Introduction:

This document outlines the SPIRE instrument requirements on the detector cooling system.  The cooling system comprises the 3He sorption cooler; the thermal links between the cooler cold tip and the detectors; the associated thermometry and any active temperature control circuitry.  Most, but not all, of the requirements apply to the cooler.

The requirements set out here come from discussions between B. Swinyard, L. Duband, B. Collaudin and M. Griffin.  Any comments or queries should be addressed to B. Swinyard (B.M.Swinyard@rl.ac.uk)

Provisional Instrument Requirements:

| Temperature at the detectors | Nominal 300 mK |
| --- | --- |
| Operating temperature control | Desirable to be able to vary the temperature of the detectors up to 320 mK and below 300 mK if this is permitted by the temperature drop across the thermal link. The evaporator cold tip temperature can be varied by heating the sorption cooler. Electronic control shall be provided to do this in the flight electronics. |
| Temperature drop across thermal link between detectors and evaporator cold tip | Maximum of 25 mK |
| Temperature drift | The temperature of the evaporator cold tip should not drift by more than 10 mK/h |
| Temperature fluctuations at the evaporator cold tip | No more than 150 nK Hz-1/2 in a frequency band from 0.1-100 Hz. |
| System low frequency temperature stability with active temperature control | TBD nK at 0.015 Hz at a maximum power dissipation of TBD W |
| Heat lift at detectors | Minimum of 10 W at 300 mK |
| Hold time | Minimum 46 hours |
| Recycle time | Maximum 2 hours |
| Thermal interface | Pumped liquid helium tank at 1.8 K for both sorption pump and evaporator |
| Thermal load onto He bath during cold operation | Maximum 1 mW |
| Time averaged thermal load onto He bath for 48 hour cycle | Maximum 3 mW (includes 20% margin) |
| Mass – including support structure | 0.6 kg (includes 20%) margin (this will be revisited if more mass is required to mount the cooler from 4-K) |
| Maximum envelope | 200x100x100 mm |
| Mechanical interface | Preferred interface is with the instrument 4-K structure – sketches below indicate how this might be achieved. |
| Preferred orientation | Horizontal with long axis along S/C Y-axis and evaporator at –Y end (see sketch) |
| Thermometers | Thermometers shall be provided on the cooler as set out in the table below. The absolute temperature measurement on the evaporator cold tip shall be 0.5% (<1.5 mK) with a resolution of TBD mK. Thermometers of the same specification shall also be provided on each detector array. |
| Sorption pump heater | The baseline design has a heater resistance of 400  implying a current of up to 20 mA for recycling. It is desirable that this heater resistance is increased so that the allowable resistance of the cryo-harness wiring can, in turn, be increased. The maximum resistance of the heater that can be driven by 28 V is about 5 k |
| Gas gap heat switches | It is noted that these are a potential single point failure in the instrument operation. Provision of some redundancy (i.e. doubling them up) is desirable but not at the expense of severe limitations on the cooler performance. |
| Ground Operation | The cooler must be capable of full operation on the ground, including recycling, when the instrument is in its normal orientation i.e. +Y horizontal and +X vertical and pointing skyward. Further it must be capable of operating with the instrument rotated to up to 90o about the S/C Y-axis (see sketch) |

|  | Instrument: SPIRE 4.3-K to 300-K interface | Name | No. of Cond. | No. of shields | Max. allowed Res. | Current | Duty Cycle | Max. Line Volt | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ID | Signal definition |  |  |  | (() | (A) | (t*T) | (V) |  |
| 14 | Pump heater (main) | PH_M | 2 | 0 | 10 TBC | 1.4E-2 | 0.014 | TBD | Br. AWG38 |
| 15 | Pump heater (red.) | PH_R | 2 | 0 | 10 TBC | 0.0E+0 | 0 | TBD | Br. AWG38 |
| 16 | Pump therm. (main) | PT_M | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 17 | Pump therm. (red.) | PT_R | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 18 | Evap. therm. (main) | ET_M | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 19 | Evap. therm. (red.) | ET_R | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 20 | Pump heat SW heater (main) | PHSWH_M | 2 | 0 | 10 TBC | 2.0E-3 | 0.96 | TBD | Br. AWG38 |
| 21 | Pump heat SW heater (red.) | PHSWH_R | 2 | 0 | 10 TBC | 0.0E+0 | 0 | TBD | Br. AWG38 |
| 22 | Evap. heat SW heater (main) | EHSWH_M | 2 | 0 | 10 TBC | 2.0E-3 | 0.04 | TBD | Br. AWG38 |
| 23 | Evap. heat SW heater (red.) | EHSWH_R | 2 | 0 | 10 TBC | 0.0E+0 | 0 | TBD | Br. AWG38 |
| 24 | Pump heat SW therm. (main) | PHSWT_M | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 25 | Pump heat SW therm. (red.) | PHSWT_R | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 26 | Evap. heat SW therm. (main) | EHSWT_M | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
| 27 | Evap. heat SW therm. (red.) | EHSWT_R | 4 | 1 | 1000 | 1.0E-5 | 1 | TBD | SST AWG38 |
|  | TOTAL |  | 44 | 8 |  |  |  |  |  |

Wiring table from IID-B