Calibrators

Subsystem Development Plan

Document Ref.: SPIRE-QMW-PRJ-000453

Issue: 2.0

Prepared by: Pete Hargrave

Alex Ellery

Last Modified on: 19 December 2000

Approved by:

Distribution list

| SPIRE-Project |  | Ken King |  |
| --- | --- | --- | --- |
|  |  | Colin Cunningham |  |
|  |  | Bruce Swinyard |  |
|  |  | Matt Griffin |  |
| QMW |  | Peter Ade |  |
|  |  | Pete Hargrave |  |
|  |  | Alex Ellery |  |
|  |  | Raul Hermoso |  |

Update history

| Date | Version | Remarks |
| --- | --- | --- |
| 03 April 2000 | Draft 1.0 | First draft. No detailed design available. |
| 14 July 2000 | Issue 1.0 | First issue |
| 19 December 2000 | Issue 2.0 | Heavily revised plan based on new model philosophy |

Table of Contents

1.	Scope	4

2.	Documents	4

2.1.	Applicable documents TBW – pay no attention to those listed	4

2.2.	Reference documents TBW	4

2.3.	Glossary	4

3.	Subsystem Description	5

3.1.	Description	5

3.1.1.	PCAL	5

3.1.2.	SCAL	6

3.2.	Requirements	7

3.3.	Interfaces	7

3.4.	Deliverables	8

4.	Constraints	8

4.1.	Technical constraints	8

4.1.1.	PCAL performance requirements	8

4.1.2.	SCAL performance requirements	8

4.2.	System constraints	9

4.2.1.	PCAL system requirements	10

4.2.2.	SCAL system requirements	10

4.3.	Calendar constraints	11

5.	Work Description	11

5.1.	Model philosophy	11

5.1.1.	PCAL	11

5.1.2.	SCAL	11

5.2.	Organisation	12

5.3.	Preliminary design phase	12

5.3.1.	Thermal source testing & development	12

5.3.2.	Modeling - SCAL	13

5.3.3.	Prototyping – PCAL & SCAL	13

5.3.4.	Testing – PCAL & SCAL	13

5.4.	Procurement of long lead-time components	13

5.5.	Detailed design phase	14

5.6.	CQM manufacture & test	14

5.7.	PFM & FS manufacture & test	14

6.	Risk analysis	15

7.	Verification plan	16

7.1.	PCAL	16

7.2.	SCAL	17

8.	Development calendar & schedule	18

9.	Calibrators sub-system schedule	18

# Scope

This document describes the development plan for the FIRST/SPIRE calibrators subsystem.

# Documents

## Applicable documents TBW – pay no attention to those listed

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD1 |  |  |  |  |
| AD2 | Interface control documents – Photometer calibrator/chopper Spectrometer calibrator/structure | P.Hargrave, C.Cunningham, B.Winter | SPIRE/ICD/1.5/1.6.1 SPIRE/ICD/1.1/1.6.2 |  |
| AD3 | SPIRE instrument – Calibrators subsystem specification document | P.Hargrave |  | April 2000 |

## Reference documents TBW

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| RD1 | Instrument Requirements Document | B.M.Swinyard | SPIRE-RAL-PRJ-000034 Issue 0.30 | May 2000 |
| RD2 | Instrument development plan |  |  |  |

## Glossary

| AD Applicable Document | FS Flight Spare |
| --- | --- |
| ATC Astronomy Technology Centre | FTS Fourier Transform Spectrometer |
| CBB Cryogenic Black Body | GSFC Goddard Space Flight Centre |
| CEA Commissariat a l’Energie Atomique | LAS Laboratoire d’Astronomie Spatiale |
| CDR Critical Design Review | MGSE Mechanical Ground Support Equipment |
| CNES Centre National des Etudes Spatiale | MSSL Mullard Space Science Laboratory |
| CoG Centre of Gravity | NA Not Applicable |
| CQM Cryogenic Qualification Model | OGSE Optical Ground Support Equipment |
| DDR Detailed Design Review | PCAL Photometer CALibrator |
| DESPA Department des Etudes SPAtiales | PFM ProtoFlight Model |
| DM Development Model | RAL Rutherford Appleton Laboratory |
| DRCU Digital Read-out and Control Unit | RD Reference Document |
| EGSE Electrical Ground Support Equipment | SCAL Spectrometer CALibrator |
| FIRST Far InfraRed Space Telescope | WE Warm Electronics |
| FPU Focal Plane Unit |  |

# Subsystem Description

The calibrators subsystem comprises the photometer calibrator (PCAL) and the spectrometer calibrator (SCAL).

## Description

### PCAL

The purpose of the photometer calibrator is to provide a repeatable signal for monitoring of detector health and responsivity for ground testing and in-flight operation. It is NOT an absolute calibrator, but may be useful as part of the overall calibration scheme. The baseline design consists of a thermal source inside an integrating cavity, the body of which will be at 4-K. The cavity will have a light pipe output with a 1-mm diameter aperture, as shown in figure (1). The calibrator will be located behind the beam steering  mirror (M4) at an image of the telescope secondary mirror, as shown in figure (2).

The fraction of M4 area obscured will be 0.2%. The limit on the calibrator aperture is set by the ratio of the telescope secondary to primary mirror diameters.

Figure 1 Schematic of photometer calibrator

Figure 2 Location of photometer calibrator (PCAL)

### SCAL

The purpose of the spectrometer calibrator is to null the telescope emission by mimicking its spectrum and brightness in the second input port of the FTS. The telescope is assumed to be at 80-K with overall wavelength-independent emissivity ( = 0.04. The overall emissivity of the system is assumed to be uncertain by a factor of 2 (actual value will not be known before launch). The baseline design, shown in figure (3), is the use of a heated black plate (or possibly a metallised film), together with a “hot” source in an integrating cavity with light pipe, to uniformly illuminate the pupil. A neutral density filter may be used to dilute the emission. The calibrator will be located at the second input port to the FTS, at an image of the telescope pupil (diameter = 30 mm) as shown in figure (4).

Throughout this document, the heated plate will be referred to as SCAL-flood, and the “hot” PCAL type source shall be referred to as SCAL-point.

Figure 3 Conceptual design of spectrometer calibrator

Figure 4 Location of spectrometer calibrator

## Requirements

Covered in 4.1 (tech. constraints)

## Interfaces

List of interfaces & reference to appropriate ICDs (to be written?)

## Deliverables

In the table below, the dates in black are those imposed by the project, and the green/red dates are estimated delivery dates. Green indicates an early/on-schedule delivery, while red indicates a possible late delivery.

Table 1 List of deliverables for each instrument model

| Item | STM | CQM | PFM | FS |
| --- | --- | --- | --- | --- |
| PCAL(1) | Mass estimate only – date TBD | 01/03/02 | 28/04/03 | 17/10/03 |
|  |  | 19/02/02 | 29/04/03 | 01/10/03 |
| SCAL(2) | 10/4/02 | 1/10/02 | 1/10/03 | 10/11/04 |
|  | 29/10/01 | `28/05/02 | 19/08/03 | 07/01/04 |

- These will be delivered to ATC

- These will be delivered to RAL

# Constraints

## Technical constraints

The performance requirements for PCAL and SCAL are listed below, and are extracted from RD1.

### PCAL performance requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALP-R01 | Nominal operating output | Equivalent to T=40 K for 200<<700 mm |
| IRD-CALP-R02. | Operating range | 4-80 K for 200<<700 mm commandable in 256 (TBC) steps |
| IRD-CALP-R03 | Equivalent obscuration of aperture through BSM mirror | <0.2%. Actual size is referred to the telescope secondary mirror image at the position of the beam steering mirror. |
| IRD-CALP-R04 | Speed of response | Requirement 150 ms. Goal 30 ms |
| IRD-CALP-R05 | Repeatability | RMS better than 1% over 20 operations. Drift less than10% over lifetime of the mission. |
| IRD-CALP-R06 | Operation | Nominally once per hour for no more than 10 seconds. |
| IRD-CALP-R07 | Frequency | Continuously or pseudo continuously variable between 0 and 5 Hz. |

### SCAL performance requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALS-R01 | Radiated spectrum | Null the central maximum to accuracy of 5% (goal 2%) [TBC] Replicate the dilute spectrum of the telescope to an accuracy of better than 20% (goal 5%) [TBC] over 200-400 mm. |
| IRD-CALS-R02 | Beam pattern | Replicate the appropriate beam pattern at the second input port pupil image |
| IRD-CALS-R03 | Adjustability | Zero - maximum in 256 steps |
| IRD-CALS-R04 | Uniformity | The uniformity of the intensity from the calibration source across the second input port pupil image shall be better than TBD% |
| IRD-CALS-R05 | Repeatability and drift | The output intensity of the calibration source shall drift by no more than 1% over one hour of continuous operation. The absolute change in the output intensity of the source shall be no more than 15% over the mission lifetime |
| IRD-CALS-R06 | Operation | The calibration source shall be capable of continuous operation for periods of up to 2 hours with no loss of operational performance. |
| IRD-CALS-R07 | Number of operations | The calibration source shall be capable of up to 12000 operational cycles |

## System constraints

The system requirements for PCAL and SCAL are listed below, and are extracted from RD1.

Additional overall system constraints are:-

- SPIRE orbital lifetime = 4.25 years

- PCAL volume envelope

- PCAL mass (including margin) = TBD (info. from ATC needed)

- PCAL vibration level = TBD (from Berend)

- PCAL shock level = TBD (from Berend)

- SCAL volume envelope (see accompanying table -TBD)

- SCAL mass (including margin) = TBD (info. from ATC needed)

- SCAL vibration level = TBD (from Berend)

- SCAL shock level = TBD (from Berend)

- Cleanliness = class 10,000??

- Transit loads = ??

- Storage = 3 years

- Bake-out temperature on AIV integration = 80(C for 48 hours

### PCAL system requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALP-R08 | Interface | The calibrator will be integrated into the beam steering mechanism. |
| IRD-CALP-R09 | Volume envelope | 30 x 15 x 10 mm |
| IRD-CALP-R10 | Thermal isolation | The temperature of the surrounding structure (including the beam steering mirror) shall rise by no more than 1 K after 10 seconds when the calibrator is operated unmodulated at nominal power output. |
| IRD-CALP-R11 | Operating temperature | 4-K |
| IRD-CALP-R12 | Cold power dissipation | Less than 2 mW when operated unmodulated at nominal power output. |
| IRD-CALP-R13 | Warm power dissipation | Less than TBD W when operated unmodulated at nominal power output |
| IRD-CALP-R14 | Operating voltage | Less than 28 V at input power level of 5 mW |
| IRD-CALP-R15 | Redundancy | Cold redundancy for the thermal source |

### SCAL system requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALS-R08 | Operating voltage | No more than 28 V DC |
| IRD-CALS-R09 | Power dissipation in the focal plane | No more than 5 mW with a goal of 2 mW |
| IRD-CALS-R11 | Envelope | 50x50x70 mm (TBC) |
| IRD-CALS-R12 | Thermal isolation | The surrounding structure of the calibrator shall rise in temperature by no more than TBD K after one hour of continuous operation |
| IRD-CALS-R13 | Operating temperature | 4 K |
| IRD-CALS-R14 | Redundancy | Fully redundant systems shall be provided for the active elements. |

## Calendar constraints

The major SPIRE project dates are (RD2):-

| PDR | 26/27 Jun 2000 |
| --- | --- |
| DDR – Interface Review | 28/29 Nov 2000 |
| DDR – Calibrators | Jan 2001 (TBC) |
| STM SCAL delivery to RAL | 10 Apr 2002 |
| STM PCAL delivery to ATC | 21 Sept 2001 |
| CQM PCAL delivery to ATC | 01 Mar 2002 |
| CQM SCAL delivery to RAL | 01 Oct 2002 |
| CDR | 3 Feb 2003 |
| SPIRE CQM delivery to ESA | Apr 2003 |
| PFM PCAL delivery to ATC | 28 Apr 2003 |
| PFM SCAL delivery to RAL | 01 Oct 2003 |
| FS PCAL delivery to ATC | 17 Oct 2003 |
| FS SCAL delivery to RAL | 10 Nov 2004 |
| SPIRE FS delivery to ESA | Dec 2005 |
| FIRST launch | 2007 |

# Work Description

## Model philosophy

The model philosophy is compliant with the SPIRE project.

### PCAL

The following models will be produced:-

- PCAL DM – several devices will be built and evaluated to produce a developmental model (DM) which will be tested and form the basis of our CQM design

- PCAL STM – This will be just a mass dummy. This may be provided by the ATC (TBD)

- PCAL CQM – This device will be provided to the ATC with limited functionality (no redundancy) to be incorporated into the BSM assembly for CQM testing

- PCAL PFM – This is the flight device

- PCAL FS – This is the flight spare device

### SCAL

The following models will be produced:-

- SCAL DM - several devices will be built and evaluated to produce a developmental model (DM) which will be tested and form the basis of our CQM design

- SCAL STM – This device will have limited functionality, but will be thermally and structurally representative of the flight device. This will be delivered to RAL for STM tests.

- SCAL CQM - This device will be provided to RAL with limited functionality (no redundancy) to be incorporated into SPIRE for CQM testing

- SCAL PFM – This is the flight device

- SCAL FS – This is the flight spare

## Organisation

Inputs, outputs & responsibilities.

- QMW to test candidate Beeman devices for suitability

- QMW to design, manufacture and test alternative options in parallel in case Beeman devices fail to conform to requirements

- QMW to design, manufacture and test the PCAL and SCAL structure

- QMW to integrate calibrators to calibrator housing

- QMW to coordinate with ATC as PCAL is an integral part of the BSM assembly

- QMW to ensure qualification, acceptance and calibration of PCAL and SCAL

- QMW responsible for transport of SCAL to RAL and PCAL to the ATC

- Calibrators to be integrated to FPU with joint responsibility between QMW, RAL, MSSL and ATC

## Preliminary design phase

The PDR freezes the technical specifications/requirements, while the interfaces are frozen at the interface review.

### Thermal source testing & development

Thermal sources for PCAL and SCAL-point are currently being developed & tested at QMW.  These are:-

- GSFC devices – Tests have shown that these devices do not meet the requirements of SPIRE in terms of power dissipation vs. photometric power. It is unlikely that these devices will be investigated further.

- Haller-Beeman Associates – Two devices have been tested at QMW so far, and both meet the SPIRE requirements for optical power vs electrical power dissipation. However, they do not yet meet the speed requirement, but QMW is working with Jeff Beeman to improve this in the next batch of devices by making some simple modifications. These devices have already been qualified for flight on MIPS.

- QMW devices – QMW are planning to develop thermal sources in-house. These are currently in the modeling phase.

### Modeling - SCAL

A MathCad model has been produced for simple thermal/mechanical analysis of design concepts. The thermal analysis is compared to results from ThermXL (Alstom technology).This is used as the first step in the design process.

Detailed thermal analysis of candidate designs will then be carried out by RAL using ESATAN. Detailed mechanical FE analysis will be carried out at QMW using the IDEAS Master Series CAD package.

### Prototyping – PCAL & SCAL

Prototype devices will be built to verify design assumptions and highlight problem areas.

### Testing – PCAL & SCAL

All prototype devices will undergo the following tests as a minimum:-

- Thermal tests

  - Device temperature vs power to SOB/structure (i.e. thermal conductance of supports & wires from heated element to 4K)

  - Thermal time constants

- Electrical tests

  - Impedance vs device temperature

  - Device temperature vs applied power

- Mechanical tests

  - Resonant frequency (warm)

- Photometric tests

  - Emitted in-band power vs applied power

Prototype devices will evolve into developmental models of PCAL and SCAL which will then undergo more rigorous tests, which include all of the above as well as repeated thermal cycling and accelerated lifetime tests.

One prototype device each of SCAL-flood and PCAL will also be tested to electrical and mechanical failure.

## Procurement of long lead-time components

Once a preliminary design has evolved, long lead-time items will be procured. These items are likely to comprise mainly of space qualified thermometers, heaters and connectors.

Note – need to get list of CPP items to Ken

## Detailed design phase

The DM SCAL will be sent to RAL as the STM SCAL for structural & thermal tests in the SPIRE STM. The results of these tests will be taken into account for any re-design needed for the CQM SCAL.

Results of tests on the DM PCAL will be used for the detailed PCAL CQM design.

The schedule allows the DM PCAL to be sent to ATC as the STM PCAL, rather than just a mass dummy.

## CQM manufacture & test

The detailed design is presented at the Detailed Design Review. The DDR must have happened before CQM manufacture can begin. To verify the design, a complete qualification and lifetests are to be conducted.

At least two models for each of the photometer and spectrometer calibrator are necessary. The two models are the development model (DM) and the cryogenic qualification model (CQM). The final versions of the PCAL-DM and SCAL-DM are deliverable for the STM to ATC and RAL respectively.

PCAL-CQM and SCAL-CQM are deliverable.

The DM models are used to qualify the design, to conduct lifetests, and as active components (rather than mass dummies) in the STM.

The CQM models are to be qualified but do not undergo lifetests.

The DM and the CQM include all the functions of the flight design, except redundancy (TBD).

The design verification tests include:-

- Verification of the basic mechanical parameters (Mass, stiffness, resonance frequencies).

- Performance verification.

- Qualification tests .

- Lifetests.

After the PCAL and SCAL CQM delivery, the SPIRE CQM is tested at project level.

The results of the qualification tests are to be presented at the SPIRE CDR which is the start point of the PFM and FS manufacture.

Then, the SPIRE CQM is delivered to ESA for cryogenic tests of the FIRST FPU.

## PFM & FS manufacture & test

Following the lifetests and SPIRE CQM tests, some modifications may have to be implemented in the design.

The design changes are to be implemented in the flight design and be validated using the DM.

The PCAL and SCAL PFM’s are then manufactured and undergo the acceptance tests and performance verification.

The FS models are duplicates of the PCAL and SCAL PFM and are manufactured at the same time as the PFM. The PCAL and SCAL FS models undergo the acceptance tests and performance verification after the PFM delivery.

# Risk analysis

| Risk | Likelihood | Impact | Preventative Action | Notes |
| --- | --- | --- | --- | --- |
| Beeman device failure | Low | No photometer calibrator. Spectrometer calibrator performance severely compromised. | None for PCAL. The FTS calibrator also has a heated black plate or metallised film and so an element of redundancy is inherent in the design although performance may be compromised. | Failure of the photometer calibrator would not lead to loss of photometer science capabilities but would prevent ground checking of detectors with modulated photometric signals when the cryostat is closed. It would compromise the ease and convenience of instrument monitoring and calibration in flight, resulting in less efficient operation. The development programme may not result in device with low enough power dissipation for required radiated power. This could restrict the frequency of operation in flight or require an undesirable reduction in the operating signal levels. Failure of the FTS calibrator would lead to inability to null the telescope offset with significant impact on dynamic range – a severe impact. |
| FTS calibrator plate has too high electrical power dissipation for the required radiated power | Low | Increased FPU dissipation or restricted FTS operation (decreased dynamic range) | Qualification testing | Heat dissipation on the FPU could become unacceptable |
| Delay in thermal source testing/prototyping | Medium | Delay in PCAL DM prototyping and testing | None – insufficient human resources | Delay in PCAL CQM delivery to ATC |
| Delay in SCAL DM prototyping | Medium | Delay in SCAL DM testing | None – insufficient human resources | Delay in SCAL STM and CQM delivery to RAL |

# Verification plan

The verification plan must be compliant with the project verification plan [Refs – AD??, RD??].

300K vibrations are conducted at QMW.

Cryovibrations are conducted at RAL (TBC).

Vacuum cycles, soak cycles, thermal cycles are conducted at QMW.

Lifetime tests are conducted at QMW.

EMI/EMC tests are conducted at TBD.

Microphonics tests are conducted at TBD.

Performance tests are conducted at QMW.

In the tables below,

| X = | a real test is conducted |
| --- | --- |
| A = | an analysis is conducted |
| NA = | Non applicable |

## PCAL

|  | PCAL-DM | PCAL-CQM | PCAL-PFM | PCAL-FS |
| --- | --- | --- | --- | --- |
| Mass measurement | X | X | X | X |
| CoG measurement | X | X | X | X |
| Spectral measurements | X | X | X | X |
| Absolute photometric calibration | X | X | X | X |
| Cross calibration with CBB sources | X | X | X | X |
| Power consumption | X | X | X | X |
| Vibrations 300K | X | X | X | X |
| Vibrations 4K | X | X | X | X |
| Thermal/Vacuum cycle | X | X | X | X |
| Accelerated lifetime (12,000 operations) | X |  |  |  |
| Radiation tolerance** | TBD | TBD | TBD | TBD |
| Microphonics | N/A | N/A | N/A | N/A |
| EMI / EMC | X | A(*) | X(*) | A(*) |

(*) : EMI/EMC tests are to be conducted on the PFM only if design changes have occurred.

(**) : The radiation tolerance is verified by analysis only, taking into account the materials involved.

## SCAL

|  | SCAL-DM | SCAL-CQM | SCAL-PFM | SCAL-FS |
| --- | --- | --- | --- | --- |
| Mass measurement | X | X | X | X |
| CoG measurement | X | X | X | X |
| Spectral measurements | X | X | X | X |
| Absolute photometric calibration | X | X | X | X |
| Cross calibration with CBB sources | X | X | X | X |
| Power consumption | X | X | X | X |
| Vibrations 300K | X | X | X | X |
| Vibrations 4K | X | X | X | X |
| Thermal/Vacuum cycle | X | X | X | X |
| Accelerated lifetime (12,000 operations) | X |  |  |  |
| Radiation tolerance | TBD | TBD | TBD | TBD |
| Microphonics | N/A | N/A | N/A | N/A |
| EMI / EMC | X | A(*) | X(*) | A(*) |

(*) : As EMI/EMC is verified on the CQM, no further verification are conducted on the subsequent models.

(**) : The radiation tolerance is verified by analysis only, taking into account the materials involved.

# Development calendar & schedule

The major project milestones pertinent to the calibrators sub-system are:-

STM SCAL to RAL		10/04/02

CQM PCAL to ATC		01/03/02

CQM SCAL to RAL 	01/10/02

PFM PCAL to ATC		28/04/03

PFM SCAL to RAL		01/10/03

FS PCAL to ATC		17/10/03

FS SCAL to RAL		10/11/04

Internal milestones and delivery dates are:-

Candidate thermal source tests complete	4/4/01

BSM/PCAL interface frozen			5/4/01

SCAL STM delivery to RAL			29/10/01

CQM PCAL to ATC					19/02/02

CQM SCAL to RAL					06/05/02

PFM PCAL to ATC					29/04/03

PFM SCAL to RAL					19/08/03

FS PCAL to ATC					01/10/03

FS SCAL to RAL					07/01/04

# Calibrators sub-system schedule

The schedule shown is realistic, with a lot of margin built into estimated task durations. However, it should be noted that this is extracted from the overall QMW SPIRE schedule (SPIRE_QMW_ALL_v1.2.mpp), which at present has a lot of tasks running in parallel. QMW need to go through this schedule and allocate resources to tasks, which may cause the calibrator schedule to change slightly to accommodate other tasks, but the delivery dates shouldn’t be affected to any great degree. This document will be updated once the resource allocation exercise has taken place.

# Text boxes

Light pipe

Thermal source

Integrating cavity

BSM

Draft

For approval