|  | FIRST SPIRE | Ref : SPIRE-QMW-PRJ-000453 Author : P. Hargrave | Page : 1 / 1 Date : 14 July 2000 |
| --- | --- | --- | --- |
| SPIRE Calibrators subsystem specification |

Issue 1.0

Distribution List :

| C.Cunningham | ATC |  |  |
| --- | --- | --- | --- |
| B.M.Swinyard | RAL |  |  |
| K.King | RAL |  |  |
| M.Griffin | QMW |  |  |
| P.Ade | QMW |  |  |
| P.Hargrave | QMW |  |  |
| S.H.Moseley | GSFC |  |  |
| M.Jhabvala | GSFC |  |  |

_________________________________________________________________________________

Queen Mary & Westfield College, Astrophysics Laboratory, Mile End Road, London E1 4NS

Tel. +44 (20) 7415 3757  Fax. +44 (20) 8980 0986

Update

| Date | Version | Remarks |
| --- | --- | --- |
| 03 April 2000 | Draft 1.0 | First draft. No detailed design available. |
| 14 July 2000 | Issue 1.0 | First issue |

| Host system | Windows NT 4.0 |
| --- | --- |
| Word Processor | Microsoft Word 2000 |
| File | F:\Projects\First\Spire\Management\Development_plans\Spire Filters subsystem specification.doc |

Table of contents

1.	Scope of the document	5

2.	Documents	5

2.1.	Applicable documents	5

2.2.	Glossary	5

3.	Subsystem description	6

3.1.	Photometer calibrator	6

3.2.	Spectrometer calibrator	6

3.3.	Thermal sources	8

1.1.1.	Modified ISO style sources	8

1.1.2.	SIRTF/IRAC type devices	8

4.	Performance & System requirements	9

4.1.	Photometer calibrator (PCAL) performance requirements	9

4.2.	Photometer calibrator (PCAL) system requirements	9

4.3.	Spectrometer calibrator (SCAL) performance requirements	10

4.4.	Spectrometer calibrator (SCAL) system requirements	10

5.	Technical requirements	10

6.	Operation	11

6.1.	Reliability	11

6.2.	Lifetime	11

6.3.	Environment	11

6.4.	Alignment	11

6.5.	Stability	11

6.6.	EMI	11

6.7.	Attitude disturbance	11

6.8.	Safety	11

6.9.	Operating modes	11

6.10.	Telemetry	11

6.10.1.	Telecommands	11

7.	Design & construction	12

7.1.	Interface requirements	12

7.2.	Reference axes	12

7.3.	Design philosophy	12

7.4.	Configuration	12

7.5.	Cleanliness	12

7.6.	Venting	12

7.7.	Accessibility and servicing	12

7.8.	Materials	12

7.9.	Finishes	12

7.10.	Processes	12

7.11.	Standard & commercial parts	12

7.12.	Standard specification	12

7.13.	Part identification	12

7.14.	Interchangeability	12

7.15.	Electrical	12

7.16.	High energy particle shielding	12

7.17.	Resource budget	12

8.	Test & Calibration	12

8.1.	Functional test	12

8.2.	Performance test	12

8.3.	Verification requirements	12

8.4.	Mechanical environment test	13

8.5.	System dynamic test	13

8.6.	Thermal vacuum performance test	13

8.7.	Life test	13

9.	Logistical requirements	13

9.1.	Handling equipment & procedures	13

9.2.	Transit equipment & procedures	13

9.3.	Protection equipment & procedures	13

9.4.	Ancillary equipment & procedures	13

# Scope of the document

This document describes the specifications for the photometer and spectrometer calibration sources for SPIRE. Detailed designs for these components do not exist at present, so this document, in it’s present draft, is limited to a presentation of conceptual designs. This document covers (or will cover):-

- The instrument requirements on the calibrators

- The performance specifications and design drivers for the photometer and spectrometer calibrators

- The manufacturing procedures and QC/QA

- Risk analysis and testing procedures

The issue of calibrator mounting and integration is covered by document AD2, and work breakdown and scheduling is covered by document AD3

# Documents

## Applicable documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD1 | Instrument Requirements Document | B.M.Swinyard | SPIRE-RAL-PRJ-000034 Issue 0.30 | May 2000 |
| AD2 | Interface control documents – Photometer calibrator/chopper Spectrometer calibrator/structure | P.Hargrave, C.Cunningham, B.Winter | SPIRE/ICD/1.5/1.6.1 SPIRE/ICD/1.1/1.6.2 |  |
| AD3 | SPIRE instrument – Calibrators subsystem development plan | P.Hargrave |  | April 2000 |

## Glossary

| AD | Applicable Document |  |  |
| --- | --- | --- | --- |
|  |  | MGSE | Mechanical Ground Support Equipment |
| CDR | Critical Design Review | MSSL | Mullard Space Science Laboratory |
|  |  | NA | Not Applicable |
| CoG | Center of Gravity | OGSE | Optical Ground Support Equipment |
| CQM | Cryogenic Qualification Model | PFM | ProtoFlight Model |

# Subsystem description

At the time of writing this draft, the calibrator subsystem is still under development.

## Photometer calibrator

The purpose of the photometer calibrator is to provide a repeatable signal for monitoring of detector health and responsivity for ground testing and in-flight operation. The calibrator is not required to provide absolute calibration, but may be useful as part of the overall calibration scheme. The baseline design consists of a thermal source inside an integrating cavity, the body of which will be at 4-K. The cavity will have a light pipe output with a 1-mm diameter aperture, as shown in figure (1). The calibrator will be located behind the beam steering  mirror (M4) at an image of the telescope secondary mirror, as shown in figure (2). The calibrator will form an integral part of the beam steering mirror subsystem.

The fraction of M4 area obscured will be 0.2%. The limit on the calibrator aperture is set by the ratio of the telescope secondary to primary mirror diameters.

Figure 1 Conceptual design of photometer calibrator

Figure 2 Location of photometer calibrator (PCAL)

## Spectrometer calibrator

The purpose of the spectrometer calibrator is to null the telescope emission by mimicking its spectrum and brightness in the second input port of the FTS. The telescope is assumed to be at 80-K with overall wavelength-independent emissivity ( = 0.04. The overall emissivity of the system is assumed to be uncertain by a factor of 2 (actual value will not be known before launch). The baseline design, shown in figure (3), is the use of a heated black plate (or possibly a metallised film), together with a “hot” source in an integrating cavity with light pipe, to uniformly illuminate the pupil. A possible alternative design is a thermal source with Winston Cone optic to present uniform illumination across the pupil as shown in figure (4). A neutral density filter will be used to dilute the emission. The calibrator will be located at the second input port to the FTS, at an image of the telescope pupil (diameter = 30 mm) as shown in figure (5).

Figure 3 Baseline concept for spectrometer calibrator

Figure 4 Alternative design of spectrometer calibrator

Figure 5 Location of spectrometer calibrator (SCAL)

## Thermal sources

Two types of thermal source are under consideration for the photometer calibrator and for the hot source for the spectrometer calibrator.

### Modified ISO style sources

Devices based on the design of the Infra Red Labs TRS sources, as used on ISO, can be manufactured in-house with improved reliability and reduced power dissipation. These are currently under development at QMW.

### SIRTF/IRAC type devices

Micromachined Si devices have been developed by NASA/GSFC for use on SIRTF/IRAC. Further development and testing is needed for use on SPIRE. This is in progress at QMW.

# Performance & System requirements

The performance and system requirements for PCAL and SCAL are listed below, and are extracted from AD1.

## Photometer calibrator (PCAL) performance requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALP-R01 | Nominal operating output | Equivalent to T=40 K for 200<<700 mm |
| IRD-CALP-R02. | Operating range | 4-80 K for 200<<700 mm commandable in 256 (TBC) steps |
| IRD-CALP-R03 | Equivalent obscuration of aperture through BSM mirror | <0.2%. Actual size is referred to the telescope secondary mirror image at the position of the beam steering mirror. |
| IRD-CALP-R04 | Speed of response | Requirement 150 ms. Goal 30 ms |
| IRD-CALP-R05 | Repeatability | RMS better than 1% over 20 operations. Drift less than10% over lifetime of the mission. |
| IRD-CALP-R06 | Operation | Nominally once per hour for no more than 10 seconds. |
| IRD-CALP-R07 | Frequency | Continuously or pseudo continuously variable between 0 and 5 Hz. |

## Photometer calibrator (PCAL) system requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALP-R08 | Interface | The calibrator will be integrated into the beam steering mechanism. |
| IRD-CALP-R09 | Volume envelope | 30 x 15 x 10 mm |
| IRD-CALP-R10 | Thermal isolation | The temperature of the surrounding structure (including the beam steering mirror) shall rise by no more than 1 K after 10 seconds when the calibrator is operated unmodulated at nominal power output. |
| IRD-CALP-R11 | Operating temperature | 4-K |
| IRD-CALP-R12 | Cold power dissipation | Less than 2 mW when operated unmodulated at nominal power output. |
| IRD-CALP-R13 | Warm power dissipation | Less than TBD W when operated unmodulated at nominal power output |
| IRD-CALP-R14 | Operating voltage | Less than 28 V at input power level of 5 mW |
| IRD-CALP-R15 | Redundancy | Cold redundancy for the thermal source |

## Spectrometer calibrator (SCAL) performance requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALS-R01 | Radiated spectrum | Null the central maximum to accuracy of 5% (goal 2%) [TBC] Replicate the dilute spectrum of the telescope to an accuracy of better than 20% (goal 5%) [TBC] over 200-400 mm. |
| IRD-CALS-R02 | Beam pattern | Replicate the appropriate beam pattern at the second input port pupil image |
| IRD-CALS-R03 | Adjustability | Zero - maximum in 256 steps |
| IRD-CALS-R04 | Uniformity | The uniformity of the intensity from the calibration source across the second input port pupil image shall be better than TBD% |
| IRD-CALS-R05 | Repeatability and drift | The output intensity of the calibration source shall drift by no more than 1% over one hour of continuous operation. The absolute change in the output intensity of the source shall be no more than 15% over the mission lifetime |
| IRD-CALS-R06 | Operation | The calibration source shall be capable of continuous operation for periods of up to 2 hours with no loss of operational performance. |
| IRD-CALS-R07 | Number of operations | The calibration source shall be capable of up to 12000 operational cycles |

## Spectrometer calibrator (SCAL) system requirements

| Requirement ID | Description | Value |
| --- | --- | --- |
| IRD-CALS-R08 | Operating voltage | No more than 28 V DC |
| IRD-CALS-R09 | Power dissipation in the focal plane | No more than 5 mW with a goal of 2 mW |
| IRD-CALS-R11 | Envelope | 50x50x70 mm (TBC) |
| IRD-CALS-R12 | Thermal isolation | The surrounding structure of the calibrator shall rise in temperature by no more than TBD K after one hour of continuous operation |
| IRD-CALS-R13 | Operating temperature | 4 K |
| IRD-CALS-R14 | Redundancy | Fully redundant systems shall be provided for the active elements. |

# Technical requirements – to be completed

These requirements are design dependent. This section will be completed once the final designs are available.

Typical list:

- Axis reference system

- Alignment constraint

- Field view

- Dimensions

- Center of gravity, mass, inertia

- Stiffness

- Dynamic perturbations emitted by the subsystem

- Power consumption

- EMC emitted by the product

# Operation – to be completed once design finalised

## Reliability

## Lifetime

## Environment

## Alignment

## Stability

## EMI

## Attitude disturbance

## Safety

## Operating modes

## Telemetry

### Telecommands

# Design & construction – to be completed

## Interface requirements

## Reference axes

## Design philosophy

## Configuration

## Cleanliness

## Venting

## Accessibility and servicing

## Materials

## Finishes

## Processes

## Standard & commercial parts

## Standard specification

## Part identification

## Interchangeability

## Electrical

## High energy particle shielding

## Resource budget

# Test & Calibration – to be completed

More detail can be found in AD3.

## Functional test

## Performance test

## Verification requirements

- Mathematical modelling

- Functional tests

- Qualification tests

- Acceptance tests

- Life tests

## Mechanical environment test

## System dynamic test

## Thermal vacuum performance test

## Life test

# Logistical requirements – to be completed

## Handling equipment & procedures

## Transit equipment & procedures

## Protection equipment & procedures

## Ancillary equipment & procedures

# Text boxes

Thermal source

Light pipe

Integrating cavity