# SPIRE Beam Steering Mirror Subsystem Specification Document   V 2.0

# Ian Pain

# DISTRIBUTION LIST

| SPIRE-Project | Ken J. King |  |
| --- | --- | --- |
|  | Bruce M. Swinyard |  |
|  | Matt Griffin |  |
| UK ATC | Colin Cunningham |  |
|  | Gillian Wright |  |
|  | Ian Pain |  |
|  | David Henry |  |
|  | Brian Stobie |  |
|  | Richard Bennett |  |
|  | Tom Paul |  |
|  | Fay Hannon |  |
| LAM | Didier Ferrand |  |
|  | Dominique Pouliquen |  |
|  | Patrick Levacher |  |

# Record of Issue

| Date | Index | Remarks |
| --- | --- | --- |
| 13.Jun.00 | 1.0 | Issue for PDR |
| 06.Nov.00 | 2.0 | Update following additional technical data. Reformatting. |

# Table of contents

SPIRE Beam Steering Mirror Subsystem Specification Document   V 2	1

Ian Pain	1

DISTRIBUTION LIST	1

Record of Issue	1

Table of contents	3

Table of Figures	5

File Description	5

1.	Scope	6

2.	Documents and Glossary	6

2.1	Applicable documents	6

2.2	Reference documents	7

2.3	Glossary	7

3.	Subsystem description	8

3.1	Outline Description of the Beam Steering mirror mechanism subsystem	8

3.2	Subsystem Description	10

3.3	Block Diagram	10

3.4	Mission profile	10

3.5	Design Assumptions	11

3.5.1	Optical Scaling Factor	11

3.5.2	Definition of Axes	11

4.	Functional Requirements	12

4.1	Performance requirements	12

4.1.1	Angular Travel - Chop Axis	12

4.1.2	Angular Travel - Jiggle Axis	12

4.1.3	Minimum Step Size	12

4.1.4	Chop Frequency	12

4.1.5	Jiggle Frequency	12

4.1.6	Holding position/ Drift Constraint	12

4.1.7	Stability	13

4.1.8	Position Measurement	13

4.1.9	Settling Time	13

4.1.10	Chop repeatability	13

4.2	Technical Specifications	14

4.2.1	Mechanical Dimensions	14

4.2.2	Operating Temperature	14

4.2.3	Thermal Isolation	14

4.2.4	Cold Power Dissipation	14

4.2.5	Warm Electronics Power Dissipation	14

4.2.6	Mirror Surface Dimensions	14

4.2.7	Mirror Surface Finish	14

4.2.8	Mirror Surface Reflectivity	15

4.2.9	Mirror Surface Emissivity	15

4.2.10	Position of Rotation Axes	15

4.2.11	Orthogonality of Rotation Axes	15

4.2.12	Fail Safe (No Drive Signal) Position	15

4.2.13	Fail Safe (Mechanical Failure) Position	15

4.2.14	Mass	15

4.2.15	Cool-down time	15

4.3	Operational Specification	16

4.3.1	Operational Safety	16

4.3.2	Lifetime	16

4.3.3	Operating modes	16

4.3.4	Jiggle Mode	17

4.3.5	Chopping Mode	17

4.3.6	Scan mapping	17

4.3.7	Stare or 'holding' mode	17

4.3.8	Combinations of Modes	17

4.4	Interface requirements	18

4.4.1	Data Outputs	18

4.4.2	Data Inputs	18

4.4.3	Self induced vibration	18

4.4.4	Stray Magnetic fields	18

4.4.5	Electro-Magnetic Compatibility	18

4.4.6	ICD's	19

4.5	Design, manufacture and test requirements	20

4.5.1	Design requirements	20

4.5.2	Electronics Card Format	20

4.5.3	Product Assurance Plan	20

4.5.4	Manufacturing and testing requirements	20

4.5.4.1	Mirror Flatness	20

4.5.4.2	Mirror Reflectivity	20

4.5.5	Parts, Material and Processes	20

4.5.5.1	General	20

4.5.5.2	Materials, Processes and Mechanical parts (MPM) approval and listing	20

4.5.5.3	Cleanliness	20

4.5.5.4	Finish	20

4.5.5.5	Outgassing	21

4.5.5.6	Material selection	21

4.5.5.7	Processes	21

4.6	Logistic requirements	21

4.6.1	Storage	21

4.6.2	Handling	21

4.6.3	Transportation	21

4.6.4	ESD precautions	21

4.7	Environmental requirements	22

4.7.1	Operating environment	22

4.7.1.1	Shock	22

4.7.1.2	Vibration	22

4.7.1.3	Vacuum Level	22

4.7.1.4	Vacuum Outgassing	22

4.7.1.5	Temperature	22

4.7.1.6	Magnetic Fields	22

4.7.1.7	Survival Temperature	22

4.7.1.8	Radiation environment	22

4.8	Verification requirements	23

4.8.1	Performance Requirements	23

4.8.2	System Requirements	23

Operating Temperature	Error! Bookmark not defined.

Thermal Isolation	Error! Bookmark not defined.

Cold Power Dissipation	Error! Bookmark not defined.

Warm Electronics Power Dissipation	Error! Bookmark not defined.

Mirror Surface Dimensions	Error! Bookmark not defined.

Mirror Surface Reflectivity	Error! Bookmark not defined.

Position of Rotation Axes	Error! Bookmark not defined.

Orthogonality of Rotation Axes	Error! Bookmark not defined.

# Table of Figures

Figure 1 Photometer Layout, BSM in green, highlighted with yellow oval	8

Figure 2 SPIRE Block Diagram showing BSM, SMEC and FSFPU	10

# File Description

Created with:	 	MS Word 97 SR-1

File: 			BSM Subsystem Specification Document v2_0.doc

# Scope

This specification defines the requirements applied to the performance, design and qualification of the Beam Steering Mechanism for the SPIRE photometer.  It includes a brief description of the proposed design to meet these requirements. The complete design description is found in RD1.

The Beam Steering Mechanism forms a sub-system of the SPIRE instrument. This specification is applicable to the CQM, the PFM and the FS.

# Documents and Glossary

## Applicable documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD1 | Instrument Requirements Document | B.M.Swinyard | SPIRE-RAL-PRJ-000034 Issue 0.30 | May 2000 |
| AD2 | Structure - BSM ICD |  | SPIRE/ICD/1.1-1.5.1 |  |
| AD3 | Deleted |  |  |  |
| AD4 | Optics - BSM ICD |  | SPIRE/ICD/1.2-1.5.1 |  |
| AD5 | Baffles - BSM ICD |  | SPIRE/ICD/1.2.2-1.5.1 |  |
| AD6 | Photometer Bolometer Arrays - BSM ICD |  | SPIRE/ICD/1.4.1-1.5.1 |  |
| AD7 | Spectrometer Bolometer Arrays - BSM ICD |  | SPIRE/ICD/1.4.2-1.5.1 |  |
| AD8 | Photometer Calibration Source - BSM ICD |  | SPIRE/ICD/1.5.1-1.6.1 |  |
| AD9 | Deleted |  |  |  |
| AD10 | On Board Software - BSM ICD |  | SPIRE/ICD/1.5.1-2.6 |  |
| AD11 | Analogue Simulator - BSM ICD |  | SPIRE/ICD/1.5.1-3.1 |  |
| AD12 | Instrument Simulator - BSM ICD |  | SPIRE/ICD/1.5.1-3.3 |  |
| AD13 | Current Optical Configuration |  | PHT 126 I |  |
| AD14 | Operating Modes for the SPIRE Instrument | B.M.Swinyard | SPIRE-RAL_PRJ 000320 issue 2.2 |  |
| AD15 | BSM Sub-system Development Plan |  | SPIRE-ATC-PRJ-003 | latest |
| AD16 | IID-A | FIRST/Planck Project Team | SCI-PT-IIDA-04624 rev 1/0 | 01/09/2000 |
| AD17 | IID-B | ESA FIRST/Planck Project Team | SCI-PT-IIDB/SPIRE-02124 rev 1/0 | 01/09/2000 |
| AD18 | SPIRE mirrors specification | K.Dohlen and D.Pouliquen | LAM.PJT.SPI.SPT.200007 Ind 1 | 6.Jun.00 |
| AD19 | MCU-BSM ICD |  | TBD |  |
| AD20 | SPIRE Product Assurance plan | RAL | SPIRE-RAL-PRJ-00017 | 5.Feb.98 |

## Reference documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| RD1 | SPIRE BSM Design Description Document | I.Pain | SPIRE-ATC-PRJ-002 | latest |
| RD2 | FIRST Telescope Specification |  | PT-RQ-04761 Issue 1/A SPIRE-ESA-DOC-000195 | January 1998 |

## Glossary

| AD | Applicable Document | QMW | Queen Mary & Westfield College |
| --- | --- | --- | --- |
| ATC | UK Astronomy Technology Centre | RAL | Rutherford Appleton Laboratory |
| BSM | Beam Steering Mirror | RD | Reference Document |
| CQM | Cryogenic Qualification Model | RMS | Root Mean Square |
| DPU | Digital Processing Unit | SMEC | Spectrometer Mechanism |
| DSP | Digital Signal Processor | TBD | To Be Decided |
| FEA | Finite Element Analysis | TBC | To Be Confirmed |
| FS | Flight Spare | LAM | Laboratoire Astronomique de Marseille |
| ICD | Interface Control Document |  |  |
| PID | Proportional – Integral - Derivative |  |  |
| PFM | Proto-Flight Model |  |  |
| MAC | Multi Axis Controller |  |  |

# Subsystem description

## Outline Description of the Beam Steering mirror mechanism subsystem

The Beam Steering Mirror mechanism subsystem (BSM) is a critical  part of the SPIRE Instrument. It is used to steer the beam of the telescope on the photometer and spectrometer arrays in 2 orthogonal directions, for purposes of fully sampling the image, fine-pointing and signal modulation.

The BSM comprises 4 main deliverables :

- The cryogenic mechanism (BSMm).

- The structural interface (BSMs).

- The warm electronics (BSMe)

- Mass and optical alignment dummies as required for SPIRE system level integration, (BSMd)

The position of the BSMm & BSMs are indicated in Figure 1.

Figure 1 Photometer Layout, BSM in green, highlighted with yellow oval

The BSMs consists of an aluminium alloy mirror, nominal diameter 30mm, machined as part of the chop axis. This is mounted orthogonally within a gimbal-type frame which provides for jiggle axis motion. The axes are suspended by flex-pivot mounts.

The BSMm is a cryogenic device with nominal temperature 4K. Nominally, the chop axis provides 2.5 ° of mirror motion at 2 Hz and the jiggle axis provides  0.5° of motion at 1 Hz.

The BSM also provides an aperture through which the Photometer Calibration Source is directed towards the detector arrays.

The BSMs provides location of the BSMm on the SPIRE optical bench, and will also provide for a light tight enclosure and structural support for harnessing and thermometry. The BSMs integrates to the SPIRE Photometer Calibration Source (QMW), a baffle (RAL…TBC) and the SPIRE optical bench (MSSL). The BSMs is a cryogenic structure with nominal temperature 4K.

The BSMe provides electrical actuators used to provide motion of the mirror. Electrical transducers are used to measure the mirror position to allow control of the mirror position. The BSMe baseline design makes use of cryogenic motors and magneto-resistive sensors used in ISOPHOT. Each axis houses a rare-earth (e.g. Cobalt-Samarium) magnet moving pole piece and is driven by a motor coil fixed to the mechanism housing/structure.

The cryogenic electronics are connected to the analogue power and amplifier electronics on the Warm Electronics (WE) by a cryogenic harness which will also feed out signal cables from thermocouples on the BSMs. The BSM operates under control of the Detector Readout and Control (FSDRC) sub-system (LAM). The BSMe will be specified and designed by the UK ATC, then manufactured by LAM in conjunction with the SMEC electronics. Integration and test will be at LAM, with support from ATC.

The BSMd may comprise several actual dummies, with at least (1) an optical dummy for initial alignment work and (2) a mass-representative model for structural vibration tests. Designs for mass and optical alignment dummies will not be specified in detail until the BSMs/BSMm design is largely complete.

## Subsystem Description

The BSM subsystem consists of the following parts.

| BSMs - BSM structural support | This forms the support for the BSM cryogenic mechanism, and attaches the cryogenic mechanism to the optical bench. It also serves as a mounting point for a baffle, light tight enclosure and the photometer calibrator. |
| --- | --- |
| BSMm - BSM cryogenic mechanism | The cryogenic mechanism comprises the mirror, chop and jiggle stages, pivots, mirror structure, motors, position sensors and thermometers. |
| BSMe - BSM warm electronics | The warm electronics provide drive signals for the motors and read the signals from the position sensors. |
| BSMd – BSM mass and alignment dummy | This is a model of the cryogenic mechanism with the correct mass and a non-moving mirror. It is used for vibration testing and optical alignment |

## Block Diagram

Figure 2 SPIRE Block Diagram showing BSM, SMEC and FSFPU

## Mission profile

TBD

## Design Assumptions

### Optical Scaling Factor

A scaling factor between the movement of the mirror and movement of the beam on the sky of 1° (at the mirror) per 50 arcsec (on the sky) is assumed. This value is derived from the SPIRE optical design (AD13).

### Definition of Axes

The FIRST spacecraft axes are defined in AD1, as follows:

- The X axis is the spacecraft boresight

- The Y axis is away from the HIFI instrument

- The Z axis is towards the sun shield.

The BSM is positioned in the SPIRE optical system such that the two rotation axes of the mirror produce nominal movements of the beam in the XY plane (around the Z axis) and in the XZ plane (around the Y axis) respectively.

The rotation axes of the BSM mirror shall be referred to as follows:

- Rotation producing a movement of the beam around the spacecraft Z axis - Chop axis

- Rotation producing a movement of the beam around the spacecraft Y axis - Jiggle axis

The mirror orientation is defined by two angles - the chop angle and the jiggle angle (c,j). The nominal zero position ( (c,j) = (0,0) ) is defined as the orientation such that the centre of the photometer detector arrays are aligned along the nominal SPIRE boresight.

# Functional Requirements

## Performance requirements

These requirements are not dependent on the particular design of the BSM.

### Angular Travel - Chop Axis

| The BSM shall allow angular movement of the mirror surface through an angle of (2.4° in the mirror chop axis. (see assumption 3.5 for corresponding on-sky motion). | IRD-BSMP-R01 IRD-B SRD-R9 |
| --- | --- |
| The minimum available chop throw shall be +/- 0.1° (see assumption 3.5 for corresponding on-sky motion). | IRD-B SRD-R9 |

### Angular Travel - Jiggle Axis

| The BSM shall allow angular movement of the mirror surface through an angle of (0.6° (TBC) in the mirror jiggle axis. (see assumption 3.5 for corresponding on-sky motion). | IRD-BSMP-R02 |
| --- | --- |

### Minimum Step Size

| The BSM shall allow movements of the mirror surface in both chop and jiggle axes in minimum increments of 0.04°. (see assumption 3.5 for corresponding on-sky motion). | IRD-BSMP-R03 |
| --- | --- |

### Chop Frequency

| The BSM shall allow movements of the mirror in the chop axis at any frequency up to 2Hz. | IRD-BSMP-R04 |
| --- | --- |
| As a goal, the BSM shall allow movements of the mirror in the chop axis at any frequency up to 5Hz. At frequencies above 2Hz, a degradation in power dissipation performance and settling time of the BSM is acceptable. | IRD-BSMP-R04 |

### Jiggle Frequency

| The BSM shall allow movements of the mirror in the jiggle axis at any frequency up to 0.5Hz. | IRD-B SMP-R05 |
| --- | --- |
| As a goal, the BSM shall allow movements of the mirror in the jiggle axis at any frequency up to 1Hz. At frequencies above 0.5Hz, a degradation in power dissipation performance and settling time of the BSM is acceptable. | IRD-B SMP-R05 |

### Holding position/ Drift Constraint

| The BSM shall be capable of moving to, and holding at, any commanded position within its range of movement, to within TBD° (RMS), for periods of up to 4 hours (TBC). Note that until this requirement is more firmly defined it is assumed that the best predictions of BSM stability will suffice – see 4.1.7 | IRD-BSMP-R06 |
| --- | --- |

### Stability

| After settling, the mirror position shall remain within 0.004° (RMS) of the mean steady state position, in the frequency range 0.03 - 25Hz. | IRD-BSMP-R07 |
| --- | --- |
| As a goal, after settling, the mirror position shall remain within 0.002° (RMS) of the mean steady state position, in the frequency range 0.03 - 25Hz. | IRD-BSMP-R07 |

### Position Measurement

| The BSM shall provide measurements of the angular orientation of the mirror surface in both chop and jiggle axes to a resolution of <0.002° (TBC), at a rate of TBD Hz. | IRD-BSMP-R08 |
| --- | --- |

### Settling Time

| The angular position of the mirror surface shall be within 0.02° of the mean steady state position in less than 0.025 sec from the application of a demand in the chop axis. | IRD-BSMP-R09 |
| --- | --- |
| The angular position of the mirror surface shall be within 0.02° of the mean steady state position in less than 0.100 sec from the application of a demand in the jiggle axis. | IRD-BSMP-R09 |
| As a goal, the angular position of the mirror surface shall be within 0.02° of the mean steady state position in less than 0.050 sec from the application of a demand in the jiggle axis. | IRD-BSMP-R09 |

### Chop repeatability

| The steady state repeatability between successive chop cycles shall be less than 0.004° (RMS). | IRD-BSMP-R10 |
| --- | --- |

## Technical Specifications

### Mechanical Dimensions

| The BSM shall fit within a volume of 130x130x30 mm (excluding mounting structure). | IRD-BSMP-R09 |
| --- | --- |
| The BSM mounting structure footprint is as defined in MSSL Interface drg. A2-5264-907 | AD2 |

### Operating Temperature

| The operational temperature of the BSM shall be at Thermal Interface Level 1 per the AD16 : nominally 4K, but in the range 3.5-6.0K . | IRD-BSMP-R10 AD16 (5.7.7.1) |
| --- | --- |
| The BSM shall be capable of operation (with reduced performance) at temperatures up to 300K. | IRD-BSMP-R10 |

### Thermal Isolation

| No part of the BSM visible to the optical path shall rise by no more than 1K from the nominal operating temperature of the surrounding structure after 1 hour of operation in any mode | IRD-BSMP-R11 |
| --- | --- |

### Cold Power Dissipation

| The average power dissipation of the BSM Cryogenic Mechanism (BSMm) and the BSM Support Structure (BSMs) shall be less than 4mW in any operating mode, when operating at the temperature defined in section 0 | IRD-BSMP-R12 |
| --- | --- |

### Warm Electronics Power Dissipation

| The average power dissipation of the BSM Warm Electronics (BSMe) shall be less than TBD Watts when chopping at 2Hz in any operating mode. | IRD-BSMP-R13 |
| --- | --- |

### Mirror Surface Dimensions

| The mirror clear diameter shall be greater than 32mm (TBC). | AD13 |
| --- | --- |
| The mirror shall include a central hole of no greater than 2mm diameter (TBC) to allow the Photometer Calibrator to be seen by the detectors. | AD13 |

### Mirror Surface Finish

| The mirror surface of the BSM shall be flat to <2(m (RMS) (TBC) and shall have a surface roughness of <10nm (RMS). | AD13, AD18 |
| --- | --- |

See also 4.5.4.1.

### Mirror Surface Reflectivity

| The reflectivity of the mirror surface of the BSM shall be >99% in the wavelength range 200 - 670 (m. | AD13, AD18 |
| --- | --- |

See also 4.5.4.2.

### Mirror Surface Emissivity

| The emissivity of the mirror surface of the BSM shall be <1% in the wavelength range 200 - 670(m. | AD13, AD18 |
| --- | --- |

### Position of Rotation Axes

| TBD | TBD |
| --- | --- |

### Orthogonality of Rotation Axes

| TBD | TBD |
| --- | --- |

### Fail Safe (No Drive Signal) Position

| When no drive signals are applied to the BSM, the mirror shall take up a position such that the mirror surface is perpendicular to the nominal (0,0) mirror position to within (0.18°. For information, this is intended to place the BSM within 50% of the beam FWHM. | ??? |
| --- | --- |

### Fail Safe (Mechanical Failure) Position

| In the event of mechanical failure on launch, the mirror shall take up a position such that the beam remains useable by the spectrometer instrument, i.e. That the mirror surface is perpendicular to the nominal (0,0) mirror position to within (1.0°. | ??? |
| --- | --- |

### Mass

| The BSM cryogenic mechanism (BSMm) and the BSM support (BSMs) shall have a combined mass of less than 1100g. | ??? |
| --- | --- |

### Cool-down time

| Upon integration to the SPIRE optical bench, the BSM shall reach a temperature of 4K within 15 hours of the commencement of cooldown. This assumes a maximum average cooldown rate of 20K/hour. | ??? |
| --- | --- |

## Operational Specification

### Operational Safety

| The BSM shall operate safely in any normal operating mode. | ??? |
| --- | --- |

### Lifetime

| The BSM shall be capable of operation in any normal mode for periods in excess of 1/6 of the nominal mission lifetime. | AD1 |
| --- | --- |
| Lifetime tests of the chop/jiggle mode will use a multiplier of 1.25x | RD16 |
| Lifetime tests of items operating <100000 times will use factors as specified in RD 16 | RD16 |

| Life time breakdown |
| --- |
| # | Parameter | IRD | Value | Note |
| OL1 | Ground Storage lifetime |  | 2 years | A guess |
| OL2 | Ground Integrated lifetime |  | 4 years | A guess |
| OL3 | Ground operational lifetime |  | 1.5 years TBC | 6 months for subsystem acceptance 6 months for SPIRE acceptance 6 months for FIRST acceptance |
| OL4 | On orbit operational (chop & jiggle mode) lifetime |  | Minimum 8.5 months cumulative time | Operating during 1/6 of the mission duration (4.25 years) |
| OL5 | On orbit de-powered operational lifetime |  | Up to 3.5 years cumulative time | TBD |
| OL6 | On Orbit powered but not chopping (servo control only) |  | Up to 3.5 years cumulative time | TBD May be required if self-damping inadequate in face of micro-vibration environment |

### Operating modes

AD14 defines the required modes of operation in detail.  There are 14 (TBC)  operational modes, but from the point of view of the BSM motion they fall into three categories :

| Jiggle | AD14 |
| --- | --- |
| Chopping | AD14 |
| Scan mapping | AD14 |

Additionally holding the 'stare' position may require active servo control.

### Jiggle Mode

Jiggle Mode is used to optimize the sampling of the detector arrays. In this mode, the mirror is stepped in small angular increments.

Jiggling may be done in any angular orientation (i.e. jiggling is not confined to the jiggle axis, and will usually be a combination of movements in the chop and jiggle axes).

Jiggle movements may be demanded in both chop and jiggle axes simultaneously.

The fine pointing mode is executed as a jiggle pattern, and the difference lies in the processing/use of the data.

Jiggle mode will be used with both the photometer and the FTS. When the FTS is in use no chop will be required. During a photometer jiggle operation the BSM is also required to chop.  i.e. at each jiggle position a number of chop cycles will be executed.

### Chopping Mode

Chopping Mode is used to provide removal of 1/f noise in the photometer detectors. The mirror is used to move the source of interest between two separate detectors.

Chopping is only required in one axis (the chop axis), and will only be used when the photometer is in use, not with the FTS

### Scan mapping

Scan mapping with chopping may be required.  The jiggle and chop axes are used to execute chopping either parallel to or perpendicular to the scan direction, with the frequency of chop set by the jiggle axis frequency limit.

### Stare or 'holding' mode

Dependent on spacecraft  vibration levels, the BSM may require active servo control in order to maintain the nominal 'power-off' hold position within adequate accuracy. This is TBD.

### Combinations of Modes

The following combinations of the basic modes are required.

- Chopping and Jiggling

- Removal of optical or mechanical cross coupling during Chop-only or Jiggle-only operations.

## Interface requirements

### Data Outputs

The following data is provided by the BSM.

| Data | Rate | Reference |
| --- | --- | --- |
| Chop axis position | TBD |  |
| Jiggle axis position | TBD |  |
| Thermometer data | 1Hz (TBC) |  |
| Motor voltage (or current, TBD) | TBD |  |

### Data Inputs

The following inputs are required to operate the BSM.

| Data | Rate | Reference |
| --- | --- | --- |
| Chop axis demand position | On Demand, max. 5Hz | AD19 |
| Jiggle axis demand position | On Demand, max. 1Hz | AD19 |
| Control system parameters (PID terms) | TBD, but slow | AD19 |
| Trajectory parameters (3 for each axis) | TBD, but slow | AD19 |
| Unlatch signal, if required | One-off | AD19 |

### Self induced vibration

The BSM will produce vibration in the region 0-5 Hz, with TBD harmonics and TBD amplitude.

### Stray Magnetic fields

To prevent spurious signal generation, magnetic field strength at the BSM position sensors should not exceed TBD

### Electro-Magnetic Compatibility

TBD - Grounding, Radiated emissions, etc…

### ICD's

The BSM interfaces with the following other subsystems in the SPIRE instrument. The interface to each sub-system is specified in the relevant Interface Control Document.

| Sub-System | ICD Reference |
| --- | --- |
| Structure – BSM | AD2 |
| Optics – BSM | AD4 |
| Baffles – BSM | AD5 |
| Photometer Bolometer Arrays - BSM | AD6 |
| Spectrometer Bolometer Arrays - BSM | AD7 |
| Photometer Calibration Source - BSM | AD8 |
| On Board Software - BSM | AD10 |
| FPU Simulator - BSM | AD11 |
| Instrument Simulator - BSM | AD12 |
| MCU-BSM | AD19 |

## Design, manufacture and test requirements

This section details additional requirements which are placed on the sub-system which are not necessary to meet the functional or operational requirements, but are necessary to enable the sub-system to be designed, manufactured or tested.

### Design requirements

UK Astronomy Technology Centre design procedures shall be adopted as applicable.

### Electronics Card Format

All electronics associated with the BSM shall be mounted on double eurocards.

### Product Assurance Plan

The BSM deliverables will comply with the Spire Product Assurance Plan, AD20.

### Manufacturing and testing requirements

### Mirror Flatness

To facilitate optical  laboratory testing, the mirror surface shall have a flatness of <100nm (P-V).

### Mirror Reflectivity

To facilitate optical testing laboratory testing, the mirror surface shall have a reflectivity of >80% at 633nm.

### Parts, Material and Processes

### General

The workmanship and materials used shall be, or shall be shown to be compatible in any future build, of a standard consistent with flight hardware. The number of materials, mechanical parts types, and processes shall be minimized. Materials and mechanical parts that have been successfully used in similar space applications shall be preferred. Standard processes or known processes previously used in space applications shall

be preferred. Material justification shall prove the hardware structural integrity during the design life.

### Materials, Processes and Mechanical parts (MPM) approval and listing

All MPM meeting the requirements of this document shall be approved by submitting to TBD the Declared Materials List. If these requirements are not met, a RFA shall be submitted to TBD. A Declared Materials List (DML) shall be drawn up, according to the requirements of ECSS-Q- 70 § 3.1.3. The format of PSS-01-700, issue 2, Annex B, is a suitable example. Materials with limited-life characteristics and non-metallic materials shall be highlighted. Mechanical parts shall be included in the DML as a separate group. Process details shall also be included in the DML for the appropriate material. Critical processes shall be highlighted.

### Cleanliness

The BSM shall be class 100 compatible

### Finish

Surface finish shall prevent deterioration from exposure to the environment. Aluminium surfaces shall be treated for corrosion protection coating. Mechanical parts made of stainless steel shall be passivated.

### Outgassing

PSS-01-702 shall be used as a guideline.

### Material selection

In addition to ESA guidelines noted elsewhere in the applicable documents, the following guidelines will be observed.

Structures will be a 6082 grade aluminium alloy to UK standards, and will be selected in consultation with the optical bench design body (MSSL). The BSM mirror may be specified in grade 6061-T6/T651 to ensure adequate thermal stability based on ATC practice.

Fasteners will be cryogenic grade stainless steel. Where practicable they will be of a self locking type. Fasteners may include UNS thread types where required on the grounds of availability in small sizes, or as required by ICD, otherwise will be metric thread. In cases where a self locking fastener is not available, Stycast or Scotchweld 1838 will be used in visible applications under fastener heads.

### Processes

The use of existing ESA documents shall be maximized. In particular the following documents shall be consulted:

- PSS-01-708 for soldering

- PSS-01-726 for crimping

Critical processes shall be identified and reported to TBD through a critical process list or the DML.

## Logistic requirements

### Storage

The BSM shall not suffer any performance degradation following storage in a dry nitrogen atmosphere for a period of up to 5 years.

Storage temperatures shall be in the range +5°C to +30°C.

### Handling

TBD.

### Transportation

TBD

### ESD precautions

TBD

## Environmental requirements

These requirements describe the environment the subsystem will encounter during its life.

### Operating environment

This section defines the environment for the BSM during operation. The BSM must meet all of the functional and operational requirements specified in section 4 when operating under this environment.

### Shock

TBD. As a working figure 50g static acceleration loads in all three axes are assumed.

### Vibration

| TBD |  |
| --- | --- |

### Vacuum Level

| Less than 10-4 Pa TBC |  |
| --- | --- |

### Vacuum Outgassing

| Less than TBC |  |
| --- | --- |

### Temperature

| The BSM shall perform within specification at the operational temperature cited in 4.2.2. In addition, the BSM will operate with reduced performance at temperatures from 4-300K. |  |
| --- | --- |

### Magnetic Fields

| To prevent spurious position signal input, stray magnetic fields must not exceed TBD at the chop and jiggle position sensors. |  |
| --- | --- |

### Survival Temperature

| The BSM shall remain fully operational following exposure to temperatures of 80°C for periods of up to 72 hours. |  |
| --- | --- |

### Radiation environment

| The integrated dose for silicon behind 2 mm of aluminium is estimated at 12 kRad and behind 5 mm of aluminium as 3.5 kRad. These figures will be taken as the radiation tolerance for components in the warm electronics boxes and inside the cryostat respectively (TBC). | RD2 |
| --- | --- |

## Verification requirements

The BSM Development Plan (AD15) contains more detail on the actual tests to be carried out on the BSM. The table below shows how each of the functional and operational requirements in this document is to be demonstrated, for each of the BSM models.

Test		T

Measurement	M

Analysis		A

Inspection	I

### Performance Requirements

| Reference | Requirement | CQM | PFM | FS |
| --- | --- | --- | --- | --- |
| 4.1.1 | Angular Travel - Chop Axis | T | T | T |
| 4.1.2 | Angular Travel - Jiggle Axis | T | T | T |
| 4.1.3 | Minimum Step Size | T | T | T |
| 4.1.4 | Chop Frequency | T | T | T |
| 4.1.5 | Jiggle Frequency | T | T | T |
| 4.1.6 | Holding position | T | T | T |
| 4.1.7 | Stability | T | T | T |
| 4.1.8 | Position Measurement | T | T | T |
| 4.1.9 | Settling Time | T | T | T |
| 4.1.10 | Chop repeatability | T | T | T |

### System Requirements

| Reference | Requirement | CQM | PFM | FS |
| --- | --- | --- | --- | --- |
| 4.2.1 | Mechanical Dimensions | M | M | M |
| 4.2.2 | Operating Temperature | T | T | T |
| 4.2.3 | Thermal Isolation | T | T | T |
| 4.2.4 | Cold Power Dissipation | T | T | T |
| 4.2.5 | Warm Electronics Power Dissipation | T | T | T |
| 4.2.6 | Mirror Surface Dimensions | M | M | M |
| 4.2.7 | Mirror Surface Finish | I | I | I |
| 4.2.104.2.8 | Mirror Surface Reflectivity | A | A | A |
| 4.2.9 | Mirror Surface Emissivity | A | A | A |
| 4.2.10 | Position of Rotation Axes | I | I | I |
| 4.2.11 | Orthogonality of Rotation Axes | T | T | T |
| 4.2.12 | Fail Safe (No Drive Signal) Position | T | T | T |
| 4.2.13 | Fail Safe (Mechanical Failure) Position | TBD |
| 4.2.14 | Mass | M | M | M |
| 4.2.15 | Cool-down time | T | T | T |

This page left deliberately blank