|  | FIRST SPIRE | Ref : DRAFT 1.1 Author : Dave L. Smith | Page : 1 Date : 12 June 2000 |
| --- | --- | --- | --- |
| SPIRE Test Facility Requirements Specification |  |

Distribution List :

| SPIRE-Project | Ken J. King | X |
| --- | --- | --- |
|  | Bruce M. Swinyard | X |
|  | Dave L. Smith | X |
| ATC | Colin Cunningham | X |
| QMW | Matt Griffin | X |

CCLRC Rutherford Appleton Laboratory, Chilton, Didcot, Oxfordshire OX11 0QX, United Kingdom

Tel: (44) +1235     -  Fax (44) +1235 445848

# Update

| Date | Index | Remarks |
| --- | --- | --- |

| Host system | Windows NT 4.0 SR4 |
| --- | --- |
| Word Processor | Microsoft Word 97 SR2 |
| File |  |

Table of contents

1	Scope of the document	4

2	Documents	5

2.1	Applicable documents	5

2.2	Reference documents	5

2.3	Glossary	5

3	Description of the TEST Facility	6

4	Requirements	8

4.1	Cryostat	8

4.1.1	General	8

4.1.2	Thermal	8

4.1.3	Vacuum	9

4.1.4	Optical	9

4.1.5	Electrical Interfaces	9

4.2	Clean Room	10

4.2.1	Cleanliness	10

4.2.2	Electrical	10

4.2.3	Access	10

4.2.4	Safety	10

4.2.5	Other	11

4.3	Control Room	11

4.4	Telescope Simulator	12

4.5	Sources	13

4.6	Test Facility Control System (TFCS)	13

# Scope of the document

This document describes the requirements specification of the FIRST/SPIRE Test Facility to be used for the thermal balance, cold functional tests and calibration of the CQM and PFM instruments (AD 1).

# Documents

## Applicable documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD 1 | SPIRE outline AIV flow | B.Swinyard | SPIRE/RAL/N/0020 |  |
| AD 2 | Outline Specification of the SPIRE instrument calibration facility | B.Swinyard | SPIRE/RAL/N/0058 |  |
| AD 3 | SPIRE Development Plan |  |  |  |

## Reference documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |

## Glossary

| AD | Applicable Document | TBC | To Be Confirmed |
| --- | --- | --- | --- |
| CDR | Critical Design Review | TBD | To Be Defined |
| CQM | Cryogenic Qualification Model | TFCS | Test Facility Control System |
| DDR | Detailed Design Review | WE | Warm Electronics |
| EGSE | Electrical Ground Support Equipment |  |  |
| FIRST | Far InfraRed Space Telescope |  |  |
| FIR | Far InfraRed |  |  |
| FPU | Focal Plane Unit |  |  |
| FTS | Fourier Transform Spectrometer |  |  |
| MGSE | Mechanical Ground Support Equipment |  |  |
| NA | Not Applicable |  |  |
| PFM | ProtoFlight Model |  |  |
| QMW | Queen Mary and Westfield College |  |  |
| RAL | Rutherford Appleton Laboratory |  |  |
| RD | Reference Document |  |  |
| SPIRE | Spectral and Photometric Imaging REceiver |  |  |

# Description of the Test Facility

# The AIV facility will comprise of two working areas.  A clean room to house the cryostat and optical bench, and a control room to house the instrument EGSE and control equipment for the cryostat and calibration equipment, Figure 1.  The working area around the cryostat will be class 100, and other areas in the clean room will be class 10, 000 or better.

# The SPIRE instrument will be mounted in a cryostat to simulate the thermal conditions provided by the FIRST cryostat, namely 7-11K, 4K and 2K.  External calibration sources will be viewed via a telescope simulator situated outside the cryostat at room temperature.  A 20K-blackbody source mounted in the cryostat will provide an absolute calibration reference.

# The telescope simulator is required to present the instrument with an F/5 beam to correctly represent the input from the FIRST telescope.

The control and monitoring of the calibration sources, telescope simulator and cryostat temperatures will be performed via a single test facility systems computer (TFSC), connected to the main SPIRE EGSE.

Figure 1: Fundamental elements of the SPIRE Test Facility

# Requirements

## Cryostat

### General

| R1.1 | The calibration cryostat must be able to accommodate the SPIRE instrument and associated electronics units (JFET/Filter box). |
| --- | --- |
| R1.2 | The cryostat will also house a cold blackbody source operating between 4K and 20K for use during calibration. |
| R1.3 | The SPIRE instrument must be mounted with the baseplate horizontal (I.E., +Y direction down) so that the instrument mechanisms do not work against gravity. |
| R1.4 | The distance from the Cryostat outer window to the SPIRE image plane must be no greater than 500mm. |
| R1.5 | All materials used for the cryostat and equipment contained in the vacuum environment must be suitable for vacuum and low temperature use. Recommended vacuum materials include: Stainless Steel Aluminium Copper Gold Glass and Quartz Teflon Polymide (Trade named Vespel) Viton Carbon Composites Crystalline Filter Materials (e.g. MgF, LiF etc) Solder Kynar Where materials not on this list are to be used (e.g. black paint, adhesives), their outgassing properties must conform to ESA and NASA outgassing rates as specified in |
| R1.6 | The instrument must be mounted in such a way as to allow ease of access and minimum effort during integration. |
| R1.7 | The cryostat will be painted yellow. |

### Thermal

| R2.1 | The cryostat must be able to simulate the thermal environment provided by the FIRST cryostat. The main interface temperatures are 7-11K, 4K and 2K. |
| --- | --- |
| R2.2 | At operating temperatures, the temperature of the SPIRE FPU should drift by no-more than TBD Kmin-1 |
| R2.3 | The cryostat should enable the SPIRE instrument and electronics to attain operating temperature within 24 hours, at a maximum rate of 20Khr-1. |
| R2.4 | The cryostat should allow a cryogen hold time of at least three days when the instrument is fully cooled. This will eliminate the need for an automatic liquid He delivery system, therefore reducing overall costs. The helium exhaust will not be recovered for future use. |
| R2.5 | The cryostat design should allow the instrument to be heated to TBDK before allowing the He shrouds to warm up to prevent contamination of the instrument optics. |
| R2.6 | Heat leaks must be kept to a minimum by anchoring the thermometer wires to the refrigerant tanks, and/or using low thermal conductivity wire. |
| R2.7 | Temperatures of the thermal environment must be measured during the tests. The location and accuracy of the thermometers will be defined in the thermal test plan. |

### Vacuum

| R3.1 | The vacuum system must be able to pump the chamber to 10-6mbar (TBD) within 24 hours before filling the cryogen tanks to ensure that the inside of the instrument is at least 10-4mbar (TBC). |
| --- | --- |
| R3.2 | The maximum pumping rate shall be TBD Lmin-1 to protect the instrument filters. These rates also apply for backfilling. |
| R3.3 | The pumping system must be mechanically isolated from the chamber to prevent vibrations being transmitted to the instrument optics. |
| R3.4 | The pumping system must be oil free if possible to prevent contamination of the optical surfaces. |
| R3.5 | A gate valve will be installed between the pumping system and cryostat. The valve must be closed before filling the cryogen tanks in order to prevent backflow of any lubricants into the main chamber. The gate valve must automatically close in the event of a failure of the pumping system. |
| R3.6 | A burst valve must be installed to prevent the chamber exceeding atmospheric pressure either during controlled let-up-to-air, or a sudden loss of vacuum. |
| R3.7 | The cleanliness standard and systems for the tank must conform to TBD. Mountings for TBD witness mirrors shall be provided in the calibration rig. |
| R3.8 | The vacuum pressure must be measured during pumpdown. The pressure range is expected to be 1bar to 10-8mbar. After high vacuum has been reached the pressure gauges will be switched off. |
| R3.9 | A Residual Gas Analyser (range 1-200 AMU) is required to monitor partial pressures of contaminants in the chamber during commissioning. This is not required for the actual calibration tests. |

### Optical

| R4.1 | The support structure must ensure that the optical alignment is maintained to TBD’’ when the instrument is cold. |
| --- | --- |
| R4.2 | Optical filters will be mounted at each of the cryostat interfaces between the instrument aperture and telescope simulator. The specifications of these are TBD at the time of writing. |
| R4.3 | A flip mirror will be mounted to allow the instrument to view the 20K-blackbody source. |

### Electrical Interfaces

| R5.1 | Electrical feedthroughs between the air and vacuum sides of the chamber will be provided for the instrument electronics, thermometry and cold calibration target(s). |
| --- | --- |
| R5.2 | It must be demonstrated that there is no electrical interference to the supplies of the SPIRE instrument and EGSE. |

## Clean Room

The clean room will house the calibration cryostat, optical bench supporting the sources and telescope assembly and the SPIRE instrument mounted on MGSE.  This area will only be used for instrument calibration and not for general AIT activities.  Hence only limited workspace will be available to prepare test equipment and the instrument for the calibration.

### Cleanliness

| R5.1 | The area immediately surrounding the cryostat, where instrument optics will be exposed, will be class 100. |
| --- | --- |
| R5.2 | All other areas in the clean room will be class 10,000 or better. |
| R5.3 | Airflow in the clean room must be directed away from the cryostat, where possible, such the entrance to the clean room and changing area are downstream. |
| R5.4 | The entrance to the clean room must be wide enough to allow the instrument and MGSE to pass through. |
| R5.5 | Cleanliness will be monitored throughout the SPIRE calibration period. |

### Electrical

| R6.1 | Two clean 28V power supplies must be provided for the SPIRE instrument. |
| --- | --- |
| R6.2 | There shall be adequate power points for the SPIRE EGSE, test equipment and vacuum system. |
| R6.3 | The electrical design of the facility must ensure that there are no ground loops that could pose an EMI problem. |
| R6.4 | All electrical connections in the facility must be fully documented and readily accessible. |

### Access

| R7.1 | The entrance to the clean room must be wide enough to allow the instrument and MGSE to pass through. |
| --- | --- |
| R7.2 | Due to limited space and to maintain cleanliness, access to the clean room will be restricted to a few authorised personnel. To prevent unauthorised access, a card key system will be installed. |

### Safety

| R8.1 | Oxygen monitors will be installed in the clean room that should sound a clear alarm in the event of O2 going below safe levels. |
| --- | --- |
| R8.2 | Escape from the room in the event of an emergency must be unrestricted. |
| R8.3 | A full risk assessment must be carried out and emergency procedures must be ready before the facility is used for cryogenic work. |
| R8.4 | Laser warning signs and eye goggles will be provided in accordance with RAL site regulations. |
| R8.5 | A full risk assessment must be carried out and emergency procedures must be ready before using laser equipment. |

### Other

| R9.1 | The clean room should be light tight to allow use as a dark room. |
| --- | --- |

## Control Room

| R10.1 | There must be adequate space for the SPIRE EGSE, TFCS, vacuum equipment and at least 4-5 people. |
| --- | --- |
| R10.2 | Access to the control room will be restricted by a card key system provided by the RAL AIV facility. |
| R10.3 | The entrance to the clean room will have changing facilities. |
| R10.4 | There will be sufficient power supplies for the EGSE and support equipment. |
| R10.5 | The control room will have a telephone. |
| R10.6 | Connection points will be provided to allow connection to the internet, and fax machines. |

## Telescope Simulator

Figure 2: Optical design of the Telescope Simulator

| R11.1 | The Telescope Simulator will present the SPIRE FPU with an F/5 beam over the full FOV of the instrument. |
| --- | --- |
| R11.2 | The SPIRE beam envelope must not be vingetted by the cryostat walls or simulator optics. |
| R11.3 | The Telescope Simulator will occupy an area no larger than 550mm x 1200mm |
| R11.4 | The distance from the 3rd fold mirror to the SPIRE image plane will be nominally 800mm |
| R11.5 | The position of the moving mirrors will be remotely controlled and monitored by the TFCS. |

## Sources

The following sources will be required for the SPIRE calibration tests.

| R12.1 | Hg Lamp |
| --- | --- |
| R12.2 | FIR Laser |
| R12.3 | FIR FTS |
| R12.4 | 20K-Blackbody |
| R12.5 | Reference Bolometer |

## Test Facility Control System (TFCS)

Fully automatic control of the calibration system is not necessary since the duration of the calibration tests is only a few weeks.  Development of a fully automated system would be disproportionately expensive on cost and effort, and high risk.  However remote control should be provided, where appropriate, for the vacuum system, temperature monitoring, liquid nitrogen and helium levels, and the calibration equipment.

| R13.1 | The TFCS will be used to provide central control for the test equipment, log temperatures and other test data (e.g. source levels). |
| --- | --- |
| R13.2 | There will be a TCP/IP network to allow transfer of data between the TFCS and the SPIRE EGSE. |
| R13.3 | The TFCS will be able to receive commands from the SPIRE EGSE |
| R13.4 | The TFCS will transmit telemetry packets to the SPIRE EGSE |
| R13.5 | The TFCS clock will be synchronised to the SPIRE EGSE to allow accurate matching of data. |
| R13.6 | The TFCS will be able to run in a stand-alone mode. |