SPIRE

| SUBJECT: | SPIRE Design Description |
| --- | --- |
| PREPARED BY: | Douglas Griffin |  |  |
| DOCUMENT No: | SPIRE-RAL-PRJ-000620 |  |  |
| ISSUE: | 0.1 | Date: | Wednesday, 04 April 2001 |
| APPROVED BY: |  | Date: |  |

Distribution:

All SPIRE Co-Investigators and Institute Managers

Change Record

ISSUE				DATE

1.	INTRODUCTION	6

1.1	The SPIRE instrument and its scientific programme	6

1.2	Purpose of this document	6

1.3	Scope	7

2.	INSTRUMENT OVERVIEW	8

2.1	Photometer design drivers	8

2.2	Spectrometer design drivers	8

2.3	Instrument functional block diagram	9

2.3.1	Imaging photometer	11

2.3.1.1	Optical Design and FPU layout	11

2.3.1.2	Detector arrays	12

2.3.1.3	Photometer observing modes	13

2.3.2	Fourier transform spectrometer	14

2.3.2.1	Optical design and FPU layout	14

2.3.2.2	Spectrometer detector arrays	15

2.3.2.3	Spectrometer observing modes	16

2.3.3	Helium-3 Cooler	17

3.	Instrument System Design	20

3.1	SPIRE Subsystems	20

3.2	The SPIRE Instrument as a System	26

3.3	Structural Design and FPU Integration	31

3.4	Optical Design	33

3.4.1	Common Optics and Photometer Optics	33

3.4.2	Spectrometer Optical Design	36

3.5	Straylight control	40

3.5.1	Bandpass Filtering	40

3.5.2	Baffling	41

3.6	Thermal design	45

3.6.1.1	Instrument temperature levels	45

3.6.1.2	Cryogenic heat loads	45

3.6.1.3	Temperature stability	45

3.6.2	Optical Design Drivers	46

3.6.2.1	Geometry	46

3.6.2.2	Optical throughput	46

3.6.3	Signal quality	46

3.6.3.1	Microphonics	46

3.6.3.2	EMC	46

3.7	Redundancy Scheme DKG	47

3.8	System Budgets	48

4.	Subsystem Design	49

4.1	Warm Electronics	49

4.1.1	Digital Processing Unit (HSDPU)	49

4.1.2	Detector Readout and Control Unit (HSDRCU)	53

4.1.2.1	Detector control unit (DCU)	55

4.1.3	Detector Control Unit (HSDCU)	57

4.2	RF Filters	58

4.3	JFET units	59

4.4	Bolometric Detector Arrays	61

4.4.1	Principle of  semiconductor bolometers	61

4.4.2	SPIRE bolometer performance requirements	62

4.4.3	SPIRE bolometer design and specifications	62

4.4.4	Bolometer readout electronics	63

4.4.5	Feedhorns and bolometer cavities	64

4.4.6	Bolometer array thermal-mechanical design	65

4.5	Mirrors	66

4.6	Filters, beam splitters	67

4.6.1	Photometer filtering scheme	68

4.6.2	Spectrometer filtering scheme	70

4.7	Internal calibrators	72

4.7.1	Photometer calibrator (PCAL)	72

4.7.2	Spectrometer calibrator (SCAL)	73

4.8	Beam Steering Mechanism (BSM)	74

4.9	Spectrometer Mechanism	79

4.9.1	Position encoder	81

4.9.2	Control  System and Readout	81

4.10	FPU structure	82

4.11	Helium-3 Cooler	88

4.12	Thermal staps	94

4.12.1	2-K to 4-K	94

4.12.2	300 mK	95

4.13	Harness	95

4.14	Shutter	98

5.	Instrument Operating Modes	101

5.1	Spacecraft pointing	101

5.2	Spacecraft movements during SPIRE observing	101

5.2.1	Nod	102

5.2.2	Raster	102

5.2.3	Line Scan	103

5.3	Photometer Observatory Functions	104

5.3.1	Point Source Photometry (POF1 and POF2)	104

5.3.2	Jiggle Mapping (POF3 and POF4)	105

5.3.3	Scan mapping (POF 4 and POF5)	106

5.3.4	Photometer peak-up (POF7)	106

5.3.5	Operation of photometer internal calibrator (POF8)	106

5.4	Spectrometer Observatory Functions	107

5.4.1	Point source spectrum (SOF1)	107

5.4.2	Fully-sampled spectral map (SOF2)	107

6.	Spire Sensitivity Estimation	108

7.	Criticality Analysis	110

8.	References	111

9.	List of Annexes	113

Glossary

| Term | Meaning |
| --- | --- |
| ADC | Analogue to Digital Converter |
| AIV | Assembly, Integration and Verification |
| AME | Absolute Measurement Error |
| AME | Absolute Measurement Error |
| AOCS | Attitude and Orbit Control System |
| APE | Absolute Position Error |
| APE | Absolute Position Error |
| AVM | Avionics Model |
| BDA | Bolometer Detector Array |
| BSM | Beam Steering Mirror |
| CDMS | Command and Data Management System |
| CDMU | Command and Data Management Unit |
| CDR | Critical Design Review |
| CVV | Cryostat Vacuum Vessel |
| DAQ | Data Acquisition |
| DCU | Detector Control Unit = HSDCU |
| DPU | Digital Processing Unit = HSDPU |
| DQE | Detective Quantum Efficiency |
| EGSE | Electronic Ground Support Equipment |
| EMC | Electro-magnetic Compatibility |
| EMI | Electro-magnetic Interference |
| ESA | European Space Agency |
| FCU | FCU Control Unit = HSFCU |
| FIR | Far Infrared |
| FIRST | Far Infra-Red and Submillimetre Telescope |
| FoV | Field of View |
| F-P | Fabry-Perot |
| FPU | Focal Plane Unit |
| FTS | Fourier Transform Spectrometer |
| FWHM | Full Width Half maximum |
| HOB | Herschel Optical Bench |
| HPDU | Herschel Power Distribution Unit |
| HSDCU | Herschel-SPIRE Detector Control Unit |
| HSDPU | Herschel-SPIRE Digital Processing Unit |
| HSFCU | Herschel-SPIRE FPU Control Unit |
| HSO | Herschel Space Observatory |
| IF | Interface |
| IID-A | Instrument Interface Document - Part A |
| IID-B | Instrument Interface Document - Part B |
| IMF | Initial Mass Function |
| IRD | Instrument Requirements Document |
| IRTS | Infrared Telescope in Space |
| ISM | Interstellar Medium |
| JFET | Junction Field Effect Transistor |
| LIA | Lock-In Amplifier |
| MCU | Mechanism Control Unit = HSMCU |
| M-P | Martin-Puplett |
| NEP | Noise Equivalent Power |
| NTD | Neutron Transmutation Doped |
| OBS | On-Board Software |
| OMD | Observing Modes Document |
| PACS | Photodetector Array Camera and Spectrometer |
| PCAL | Photometer Calibration source |
| POF | Photometer Observatory Function |
| PUS | Packet Utilisation Standard |
| SCAL | Spectrometer Calibration Source |
| SCUBA | Submillimetre Common User Bolometer Array |
| SED | Spectral Energy Distribution |
| SOF | Spectrometer Observatory Function |
| SPIRE | Spectral and Photometric Imaging Receiver |
| SSSD | SubSystem Specification Document |
| STP | Standard Temperature and Pressure |
| SVM | Service Module |
| TBC | To Be Confirmed |
| TBD | To Be Determined |
| URD | User Requirements Document |
| UV | Ultra Violet |
| WE | Warm Electronics |

# INTRODUCTION

## The SPIRE instrument and its scientific programme

SPIRE (the Spectral and Photometric Imaging REceiver) is one of three cryogenic focal plane instruments for ESA’s Herschel mission. Its main scientific goals are the investigation of the statistics and physics of galaxy and structure formation at high redshift and the study of the earliest stages of star formation, when the protostar is still coupled to the interstellar medium. These studies require the capabilities to carry out large-area (many tens of square degrees) deep photometric imaging surveys at far-infrared and submillimetre wavelengths, and to follow up these systematic survey observations with spectroscopy of selected sources. SPIRE will exploit the advantages of Herschel: its large-aperture, cold, low-emissivity telescope; the complete lack of atmospheric emission and attenuation giving access to the poorly explored 200-700-(m range, and the large amount of high quality observing time. Because of these advantages, SPIRE will have unmatched sensitivity for deep imaging photometry and moderate-resolution spectroscopy.

Galaxies emit a large amount (from 30% to nearly 100%) of their energy in the far infrared due to re-processing of stellar UV radiation by interstellar dust grains. The far infrared peak is redshifted into the SPIRE wavelength range for galaxies with redshift, z, greater than ~ 1. The total luminosity of a galaxy cannot be determined without an accurate measurement of its Spectral Energy Distribution (SED). The study of the early stages of galaxy evolution thus requires an instrument that can detect emission from high-z galaxies in the submillimetre, enabling their SEDs and luminosities to be derived.

Stars form through the fragmentation and collapse of dense cloud cores in the interstellar medium (ISM), and the very first stages of this process are not well known. A good understanding of this early evolution is crucial, as it governs the origin of the stellar initial mass function (IMF). Sensitive far infrared and submillimetre observations with high spatial resolution are necessary to make complete surveys of protostellar clumps to determine their bolometric luminosities and mass functions. SPIRE will also, for the first time, enable astronomers to observe at high spatial resolution the physical and chemical conditions prevailing in the cold phases of the interstellar medium and to study the behaviour of the interstellar gas and dust before and during star formation. SPIRE’s uniquely high sensitivity to very cold dust emission also makes it the ideal instrument to study the material that is ejected in copious quantities from evolved stars, enriching the interstellar medium with heavy elements. Large amounts of matter - as yet undetected - are ejected from stars before the white dwarf stage. Theories of stellar evolution, and of the enrichment of galaxies in heavy elements and dust, will be incomplete until these earlier mass loss phases are characterised and understood. Studies of star formation and of the interaction of forming and evolved stars with the ISM are also, of course, intimately related to the investigation of galaxy formation and evolution, which occur through just these processes.

These high priority programmes for Herschel require sensitive continuum imaging in several bands to carry out surveys, and a low-resolution spectroscopic mode to obtain detailed SEDs of selected objects and measure key spectral lines. Although SPIRE has been optimised for these two main scientific programmes, it will offer the astronomical community a powerful tool for many other astrophysical studies: giant planets, comets, the galactic interstellar medium, nearby galaxies, ultraluminous infrared galaxies, and active galactic nuclei. Its capabilities will remain unchallenged by the ground-based and the airborne observatories which are planned to come into operation over the next decade.

The scientific case for SPIRE is described in more detail in the SPIRE proposal (Griffin et al., 1998a)

## Purpose of this document

This document outlines the essential features of the design and operation of SPIRE. It is intended to provide an introductory account of the key features of the SPIRE system and subsystem design and operation - but not to provide a complete description: for that the reader is referred to other more detailed project documents, particularly the following:

| Document | Reference | Abbreviation |
| --- | --- | --- |
| SPIRE Scientific Requirements | SPIRE-UCF-PRJ-000064 | SRD |
| SPIRE Instrument Interface Document Part B | SPIRE-ESA-DOC-000275 | IID-B |
| SPIRE Instrument Requirements Specification | SPIRE-RAL-PRJ-000034/1.0 23 Nov. 2000 | IRD |
| Operating Modes for SPIRE | SPIRE-RAL-PRJ-000320 | OMD |
| SPIRE On-Board Software URD. | SPIRE-IFSI-PRJ-000444/ | OBS URD |
| Subsystem Specification Documents for each of the SPIRE subsystems | Detector Subsystem Specifications | SPIRE-JPL-PRJ-456 | SSSDs |
|  | FTS Subsystem Specifications |  |  |
|  | BSM Subsystem Specifications |  |  |
|  | Cooler Subsystem Specifications |  |  |
|  | DPU Subsystem Specifications |  |  |
| MCU Design Description ? | LAM/ELE/SPI/000619/1.1/ 20 Dec. 2000 |  |

Additional references are given in Section 8.

## Scope

The document concerns the design and operation of the instrument. It includes:

- the Warm Electronics (WE) which is composed of the Digital Processor Unit (DPU), the Detector Control Unit (DCU) and the FPU Control Unit (FCU) and their associated electrical harnesses;

- the Focal Plane Unit (FPU) which contains the optical components, detector arrays, 300-mK sorption cooler, mechanisms, internal calibrators, thermal straps, and various housekeeping units, and the mechanical structure which houses these elements;

- the on-board software (OBS) which allows commanding of the instrument and telemetry of the science and housekeeping data.

The design of the Herschel telescope, which SPIRE shares with the other experiments is only referenced in as much as it influences the design of SPIRE as external inputs.  The thermal straps from the Herschel cryostat to the FPU, the interface between the DPU and the spacecraft Power Distribution Unit and the Command and Data Handling System (CDMS) are only referenced only in so far as they impact on the design of the units listed above.

The various items of SPIRE ground support equipment also fall outside the scope of the document

# INSTRUMENT OVERVIEW

SPIRE contains a three-band  imaging photometer and an imaging Fourier Transform Spectrometer (FTS), both of which use 0.3-K feedhorn-coupled “spider-web” NTD germanium bolometers cooled by a 3He refrigerator.  The photometer and spectrometer are not designed to operate simultaneously. The field of view of the photometer is 4 x 8 arcminutes, the largest that  can be achieved given the location of the SPIRE field of view in the Herschel focal plane and the size of  the telescope unvignetted field of view. Three bolometer arrays provide broad-band photometry ((/(( ~ 3) in wavelength bands centred on 250, 350 and 500 (m. The field of view is observed simultaneously in all three bands through the use of fixed dichroic beam-splitters. Spatial modulation can be provided either by a Beam Steering Mirror (BSM) in the instrument or by scanning the telescope across the sky, depending on the type of observation. An internal thermal calibration source is available to provide a repeatable calibration signal for the detectors. The FTS uses novel broadband intensity beam dividers, and combines high efficiency with spatially separated input ports. One input port covers a 2.6-arcminute diameter field of view on the sky and the other is fed by an on-board calibration source. Two bolometer arrays are located at the output ports, one covering 200-300 (m and the other 300-670 (m. The FTS will be operated in continuous scan mode, with the path difference between the two arms of the interferometer being varied by a constant-speed mirror drive mechanism. The spectral resolution, as determined by the maximum optical path difference, will be adjustable between 0.04 and 2 cm-1 (corresponding to (/((  = 1000 - 20 at 250 (m wavelength.

## Photometer design drivers

The basic design of the photometer was dictated by the following consideration:

- the major scientific priorities for SPIRE require an instrument which is capable of deep mapping of large areas of sky efficiently with full spatial sampling and multi-wavelength coverage;

- observations  of point and compact sources should also be possible with good efficiency;

- the instrument should be as simple as possible for affordability, reliability  and ease of operation;

- the wavelength coverage should complement that of the PACS instrument on board Herschel and ground-based facilities operating at near-millimetre wavelengths;

These considerations led to the choice of a system with fixed dichroic beamsplitters allowing simultaneous observation of a large field of view in the three chosen submillimetre bands, with the inclusion of a beam steering mirror to allowed point source observations and small maps to be carried out efficiently.

## Spectrometer design drivers

Fabry-Perot and grating spectrometer designs were also studied for SPIRE (Griffin, 1997; Griffin et al. 1998b). The FTS was chosen for a number of reasons:

- it allows for two-dimensional imaging spectroscopy, which is not possible with the grating instrument;

- the spectral resolution can easily be adjusted and tailored to the scientific requirements of the observation;

- the detectors can be operated at 300 mK because the photon noise limited NEP is higher for the FTS, whose detectors observe broad-band, than for grating or Fabry-Perot spectrometers in which they observe in narrow-band mode - this allowed a 3He sorption cooler to be adopted for SPIRE - a considerable simplification over the 100-mK dilution cooler that would have been needed for a grating or F-P instrument;

- whilst the grating is more sensitive, at least in principle, for observations of known spectral lines, there is little difference in the sensitivities of the two options for spectral survey observations, which are of greater scientific priority for SPIRE;

- the FTS is less vulnerable to degradation in performance arising from stray light and out-of-band leaks which can be problematic with a low-background grating instrument;

- practical implementation of either a grating or an F-P instrument within the constraints of the Herschel system would have posed serious accommodation problems (grating) or risk of FPU failure due to having multiple mechanisms in series (F-P).

## Instrument functional block diagram

Figure 2-1 shows a functional block diagram of the SPIRE instrument. The subsystems are prefixed HS standing for Herschel-SPIRE. The focal plane unit is mounted on the Herschel Optical Bench as shown in Figure 2-2. It is approximately 690 x 410 x 410mm in size, and has three separate temperature stages at nominal temperatures of 4 K, 2 K (both provided by the Herschel cryostat) and 300 mK (provided by SPIRE's internal cooler). The main 4-K structural element of the FPU is an optical bench panel which is supported from the 10-K cryostat optical bench by stainless steel blade mounts. The photometer and spectrometer are located on either side of this panel. The majority of the optics are at 4 K, but the detector arrays and final optics are contained within 2-K enclosures. The 3He refrigerator cools all of the five detector arrays to 0.3 K. Two JFET preamplifier boxes (one for the photometer and one for the FTS) are attached to the 10-K optical bench close to the 4-K enclosure, with the JFETs heated internally to their optimum operating temperature of ~ 120 K. Note that the shutter is not designed for flight operation and is controlled by SPIRE EGSE.

The SPIRE warm electronics consist of a Detector Readout Unit (DCU), a FPU Control Unit (FCU) and a Digital Processing Unit (DPU). The DCU provides bias and signal conditioning for the arrays and cold readout electronics and reads out the detector signals; the FCU controls the FPU mechanisms, the 3He cooler and the internal calibrators, and housekeeping thermometers. The DPU acts as the interface to the spacecraft, including instrument commanding, and formats science and housekeeping data for telemetry to the ground.

Figure 2-1 - Functional block diagram of the SPIRE instrument

Figure 2-2 Location of the SPIRE FPU on the Herschel Optical Bench

### Imaging photometer

#### Optical Design and FPU layout

The photometer layout is shown in Fig. 2.3. The 4-K optical elements are mounted directly from the optical bench panel. The 2-K enclosure is also supported from the panel by stainless steel blades, and contains the detector arrays, dichroics, and fold mirrors. The three array modules are bolted to the outside wall of the 2-K box. Within each module, the detector arrays, feedhorns and the final filter are thermally isolated from the 2-K structure by Kevlar wires, and are cooled by a thermal strap to the 3He refrigerator (see Section TBD below). The photometer input optics are shared with the spectrometer. The separate spectrometer field of view is directed to the other side of the optical bench panel by a pick-off mirror.

Figure 2-3 Computer-generated picture of the SPIRE photometer layout (4-K cover not shown)

The photometer optical design is shown in Figure 2-4, and is described in more detail by Dohlen et al. (2000) and in Section TBD below. It is an all-reflective system except for two dichroic beam dividers used to direct the three wavelength bands onto different bolometer arrays, and various transmissive band-pass and edge filters used to reject out-of-band radiation. It is optimised to give close to diffraction-limited imaging across the whole 4 x 8 arcminute field of view. The SPIRE field of view is offset by 11 arcminutes from the centre of the Herschel telescope's highly curved focal surface. Mirror M3, which lies below the focus, receives the f/8.68 beam from the telescope and forms a pupil image of the telescope secondary at the flat beam steering mirror, M4. Mirror M5 converts the focal ratio to f/5 and provides an intermediate focus at the next mirror, M6, which re-images the aperture stop at M4 to a cold stop located at the entrance to a 2-K enclosure. M7, M8 and M9 constitute a one-to-one optical relay to bring the M6 focus to the three detector arrays. The beams for the three bands are directed onto the arrays at f/5 by a combination of flat folding mirrors and fixed dichroics set at 25º to the beam axis. M3 - M8 are at 4 K and the cold stop and all subsequent optics are at 2 K.

Figure 2-4 Imaging photometer optical design

A shutter at the entrance aperture of the instrument (just above the telescope focus) can be inserted to block the beam. This will be important for ground testing in the Herschel cryostat where the background radiation from the cryostat shields and lid will be much greater than in flight. The shutter will allow the detector performance to be tested and verified under controlled flight-representative conditions. Flight operation of the shutter is not planned.

An internal calibration source provides a repeatable signal for the bolometer arrays. It radiates through a 1-mm hole in the centre of the beam steering mirror, M4. As this is at a pupil image, the illumination is close to uniform over the arrays. The source can be modulated at frequencies up to 5 Hz, and operated at temperatures up to 80 K to give sufficient signal on the arrays, with peak power dissipation < 2 mW. The beam steering mirror is capable of chopping (2 arcminutes along the long axis of the 4 x 8 arcminute field of view, at frequencies up to 2 Hz with an efficiency of 90% and power dissipation < 2 mW. It can operate at higher frequencies with reduced efficiency and increased power dissipation. The beam steering mechanism can simultaneously chop at up to 1 Hz in the orthogonal direction by up to 30 arcseconds. Two axis motion allows "jiggling" of the pointing to create a fully sampled image of the sky with the feedhorn-coupled detectors whose diffraction-limited beams on the sky are separated by approximately twice the beam FWHM.

The SPIRE filtering scheme is designed to provide precise definition of the spectral passbands with high out-of-band rejection and maximum in-band transmission, and also to minimise the thermal loading on the 4-K, 2-K and 0.3-K stages by reflecting short-wavelength radiation. To achieve complete rejection out to UV wavelengths, four blocking filters are needed in the chain in addition to high-pass and low-pass edge filters which define the band.

#### Detector arrays

SPIRE will use spider-web bolometers with NTD germanium thermometers (Mauskopf et al. 1997; Bock et al. 1998). The bolometers are coupled to the telescope by hexagonally close-packed single-mode conical feedhorns, providing diffraction limited beams. The horn diameters are 2F( where F is the focal ratio of the final optics and ( is the wavelength: this provides a diffraction-limited beam with maximum coupling efficiency of the detector to a point source. Modelling of the complete optical train predicts FWHM beam widths of 17.1, 24.4 and 34.6 arcseconds at 250, 350 and 500 (m, respectively. The numbers of detectors in the three arrays are 139, 88, and 43 for 250, 350 and 500 (m respectively, making a total of 270 detectors for the photometer. The detector arrays are shown schematically in Fig. TBD, and a photograph of a prototype array module is shown in Fig. TBD. Each array unit has an interface to the 2-K box, with a thermal strap from the 3He cooler to the 0.3-K stage, which is supported by Kevlar strings from the 2-K level. The electrical connections to the detectors are made with Kapton ribbon cables within the array modules and with woven manganin cables between the array modules and the JFET units. The bolometers are excited by an AC bias at approx. 100 Hz, which eliminates 1/f noise from the JFETs, giving a 1/f knee for the system of less than 100 mHz. Conservative estimates of the bolometer Detective Quantum Efficiency (DQE) vary between 0.6 and 0.7 ensuring that the overall NEP will be dominated by the thermal emission from the Herschel telescope.

#### Photometer observing modes

The photometer will have three principal observing modes, as illustrated in Figure 2-6 and described below. These modes are described in greater detail in OMD.

Figure 2-5 - :  (a) layout of the three photometer arrays; (b) photograph of array mechanical prototype

Point source photometry: For photometric observations of point or compact sources, chopping will be used. There are several sets of three detectors for which the beams at the three wavelengths are exactly co-aligned on the sky, indicated by the shaded circles in Figure 2-5. By chopping through the appropriate angle (approx. 126 arcseconds), 3-band photometric observations can be carried out simultaneously with maximum efficiency. To account for the possibility of positional errors due to telescope pointing inaccuracy or imperfect knowledge of the source position, the beam steering mirror can be used to implement a seven-point mapping routine in this mode. Assuming an angular offset of 6" for the seven-point, the loss in S/N for a given integration time varies between 6% at 500 (m and 20% at 250 (m, which is a small penalty to pay for assurance that telescope pointing or source position errors do not result in an underestimate of the source flux density.

Field mapping: For mapping of regions a few arcminutes in extent, the beam steering mirror will be used to carry out a jiggle map, similar to the mode of operation of the SCUBA bolometer camera on the JCMT (Holland et al. 1999). A 64-point jiggle pattern is needed to achieve full spatial sampling in all bands simultaneously, with a step size of 9 arcseconds (half-beam spacing at 250 (m). A maximum field size of 4 x 4 arcminutes is available in this mode as the 2-arcminute regions at each end of the array will be chopped outside the field of view admitted by the photometer optics.

Scan mapping: This mode will be used for mapping large areas of sky (much bigger than the SPIRE field of view), including deep survey observations. The telescope will be scanned across the sky (at up to 1 arcminute per second, the maximum rate that the spacecraft can provide). Because of the excellent 1/f stability of the NTD detectors, the beam steering mirror does not need to be operated - signal modulation is provided by the telescope motion. To provide the necessary beam overlap for full spatial sampling over the strip defined by a single scan, the scan angle must be 14.5o with respect to one of the array axes.

The available telemetry rate of 100kbs allows all of the 270 photometer detectors to be sampled with 16-bit resolution at up to 28 Hz and the data telemetered directly to the ground with no on-board processing.

Figure 2-6 Overlaid photometer arrays and summary of photometer observing modes.

### Fourier transform spectrometer

#### Optical design and FPU layout

The layout of the FTS and its optical scheme are shown in Figure 2-7 and Figure 2-8 respectively. The main design features of the FTS are described in Swinyard et al. (2000) and in Section TBD below. It uses two broadband, high-efficiency, intensity beam splitters in a Mach-Zehnder configuration rather than the traditional polarising beam dividers. This configuration has the advantage that all four ports are separately accessible, as in the classical Martin Puplett (M-P) polarising FTS (Martin, 1982). But the throughput is a factor of two higher than for the M-P as none of the incoming radiation is rejected. This design is also insensitive to the polarisation of the incident radiation. The performance of the beam dividers and of a bench-top implementation of this design has been demonstrated (Ade et al. 1999). A thermal calibrator is located at a pupil image in the second input port of the FTS, and provides a thermal input that mimics the dilute 80-K black body emission of the telescope. This allows the large telescope background to be nulled, thereby reducing the dynamic range requirements for the detector sampling. Two band-limited detector arrays are placed in the two output ports, covering 200-300 m and 300-670 m. A single back-to-back moving roof-top mechanism serves both arms of the interferometer, with a frictionless carriage mechanism using double parallelogram linkage and flex pivots. The pick-off mirror (on the photometer side of the optical bench panel and located at the intermediate field image) directs the spectrometer field of view through a hole in the optical bench panel into the FTS side of the instrument. A 4-K pupil stop is located between the pick-off mirror and the input fold mirror. The input relay mirror focuses the beam to an intermediate image plane located just after the first beam divider, after which the beam is collimated and sent to the moving corner cube assembly. The corner cube shifts the beam and sends it towards the camera mirror, which produces an image plane just before the output beam divider. The output relay mirror focuses the beam onto the detector arrays. A pupil image is located near the final fold mirror, making this a convenient location for the entrance aperture to the 2-K enclosure. As this pupil moves when the optical path difference changes, it is not a good place for a limiting cold stop. Instead, the limiting aperture is located at the 4-K pupil plane between the pick-off mirror and the input fold mirror.

Figure 2-7 - Physical layout of the spectrometer.

Figure 2-8 - Optical design of the spectrometer.

The FTS design is optimised for the 200-400 m band. The wavelength coverage is extended to 15 cm-1 (670 m) to give access to the astrophysically important 609-m line of CI in our own and nearby galaxies, and to increase the range over which the spectral energy distribution of sources can be measured in the FTS low-resolution mode. A filtering scheme similar to the one employed for the photometer channel is used to restrict the passband of the instrument. Filters on the bolometer arrays themselves define the passband for each array.

#### Spectrometer detector arrays

The field of view of the FTS is approximately 2.6 arcminutes in diameter, and is covered by 37 hexagonally close-packed detectors in a short-wavelength array and 19 in a long-wavelength array. The detector modules will be similar to those used for the photometer, with a mechanical interface to the wall of the 2-K enclosure.

The two FTS arrays cover the 200-300 and 300-670 (m bands. The detectors and feedhorns for the short wavelength band are similar to those for the photometer 250-m channel. The long wavelength band is optimised for the 300-400 (m range. There is a degradation in point source coupling efficiency at wavelengths beyond around 400 (m due to the decreasing aperture size relative to the wavelength. The waveguide coupling the horn to the bolometer must also have a diameter large enough to transmit at 670-m, and so is overmoded at the shorter wavelengths within the band. This results in an increase in background radiation on the detectors and a broadening of the beam by about 20% compared to the diffraction limit at the lower end of the band (Caldwell et al. 2000).

The layout of the FTS arrays is shown in Figure 2-9. The detectors on the periphery of the arrays are partly vignetted by the 2.6-arcminute field of view admitted by the instrument optics (indicated by the large circles in Figure 2-9.). The short-wavelength array feedhorns are sized to give 2F pixels at 225 m and the long-wavelength horns to give 2F pixels at 389 m. This arrangement, although slightly non-optimal from the point of view of point source sensitivity at the central wavelengths of the two arrays, has the advantage that there are numerous co-aligned pixels in the combined field of view. This maximises the observing efficiency for measuring a point source spectrum together with its surrounding sky background and also provides redundancy to the spectrometer in the case of failure of a single pixel within one array.

Figure 2-9 Layout of spectrometer detector arrays

#### Spectrometer observing modes

The FTS will be operated in continuous scan mode with the mirrors moving at a constant speed of up to 0.1 cm-1, corresponding to a signal frequency range of 6 - 20 Hz. The spectral resolution can be adjusted between 0.04 and 2 cm-1 (/ = 20 - 1000 at 250 m). The maximum scan length is 3.5 cm (taking 35 seconds or more and giving an optical path difference of 14 cm). To ensure that mechanism jitter noise is well below the photon noise level, a relative accuracy of 0.1 m is required for the mirror position. The FTS calibration source will be on continuously while the spectrometer is operating, with a peak power of no more than 5 mW. For spectral mapping of extended sources, the beam steering mirror will be used to provide the necessary pointing changes between scans. The scanning mirror control system uses a digital feedback loop to provide a constant speed over the scan length, with an accuracy requirement of 1% (goal 0.5%). The position readout uses a Heidenhain Moiré fringe sensing system. The detectors are read out asynchronously with the samples time-stamped to match them to the corresponding mirror locations. No on-board processing will be done - the raw interferograms will be telemetered to the ground. The number of detectors and the available telemetry rate are compatible with an oversampling factor of two with respect to the Nyquist sampling rate of 40 Hz (sampling at approx. 80 Hz per detector). An oversampling factor somewhat greater than this is desirable - options to achieve this include increasing the data rate, decreasing the mirror speed, sampling only a fraction of the detectors in some cases (e.g., point source observations), or a combination of these.

### Helium-3 Cooler

The 3He cooler (Duband 1997) uses porous material to adsorb or release a gas when cooled or heated. This type of refrigerator is well-suited to a space environment. Gas-gap heat switches (Duband 1995) are used the control the refrigerator and there are no moving parts. It can be recycled indefinitely with over 95% duty cycle efficiency and the lifetime is only limited by that of the cold stage from which it is run (in this case, the lifetime of the Herschel cryostat). The evaporation of 3He naturally provides a very stable operating temperature under constant heat load over the entire cycle. The cooler requires no mechanical or vacuum connections and only low-current electrical leads for its operation, making the mechanical and electrical interfaces very simple. For operation in a zero-g environment two aspects of the design of a 3He refrigerator have been addressed: the liquid confinement and the structural strength required for the launch. The confinement within the evaporator is provided by a porous material which holds the liquid by capillary attraction. For the thermal isolation and structural support of the refrigerator elements, a suspension system using Kevlar wires has been designed to support the cooler firmly during launch whilst minimising the parasitic heat load on the system. The base-line SPIRE cooler contains 6 STP litres of 3He, fits in a 200 x 100 x 100 mm envelope and weighs about 1.6 kg. Its performance has been analysed using the same methods that successfully predicted the performance of the IRTS cooler on orbit. When operated from a 1.8-K heat sink it achieves a temperature of 287 mK at the evaporator with a 10 (W heat load, a hold time of at least 46 hours and a duty cycle efficiency of 96%. The energy input to the helium tank during recycling of the fridge is about 700 Joules. The 3He cooler is a potential single point failure for the instrument. Its reliability and redundancy are under analysis, and an option with double parallel heat switches is being considered. Figure 2-10 is a photograph of a development model cooler which is similar to the SPIRE design.

Figure 2-10 Prototype 3He cooler

# Instrument System Design

In this section the SPIRE instrument is described in terms of its systems level implementation.  First an overview is given of the SPIRE instrument subsystem implementation, then the instrument is described as a number interacting systems.  Finally a more detailed description is given of the systems design approach taken in the critical areas of structural design; optical design; thermal design; electrical design and EMI protection.

## SPIRE Subsystems

The SPIRE instrument can be divided into two separate sections, the cryogenic subsystems located on the HERSCHEL Optical Bench inside the Cryostat Vacuum Vessel (CVV) and the Warm Electronics (WE) which are located outside the CVV on the HERSCHEL Service Module SVM below the CVV. A schematic diagram of the SPIRE subsystems inside the CVV is shown in .  The following features are illustrated in this Figure.

- The HERSCHEL Optical Bench (HOB) upon which SPIRE, PACS and HIFI are mounted is shown with Black hatching.

- The thermal straps between the various stages of the cryostat, cryocooler and the detector boxes are shown as Blue hatching

- The Photometer and Spectrometer Detector Boxes which are kept at 2K are shown are indicated by Blue crosses.

- The 4K FPU is indicated by Green crosshatching. This unit is sealed to prevent the ingress of against EM radiation and/or stray light. .

- Internal stay light baffles are indicated by Black hatching.

- Indicative Harness routing is shown by Magenta.

- The light rays from the Primary Telescope is shown by Red lines.

The only interface between the SPIRE subsystems inside the CVV and the WE on the SVM is via the 13 bulk head connectors (and associated harnesses) shown at the top of Figure 3-1.  The SPIRE harnesses that connect to the outside of the CVV connect to the WE on the left hand side of Figure 3-2.  This schematic shows the Detector Readout and Control Unit (DRCU), which consists of two sub-units; the Focal-Plane Control Unit (FCU) and the Digital Control Unit (DCU), together with the electrical interfaces to the HERSCHEL spacecraft.

Figure 3-3 shows a view of the FPU on the HERSCHEL Optical Bench.  Figure 3-4 and Figure 3-5 show views of the Photometer and the Spectrometer with the physical locations of the various subsystems.  Table 3-1 lists and describes the sub-systems and gives the unit names used in the IID-B. A detailed description of the design of all the SPIRE subsystems is given in §4.

Figure 3-1 Schematic diagram of the sub systems of SPIRE contained within the HERSCHEL CVV.

Figure 3-2 Schematic diagram of  the sub-systems of SPIRE contained on the HERSCHEL SVM.

The only interface between the SPIRE subsystems inside the CVV and the WE on the SVM is via the 13 bulk head connectors (and associated harnesses) shown at the top of Figure 3-1. The SPIRE harnesses that connect to the outside of the CVV connect to the WE on the left hand side of Figure 3-2. This schematic shows the Detector Readout and Control Unit (DRCU), the Focal-Plane Control Unit (FCU), the Digital Control Unit (DCU) and the electrical interfaces with the HERSCHEL spacecraft.

Figure 3-3 shows a view of the FPU on the HERSCHEL Optical Bench. Figure 3-4 and Figure 3-5 show views of the photometer and the spectrometer with the various subsystems. A detailed description of the design of all the systems of SPIRE is given in §4.

Figure 3-3 External view of the SPIRE FPU on the HERSCHEL Optical Bench c/w with photometer and spectrometer beam envelopes.

Figure 3-4 View of photometer side of the instrument with the FPU outer cover removed.

Figure 3-5 View of the spectrometer side of the instrument with the FPU outer cover removed.

A complete list of the description of the design of all the sub systems of SPIRE is given in §4. ubnj nbc oefh opie

Table 3-1 - List of the SPIRE sub-systems.

| Subsystem Name | Description | Unit | Redundancy |
| --- | --- | --- | --- |
| Structure | Focal plane unit structure to hold all cold sub-systems in the focal unit. This includes all thermometers necessary to monitor the instrument during cool down and operation. | HSFPU | No redundancy |
| Optics | All mirrors for the photometer and spectrometer channels | HSFPU | No redundancy |
| Filters | All filters; beam splitters and dichroics for the photometer and spectrometer channels The requirements on these are included with those for the optics. | HSFPU | No redundancy |
| Baffles | Straylight control baffles for the photometer and spectrometer channels | HSFPU | No redundancy |
| Cooler | 3He cooler unit cools the photometer and spectrometer detector arrays to 300 mK | HSFPU | No redundancy on hardware, full redundancy on control system and harnessing |
| Detector Arrays | Bolometer array modules for the photometer and spectrometer | HSFPU | No redundancy. Loss of a single Bolometer element is a soft failure mechanism but several hard failure mechanisms exist |
| Beam Steering Mechanism | This mechanism allows the photometer and spectrometer fields of view to be stepped or chopped across the sky. | HSFPU | No redundancy in the structure, mirror or flexure pivots. Redundancy in instrumentation, actuation and harnessing. |
| FTS Mechanism (SMECm) | The FTS moving mirrors drive mechanism and position measurement system. SMECm designates the mechanism and position encoder | HSFPU | No redundancy in the structure, mirror or flexure pivots. Redundancy in instrumentation, actuation and harnessing. |
| FTS encoder amplifier (SMECp) | SMECp the cold pre-amplifier that may be required for the encoder detectors. | HSFPU |  |
| Shutter Mechanism | A shutter is required in the instrument for ground test to allow the detectors to see the correct radiation environment. | HSFPU | No redundancy as not a flight item. |
| Photometer Calibration Source | Calibration source for photometer | HSFPU | No redundancy on light guide and reference source. Full redundancy on electronics and harnesses. |
| Spectrometer Calibration Source | Calibration source for the spectrometer | HSFPU | No redundancy on reference source. Full redundancy on electronics and harnesses. |
| RF Filter Modules | Each sub-system harness into the cold FPU must have an electrical RF filter to prevent EMI problems with the bolometers. These will be mounted in standard RF filter modules on the wall of the FPU box. | HSFPU | Fully redundant |
| Photometer JFET Box | JFET pre-amplifiers for photometer NTD germanium bolometers. | HSFTBp | No redundancy. Failure of single JFET amplifier is a soft failure mechanism |
| Spectrometer JFET Box | JFET pre-amplifiers for spectrometer NTD germanium bolometers. | HSFTBs | No redundancy. Failure of single JFET amplifier is a soft failure mechanism |
| Detector Read-out & Control Unit | Detector amplifier and digitisation chain and instrument control electronics. Conceptually this is a single unit however for accomodation reasons it will be split into two phyiscal units | HSDRC | See below |
| FPU Control Unit | Contains the electronics for the power conversion and distribution to the DRCU; for the control and read-out of the thermometers; cooler; calibration sources and the cold mechanisms | HSFCU | Full redundancy |
| Detector Control Unit | Contains the bias conditioning electronics for the bolometers arrays and JFET units and the lock in amplifiers and readout electronics for all the detector arrays. | HSDCU | No redundancy on the Lock-in Amplifiers. Full redundancy on DAQ. No redundancy on Bias generators but failure of a single Bias Generator propagates only to a portion of the detectors (soft failure) |
| Digital Processing Unit | Instrument on board computer – forms interface to CDMS | HSDPU | Full redundancy |
| Warm Interconnect Harness | Harnesses between warm boxes | HSWIH | Full redundancy |
| On Board Software | All on board software that controls the function of the instrument. This is all contained in the DPU | HSOBS | Implicit full with full redundancy of the HSDPU |
| FPU Simulator | A set of electronic components, either passive or active, that mimics the analogue response of the FPU sub-systems to the warm electronics. | HSFPS | No redundancy as not a flight item |
| DRCU Simulator | A set of interface hardware and computer software that mimics the response of the DRCU and FPU to the DPU and on board software. | HSDRS | No redundancy as not a flight item |

## The SPIRE Instrument as a System

The SPIRE instrument can be viewed not just as a series of physical sub-systems but also as a series of interacting systems.  Figure 3-6 is a system topology of the SPIRE instrument that attempts to break the SPIRE instrument into a number of systems areas with over lapping areas of interest. Table 3-2 expands on Figure 3-6 and gives details of what each system area represents; the issues to be addressed under each system area; the physical components that can be associated with each system and what methods of analysis and verification we intend using to ensure that each area is properly considered in the implementation of the SPIRE instrument

Although Figure 3-6  and Table 3-2 are a very much-simplified view of the systems interactions in the SPIRE instrument, they do serve to illustrate some important points about the system level design of the SPIRE instrument:

The Radiation Detection System – going from the cold detector arrays through to the digitised signals from the DRCU, is at the very heart of the instrument.  All systems issues ultimately come back to ensuring that the detection system can operate correctly and without undue interference.

The Electrical system and Structure System can be seen as the “glue” that bonds the instrument together into a single unit.  These must be right before the instrument will work as a single entity.

The EMC/EMI protection system touches on virtually every aspect of the instrument design.  The issues raised by consideration of the EMC/EMI must always be taken into account in design and implementation of virtually every physical sub-system in the instrument

Figure 3-6 – Simplified view of the SPIRE instrument as broken into “systems.”

Table 3-2 -  Description of the SPIRE systems

| System | Description/Issues | Sub-systems | Design analysis Tools | Design verification methods |
| --- | --- | --- | --- | --- |
| Structural | To ensure that the SPIRE instrument is mechanically compatible with the FIRST system and capable of withstanding the launch environment Mechanical frequency response Ability to withstand launch environment Mechanical interface with FIRST system Instrument level integration Sub-system mechanical interfaces | Primarily instrument Structure and JFET enclosures Interfaces to all cold FPU sub-systems | CAD FEM | Prototype material testing STM/CQM instrument model vibration tests CQM system level integration |
| Opto-mechanical | To ensure that only the legitimate optical radiation reaches the radiation detection system and does so in a manner that fulfils the instrument requirements Optical design Optical interface to FIRST system Straylight Instrument optical performance Integration and alignment Sub-system optical interfaces | Structure Optics Filters Calibration Sources Detector Arrays Baffles SMEC BSM | Synopsis ASAP APART Feedhorn model (Gaussian Mode analysis; HFSS) | Component testing (filters etc) Optical alignment Instrument level tests |
| Thermo-mechanical | To ensure that the different parts of the instrument run at the correct temperature and that the instrument functions at the correct temperature according to requirements for all defined instrument operating and environmental conditions Thermal performance under all operating conditions Thermal interface to FIRST system Sub-system thermal interfaces Sub-system thermal control | Structure Cooler Thermometry Temperature Control JFET Amplifiers JFET Enclosure Filters Thermal straps SCU | ESATAN model Other computer models | Prototype sub-system tests (cooler; cooler plus strap etc) STM/CQM sub-system cold tests Instrument level STM cold tests Instrument level CQM cold tests System level CQM cold tests |
| Electro-mechanical | To ensure that the moving parts of the instrument meet the instrument requirements; do not unduly influence the operation of other parts of the instrument and that the instrument can operate according to requirements in the micro-vibration environment expected in the FIRST satellite Micro-vibration environment Mechanism control Harness mechanical frequency response and routing | FPU Harnesses Detector arrays SMEC BSM Shutter JFET Amplifiers Cryostat cold harness Cryostat warm harness MCU Shutter electronics | Dynamical analysis model (DSPACE?) at sub-system level only FEM??? | Prototype sub-system tests Instrument level STM cold tests (?) Instrument level CQM cold tests System level CQM cold tests |
| Radiation Detection | To ensure that the radiation transmitted by the opto-mechanical system is efficiently detected and converted into digital signals without excess noise or contamination from other electrical signals. Detector performance versus environment (temperature; photon background; micro-vibration; EMC) JFET Amplifier performance versus environment (ditto) Harness performance Detector sub-system interface compatibility – thermal; electrical; mechanical End-to-end system performance | Detector Arrays Thermal Straps Temperature Control Cooler FPU Harnesses RF Filters JFET Amplifiers Cryostat cold harness Cryostat warm Harness DCU | Mathcad Models System analysis | Prototype cold units in representative environment with representative electronics STM sub-system cold units for thermal and environmental test CQM sub-system end to end test CQM instrument level end to end test CQM system level end to end test |
| EMI/EMC protoection | To ensure that no radiofrequency EM radiation enters the radiation detection system from any source within the FIRST system. Also that the SPIRE instrument does not emit any radiofrequency EM radiation that might influence the operation of any part of the FIRST system EMC susceptibility and emission – radiated/conducted Electrical grounding Faraday cage integrity and performance RF filter performance Harness performance Power supply cleanliness Digital/analogue separation | Structure FPU Harness RF Filters JFET Box (FSFTB) Cryostat cold harness Cryostat warm harness DRCU (FSDRC) | Systems Analysis SPICE model | EM and QM electronics units as sub-system with simulator and EMC tested CQM instrument level testing CQM system level testing |
| Electrical | To ensure that the SPIRE instrument is electrically compatible with the FIRST system and that the different parts of the SPIRE instrument are mutually electrically consistent with each other Electrical interface to FIRST system Power supply distribution and control Sub-system electrical interfaces Wiring tables Analogue to digital interfaces Digital to digital interfaces | DRCU (FSDRC) SPIRE Warm harness (FSWIH) DPU (FSDPU) S/C PDU S/C Warm harness DRCU Simulator FPU Simulator | Systems analysis | EM and QM electronics units tested as sub-system with simulator(s) CQM Instrument level testing AVM and CQM system level testing |
| Instrument control and communication | To ensure that the SPIRE instrument communicates with the FIRST system; that the different parts of the SPIRE instrument are mutually consistent with the operations concept and that the instrument operates safely and to requirements in all operational modes Data interface to FIRST system Operating mode definition Instrument commanding definition On board software definition Sub-system operational and control interfaces Sub-system data interfaces | DRCU (FSDRC) SPIRE warm harness (FSWIH) DPU (FSDPU) S/C CDMS FPU Simulator DRCU Simulator | Systems analysis Software simulators | EM and QM electronics units tested as sub-system with simulator(s) CQM Instrument level testing AVM and CQM system level testing |
| Instrument data processing | To ensure that the data produced by the SPIRE instrument are compatible with the requirements of the FIRST system and are processed into the required data products Interfaces to the ICC Data product definition Data processing definition Sub-system data processing interfaces Observing mode data processing interfaces | DPU (FSDPU) DRCU Simulator FPU Simulator ICC | Systems analysis Software simulators | Data sets produced by simulators EM and QM electronics units tested as sub-system with simulator(s) produces data sets Instrument level CQM tests for observation verification and producing data sets System level AVM and CQM tests for end to end verification |
| Calibration | To ensure that the data produced by the instrument can be converted into meaningful physical units to allow the correct operation of the instrument in all modes and the processing of the instrument data into the required data products Observing mode calibration definition Ground commissioning and calibration plan Flight commissioning and calibration plan Instrument to ground facility interfaces Ground facility definition Ground based observing programme definition | Photometer Calibrator Spectrometer Calibrator DPU (FSDPU) ICC | Systems analysis Instrument performance models | Prototype sub-system tests CQM instrument level performance verification Ground based observing programme |

## Structural Design and FPU Integration

We have already discussed the need to have various temperature zones with the SPIRE instrument.  This, combined with the need for two essentially separate instruments in the SPIRE instrument, has dictated the design approach to be taken for the SPIRE structural design.  Figure 3-7 shows the conceptual design of the SPIRE instrument structure.  A single stiff optical bench is used to mount all the subsystems and optical components, including two detector boxes that are thermally isolated from the optical bench on stiff space frames.  On one side of the bench the components for the common entrance optics and the photometer channel are mounted, and on the other the components for the spectrometer channel.  Each side of the optical bench has a cover that forms a structural “monocoque” element in the design.  The integrated instrument box is mounted from the Herschel optical bench via three thermally isolating supports.  One of these is directly mounted from the SPIRE optical bench and forms a fixed reference point, the other two are mounted from the two covers and are bipods with flexibility in one direction to allow for any differential thermal contraction during system cool down.

The instrument covers also form both a straylight shield to protect the instrument from the ambient thermal radiation environment in the Herschel cryostat and an RF shield to protect the detectors from any radiated EMI.  All sub-system wiring entering the instrument box must pass through passive RF filters mounted in boxes from the SPIRE optical bench on the spectrometer side.  When the cover is integrated with the optical bench the RF filter boxes will be sealed to the cover.  The exception to this are the harnesses for the detectors themselves that connect the bolometer arrays to the externally mounted JFET units.  These are filtered within the JFET units and then pass to the instrument box via a drilled plate hard mounted to the SPIRE optical bench.  The wiring harnesses therefore form part of the RF shield therefore and careful attention must be paid their electrical shielding.

In addition to sealing the instrument box against RF, it must also be sealed against the possibility of stray optical radiation entering via routs other than the legitimate path defined by the telescope and SPIRE optical elements.  To this end the thermal straps that must broach the covers to connect the sorption cooler and the detector boxes directly to the Herschel helium tank at 1.7 K must pass through light baffles.

The outline integration sequence for the SPIRE instrument FPU is as follows (see the SPIRE AIV Plan for more details):

- The 300 mK thermal bus bars (see section on thermal design below) are fitted into the detector boxes

- The optical elements that go into the detector boxes are fitted and their alignment verified (see Figure 3-1 for which elements are fitted into the boxes)

- The detectors are fitted in to the detector boxes and connected to the 300 mK busbar

- The “baseplates” of the detector boxes and the detector harnesses are fitted to the optical bench panel

- The mirror and filter mounts are fitted to the optical bench panel

- The positions of the interfaces between the mounts and the optical elements are verified by 3-D mechanical metrology (see section on alignment below)

- The optical elements are fitted and the detector boxes are fitted to their baseplates

- The rest of the subsystems that mount to the optical bench are fitted together with their harnesses

- The covers are placed onto the optical bench panel

- The instrument FPU can be fitted to the Herschel optical bench (or the test cryostat cold plate)

- The JFET units are fitted to the Herschel optical bench (or cryostat cold plate) and the harnesses from the detectors connected

Because the instrument FPU does not stand alone before the covers are fitted some MGSE is envisaged that will hold the optical bench panel and allow access to all sides during integration.

Figure 3-7 - Conceptual design of the SPIRE instrument structure. Once the 4-K covers are integrated the instrument box forms a straylight and RF shield

## Optical Design

### Common Optics and Photometer Optics

The 3.5 m FIRST telescope will be either a Cassegrain or Richey-Chretien system.  In either case it will provide a well-corrected image at a focal ratio of F/8.68 (IID-A).  The low focal ratio of the primary mirror (F/0.5) causes the telescope focal surface to be highly curved.  SPIRE uses an off-axis part of the telescope FOV and its object surface is therefore tilted with respect to the central (gut) ray. Figure 3-8 shows the FIRST telescope with, to scale, the SPIRE photometer.

At an early stage in the design process several decisions were taken on the system design of the instrument that were to be design drivers for the optical design:

- The spectrometer and photometer are to have separate fields of view on the sky and no beam switching mechanism is to be used

- The largest field of view available to SPIRE should be used for the photometer channel – this is 4x8 arcmin rectangular.

- Whatever is left should be available for the spectrometer – this is 2.6 arcmin circular

- An ability to steer the beam around the sky to allow chopping and spatial sub-sampling is required

- The beam steering mechanism is to be available for both the photometer and spectrometer channels – i.e. it is to be placed ahead of any division of the fields into the separate instrument channels

- The secondary of the telescope is to be used as the final exit pupil of the instrument and a cold aperture stop close to the detectors is required

- No compromise is made in the photometer optical design to accommodate the spectrometer channel.

- Imaging in the Spectrometer is not a key requirement; the primary science driver is for point sources viewed on axis.

Figure 3-8 Ray diagram of the FIRST telescope with the SPIRE instrument to the same scale. The primary mirror has a diameter of 3.5 m and the SPIRE instrument is about 500 mm wide.

The SPIRE optics are therefore divided into a common portion followed by the photometer and spectrometer optical chains.  The common optics form an image of the secondary at the location of the beam steering mirror and provide a well corrected intermediate focus to allow the separation of the spectrometer and photometer beams into different parts of the instrument box

In addition the combination of the common optics and the rest of the photometer optical train has to provide;

- a well-corrected, flat focal surface, perpendicular to the gut ray, at F/5 at the proposed location of the detector arrays – i.e. correction for the highly tilted focal plane of the telescope;

- sufficient back-focal length (BFL) to allow separation of the photometer beam into three spectral bands using dichroics;

- a non-aberrated image of the telescope secondary at a physical location as close to the photometer detectors as possible and physically accessible to allow a real stop to be placed there.

All these functions are assured by the SPIRE common and photometer optical design as shown in Figure 3-9.  M3 is an off-axis ellipsoid projecting an image of the telescope secondary (M2) onto M4.  This image is well-corrected and in focus at the centre of the M4 so that the pupil image at the cold stop stays fixed during chopping and beam steering.  M4 is a flat mirror whose orientation is adjustable in flight to permit 2¢ chopping in the sagittal plane, allowing off-field chopping for a 4¢ x 4¢ sub-field, and 30¢¢ beam steering motion in both the tangential and sagittal planes to obtain fully Nyquist sampled images.  The toric M5 mirror reimages the focal plane onto M6 and converts the focal ratio of the beam from f/8.68 to f/5.  While M3 and M4 are common for both photometer and spectrometer, the two systems separate at M6.  The photometer M6 is toric and sends the beam into an Offner-type relay system consisting of three spherical mirrors: M7 (concave), M8 (convex), and M9 (concave). The tilt angles, separations and curvatures of these mirrors provide enough free variables to satisfy the need for well corrected flat image planes; a long back focal length and an image of the telescope secondary mirror close to the detectors.  In particular, an easily accessible pupil image is provided between M8 and M9 in which the cold stop is located.  The physical separation between the 4K cavity and the 2K cavity (detector box) into which the detectors are mounted, is made here so that the entrance to the cold environment containing the detectors is as small as possible to restrict the possibility of straylight hitting the detector arrays.

Figure 3-9 - Ray diagram of the SPIRE instrument showing the ray paths for three points in the tangential plane, centre and extremes of the FOV (2().  Also shown here are the locations of the straylight baffles and optical bandpass filters

Figure 3-10 - The beam folding within the photometer detector box.  The dichroics are low pass devices so DET1 is the shortest wavelength array; DET2 the medium wavelength array and DET3 the long wavelength array.

The detector box contains M9, the last powered mirror of the optical train, and dichroics and beam folding mirrors distributing the light between three detector arrays, covering the short (200-300 µm), medium (300-400 µm) and long (400-670 µm) wavelength bands. An edge filter at the cold stop minimizes stray radiation entering the cold box and band-pass filters in front of each detector array ensures the spectral limitation of each band. Figure ?? shows a 3D view of the photometer detector box optics.

Figure 3-11 shows geometrical spot diagrams across the photometer FOV. The final focal surface is flat and perpendicular to the gut ray, and the exit pupil is close to telecentric. With a theoretical Strehl ratio better than 0.986 at 250 µm the system leaves headroom for manufacturing tolerances and the field distortion is below 1.1%.

Figure 3-11 - Geometric spot diagrams across the 4( x 8( SPIRE photometer FOV. The spots are plotted in their actual positions and to scale. The concentric circles around the central spot have diameters 3.0, 4.3, and 6.1 mm and indicate the Airy disk size at 250, 350, and 500 µm, respectively. With a maximum RMS wavefront error of 4.7 µm, the theoretical Strehl ratio is better than 0.99 anywhere in the FOV at 250 µm. A slight distortion is observed, corresponding to 6( or 1.1% of the FOV diagonal. The average focal ratio is F/4.9.

### Spectrometer Optical Design

An imaging FTS has been chosen for SPIRE rather than a grating-based solutions because of its superior imaging capability, lower stray-light sensitivity, and variable spectral resolution (RD3).  Among a large number of possible interferometer concepts, three were chosen for a final comparison, see Figure ??.  In each case, two separate, band-limited detector arrays are required to divide the 200-670 µm band into two sub-bands: 200-300 µm, and 300-670 µm.  Covering the entire spectrum with a single array would be too constraining with respect to spatial sampling and detection efficiency.

The Martin-Puplett interferometer (Fig. ?? a) offers a good and robust solution to this problem. Based on the use of three polarisers, components which can provide excellent efficiency over a broad band in the far infrared region, and roof-top mirrors, it provides two input and two output ports with a minimum of complexity. When the polarisers are properly oriented, the roof-top mirrors switch the polarisation of the beams so that 100% of the light incident upon P2 is transmitted towards the detectors. One polarisation of the incident light is lost at P1 however, reducing the optimal efficiency of this concept to 50%. P3 is required to analyse the interfering beams, sending complementary interferograms towards the two detectors. Usually, both detectors would see the entire band, hence detecting all the light incident onto P3, but in our case the spectrum would have to be divided into two by band-pass filtering each detector.  This loses another 50%, reducing the theoretical efficiency to 25%.

Figure 3-12 Three possible interferometer concepts for the SPIRE spectrometer: (a) Martin-Puplett polarising interferometer, (b) classical Michelson interferometer, and (c) Mach-Zehnder type dual beam interferometer. BB: blackbody source, RT: roof-top mirror, P: polariser, F: filter, D: detector, CC: corner-cube reflector (could also me mirrors or roof-tops), BS: beamsplitter, Dic: dichroic beam divider, M: mirror, BCC: back-to-back corner cubes (or roof-tops).

As a second option, we considered a simple Michelson interferometer as shown in figure ?? (b).  This option was made possible thanks to a new development of 50/50 beamsplitters (Reference to PARAs paper on beam splitters), providing greater than 90% efficiency (4RT) over the entire SPIRE band.  No output polariser is required in this case and it can be replaced with a dichroic beam divider, offering a theoretically loss-less channel separation. There is of course a 50% loss at the beamsplitter since half the incident radiation is sent back out through the telescope. This configuration is still twice as efficient as the previous one.  Its main drawback is the lack of a second input port, required for balancing off the telescope background radiation.

The preferred solution is shown in figure ??(c). Rather more complex than the former options, it provides both a second input ports and a 50% theoretical efficiency.  The concept is based on a Mach-Zehnder interferometer with its arms folded in order to avoid beam shearing during scanning of the optical path difference (OPD) and uses two 50/50 beamsplitters. If the detectors could be used over the entire spectral range, this concept would provide 100% efficiency, but the requirement for two separate bands imposes a 50% channel separation loss as in the Martin-Puplett case.  The folding allows the optical path of both arms to be changed simultaneously with a single scanning mechanism, hence doubling the available resolving power for a given mirror-moving mechanism.  A resolving power of 1000 at 250 µm, requiring a maximum OPD of 125 mm, is therefore obtained with a lopsided movement from -3 mm to +31 mm. The lowest resolving power, R= 20, is achieved using a double sided scanning of 0.6 mm.

One of the difficulties encountered in the optical design of the interferometer concept of figure ??(c) was the long distance between separation (at BS1) and recombination (at BS2) of the beams. Due to the FOV dependent beam spread, the size of the beam splitters and collimating and camera optics became prohibitive.  Also, it was difficult to find space for the scanning mechanism.  To improve the situation it was decided to move collimator and camera optics to within the interferometer by making the four mirrors M1, M2, M1(, and M2( of figure ??(c) powered.  This is not without disadvantages, since at non-zero OPD, the two arms do not see the same optical system.  A differential aberration analysis is therefore necessary.  Keeping to a strict scheme of symmetry ensures minimal aberrations in the system, and the only residual aberration of some concern is differential distortion giving a lateral separation between the images of a point source at the edge of the FOV.  The induced contrast reduction is not negligible but small compared with other sources, notably alignment tolerances

Figure 3-13 shows the ray diagram of the upper half of the spectrometer.  The lower half has the same optical design. After reflection from the common mirrors M3, M4, and M5, the spectrometer beam is picked off by the toric M6s and sent out of the plane of the photometer system.  The flat M7s redirects it into a parallel plane, separated by 170 mm from the photometer plane.  The input relay mirror (Rin) focuses the beam to an intermediate image plane located just after the first beam splitter, after which the beam is collimated (Coll) and sent vertically towards the corner cube assembly.  The corner cube, modelled by non-sequential raytracing, shifts the beam and sends it up towards the camera mirror (Cam).  Symmetrical with the collimator, the camera focuses the beam to an image plane just before the output beam splitter.  The output relay mirror (Rout) focuses the beam onto the detector arrays.  To accommodate the components within the available volume, a fold mirror is needed to take the beam out of the plane again.  The input and output relays are toric in order to control astigmatism and image anamorphism.  A slight asymmetry in the input and output relays is introduced in order to adjust the final focal ratio.  The collimator and camera mirrors are spherical.

A pupil image is located near the final fold mirror, making this a convenient place for the entrance hole in the 2K enclosure.  This pupil moves as the OPD changes, however, so it is not appropriate for a limiting cold stop. Instead, a limiting aperture is placed in another pupil image at 4 K located between M6s and M7s.

Figure 3-13 Raytracing diagram of the upper half of the SPIRE spectrometer. The symmetrical lower half is generated by reflection about the plane containing the two beam splitters. The location of the cold stop and the bandpass filters are also indicated.

Figure 3-14 shows spot diagrams for the spectrometer.  Clearly the imaging performance is not quite as good as that of the photometer, the spots reflect the image quality in the intermediate focal plane at M6s.  Since the planar symmetry is lost, it is very difficult to improve on this.  However, the astigmatism has been brought to zero at the centre of the FOV and a good balance of aberrations over the rest of the FOV has been achieved by introducing a 3.8° rotation of the output relay mirror around its normal.  The worst rms wavefront error is 6.6 microns, giving a Strehl ratio at 250µm of 0.97.  Apart from a slight rotation, the image suffers from a distortion of up to 9((, corresponding to 6% of the FOV diameter.

Figure 3-14  - Geometric spot diagrams at the centre, half field, and full field of the 2.6( diameter spectrometer FOV. The spots are plotted in their actual positions and to scale. The concentric circles around the central spot have diameters 3.7 and 6.7 mm indicating the Airy disk size at 300 and 550 µm, respectively. With a maximum RMS wavefront error of 6.6 µm, the theoretical Strehl ratio is better than 0.97 anywhere in the FOV at 250 µm. Distortion corresponding to 9(( or 6% of the FOV diameter is observed. The average focal ratio is F/4.9.

One feature of the optical design of the spectrometer is that the final re-imaging optics are non-telecentric – that is the beams at different parts of the field of view are not parallel (or nearly so), unlike in the photometer design.  This feature is forced by the lack of physical space available to the instrument.  Whilst with most detection systems this would not make very much impact, with the feedhorn arrays this causes some loss of signal as the detector “beams” only illuminate a portion of the pupil for off axis detectors.  Figure 3-15  illustrates the situation.

Figure 3-15 - Sketch of the final optical beams onto the spectrometer focal planes.  There is a real pupil image of ~20 mm at about 100 mm (check these distances) from the focal plane which has a physical diameter of 12.2 mm. Off axis images therefore come onto the array at an angle of up to 3.5 degrees – shown by the dashed line.  The detectors themselves, because of the feedhorns, can only accept a certain range on input angles and, as they are configured to stare straight ahead, they only partially illuminate the pupil – shown in pink.  Whilst this only causes a small loss of signal at the centre of the FTS mirror movement, at large mirror displacements the pupil images from the two interferometer arms shear past each other and the loss in fringe contrast will be greater.

## Straylight control

Straylight control is defined as the reduction of any unwanted optical power falling onto the detectors in an instrument to as low as practically possible.  This includes both the removal of out of spectral band optical power and the removal of in band optical power from sources outside of the field of view.  These might be radiating surfaces within the structure of the instrument or satellite directly viewed by the detectors or seen via reflection or diffraction from other parts of the optical chain and structure.

### Bandpass Filtering

Any surface with a temperature greater than ~10-15 K will emit radiation in the detection band of the SPIRE instrument.  The Herschel telescope will be at a temperature of about 70-80 K with an effective emissivity of 3-4%.  There will, therefore, be a relatively large photon flux falling on the SPIRE instrument entrance aperture from the telescope alone.  The bolometer detectors are sensitive to optical power at all wavelengths, therefore the first task in managing the amount of unwanted optical power falling on the detectors themselves is to limit the spectral band of the radiation falling on the detectors.   This is achieved by a series of optical filters as described in the SPIRE Filter Specification Document.  In practice a single passband filter does not have enough out of band rejection on its own and four filters are used strategically placed along the optical path to reduce the out of band radiation entering each part of the instrument structure.  The overall passband provided by these filters is shown in figure ??; the physical locations of the filters are shown in figures ?? and ?? above.

Figure 3-16 - Example showing the definition of a 350 micron photometer channel using real filters. The passband ((/(() in this case is 3.

### Baffling

Controlling unwanted optical power that is in the spectral band of the detectors is effected in the SPIRE instrument by having a series of compartments within the instrument separated by physical barriers or baffles.  The overall concept for the photometer is shown in figure ?? and for the spectrometer in figure ??.  The optical design, as described above, is made so that the entrance to the instrument itself is at the field plane of the Herschel telescope and a physical aperture plate is placed here together with an optical filter that rejects most of the out of band radiation.   The aperture plate is made just large enough to allow the beams generated by the detectors to pass out of the instrument with the minimum vignetting.  In fact the beams will be chopped and jiggled around the sky so the aperture plate is slightly oversized compared to the detector footprint on-axis –figure ?? (also see Definition of a combined focal plane plate for the SPIRE instruments).

Figure 3-17  Chopped and jiggled footprints of the photometer and spectrometer arrays at the focal plane of the Herschel telescope.  The solid line shows the boundary used to define the physical stop that will be placed at the entrance to the SPIRE instrument.

Figure 3-18 - Overview of the straylight baffling scheme for the photometer

Figure 3-19 - Possible baffling scheme for the spectrometer

The photometer detectors are mounted to a sealed box that is kept at the<2 K by directly strapping to the Herschel helium tank.  An image of the secondary mirror of the telescope is made at the entrance to this box and a physical aperture placed around at exactly the size of the image – see figure ??.   In figure ??a the M4 beam-steering mirror in its neutral position, in (b) it is tilted by 2.17°, changing the instrument pointing by 2( in the sagittal plane.  Assuming the nominal M2 image to run through the centre of gravity

| (a) | (b) |
| --- | --- |

Figure 3-20 - Pupil spot diagrams obtained by tracing rays from 15 positions in the FOV through 16 points along the rim of the telescope pupil to the cold stop

of each spot, we may measure the radial pupil error R for each point in the FOV at each point along the pupil edge.  A useful measure of pupil aberration is relative pupil displacement R/R, found to be less than 5% both in the chopped and unchopped configuration.  The good control of the geometric image of the telescope secondary means that the primary source of unwanted radiation from sources outside of the field of view from any part of the Herschel cryostat; telescope or sunshade is that passing into the instrument via reflection or diffraction from the telescope secondary mirror.  Control of the straylight from such sources must carried out by the system level straylight control design. A further baffle is implemented in the photometer optical path between M7 and M8 to prevent the direct illumination of the cold box aperture by sources within the instrument box itself.  The second bandpass filter is mounted at this point.

The straylight control within the spectrometer is more problematic than in the photometer as the moving mirror means that the final image of the telescope secondary mirror both shears and moves along the optical axis.  Instead of placing the pupil at the entrance to the cold detector box, therefore, it is placed at an image of the secondary made by the Fabry mirror SM6 at the field plane where the spectrometer and photometer fields are separated.  A physical stop is placed at this pupil image, which is at 4 K.  A filter is also placed at the stop.  As this stop is only ~25 mm diameter and, apart from holes for cables and thermal straps is the only aperture in the optical bench, the spectrometer part of the instrument box is very well baffled against radiation entering the instrument through the common entrance aperture.  Two other baffles will be placed in spectrometer; one to prevent stray radiation from the calibrator reaching the detectors and one to prevent radiation from the mirror mechanism reaching the detectors – no filters will be placed on these baffles.  The final baffle in the spectrometer is at the two entrances to the cold detector box where another filter will be placed as in the photometer.

The straylight model of the instrument has been constructed using the BRO APART code.  A “reversed” Code V model has also been constructed to generated the optical paths from the detectors out through the system to space.  These are the solid model beams shown in the general views of the instrument.  By reversing the optical model one can determine more easily what it is possible for each part of the field of view to directly “see” within the Herschel system.  The APART code is used to determine which straylight paths reach to detectors via multiple scattering and diffraction and what level of radiation falls onto the detectors from those paths given the physical construction; temperatures and scattering properties of the real system.

### Diffraction Limited Optical Analysis

One final aspect of the optical design that influences both the amount of straylight falling onto the detectors and the optical performance of the instrument in terms of both throughput and image quality, is the diffraction limited nature of the SPIRE optics.  For the SPIRE instrument the sizes of the optical elements and stops within the system are significant compared to the longest wavelength radiation that has to pass through the system – (~0.7 mm compared to optical stops of 20-25 mm in some cases.  Whilst use of the feedhorns on the detectors makes the diffraction limited nature of the optical design tractable by effectively acting as spatial filters and reducing sidelobes to a minimum, great care must still be exercised in the optical design to ensure that truncation of the beam from the detectors only occurs at the desired apertures.  That is we wish to limit the effective throughput of the instrument only at the telescope secondary and its image within the instrument optical train.  Some truncation at the edges of the field of view is inevitable at the field stop.  All other optical components in the SPIRE instrument must be sufficiently oversized to allow the beams from the detectors to pass un-truncated at each and every point in the field of view.

The most basic rule of thumb that can be employed in taking into account the diffraction limited beam is to oversize all components by 20% of the geometrical footprint of the beam at the location of the element (see figure ??).  This is complicated by having a wide field of view as all beams from each detector must be amalgamated to make a single “instrument beam” before the oversizing is determined.  This has been successfully done for all elements in the photometer optical train.  In the spectrometer the situation is more difficult due to the moving mirror.  The spectrometer components have been oversized to the physical limits possible but some truncation will still occur for some parts of the field of view at the limits of the mirror travel.

The BRO ASAP program has been used to ensure that the “20% rule” is indeed sufficient given the real Gaussian modes generated by the feedhorns and to determine what influence the truncation of the Gaussian modes has on the point spread function of the SPIRE instrument.  Figure (which I will find get from MEC) shows the comparison between the theoretical point spread functions for the Herschel telescope and those calculated for the expected Gaussian modes for the SPIRE wavelength bands.

One further aspect of the diffraction limited design of the SPIRE instrument is the change in focal length of the instrument optics with wavelength due to the quasi-optical nature of the relatively low frequency (long wavelength) radiation to be detected (Goldsmith or Lesurf as references).  This can be analysed using the ASAP code (Martin’s many notes and Munich paper) and the positions of the real foci found with respect to those given by the geometrical optical design.  The detectors will be displaced along the optical axis at the position that gives the best coupling between the radiation field from the telescope and optics and the Gaussian mode generated by the feedhorns.

Figure 3-21 - Illustration of the 20% oversize rule and the comparison between 
component margin and instantaneous beam footprint.

Note: this is much too complicated to explain in less than many pages of dense text and mathematics – I suggest we leave the interested reader to go to Martin’s paper or tech notes…..

### Optical Alignment

Alignment of the SPIRE instrument will be performed according to a philosophy based on high-precision machining and pre-assembly 3D measurements and a program of optical alignment checks during and after assembly. Nominally no adjustments will be necessary but, if a serious misalignment is detected, its compensation will be possible by re-machining of the M6 mirror stand. Alignment of the instrument with respect to the telescope axis and pupil will be performed using a Herschel optical bench simulator consisting of a set of reference mirrors accurately located with respect to the instrument interface points. Verification of image quality and internal alignment stability will be effectuated using a set of alignment tools (sources, reticules, theodolites etc) mounted in strategic positions in the optical train (object plane, cold stop, image plane).  The detailed alignment procedures are described in the SPIRE Alignment Plan. Details of the calculation of the optical alignment budget are given in the SPIRE Optical Error Budgets document, here we give an overview of the requirements on the alignment.

An optical sensitivity study shows that with an alignment tolerance of 0.1 mm and 1( applied to all the mirrors in the photometer, the alignment-related relative pupil displacement will be less than R/RI = 4.1%. This compares favourably with the contribution from telescope alignment errors, budgeted to R/RT = 6%, and the theoretical design contribution of R/RD = 5%.  The total instrument budget is estimated by square-summing of the random alignment errors and summing of the deterministic design error:

,

giving a total of 12.3%.  For an exactly sized pupil this gives a loss of 8% in telescope transmission factor.

For the spectrometer, the predominant alignment criterion is interferometer contrast, calculated from the misalignment-induced lateral separation of the interfering images using the van Cittert-Zernike theorem. Note that this only concerns mounting tolerances of the fixed optical components within the interferometer (beamsplitters and collimator/camera mirrors) since the interferometer design with its back-to-back corner cube reflector leaves the interferogram contrast insensitive to errors in the scanning movement. Again, tolerances of 0.1 mm and 1( have been found appropriate, offering a contrast in the interferogram of 87%. Including mirror surface quality and differential aberrations a total contrast greater than 80% is expected.

## Thermal design

The thermal design of the SPIRE instrument is predicated on the following constraints:

- The need to reduce the temperature of the structure; mechanisms and optics to below a temperature where they can radiate significant optical power onto the detectors

- The need to get the bolometers to 300 mK or less and keep them at this temperature in a stable manner for all instrument observing modes

- The need to minimise the thermal load into the Herschel cryostat to make the mission last as long as possible given the amount of liquid helium available

### Instrument temperature levels

Given these drivers and the structural layout of the SPIRE instrument, three basic temperature levels have been defined within the SPIRE FPU.  The various optical components (mirrors, filters, beam splitters etc.) and other instrument sub-systems (structure, mechanisms, harnesses etc.) must not be allowed to radiate significant IR power in these wavelengths. The temperature of all components that can either directly or indirectly (i.e. via multiple reflections) irradiate the detectors must therefore be below 5 K.  The main structure of the instrument and most of the optics are therefore held at 4 K (termed Level 1 of the cryostat); the structure that houses the detectors is held at 2 K (Level 0), and, in order to achieve a sensitivity matched to the photon noise limit from the telescope thermal background, the detectors must be at 300 mK – this temperature is generated internally by the sorption cooler which is mounted from the 4-K optical bench.  The detector JFET amplifiers must run at a temperature of ~120 K so they are mounted on silicon nitride membranes which isolate them thermally from the surrounding structure.  These membranes are then housed in a JFET units that is hard mounted to the Herschel optical bench at the cryostat level 2 temperature of between 9 and 12 K.

Figure 3-22 - Summary of the thermal analysis model for the subsystems within the CVV.

### Cryogenic heat loads

To maximise the life of the mission and therefore the quantity of science data, the rate of consumption of helium must be minimised.  This is done by (i) minimising the power dissipated in the various sections of the instrument, and (ii) by minimising the flow of heat from warmer stages to stages requiring lower operating temperatures. Stringent thermal budgets have been applied to all dissipating elements in the focal plane, and to the mechanical mounts that support the FPU from the Herschel optical bench; the mechanical mounts for detector cold boxes from the SPIRE optical bench and the support system for the various elements that have to be maintained at 300 mK.  The electrical harnesses for the mechanisms and the detectors also contribute significantly to the thermal loads and these will have to be carefully designed to ensure that they meet both the electrical and thermal requirements placed upon them.  The heat flows within the instrument are represented in an ESATAN model and shown schematically in figure ??.  A detailed description of the SPIRE thermal mathematical model is found in the SPIRE Thermal Configuration Control Document.

### Temperature stability

The temperature of the bolometric detectors must be highly stable to achieve the required sensitivity, and there are corresponding requirements on the thermal stability of the instrument itself.  A detailed analysis of the effects of temperature drift on the sensitivity of the instrument in various operating modes has been conducted (Jamies note) and shows that the detector temperature must have a stability of 0.6-1.2  in the detection band of the detectors and, further, must drift by no more than 1 mK per hour over the course of an observation.  There is a very weak direct dependence on the temperature stability at the 1.7 K level and the corresponding figures for the 4-K level are, for the case of drift scanning to detect point sources, 5  and 0.3 K per hour.

## EMC

The Bolometer detectors act effectively as Light Power ( Voltage transducers. Any photon from any portion of the Electro-magnetic spectrum absorbed by a detector element will therefore generate a spurious signal. Over and above this, any spurious voltage generated across the detector by conducted EMI will be dissipated via ohmic heating of the detector element. The temperature rise of the detector elements via this heating will also generate a spurious signal. It is critical therefore that the detector elements be protected from EMI via an effective RF shield and a grounding scheme that prevents ground loops and protects the sensitive parts of the system from injected EMI.

### Signal quality

The signal; gain and noise figure expected at each portion of the detection chain is shown schematically in figure ??.  The basic signal level is of order of a few mV at the detector and, after current amplification, at the output of the JFETs.  The voltage equivalent noise at the detectors is of order 25 nV/(Hz so all amplification and digitisation stages must contribute insignificant noise compared to this.  The value of the basic detection noise also represents the sensitivity of the system to injected EMI.

### Grounding and RF Shield

In order to maintain the very low noise required at the detector it is necessary to create an RF shielded environment around the detectors and to filter all wiring entering the enclosure.  In the SPIRE instrument the RF shielded enclosure is created using the FPU box and the JFET boxes together with harnesses running between them – this is shown in yellow in figure ??.  The only significant aperture into the RF shield will be the optical entrance aperture of the instrument itself.  Whether there is sufficient attenuation of RF frequencies by the apertures along the optical path of the instrument remains to be seen.  If tests prove that that further attenuation is required, a “chicken wire” type RF filter will be fitted to the instrument entrance aperture.

The grounding of the SPIRE electrical system is complex, even more complex than shown in figure ??.  The basic grounding scheme treats the two halves of the instrument, the photometer and spectrometer, as entirely separate electrical systems with their own power supplies and grounding.  The analogues grounds of the two systems can be brought to a unipoint connection to the system ground through the thermal straps to the cooler – as shown in figure ??.  However, there is a separate ground and power supply for the digital part of the detector control electronics (not shown in figure ??) which also has to be connected to system ground.  There is some concern that this may lead to injection of noise into the analogue system via the system ground.  Another possibility therefore is to isolate the RF shield at the cold end and connect to system ground in the warm electronics.  Both these possibilities are open to us providing the mechanical structure of the SPIRE instrument is electrically isolated from the cryostat.  This will be implemented by electrical isolation washers on the feet of the FPU mounts and sapphire plates on the thermal straps between SPIRE and the Herschel cryostat.

Figure 3-23 - Schematic representation of the SPIRE signal chain showing the noise sources and gains through the system

Figure 3-24 - Outline grounding diagram for the SPIRE instrument – this is subject to change and is shown for illustrative purposes only.

Figure 3-25 - Spice model of the detector driving electronics used for EMC analysis.

### Microphonics

The impedance of the Bolometer detector elements are relatively high, being in the order of 5M The high impedance of the detection circuit is reduced by JFET transistors located outside the FPU which convert it to approximately 7 k. Any differential charge that appears on the harness between the detectors and the JFETs will cause a significant voltage to be induced across the detector element.

Charge can build up on the conductors due to very small amplitude vibrations of the harnesses. This phenomenon is called micro-phonics or phonics. The physical source of the charge build up can be minute changes in the wire-ground plane capacitance or the physical solicitation of asperities on the conductor and insulator interface. The magnitude of microphonic induced noise increases with the magnitude of vibration of the harness and the support structure.

As the harness connecting the bolometer elements to the JFET units represents the most vulnerable part of the detection system it will be strapped down to the structure to give a minimum resonant frequency of any part of the harness of 1 kHz.  In practice this means running the harnesses via the cold detector box supports, onto the optical bench and along the FPU covers.  The harnesses will be strapped to the structure about every 1 cm.  This routing is not optimum for the electrical requirements as the capacitance of the harness will be near to 100 pF, meaning that the bias frequency may have to be lower than initially desired to prevent roll off.  However, the need to prevent induced noise from microphonics is felt to be an overriding concern as the bias frequency can be varied sufficiently in the electronics to optimise the performance of the detection system.

## System Level Criticality

A top level analysis has been conducted into the effects of a failure or partial failure of one of the SPIRE subsystems (Assessment of System Level Failure Effects for SPIRE).  In this analysis the following failures are shown to be mission critical – i.e. a failure of one of these sub-systems will cause major loss of scientific capability for the SPIRE instrument:

- Total loss of the cooler

- Structural failure in the 300 mK system leading to thermal short

- Total loss of the photometer long wavelength array

- Total loss of either spectrometer array

- Total loss of the FTS mirror mechanism

All other sub-system failures will lead to a greater or lesser degree of loss of performance and difficulty of operation, but they do not lead to a total failure of either the photometer or spectrometer scientific goals.  The redundancy and reliability of these sub-systems will be addressed as a first priority.

For most sub-systems cold redundancy can be provided to ensure a high probability of avoiding total failure in any part of the sub-system implementation (see section 3.8.1).  However in some cases this is not possible, for instance there will not be multiple detector arrays and only a single cooler will be fitted.  In the case of the detectors reliability is achieved by having many pixels arranged in blocks for the purposes of power supplies and multiplexing into A-to-D converters etc.  In the case of the cooler, and the 300 mK thermal architecture, either large safety margins will be implemented backed up by testing or “soft” failure modes will be designed to prevent dead thermal shorts in the event of structural failure.

An additional method of providing operational reliability is to define backup operational modes for the sub-systems and instrument.  The following instrument backup operating modes are required in event of sub-system or system failure:

- More frequent cooler recycling including the possibility of autonomous recycling under control of the DPU alone.

- Slow chop mode in the event of partial BSM failure

- Open loop BSM control using commanded current to the actuators

- Single axis BSM operation

- Slow scanning of FTS mirrors

- Step and look operation of the FTS in conjunction with the BSM

- Open loop operation of the FTS mechanism by commanding the current to the actuator

- DC operation of photometer calibrator this will allow V-I’s on detectors under different loadings for calibration

- Selection of smaller numbers of detectors from photometer arrays in event of telemetry bandwidth problems

- Selection of smaller number of spectrometer detectors in event of problems with telemetry bandwidth and/or loss of spectrometer calibrator

## Redundancy Scheme

The general design philosophy of the instrument, as far as is possible, is that the total failure of a single sub-system will does not lead to the total loss of instrument operations. In order to achieve this, SPIRE has been designed with both a prime and a redundant side to the instrument. There is no electrical cross strapping between these two sides of the instrument, except at the SPIRE/Herschel data interface where both the Prime and Redundant HSDPU subsystems are each connected to the Prime and the Redundant Herschel MIL-STD-1553 data buses (See Figure 3-2).  To normally switch between the Prime and Redundant sides of the instrument, the Spacecraft sends appropriate commands on the IEEE 1553 bus to the HSDPU to firstly shut down the HSFCU and then to prepare itself for shut down. Once the HSDPU is ready, the two prime spacecraft level LCLs in the HPDU that power the HSDPU and HSMCU are unlatched and the Prime side is then shut down. The two redundant LCLs are then latched and the Redundant HSDPU and HSFCU are powered up.  Due to impracticalities,  some systems (for example, the Optical System) are not duplicated. In these cases, they can be controlled by either the Prime or Redundant side of the instrument.  Importantly, in the signal detection subsystems, there is no redundancy  in the detectors, the JFETs and the Lock-in Amplifiers. The specific redundancy scheme adopted for each sub-system is described in Table 3-1.  Figure 3-2 illustrates the redundancy in the Warm Electronics.  Figure 3-1 shows the redundancy in the FPU subsystems. It can be seen that the hermetic connectors 10 and 11 on the CVV wall are harnessed to the Prime Filter Boxes (and from there the prime subsystems) Connectors 12 and 13 are connected to the redundant Filter Boxes and redundant subsystems.

## System Budgets

A summary of the various budget allocations to SPIRE are summarised below in Table 3-3 .

Table 3-3 - SPIRE Budget allocations

| Item | Budget Allocation |
| --- | --- |
| Focal Plane Mass Budget | 57.6kg |
| SVM Mass Budget | 30kg |
| FPU ( Stage 0 Thermal Load | 10 mW average | 100 mW peak |
| FPU ( Stage 1 Thermal Load | 25 mW average | 100 mW peak |
| JFB ( Stage 2 Thermal Load | 33 mW average | 200 mW peak |
| WE dissipation to SVM | 86W |
| Data transmission from HSDPU to CMDS | 100 kbps |

# Subsystem Design

## Warm Electronics

The Warm Electronics are located on the Service Module of Herschel below the Cryostat. The temperature of the SVM is conditioned to remain between -15°C and +45°C during operation of the equipment. The principal function of these subsystems is to (i) act as the back end for the instrumentation contained in the CVV, and (ii) to interface SPIRE with the Herschel spacecraft. This is illustrated in Figure 4-1.

Figure 4-1 - Schematic representation of the Warm Electronics.

### Digital Processing Unit (HSDPU)

The DPU interfaces with the Detector readout and Control Unit subsystem (DRC) and with the spacecraft telemetry, telecommand (the on board CDMS, Command and Data Management system) and power systems. Figure 4-2 is a high level block diagram showing the electric and data interfaces of the DPU to the spacecraft and the DRC. The interface with DRC will be composed by three high speed data links for science data collection and one low speed, bi-directional serial bus with three output buffers, for command transmission and housekeeping data collection The interface with the spacecraft has a baseline data rate of 100 kbps, with burst mode transmission up to 300 kbps. This interface is compliant with the MIL-STD-.1553B standard, with the DPU acting as a remote terminal and the spacecraft CDMS acting as the bus controller.

DC/DC Converter: The DC/DC converter is a SMPS running off a spacecraft level clock signal at 131.072 kHz which powers the DPU board. The input voltage provided by the s/c power bus at 28V and the nominal output voltage provided to the 5V required by the DPU board. The average power delivered to the

CPU Board: A schematic representation of the information flow within the DPU is shown in Figure 4-3.  On the top left hand side of the Figure, the DPU “Command handling” electronics receives and interprets commands from the CDMS. Upon successful receipt of these commands, an acknowledge message is sent back to the bus. The high level commands from the CDMS are processed by the DPU to provide digital driver commands to the DRC (which acts the analogue driving electronics for the cryogenic part of the instrument) and are in general executed in real time. It is the task of the spacecraft to determine the sequence and absolute timing of the commands to be passed to SPIRE. These commands are passed to the DRC via the “Commanding” block in the lower left hand corner of Figure 4-3. Science frames and housekeeping data from the instrument are passed to the “Data Collection” block via the DRC. The housekeeping data (such as instrument temperature, control currents etc.) are monitored by the “Autonomy” block. In the event of a fault condition in the DRC or in the cryogenic section of the instrument, commands are passed to the DRC to put the instrument into safe mode. Notification of a fault condition is also passed to the CDMS via the “Packetisation Block” in the top right corner of Figure 4-3. The CDMS then determines if the instrument should be turned off and unlatches the LCL for the DRC and the DPU as necessary.  All data passed from the DPU to the CDMS is formatted according to the ESA Packet Utilisation Standard.

The core microprocessor of the DPU is the TEMIC TSC 21020 which is the Analog Devices 21020 Digital Signal Processor adapted for space applications by TEMIC. The CPU board will be based on this chip (20 MHz clock), with:

- an appropriate timer for time management and synchronisation purposes

- a watch-dog system that can be hardware disabled through jumper;

- Rad Tolerant memories and components

Figure 4-4 shows the block diagram for the CPU board. There are three high speed, unidirectional  synchronous serial 1355 links with a 1 MHz clock between the DRC and the CPU board. These three links upload the science data from the DRC. Three serial synchronous buses are used to interface the control electronics of the focal plane unit subsystems: the bus will be used to transmit commands and receive responses or to control and receive housekeeping information. The baseline clock speed is 0.2 MHz.  Figure 4-5 shows the memory organisation of the CPU and the interfaces between the DPU and the spacecraft and DRC with the size of the various types of memories. The CPU directly accesses 1 MB of EEPROM memory where the boot sequence software resides. It also accesses 2MB of SRAM for program data.

Figure 4-2 High level DPU block diagram.

Figure 4-3 - Conceptual diagram showing the flow of information within the DPU

Figure 4-4 – CPU board block diagram.

Figure 4-5 - DPU internal memory

A SMPS within the DPU converts the 28V power from the spacecraft main power bus. The switching frequency is controlled by a spacecraft level clock signal. This clock signal synchronises all the SMPSs connected to the main power bus and avoids the generation of conducted EMI harmonics

Onboard Software: The CPU board carries (on the PROMs) the basic software to drive and control the DPU.  The tasks the OBS carries out are schematically shown in Figure 4-6.

Figure 4-6 SPIRE DPU On Board Software Context Diagram

The main functions of the OBS are:

- Acceptance of instrument commands from CDMS;

- Execution of predefined commanding sequences;

- Instrument health/status monitoring;

- Implementation of pre-defined procedures on detection of instrument anomalies: the instrument shall be able to adjust parameters and/or switch operating mode and/or activate subsystem redundancy when an anomaly occurs.

- Science data acquisition and packetisation;

- HK data packetisation;

- Transmission of data (science, HK, events and TC verification) from the instrument to the CDMS;

In addition, the OBS shall provide the following SW oriented functions:

- The ability to load, via TCs, replacement and/or additional SW (patches, tables, command sequences, TBC);

- Self test and SW verification facilities;

- Possibility to load and dump part of DPU/ICU memory;

- Possibility to write and check EEPROM: possibility to inhibit these functions during flight operations.

### Detector Readout and Control Unit (HSDRCU)

The DRC is an electronic unit housed into two boxes: the FCU and the DCU connected between the FPU and the DPU.  The DRCU includes the front-end electronics of the following sub-systems:

• Detectors,

• Fourier Transform Spectrometer mechanism,

• Beam Steering Mirror,

• Cooler,

• Photometer and spectrometer calibration sources,

• Thermometry & Analog Housekeeping,

The DRCU comprises 4 physical sub-units which are:

• the DCU which includes the detector control, biasing, readout and digitization electronics;

• the MCU includes the Mechanisms (FTS+BSM) control electronics;

• the SCU includes the Sub-system Control Electronics; and

• the PSU which takes power from the spacecraft and converts it to the required voltage.

These components are schematically represented in Figure 4-7.

Figure 4-7 - Schematic block diagram of the DRCU including the interfaces to the DPU and the FPU.

#### Detector control unit (DCU)

The DCU acts as the interface between the analogue signals coming from the detectors and the digital DPU. A schematic drawing of the DCU together with the detector train components in the FPU and the interfaces with the DPU and FCU is shown in Figure 4-8

Figure 4-8 - Schematic representation of the DCU including it's interfaces with the FPU, DPU and FCU.

Detector Bias Generators: As described in §4.4.4, the detectors are differentially biased with an AC voltage which is provided by the Bias Generator board. Due to the criticality of this component of the detector read out electronics, there is both a prime and redundant

Figure 4-9 - Back plane connections to the photometer LIA cards.

In order to limit the mechanical properties of this unit it is now divided into two boxes which are respectively the DCU and the FCU including the MCU, the SCU and PSU.

The SCU function is manifold: it is in charge of interfacing FPU sub-systems such as the

cooler heaters, the calibrators and the thermometry sensors plus housekeeping parameters

with the DPU through digital interfaces. This function covers the sensors biasing, the signal

amplification and the digitisation of the Analog parameters. The SCU is also in charge of

providing all the sub-systems with secondary power supply lines from the S/C 28 V interface

by means of the PSU (a set of DC/DC converters). This PSU is associated to a supply current

monitoring capability for safety purpose. As illustrated below the DRCU has electrical

interfaces with:

• the FPU

• the DPU

• the PDU (S/C)

• the EGSE (for shutter ground test only)

The DRCU is composed of two subsystems; the Detector Control Unit (DCU) and the FPU Control Unit (FCU) which are both housed in a single box. Each system has analogue interfaces with the systems in the CVV and digital interfaces with the DPU for the transfer of scientific data, telemetry and commands/acknowledge packets. There is no electrical connection between the DCU and FCU.

A common SMPS (with several outputs) is part of the DRCU box. The secondary output lines of this SMPS supply the sub-systems via distribution electronics whose function it is to monitor the current consumption of each sub-system and enables switching off in case of over current.

Electrical interfaces with the FPU are located in the front cover of each board while interface with DPU are located in the rear cover of the DRCU box : this configuration minimizes wiring length and risk of cross-talk of sensitive analogue signals of the FPU/DRCU interface.

### Detector Control Unit (HSDCU)

The Detector Control Unit controls the Detectors of the photometer and the spectrometer, provides the correct detector bias current, conditions the output from the detectors and converts the analogue signals into digital signals.

This is achieved with three different cards mounted in the DCU, viz. the LIA Card, DAQ+I/F Card and the Bias Card. Each LIA Card takes 32 differential signal inputs from the detectors and preamplifies the positive and negative signals with four OP-400 operational amplifiers. They are then fed into a DG303A CMOS Analogue switch

| Number of channels | 48 96 144 24 48 | PLW BDA PMW BDA PSW BDA SLW BDA SSW BDA |
| --- | --- | --- |
| Signal AC amplitude | 10 mV | Dark condition |
| Signal DC level | 5 mV | JFET VOSmax |
| Common mode offset | 1 V | DC |
| Noise allocation | 7 nVrms/rt(Hz) 0.64pArms/rt(Hz) | 0.05 to 25 Hz |
| Input capacitance | < 100 pF |  |
| Input impedance | > 1 M |  |
| Base band signal bandwidth | 0.03 to 5 Hz 0.03 to 25 Hz 5 Hz | Photometer Spectrometer Thermometry |
| Input noise impedance | > 7k |  |
| Common mode rejection | - 60 dB | 30–300 Hz |
| Interface Type | balanced signal + shield |  |

## RF Filters

The Bolometer detectors in the spectrometer and the photometer are sensitive to all frequencies of electromagnetic radiation. Although the BDA feedhorns act as effective high pass filters, it is essential that emissions from within the FPU unit be minimised to eliminate spurious sources of signal. The harnesses that pass from the WE on the SVM to the FPU can act as receiving antennas. The voltages received by the harnesses outside the cryostat could then possibly retransmit EM radiation within the FPU and corrupt the signals from the BDA.

To eliminate this possible source of  noise on the detector signal, low pass RF filters protect all harnesses passing into the FPU. Three basic types of harnesses pass into the FPU; (i) Bolometer signal wires, (ii) Bolometer bias wires, and (iii) FPU sub-system drive and instrumentation wires (eg. Coil drive current wires, thermistor readout signals etc.). These three types of harnesses can be identified on the SPIRE system diagram, Figure 3-1. The Bolometer signals are filtered within the JFET boxes and will described separately below in §4.3. The Bolometer bias harnesses and the subsystem harnesses are conditioned by sets of filter boxes.

There are six filter boxes for the the subsystems are located on the spectrometer side of the FPU (see Figure 3-5). They have two 37 way MDM socket connectors on the WE side of the box connected to two 37 way MDM plug connectors on the subsystems side of the box. Each pin on WE side of the box is connected to an equivalent pin on the detector side via a X (TBD) pole passive filter. There are therefore 444 individual filters allocated to the subsystems. The filter rejects all frequencies from 500 MHz to 10 GHz at –60 dB. The bode diagram for the filter is shown in Figure 4-10.

Figure 4-10 - Bode plot for the filter units.

Figure 4-11 - Schematic representation of the Filter Box mounting technique showing effective limit of the FPU RF shield.

The filter units are attached to the SPIRE optical bench by a blade support structure as shown in Figure 4-10. As the spectrometer cover is placed over the SPIRE optical bench, the outer cover meets with the filter units. The filter boxes are then screwed tightly to the outer cover with a 30 mm bolt spacing and become a part of the RF tight FPU enclosure. The effective outer extent of the FPU RF tight enclosure is shown on Figure 4-10 by the green dashed line. The required mechanical rigidity of the filter boxes come from the bolts attaching them to the outer cover. The first natural frequency of the filter boxes is above the required 200 Hz.

## JFET units

The functionality of the JFET transistors in the read out electronics of the BDAs is adequately described in §4.4.4 below.  The JFET boxes are described here as a sub-system in themselves.

There are two sets of JFET boxes, one set drives the signals from the photometer detectors and is located

Thermal Design: To attain the noise performance the individual JFET devices need to operate at around 110K. The

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-FTB-R01 | Amplifier noise | Requirement better than 10 nV Hz-1/2 over a bandwidth of 100 to 1400 Hz Goal 7 nV Hz-1/2 |  |
| IRD-FTB-R02 | RF rejection | The RF filters, as fitted in the box and with the correct harness, connectors and back-shells; shall reject all frequencies from 500 MHz to 10 GHz at –60 dB. |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-FTB-R04 | Envelope | The JFET/Filter box shall be no more then 300x100x100 mm (TBD) |  |
| IRD-FTB-R05 | Dissipation | The dissipation of JFET amplifiers shall be heat sunk to the level 2 cryostat stage. The dissipation shall be no more than 33 mW (TBC) average for all operating modes with the level 2 temperature at 11 K. Any change in the system design – i.e. changes in the cryostat specification - will lead to a revision of this requirement. |  |
| IRD-FTB-R06 | Operating temperature range | The JFET amplifiers and RF filters shall be capable of operating in with the temperature of the mounting point of the box in the range 4 to 300-K |  |
| IRD-FTB-R07 | Mechanical Interface | The FTB shall mount directly to the FIRST optical bench. |  |
| IRD-FTB-R08 | Nominal operating temperature | The JFET amplifier and RF filter performance requirements shall be maintained with the temperature of the mounting point of the box within the range 4 to 20 K. |  |
| IRD-FTB-R09 | First natural frequency | The first eigenfrequency of the FTB on its mounts shall be greater than 100 Hz (TBC) with a gaol of > 150 Hz |  |

•4 interface bolt holes at ESA side of rack

•1 interface bolt hole at each side of JFET module

•Rack will have a front panel and back panel

•side panel for shear stiffness

•3 JFET modules per rack

•1 RF-filter module per rack

•JFET module rack will be electrically isolated from Herschel optical bench

•JFET module rack will be thermally strapped to level 2 feed line

•JFET module rack thermally isolated from Herschel optical bench?

## Bolometric Detector Arrays

### Principle of  semiconductor bolometers

The SPIRE detectors are semiconductor bolometers. The basic principles of operation are illustrated in Figure 4-10.

Figure 4-12 Principles of bolometer operation

The bolometer comprises a low heat capacity, the absorber, designed to absorb the incident submillimetre radiation. The absorber, is coupled to a heat sink at a fixed temperature To by a thermal conductance, G. Absorbed radiant power, Q, is thermalised in the absorber resulting in an increase in temperature over the equilibrium value in the absence of illumination. A semiconductor thermometer is attached to the absorber. A bias current, I, is passed throughout the thermometer, and the corresponding voltage across the thermometer is measured. The bias current dissipates electrical power P = I2R, which heats the bolometer to an operating temperature T, slightly higher than To. The main performance parameters for bolometric detectors are:

Responsivity:	S = dV/dQ

Noise Equivalent Power	NEP = a(4kTo2G)1/2 where k is Boltzmann's constant and a is a constant 
of order 1.

Time constant:	( ~ C/G where C is the heat capacity of the bolometer.

For a full account of the theory and practice of semiconductor bolometers, see, for example, Mather (1982), Griffin & Holland (1988), Richards (1998).

### SPIRE bolometer performance requirements

The combination of good sensitivity and speed of response requires low-temperature operation. The ultimate limit to the sensitivity of the SPIRE instrument is determined by the thermal background power from the telescope. The background power incident on a SPIRE detector is typically a few pW , almost all of which is from the telescope. This corresponds to a background photon rate of ~ 1010 photons per second. Note that this is enormously larger than in the case of optical observations. The accurate subtraction of this thermal background is essential if the much fainter astronomical signals are to be measurable. Furthermore, statistical fluctuation in the arrival rate of these background photons creates a fluctuating noise power which represents a fundamental thermodynamic limitation to the sensitivity. For SPIRE, the associated photon-noise limited NEP, NEPph, is a few x 10-17 W Hz-1/2. In order to achieve photon noise-limited performance, the inherent NEP of the detector system must be comparable to or lower than this. The noise of the bolometer itself is a combination of Johnson noise and phonon noise (due to the quantised flow of thermal energy from the bolometer to the heat sink).

The instrument performance models described in Section TBD contain detailed calculations of the background power levels and sensitivity for SPIRE. In addition, the operating modes require time constants of 30 ms for the photometer and 16 ms for the FTS. With current bolometer technology, these sensitivity and time constant requirements can be met by detectors operating at a temperature of around 0.3 K, cooled by a 3He refrigerator.

### SPIRE bolometer design and specifications

SPIRE will use arrays of semiconductor bolometers developed at Caltech/JPL. Figure 4-11 illustrates the basic detector design. The absorber is a "spider-web" of metallised silicon nitride. This appears as a filled planar absorber to the submillimetre radiation which has a wavelength much longer than the grid spacing. The spider-web structure has high mechanical strength and a low filling, factor providing low heat capacity and immunity to glitches that can be caused by ionising radiation. The diameter of the spider-web is typically a few times the wavelength to be detected. The thermometer is a small (20 x 100 x 300 (m) crystal of Neutron Transmutation Doped (NTD) germanium. This material is highly suited for use in low temperature bolometers as its thermal behaviour closely approaches that of an ideal thermal device, the material displays very low 1/f noise, and the manufacturing processes are highly repeatable and reliable. A wide range of NTD materials is available to tailor the impedance of the device to the desired range for the chosen operating temperature. Large arrays of bolometers can now be made, as illustrated by the array wafer shown in Figure 4-11.

Figure 4-13 - Left: individual spider-web bolometer. Right: large-format array wafer used as a 350-(m SPIRE prototype array. The absorbers have a 0.725-mm diameter with a grid spacing of 72.5 (m. The filling factor is 8%. The absorber is suspended by five 5-(m wide, 240-(m long support legs. The thermistors are placed to one side of the absorber and read out with two leads deposited on a single, 18-(m wide support member.

### Bolometer readout electronics

The resistance of the detector at the operating point is typically 5 M. Lower values make it difficult to avoid being preamplifier noise limited while higher values can pose problems with electromagnetic pick-up and sensitivity to microphonic disturbance. The detector noise is typically 10-20 , and the noise of the readout amplifier must be of this order or less. With current transistor technology, this requires the use of silicon JFETs, which must operate at a temperature of around 100 K, and must also be located as close as possible to the detectors. This is why SPIRE employs JFET amplifier modules as part of its cold FPU. The essential features of the readout electronics are shown in Fig. TBD. The bolometer is biased (heated to its optimum operating temperature of around 1.3To) by a sinusoidal current bias at a frequency of ~ 100 Hz, applied via the 10-M load resistors. The bias excitation is much faster than the thermal time constant, so that bias itself does not produce a temperature modulation. This is preferred over DC bias as it up-converts the signal information to the bias frequency, getting well above the 1/f noise knee of the JFET readout amplifiers. With this arrangement, because of the inherently low 1/f noise of the bolometers, the1/f noise knee of the system can be very low (less than 0.1 Hz).

Figure 4-14 Bolometer bias and cold readout circuit

Model U401 Silicon JFETS are used to read out the SPIRE detectors, mounted in groups of 24 on silicon nitride membranes which provide high resonant frequency and low thermal conductance. The contract traces and JFET source resistors are lithographed on the membranes. The JFET modules for the photometer and spectrometer are mounted in separate enclosures (illustrated in Figure 4-13) on either side of the FPU.

Figure 4-15 SPIRE JFET module design

### Feedhorns and bolometer cavities

The size of a composite bolometer (typically mm diameter) is small compared to the telescope diffraction spot (in the case of SPIRE, with f/5 final optics, the spot size is 2.44(F = 12.2 m at 500 m) and the first-pass absorption efficiency of the absorber is only ~ 50%. In order to couple the telescope beam onto the detectors in the array, conical feedhorns are used, with a short section of waveguide at the end of the horn to feed the radiation into a cavity containing the bolometer, as shown schematically in Figure 4-14. The waveguide acts as a low-pass filter as it does not propagate radiation of free space wavelength > TBDa where a is the radius. For high efficiency, the absorber is located in the centre of the cylindrical cavity with l/4 spacing between the front of the cavity and a reflecting back-short at the back. The feedhorns are packed in a hexagonal arrangement in the focal plane to fit as many as possible into the area available.

Figure 4-16 - Feedhorn and cavity design

Maximum feedhorn aperture efficiency of 70-80% is achieved for a horn diameter close to 2F( (Murphy et al, TBD), corresponding to a beam spacing on the sky ~ 2(/D, where D is the telescope diameter. The horn restricts the detector field of view, giving a tapered (near Gaussian) illumination of the telescope primary mirror (with an edge taper of approximately 8 dB in the case of SPIRE). Whilst the horns are close-packed in the focal plane, their beams on the sky do not fully sample the image unless the horn diameter is = 0.5F(. Several separate telescope pointings are therefore needed to create a fully-sampled image. For the 2F( horns 16 pointings are required in principle, as illustrated in Figure 4-15. In the case of SPIRE, the step size is dictated by the shortest wavelength channel (250 m) and the number of steps is dictated by the longest wavelength channel (500 m) so that a 64 point jiggle map is needed to achieve simultaneous full spatial sampling in all photometer bands.

Figure 4-17 A 16-point “jiggle pattern” is needed to achieve a fully sampled map with 2F( feedhorns (for hexagonal packing the jiggle pattern is slightly different but 16 steps are still required).

The main advantages of feedhorn arrays are:

- maximum efficiency for detection of a point source with known position;

- well understood horn properties, allowing good control of the beam and reliable design;

- good stray light rejection - the bolometer field of view is restricted to the telescope;

- good rejection of EM interference - the horn plus integrating cavity act as a Faraday enclosure;

- minimum number of detectors for a given total field size.

The main disadvantages are that the observing modes are complicated (jiggling or scanning) and that the efficiency for mapping is a less than the ideal (Griffin, 2000). This means that the full collecting area of the telescope is not used with maximum efficiency.

In order to collect the radiation from a point source with good efficiency, the detector through put must be of the order of (2. The diffraction-limited beamsize of a telescope of diameter D is approximately ( ( 1.2(/D. The corresponding solid angle on the sky is (sky = (((/2)2. The throughput (the product of area and solid angle, which is conserved in an ideal optical system) is therefore

Atel(sky ( [(D2/4][((1.2(/2D)2] ( (2.

The SPIRE photometer uses single-mode feedhorns - the circular waveguide section allows propagation of only one mode (TE-TBD). The corresponding throughput (defined as the area-solid angle product) is exactly (2.

### Bolometer array thermal-mechanical design

The design of the SPIRE array units (called Bolometer Detector Arrays, BDA's) is illustrated in Figure 4-16. Each BDA unit is split into a 2-K and a 300-mK section. The 300-mK section contains the feedhorns, waveguides, bolometer array wafer and the supporting structure. The 2-K detector structure forms the mechanical interface between the 300-mK structure and the photometer or spectrometer 2-K enclosures. The enclosures are connected via thermal straps to the 1.7-K stage of the Herschel cryostat. The heat load from the 2-K structure to the 300-mK structure of the BDA is minimised by suspending the 300-mK section from the 2-K structure with two pre-tensioned 3000-denier Kevlar cords. Kevlar has high mechanical strength in tension and a very low thermal conductivity at this temperature, and forms gives good thermal isolation between the assemblies. The estimated heat load is less than TBD-(W per array. The high mechanical rigidity of the Kevlar cord yields a high resonant frequency (> 200Hz).

Figure 4-18  Bolometer Detector Array module

Table 4-1 lists the main design and performance parameters for the SPIRE detectors (update from SSSD):

Table 4-1 main design and performance parameters for the SPIRE detectors

| Yield | > 90% |
| --- | --- |
| Operating Resistance | ~ 5 M( |
| Time constant | < 30 ms (photometer) < 16 ms (spectrometer) |
| Thermal conductance at To | ~ pJ K-1 |
| Time constant | 11 ms |
| Heat capacity at To | 1 pJ K-1 |
| Responsivity | 0.4 GV W-1 |
| Detector noise | 14 nV Hz-1/2 |
| Efficiency of bolometer-cavity-feedhorn combination | > 0.6 |
| BDA total mass | 0.6 kg |
| BDA load on 3He cooler | < 3 (W (goal 1.6) |

## Mirrors

The basic design of the mirrors is the same as the one used for ISO-LWS. The mirrors are made from Al-6061. They all have a standard interface with the structure, i.e. an M8 screw and a pin. Each mirror is machined in a single block of aluminium (diamond cutting). The screw part of the attachment exerts pressure only on the shoulder part of the mirror, avoiding deformation of the optical surface. The mount of each mirror is located on the optical bench by means of a pin. This pin ensures that in case of dismounting of the mirror it will be mounted again in the same position. During integration of the mirrors in the SPIRE structure, the mirrors are mounted on brackets which are provided by MSSL. The figure below gives a drawing of the CM5 mirror, which is the heaviest one.

## Filters, beam splitters

Filters are placed at a number of positions within the SPIRE FPU, serving the following functions:

- definition of the spectral passbands;

- minimisation of the thermal loading on the 3He fridge, 2-K, and 4-K stages by rejecting short wavelength thermal energy;

- minimisation of stray light getting to the detectors;

- maximisation of the in-band spectral transmission.

The filters are implemented as combinations of capacitative and inductive grids formed by evaporation of copper on thin (~ 2 (m) mylar or polypropylene substrates (Lee et al., TBD). Several grids are used in each filter. The grids must be maintained at accurate spacings. This is achieved either by the use of annular rings for vacuum-gap filters or solid dielectric spacers for “hot-pressed” filters. The latter are formed by pressing together a stack of grids and spacers while heating the combiation. The result is a solid and robustly self-supporting filter.

The SPIRE filter chains are designed on the pessimistic assumption that the telescope surface is perfect even out to UV wavelengths. For complete blocking out to UV, 3 or 4 blocking filters are needed in the chain

The filters can withstand multiple cycling to low temperature, and are robust with respect to low temperature vibration. They are unaffected by UV and high energy particle irradiation and have a very low outgassing rate. They are based on technology developed and implemented in the ISO Long Wavelength Spectrometer and also used in numerous ground-based submillimeter instruments. The filters are interfaced to the structure by simple 3-point mounts to ensure that flatness of the grids is not affected by mechanical stress on the support rings.

### Photometer filtering scheme

The photometer filtering and dichroic scheme is summarised in Table 4-2. The two dichoics allow the spectral bands to be directed to the three arrays with high efficiency

Table 4-2 Photometer filtering scheme [Table to be updated by Peter Hargrave]

| Component ID | Temp (K) | Location | Type | Edges (cm-1) | Clear aperture (mm) | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| CFIL1 | 11 | TBD | Edge filter | 10% 105 50% 100 90% 95 | 70 x 150 |  |
| CFIL2 | 4 | TBD | Edge filter | 10% 84 50% 80 90% 76 | 70 x 150 |  |
| PFIL3 | 2 | TBD | Edge filter | 10% 74 50% 70 90% 66 | 45 dia. |  |
| PDIC1 | 2 | TBD | Low Pass Dichroic | 10% 40 50% 37 90% 34 | 90 dia. TBC | Transmits long ( |
| PDIC2 | 2 | After PDIC1 | Low Pass Dichroic | 10% 27 50% 25 90% 23 | 90 dia. TBC | Transmits long ( |
| PFIL4-S | 0.3 | At SW array | Low pass edge High pass edge | 53 38 | 40 dia. |  |
| PFIL4-M | 0.3 | At MW array | Low pass edge High pass edge | 36 25.7 | 40 dia. |  |
| PFIL4-L | 0.3 | At LW array | Low pass edge High pass edge | 24 17 | 40 dia. |  |
| PFIL5-S | 0.3 | At SW array | Edge filter | 10% 70 50% 66 90% 62 | 40 dia. |  |
| PFIL5-M | 0.3 | At MW array | Edge filter | 10% 46 50% 43 90% 40 | 40 dia. |  |
| PFIL5-L | 0.3 | At LW array | Edge filter | 10% 27 50% 26 90% 25 | 40 dia. |  |

Figure TBD shows a prototype photometer 350-(m band filter chain transmission profile on linear and logarithmic plots. The optical and UV transmissions of the individual elements have been measured separately and multiplied together to derive the overall transmission. The horizontal lines on the logaritmic plot indicate the out-of-band rejection requirements, which are comfortably met by the filter chain.

Figure 4-19 - Transmission characteristics of prototype 350-(m photometer filter chain

The SPIRE dichroics are implemented as capacitative low-pass edge filters placed at an angle to the beam. To maintain the sharpness of the edge, the angle of incidence is set at < 25o. Figure TBD shows typical transmission and reflection profiles for such a dichroic.

Figure 4-20 - Performance of a prototype 15-cm-1 dichroic (red: reflection; blue: transmission)

### Spectrometer filtering scheme

The spectrometer filtering and beam-dividing scheme is summarised in Table 4-3. The first two filters are shared with the photometer, and the remaining ones are placed over the arrays. The beam dividers are designed to have near 50% transmitivity (T) and reflectivity (R) over the FTS spectral range. The beam divider efficiencies as viewed by the two detector arrays are 4RT and R2 + T2, both of which should be close to 100%.

Table 4-3 Spectrometer filtering scheme [Table to be updated by Peter Hargrave]

| Component ID | Temp (K) | Location | Type | Edges (cm-1) | Clear aperture (mm) | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| CFIL1 | 4 | TBD | Edge filter | 10% 105 50% 100 90% 95 | 70 x 150 | Shared with photometer |
| CFIL2 | 4 | TBD | Edge filter | 10% 84 50% 80 90% 76 | 70 x 150 | Shared with photometer |
| SBDI1 | 4 | TBD | Beam divider | 15 to 60 > 90% 4RT | 36 dia. |  |
| SBDI2 | 4 | TBD | Beam divider | 15 to 60 > 90% 4RT | 36 dia |  |
| SFIL3-A SFIL3-B | 2 | TBD | Edge filter | 10% 74 50% 70 90% 66 | 40 dia. |  |
| SFIL4-S | 0.3 | At SW array | Low pass edge High pass edge | 50 33 | 15 |  |
| SFIL4-L | 0.3 | At LW array | Low pass edge High pass edge | 33 None | 15 |  |
| SFIL5-S | 0.3 | At SW array | Edge | 10% 70 50% 66 90% 62 | 15 |  |
| SFIL5-L | 0.3 | At LW array | Edge | 10% 46 50% 43 90% 40 | 15 |  |

Figure TBD shows the measured efficiency of a prototype beam divider with the SPIRE wavelength range indicated by the shaded region.

Figure 4-21 - Measured efficiency of FTS beam divider prototype

## Internal calibrators

### Photometer calibrator (PCAL)

The purpose of the photometer calibrator is to provide a repeatable signal for monitoring of detector health and responsivity for ground testing and in-flight operation. It is not an absolute calibrator, but may be useful as part of the overall calibration scheme. The baseline design consists of a thermal source inside an integrating cavity, the body of which is at 4 K. The cavity will have a light pipe output with a 1-mm diameter aperture. PCAL will be located behind the beam steering mirror (M4) at an image of the system pupil (telescope secondary mirror). The fraction of M4 area obscured is 0.2%, and does not result in any loss of signal as the obscured central area is within the region of the pupil obscured by the holw in the primary mirror. . The PCAL unit is located in within the Beam Steering Mirror housing, and shares the BSM wiring harness.

Because all of the detectos in the photometer arrays view the pupil with near equal efficiency, PCAL produces a very uniform illumination over the arrays (this is an advantage but not a requirement). It is envisaged that the calibrator will be operated in flight at regular but not frequent intervals (once per hour or more). Operation of PCAL requires the BSM to be switched off and the telescope pointing to be fixed, so that there are no sources of detector power modulation except the calibrator itself. A predetermined current excitation sequence will be applied over a period of ~ 10 seconds, and the corresponding detector signal measured.

The requirement for the brightness of PCAL is that it provide the equivalent of a unit emissivity black body temperature of 40 K or more over its 1-mm emitting area. PCAL occupies (1/30)2 of the 30-mm diameter pupil area. The power level that it produces at thte detectors can be compared to that due to the telescope (80-K 4% emissivity) by the following simple calculation. With a temperature of 40 K and unit emissivity, it will produce a power at the detector of (40/80)(1/0.04)(1/30)2 = 1.4% of the telescope background power. The latter is typically 3 pW, so the signal level provided by PCAL is around 4 x 10-14 W. With a detector NEP of a few x 10-17 W Hz-1/2, this provides a very large instantaneous S/N.

The TBD ms time constant requirement for PCAL is that it be comparable to or faster than the photometer detectors. A power dissipation requirement of < 2 mW has been adopted to ensure that PCAL operation makes a negligible contribution to the average load of the SPIRE FPU on the Herschel helium tank, and that local heating of the environment will not result when it is operated.

The baseline PCAL device is essentially an inverse bolometer, developed by J. Beeman of UC Berkeley for use in the SIRTF MIPS instrument. Figure TBD shows a photograph of a SPIRE PCAL prototype.

The device is similar in principle to the one used in the ISO Long Wavelength Spectrometer (but with lower power dissipation). A thin 1 x 1 mm dielectric substrate is coated with a metallic film and suspended from its 4-K copper housing by thin nylon threads. Electrical contacts allow a current to be passed through the metal film, heating it up to a temperature of 40 K or more. A photographs of a prototype device is shown in Figure TBD. A typical resistance is 300 (, requiring ~ 2.5 mA drive current for 2 mW power. There is a 1% stability requirement on the constant-current drive, which translates to a comparable stability for the radiant power output.

Figure 4-22 Photograph of prototype PCAL device

### Spectrometer calibrator (SCAL)

The signal output from one of the FTS detectors represents the difference between the spectra presented at the two input ports. When the FTS is at its zero path difference position, all frequencies are in phases and the signal is at its maximum. The astronomical signals will invariably be much smaller than the thermal background from the telescope. To minimise the required dynamic range for detector signal measurement, it is desirable to replicate the telescope spectrum in the second input port. This is the purpose of SCAL, which is designed to null the telescope emission to within 20% over the 200-400 (m range by minimising its spectrum and brightness in the second input port of the FTS. It is located at the second input port to the FTS, at an image of the telescope pupil (diameter = 30 mm). The telescope is assumed to be at 80 K and to have an overall emissivity of 4%. It is assumed that the overall emissivity of the system is uncertain by a factor of two, and that the actual value will not be known before launch, and SCAL must therefore have sufficient adjustability to accommodate this uncertainty. The baseline SCAL design is shown in Figure TBD. SCAL incorporates two active elements:

- a device similar to PCAL, designed to operate at a high temperature, filling a small fraction of the pupil area, which allows the responsivity of the spectrometer detectors to be monitored in the same way as for the photometer;

- a large area heated plate to provide uniform illumination of the pupil and to replicate the dilute 80-K telescope spectrum.

The heated plate is supported on a thermally isolating tripod using Torlon struts. Redundant 4-wire Cernox thermometers will be used to monitor the temperature of the plate. The unit has a mass of < TBD gm and an allowed volume envelope of 50 50 70 mm. It interfaces directly to the 4-K structure.

Figure 4-23 - Schematic diagram of SCAL. The diameter of the heated metal plate is appro. 30 mm.

In order to replicate the spectral shape of the telescope emission, SCAL must operate in the Rayleigh-Jeans part of the black body spectrum even at the longest wavelength covered by the FTS. This requires a temperature in excess of ~ 40 K. For a source of this temperature, filling the pupil, the emissivity must be considerably less than unity (approximately 8% if the telescope is at 80 K and has 4% emissivity). Two options are under consideration for achieving the desired SCAL heated plate emissivity:

- a neutral density filter may be placed in front of the device;

- the plate may be coated with a suitable material to achieve the desired emissivity.

SCAL is required to have stable emission over the timescale of an FTS observation. The radiant output is required to be stable to within 1% over a period of > 1 hr. A long time constant is therefore desirable. A PID controller, implemented in software is baselined to control the SCAL temperature, and the device shall be designed to have a long intrinsic time constant. However, the cool-down time should not be excessive to prevent emission persisting after the FTS has been switched off (this could, for instance, create stray light problems for photometer observations ). SCAL will be designed to have a cool-down time of ~ 15 minutes. The required SCAL power dissipation when switched on is < 5 mW with a goal of 2 mW.

Thermal testing of SCAL prototypes using a 2-mm thick aluminium plate indicate that an operating teperature of 80 K for the heated plate is achievable with power input of ~ 2 mW, comfortably within the requirement. A modelling and prototyping programme is continuing to optimise the device time constant and to control accurately the emissivity of the plate.

Note that failure of SCAL does not result in loss of the FTS, but the performance for low resolution spectrophotometery would be degraded.

Further details in of the PCAL and SCAL device may be found in the SPIRE Calibrators SSSD.

## Beam Steering Mechanism (BSM)

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-BSMP-R01 | Maximum throw in chop axis | The BSM shall move the imaged field of view of the detectors by a maximum of 2 arcmin on the sky in the Y axis of the satellite |  |
| IRD-BSMP-R02 | Maximum throw in jiggle axis | The BSM shall move the imaged field of view of the detectors by a maximum of 30 arcsec (TBC) in the Z axis of the satellite |  |
| IRD-BSMP-R03 | Minimum step in both axis | The minimum step size in either chop or jiggle axes shall be 2 arcsec |  |
| IRD-BSMP-R04 | Frequency of chop | The chop frequency in the chop axis shall be continuously variable or selectable in 16 steps from 0 to 2 Hz for nominal operation and power dissipation. The chop frequency should be capable of reaching 5 Hz with increased power dissipation and settling time. |  |
| IRD-BSMP-R05 | Frequency of jiggle | Goal of 1 Hz. Requirement of at least 0.5 Hz |  |
| IRD-BSMP-R06 | Holding Position | The BSM shall be capable of moving to and holding indefinitely at any commanded position within its range of movement |  |
| IRD-BSMP-R07 | Stability | The angle on the sky must not vary by more than 0.2 arcsec r.m.s (TBC) over 60 sec at the commanded mirror position in the frequency range 0.03 to 25 Hz. |  |
| IRD-BSMP-R08 | Repeatability in successive positions | 0.2 arcsec |  |
| IRD-BSMP-R09 | Position Measurement | The absolute knowledge of the mirror position shall be equivalent to less then 0.1 arcsec (TBC). |  |
| IRD-BSMP-R10 | Duty Cycle | The mirror shall settle to within 1 arcsec of its commanded position in less than 25 milliseconds in the chop axis. (90% duty cycle for 2 Hz) Goal is 50 milliseconds for the jiggle axis equivalent to 90% duty cycle. |  |

The BSM is designed to allow the image of the sky to be positioned and moved across the detector arrays in a controlled manner. Its specifications are fully described in the BSM SSSD.

The BSM is mounted on the Photometer side of the Spire Optical Bench, with the moveable mirror coincident with an image of the system pupil stop, the Herschel secondary mirror. The BSM must be able to steer the FoV in two independently controlled orthogonal axes. Rotations of the mirror around the Y-axis of the satellite represent “chopping” motions of the FoV while rotations about the Z-axis represent “jiggling” motions. Various observation modes of the telescope call for simultaneous rotations in both axes.

The key specifications of the BSM are summarised in Table TBD.

|  | Chop stage | Jiggle Stage |
| --- | --- | --- |
| On-sky rotation | ( 2 arcminutes | ( 30 arcseconds |
| Mirror rotation | ( 2( 24’ | ( 0( 36’ |
| Frequency | 2 Hz (design), 5 Hz (goal) | 0.5 Hz (design), 1 Hz (goal) |
| Power dissipation |  |  |
| Mass |  |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-BSMP-R11 | Volume envelope | The BSM shall fit within a volume of 130x130x30 mm (TBC) not including its bracket. |  |
| IRD-BSMP-R12 | Operating temperature | Nominal operating 4 K. The mechanism shall be capable of operating in a temperature range of 4-300 K |  |
| IRD-BSMP-R13 | Thermal isolation | The beam steering mirror structure or mirror temperature shall rise by no more than 1 K (TBC) from the nominal temperature of the surrounding structure after one hour operation in any mode. |  |
| IRD-BSMP-R14 | Cold power dissipation | The power dissipation at 4 K shall be no more than 4 mW (TBC) in any operating mode. |  |
| IRD-BSMP-R15 | Warm power dissipation | The power dissipation in the warm electronics shall be no more than TBD W when chopping at 2 Hz in any operating mode. |  |

The BSM is composed of the following components:

- Mechanism (BSMm)

- Mirror and Photometer calibration source (P-Cal)

- Flex-pivots

- Gimbal frame

- Magnetic actuators

- Position sensor

- Stray light baffle

- Thermometer

- Structural interface to the SPIRE Optical Bench (BSMs)

- Warm electronics (BSMe)

- Launch-lock device

Figure 4-24 Beam Steering Mechanism: A. P-Cal, B. Chop stage motors, C. Jiggle Stage Motors, D. Chop Stage Flexure Pivots, E. Gimbal frame, F. Mirror surface, G. Chop Stage Magneto-resistive position sensor, H. BSM Structure, I. Electrical connectors. Not shown; stray light baffle.

Mirror: The mirror has a diameter of 32 mm and is machined from a monolithic block of Al-6160. The front reflecting surface is machined flat to within 1 (m rms and polished to within 10 nm rms. The mirror has a reflectivity in excess of 99% and an emissivity of less than 1% in the wavelength range 200-670 (m. Perpendicular stiffening ribs machined into the rear surface of the mirror minimise the mass while retaining sufficient in plane stiffness. There is a 1.5-mm diameter (TBC) hole at the centre of the of the mirror, through which PCAL, the photometer calibration source (see section TBD), radiates when operated. This hole is conjugate with the central obscuration of the Herschel telescope. The PCAL source is mounted behind the BSM mirror with a short length of light-pipe coupling the radiation to the aperture at the centre of the mirror.

Flexure Pivots: The Flexural Pivots allow rotation of the mirror while constraining translational degrees of freedom. Lucas 5010-800 Flexure Pivots are used for the chop axis while Lucas 5010-600 Flexure Pivots are used for the jiggle axis. These pivots have high radial stiffness, which virtually eliminates unwanted rotation of the mirror due to bearing flexibility. They do not have minimal static friction to overcome which would greatly complicate the design of the control system and possibly increase the power demand on the actuator. The torsional spring rate of these pivots is 5.02 Nm/( for the 5010-600 and 0.631 Nm/( for the 5010-800. This restoring force ensures that the mirror returns to the neutral position in the event of power failure to the BSM.

Gimbal Frame: The gimbal frame is fabricated from Al-6061. It houses the

Balancing Mass

Moment of inertia

Imbalance forces

Magnetic Actuators: The magnetic actuators are located at the edge of the . . . They consume X mW TBD. They produce a torque of X Nm.

Response/rise time

Diagram of specified waveform

Chopping and jiggling frequency

Position sensors: The position sensors are magneto-resistive devices. As the position of the gimbal frame moves relative to the position of the sensor, the resistance of the element changes in proportion to the displacement. The sensor is connected in a Wheatstone bridge configuration driven by a constant current source. Hence, the voltage measured across the bridge is proportional to the displacement of the mirror.

Stray light baffle: Parasitic losses in the actuators, electronics and wiring may cause the temperature of some of the items mounted on the BSM to rise above the 4 K temperature of the Level-1 enclosure. Any radiation emitted from these items must be shielded from the detector arrays. This is achieved through a baffle made from 0.25-mm thick (TBC) Al-6160 (TBC) placed over the components and around the mirror. The baffle fits tightly around the mirror allowing full movement without interfering with the primary beam.

Thermometer: The temperature of the BSM is monitored by . . . .

Structural interface to SPIRE Optical Bench: The mirror needs to be accurately mounted on the SOB.

Mounting accuracy and stability

Vibration

Electrical connector: A single TBD-pin connector accommodates all electrical connections to the BSM and the PCAL device.

Launch-lock device: During the launch (and operation) of Herschel, failure the BSM Flexure Pivots could result in the mirror being fixed in an undesigned off-axis position. The Spectrometer functionality is particularly susceptible to this failure mode - if the BSM gets stuck at one extreme of its angular range, then the field of view of the FTS detectors could be projected outside the SPIRE field of view resulting in loss of the spectrometer. To prevent this from happening, a launch lock device is incorporated into the BSM. This device is a solenoid-driven locking device that reduces the maximum travel range of the Chop and Jiggle axes. During launch or in the event of failure of one or more of the Flexure Pivots, the launch lock solenoids can be energised, constraining the mirror to point ± 1”(TBC) in each axis of the telescope boresight. The state of the launch lock solenoid is monitored by a prime and redundant micro-switch.

Control system: The BSM has a prime and redundant electrical interface with the FCU to provide the necessary feedback control of the mirror pointing angle. The position of the mirror is controlled via a PID scheme which is part of the OBS of the DPU. The control scheme treats each rotational axis of the mirror independently of the other. More detail on the hardware implementation of the actuator control and power system is described in Section TBD.

## Spectrometer Mechanism

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-SMEC-R01 | Linear Travel | Assumed folding factor of 4 for baseline design and single sided interferograms with short travel beyond zero path difference for phase correction. Total OPD required 14 cm. Maximum mirror travel required (wrt ZPD position): – 0.32 to +3.2 cm |  |
| IRD-SMEC-R02 | Minimum movement sampling interval | Short wavelength band minimum measurement interval of 5 m is required (equivalent to 20 m OPD) For long wavelength band the requirement is 7.5 m (equivalent to 30 m OPD) |  |
| IRD-SMEC-R03 | Sampling step control | The measurement interval must be variable between 5 and 25 m. |  |
| IRD-SMEC-R04 | Scan length | The system shall be capable of starting and stopping a scan at any position within the required scan range |  |
| IRD-SMEC-R05 | Dead-time | A goal is to have a dead-time of no more than 10% per scan when taking data at resolution of 0.4 cm-1 |  |
| IRD-SMEC-R06 | Mirror velocity | For assumed detector response of 20 Hz the maximum required rate of change of the OPD is 0.4 cm s-1. Required max. mirror velocity 0.1 cm s-1. A capability to have mirror velocity of 0.2 cm s-1 is desirable and is set as a goal. |  |
| IRD-SMEC-R07 | Velocity control | The mirror velocity should be selectable from 0.02 to 0.1 cm s-1 – or 0.2 cm s-1 if the goal performance is achieved. |  |
| IRD-SMEC-R08 | Velocity stability | The mirror velocity shall be within 0.001 cm/s r.m.s. within a band width of 0.03 to 25 Hz over the entire scan range. The velocity from scan to scan shall not vary by more than 1% over a period of 24 hours under nominal operating conditions. |  |
| IRD-SMEC-R09 | Position measurement | Required OPD position accuracy is 1/50 of the smallest step size. Simulation confirms that this adds minimal system noise to the resultant interferogram. Required mirror position measurement accuracy 0.1 m over +- 0.32 scan range and 0.3 m thereafter. |  |
| IRD-SMEC-R10 | Sampling frequency | The position is sampled at the frequency required for the short wavelength array – i.e.(mirror velocity)/(measurement step size for short wavelength array) |  |
| Requirement ID | Description | Value | Source |
| IRD-SMEC-R11 | Maximum thermal load onto 4 K during cold operation – mechanism and cold position measurement system. | Under zero-g maximum TBD mW Under 1-g maximum TBD mW |  |
| IRD-SMEC-R12 | Maximum envelope | TBD |  |
| IRD-SMEC-R13 | Thermometers | At least one thermometer shall be provided on the FTS mechanism. The temperature range of the thermometer shall be 2 to 20 K (TBC). The absolute temperature measurement accuracy shall be 5% (TBC) with a resolution of TBD mK. |  |
| IRD-SMEC-R14 | Ground Operation | The mechanism and position measurement system must be capable of full operation on the ground when the instrument is in its normal orientation in the test facility cryostat. |  |

The spectrometer mirror MEChanism subsystem (SMEC) controls the movement of the rooftop mirrors inside the SPIRE spectrometer. The critical performances of SMEC [AD1] are the mirror velocity and its stability, the mirror movement around its travel axis and the required accuracy of the mirror position measurements. The SMEC is made of 3 main parts :

The cryogenic mechanism (SMECm)

The SMECm is the mechanism which moves the mirrors, The mirrors are considered a part of SMECm. Basically, the movement of the mirrors is a translation only, obtained through the action of a linear actuator. The position of the mirrors along their travel is measured by an incremental optical sensor coupled to a rule. The mechanical design is based on the GSFC design for a balloon project. A base plate supports the mechanism and is mounted on the SPIRE Optical Bench (Structure). On the baseplate are mounted the fixed part of the actuator and the optical encoder. A moving plate supports the rooftop mirrors and the rule for the optical encoder. The base plate and the mirror moving plate are linked by four "legs". Each leg has two arms, one arm linking the base plate and the intermediate moving plate, the other linking the intermediate moving plate and the mirror moving plate. The articulations at both extremities of the legs use flex pivots. The stiffness of the articulation is very low to keep the actuator power consumption within the specified limits. Consequently, for SPIRE, a launch latch item is added to allow the mechanism to sustain the launch vibrations without damage. The latch is placed between the baseplate and the mirror moving plate.

The preamplifier (SMECp)

Due to the loss of current on the output signals of the optical encoders when they are cooled down to 4K, a preamplifier is necessary (TBC). The electronic components (JFETs, etc..) are implemented on a card integrated in a separate box on the SPIRE structure. The temperature of the components is set around 100K.

The Mechanisms Control Unit (MCU)

The SMEC warm electronic comprises the SMEC analog electronic i.e. the power amplifier for the actuator, the position acquisition electronics and interface unit and the limit position sensors conditioning. The main and redundant circuitry are implemented on one board. This board is a part of the Mechanism Control Unit (MCU) which comprises the SMEC board, the BSM board and the Multi Axis Controller (MAC) board which houses the interface electronics (FPGA) and the mechanism control electronics (DSP based). The MAC board takes charge of the commands sent by SPIRE Digital Processing Unit (DPU) via the Digital Read-out Control Unit (DRCU), controls the movement of the mechanisms via the relevant analog board, delivers the position measurements and transmits the housekeeping data to SPIRE DPU via the DRCU. The power is provided by the DRCU in which the MCU is integrated.

The requirements on the mirror mechanism are that it does not shift or tilt the beam during a scan of –0.3 to 3.5 cm and that the position at any point in the scan range can be reconstructed with an accuracy of 0.1 m.  The latter is set by the requirement that the achieved signal to noise in the spectrum is not dominated by position jitter, at least in the case of a compensated signal (see discussion in reference 4 and section 7).  A prototype mechanism has been constructed at Goddard Spaceflight Center (GSFC) that is based on a double parallelogram system with a frictionless gear to ensure that the two parallelogram structures move together.  Figures 8 shows the side and end views of the prototype mechanism with a commercial actuator fitted and a dummy corner cube mirror.

### Position encoder

For the position measurement system, a Heidenhain LIP interferential linear encoder11 is preferred over a linear variable differential transformer (LVDT) transducer for its better accuracy.  Cryogenic tests indicate that the Heidenhain encoder may be used at liquid helium temperatures with only minor modifications to the optical head.  The optical encoder will have a resolution of 0.1 m or better and gives a “pulse” as each position is passed rather than an absolute distance measurement.  The encoder will be fitted to the outside of the mirror carriage.  Measuring the position at a distance from the signal beam imposes stringent requirements on the “nodding” movement of the mirror mechanism, possibly stricter than those imposed by the optical requirements of the spectrometer.  The actual value of the requirement will depend on the final configuration of the flight optical encoder and the final carriage design.

### Control  System and Readout

The baseline design of the spectrometer is that it should be operated in “rapid scan” mode with a constant mirror velocity and the time at which each position is passed recorded.  The mirror movement will be controlled by a digital feed back loop that uses the time between successive pulses from the optical encoder to control the speed of the mirrors.  In order that the mirror velocity jitter in the signal detection band (essentially 0.03-25 Hz) does not affect the achieved signal to noise, the feedback loop will be operated at about 2 kHz.  The digital feedback loop will be implemented using a digital signal processor, most likely the TEMIC TSC21020 type that is already baselined for the instrument on board computer.  The detectors could be read out in one of two methods given the encoder type we plan to implement:

Position sampling with synchronised detector readout.  In this scenario the mirror position pulse is used to trigger the readout of the detectors.  The time taken to readout the detector array will need to include any possible uncertainty in the time interval between positions due to the uncertainty in the mirror velocity.  This method has the advantage that each frame will be recorded at a known mirror position within the uncertainty of the mirror velocity from one position sample to the next.  This means the interferograms can be safely co-added with no need to interpolate the data onto a common position grid.  The disadvantage of this method is that it requires both a synchronisation pulse from the mirror drive electronics to the detector read-out electronics (and therefore makes the electronics more complicated) and a higher readout rate for the detectors.

Time sampling.  Here the time of each pulse from the position sensor is recorded and the detectors are independently sampled; with perhaps a start pulse at the start of the scan to synchronise the readout.  The advantage of this scenario is that it is simpler to implement in the electronics and will be less prone to operational difficulties associated with timing between the position readout and detector readout.  The disadvantage is that, because of the uncertainty in the mirror velocity and any hysteresis or scan to scan variability in the positioning of the mirror, the detector readouts will have to be interpolated onto a common position grid before co-addition of the interferograms.  In practical terms this will probably prevent on board co-addition of the interferograms and the mirror scan speed will therefore have to be adjusted to allow the data sampling rate to fit within the 100 kbit/second telemetry bandwidth available for the FIRST instruments.

After an assessment of the implementation of the two methods in the electronics system for the SPIRE instrument, it was decided to adopt the time sampling method as the default operating mode.  The impact of this on the performance of the instrument has been simulated and is reported in section 7.

## FPU structure

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-STRC-R01 | Alignment of the instrument w.r.t. the FIRST optical axis | The SPIRE common structure shall allow the alignment of the instrument and telescope optical axes to within +-2.5 (TBC) mm lateral and +- TBD arcmin rotational about any axis. |  |
| IRD-STRC-R02 | Attenuation of RF by 4-K Common Structure covers | The covers as fitted on the instrument will attenuate all frequencies lower than 8 GHz by TBD dB. (Needs confirmation May 2000) |  |
| IRD-STRC-R03 | Items requiring support from the 4-K Common Structure | Photometer and common sub-systems Photometer 4-K optics Photometer filters 4-K Thermal Strap 3He Cooler 4-K Baffles All sub-system harnesses BSM Mechanism and structure Shutter mechanism and mount Photometer 2-K enclosure Spectrometer All spectrometer optics Beam splitters Mechanism structure Mechanism motor Calibration source and mount Spectrometer 2-K enclosure |  |
| IRD-STRC-R04 | Optics and associated sub-system alignment | The common structure shall be capable of maintaining the alignment of the photometer and spectrometer optics and associated components (i.e. filters; 2-K enclosures; BSM etc) to within the specifications given in RD7 both at room temperature and during cryogenic operation. |  |
| IRD-STRC-R05 | Surface finish of the 4-K Common Structure cover | The inside and outside of the box shall have a finish with a low emissivity. At least =0.2. Some parts of the structure walls may be blackened as part of the straylight control. |  |
| IRD-STRC-R06 | Pumping port | The total effective pumping conductance of the 4-K enclosure must be greater than or equal to 7.8 l/s (TBC) |  |
| IRD-STRC-R07 | Thermometry | The structure subsystem shall provide thermistors and associated wiring to allow the temperature of critical parts to be monitored during in-flight operations. |  |
| IRD-STRC-R08 | Attenuation of radiation from cryostat environment | Requirement <2x10-5 (TBC) To illustrate this, the requirement is the equivalent of a ~4 mm diameter hole in a total area of the box cover of 1 m2 (TBC) |  |
| IRD-STRC-R09 | First natural frequency of the instrument assembly | The first eigenfrequency of the integrated instrument assembly shall be greater than 100 Hz (TBC) with a goal of greater than 120 Hz |  |
| IRD-STRC-R10 | Instrument mechanical interface | The mechanical interface of the instrument will be directly to the FIRST optical bench and the instrument will be in direct thermal contact at that interface. |  |
| IRD-STRC-R12 | Grounding | All parts of the SPIRE structure shall be electrically connected one to another. Resistance to be no more than 0.1 TBCbetween any two parts of the structure |  |
| IRD-STRC-R13 | Electrical isolation from FIRST | All parts of the SPIRE structure shall be electrically isolated from the FIRST optical bench and cryostat. Resistance to be greater than TBD  |  |
| IRD-STRC-R14 | Thermal isolation | The conductance from the level 2 to level 1 stage is required to be no more than 6 mW (TBC) assuming level 2 is 9 K and level 1 is 4 K. |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-STRP-R01 | Items requiring support | The photometer 2-K structure shall support: Photometer 2-K optics; dichroics and filters Detector array modules; Detector thermal straps |  |
| IRD-STRP-R02 | Optics and filters alignment | The 2-K photometer structure shall be capable of maintaining the alignment of the photometer 2-K optics; filters and dichroics to within the requirements set out in RD7 at room temperature and during cryogenic operation. |  |
| IRD-STRP-R03 | Array module alignment | The 2-K photometer structure shall be capable of maintaining the position of the detector array modules to within the requirements set in RD7 about any axis during cryogenic operation of the instrument. |  |
| IRD-STRP-R04 | Surface finish | The outside of the box shall have a finish with a low emissivity. At least =0.2 The inside of the box shall have a low reflectivity finish on all non-optical surfaces. |  |
| IRD-STRP-R05 | Pumping port | The total effective pumping conductance of the 2-K box must be greater than or equal to 5.6 l/s (TBC) |  |
| IRD-STRP-R06 | Attenuation of radiation from 4-K environment | Requirement 5x10-7; goal is 5x10-8 (TBC) To illustrate this, the requirement is the equivalent of a 0.5 mm diameter hole in a total area of the box cover of 0.5 m2 This, of course, excludes the hole that lets the beam in. |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-STRP-R07 | First natural frequency | The first eigenfrequency of the photometer 2-K structure on its mounts shall be greater than 100 Hz (TBC) with a gaol of > 150 Hz |  |
| IRD-STRP-R09 | Thermal isolation | The conductance of from the 4-K to 2-K stage shall be no more than 0.75 mW (TBC) |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-STRS-R01 | Items requiring support | The spectrometer 2-K structure shall support: Cold Stop filters Fold mirrors |  |
| IRD-STRS-R02 | Optics alignment requirements | The spectrometer 2-K structure shall be capable of maintaining the alignment of the spectrometer 2-K optical components to within the requirements set out in RD7 |  |
| IRD-STRS-R03 | Surface finish | The outside of the box shall have a finish with a low emissivity. At least =0.2 The inside of the box shall have a low reflectivity finish on all non-optical surfaces. |  |
| IRD-STRS-R04 | Pumping port | The total effective pumping conductance of the 2-K box must be greater than or equal to 5.6 l/s (TBC) |  |
| IRD-STRS-R05 | Attenuation of radiation from 4-K environment | Requirement 5x10-7; goal is 5x10-8 (TBC) To illustrate this, the requirement is the equivalent of a 0.5 mm diameter hole in a total area of the box cover of 0.5 m2 This, of course, excludes the holes that let the beams in. |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-STRS-R06 | First natural frequency | The first eigenfrequency of the spectrometer 2-K structure on its mounts shall be greater than 100 Hz (TBC) with a gaol of > 150 Hz |  |
| IRD-STRS-R08 | Thermal isolation | The conductance of from the 4-K to 2-K stage shall be no more than 0.25 mW (TBC) |  |

The structure has three main requirements:

- To mechanically support and positively locate the components of the photometer and the spectrometer;

- To provide the correct thermal isolation between the various temperature stages of the instrument; and

- To provide adequate EM shielding and stray light shielding to the photometer and spectrometer such that the detectors see no spurious sources of EM power.

Requirements on the natural frequecy and 2 are conflicting!

The structure is comprised of three main components: The Spire Optical Bench (SOB), The photometer Cover and the spectrometer Cover. The Spire Optical Bench is machined from a single block of Al-60601 approximately 675mm long (aligned with the Spacecraft Z axis, See Figure 2-2) by 410mm high (aligned in the X axis) and 50 mm thick. All the optical components of the spectrometer are located on the +Y side of the SOB while all the components of the photometer are located on the –Y side. The mass of the SOB is minimised by machining ribbed pockets (nominally 70x70mm) on both sides. The nominal thickness of the ribs is 3mm and the centreline plane of the bench is 3mm (TBC) as shown in Figure 4-23-A.

Figure 4-25 Details of  (A) SOB lightweighting and optical component interfaces and (B) cover joint design.

The optical components are located on the bench by hollow dowel pins and 8-32 UNC screws provides the mechanical attachment. The SOB is manufactured and toleranced in such a way so that the position of the components located on the bench are in the correct position when the bench has undergone the contraction as is cooled from room temperature to its’ nominal operating temperature of 4K.

Figure 4-26 - FPU/HOB interface strut.

The ribbed design of the SOB gives it excellent bending stiffness but it will have a relatively low torsional stiffness. The required torsional stiffness is achieved through the photometer and spectrometer covers. These Covers are created by screwing a series of lightweighted 12mm thick plates together in the manner shown in Figure 4-23-B. The integration plan for the FPU structure sees the covers being loosely assembled with the screws not being fully tightened. The covers are then screwed tightly onto the SOB and the cover screws fully tightened thus forming a monocoque structure that will not be disassembled. Access to the SOB with integrated optical components will be done by removing either the photometer or the spectrometer covers. The remaining cover and optical bench will give the required stiffness to the structure. The two covers and the SOB are machined from Al-6082. After final machining, it is passivated with Aluchrome.

The FPU structure is connected to the Herschel Optical Bench via three feet structures. These feet have to:

- Proved a rigid connection between SPIRE and Herschel so that the first natural frequency of the integrated instrument is higher than 100Hz and prescribed in the IRD.

- Provide adequate thermal isolation between the 10K HOB and the FPU structure so that the

- Allow for contraction

Figure 4-27 - View of the Lakeshore, Cernox SD hermetically sealed RTD thermometers.

Figure 4-28 Interface drawing for the FPU structure within the CVV showing the positions of the photometer and spectrometer JFB, and the 2K strap feed throughs.

| Unit | Location | Acronym | Sensor Type | Temp. Range | Resolution over Range | Prime /Red |
| --- | --- | --- | --- | --- | --- | --- |
| FSFPU | PHOT Level 0 box | T_PL0_1 | CX-1050 | 1 K>10 K | 2 mK | P |
| FSFPU | SPEC Level 0 box | T_SL0_1 | CX-1050 | 1 K>10 K | 2 mK | R |
| FSFPU | SMEC Mechanism | T_FTS_1 | CX-1050 | 3 K>50 K | 10 mK | P |
| FSFPU | SMEC Mechanism | T_FTS_2 | CX-1050 | 3 K>50 K | 10 mK | R |
| FSFPU | SMEC/SOB Interface | T_FTS_3 | CX-1070 | 3 K>300 K | 100 mK | P |
| FSFPU | SPEC Calibrator | T_SCAL_1 | CX-1050 | 10 K>80 K | 5 mK | P |
| FSFPU | SPEC Calibrator | T_SCAL_2 | CX-1050 | 10 K>80 K | 5 mK | P |
| FSFPU | SPEC Calibrator | T_SCAL_3 | CX-1050 | 10K>80K | 5 mK | R |
| FSFPU | Cooler Pump | T_CPMP_1 | CX-1050 | 3 K>100 K | 25 mK | P |
| FSFPU | Cooler Pump | T_CPMP_2 | CX-1050 | 3 K>100 K | 25 mK | R |
| FSFPU | Cooler Evaporator | T_CEV_1 | CX-1030 | 0.2 K>5 K | 1 mK | P |
| FSFPU | Cooler Evaporator | T_CEV_1 | CX-1030 | 0.2 K>5 K | 1 mK | R |
| FSFPU | Cooler Pump heat sw. | T_CPHS_1 | CX-1050 | 1 K>50 K | 10 mK | P |
| FSFPU | Cooler Pump heat sw. | T_CPHS_2 | CX-1050 | 1 K>50 K | 10 mK | R |
| FSFPU | Cooler Evap. Heat sw. | T_CEHS_1 | CX-1050 | 1 K>50 K | 10 mK | P |
| FSFPU | Cooler Evap. Heat sw. | T_CEHS_2 | CX-1050 | 1 K>50 K | 10 mK | R |
| FSFPU | Shutter Vane | T_SHUT_1 | CX-1070 | 3 K > 100 K | 25 mK | N/A |
| FSFPU | BSM Mechanism | T_BSM_1 | CX-1050 | 3 K>20 K | 10 mK | P |
| FSFPU | BSM Mechanism | T_BSM_2 | CX-1050 | 3 K>20 K | 10 mK | R |
| FSFPU | BSM/SOB Interface | T_BSM_3 | CX-1070 | 3 K>300 K | 100 mK | R |
| FSFTBs | SPEC JFET box | T_FTBS_1 | CX-1070 | 3 K>100 K | 25 mK | P |
| FSFTBs | SPEC JFET box | T_FTBS_2 | CX-1070 | 3 K>100 K | 25 mK | R |
| FSFTBp | PHOT Filter box | T_FTBP_1 | CX-1070 | 3 K>100 K | 25 mK | P |
| FSFTBp | PHOT Filter box | T_FTBP_2 | CX-1070 | 3 K>100 K | 25 mK | R |

## Helium-3 Cooler

The cooling of the SPIRE and PACS detectors down to 300 mK will be effected by a helium three sorption cooler. This sub-Kelvin sorption cooler provides a wide range of heat lift capability at temperature below 400 mK. It relies on the capability of porous materials to adsorb or release a gas when cyclically cooled or heated. Using this physical process one can design a compressor/pump which by managing the gas pressure in a closed system, can condense liquid at some appropriate location and then perform an evaporative pumping on the liquid bath to reduce its temperature. Helium sorption refrigerators have no moving parts, are vibrationless and can be designed to be self contained and compact with a high duty cycle efficiency.

As shown on the figure hereafter the cooler is basically made of 6 components designated as a sorption pump SP and an evaporator EV connected via a pumping line to a thermal shunt TS comprising an heat exchanger, two gas gap heat switches HSP and HSE respectively connected to the sorption pump (SP) and evaporator (EV), and finally a support structure SST. SP, EV, TS and the pumping line are assembled to form a unique component which is the actual “heart” of the cooler. This component is held within SST, which provides firm mechanical support (launch environment) while minimising any parasitic conductive load on the cooler (low temperature environment).

Heat switches are then required for operation of the cooler. The two switches are used to control the temperature gradient; during the condensation phase they are set such that the sorption pump can be heated to release the helium gas and such that liquid condensation occurs into the evaporator EV maintained as the coldest point. The liquid is held into EV by capillary attraction inside some porous material : both the surface tension and the vapour pressure provide forces that drive and hold the liquid at the coldest point. Then the switches are set such that the sorption pump is thermally grounded to the heat sink and such that the evaporator is thermally isolated. The sorption pump provides an evaporative pumping on the liquid helium bath which temperature quickly drops to sub-Kelvin temperature.

Cryogenic switches can be based upon different physical mechanisms. Gas gap heat switches have been selected as the preferred design for the present project. Gas gap heat switch utilises concentric copper cylinders separated by a small gap which is filled with or emptied of He gas to achieve the switching action. The thermal separation between the two ends is achieved by a thin-walled tube which also provides the mechanical support. The presence or absence of gas is controlled by a miniature cryogenic adsorption pump that can be temperature regulated. Thus one of their main feature is the absence of any moving part, and consequently operation of the cooler is almost then fully controlled by three heaters.

For operation in a zero-G environment two aspects are addressed: the liquid confinement and the structural strength required for the launch. The confinement within the evaporator is provided by a porous material which hold the liquid by capillary attraction, and a suspension system using Kevlar wires is designed to firmly support the refrigerator during launch while minimising the parasitic heat load on the system.

The sorption cooler will be mounted off a 4 K plate and 1.7 K thermal paths will be provided for the heat switches and thermal shunt for the operation of the cooler. The radiative environment will be 4 K. The recycling time shall be as short as possible, compatible with the warm heat sink heat evacuation capability. This recycling time has no real impact on the design of a cooler operated from a helium bath. It only has an impact on the recycling procedure (timing, control of heat switches, peak dissipation). The sorption cooler should be able to operate in any orientation under 1g. However it is understood that during ground test some specific orientations shall be avoided because of potential convective effect. This constraint only applies during the recycling phase while the pump is heated to above 20 K. A design “plug in” type shall be favored.

- The net heat lift at 290 mK shall 10 µW at a minimum

- Inside the cooler itself the parasitic heat load to the cold tip must be minimised. As a goal

- it shall not exceed 12 µW under the nominal operating conditions (1.7 K / 4 K - AD1).

- It is desirable to be able to vary the detectors temperature up to 320 mK. Electronic control shall be provided to perform this task in the flight electronic

- The temperature of the evaporator cold tip should not drift by more than 0.1 mK/h under active temperature control

- The temperature fluctuations at the evaporator cold tip shall be no more than 10 µK/Hz 0.5 in a frequency band from 0.1-10 Hz

- The hold time shall be no less than 22 hours and shall be based on a finite number of days (a minimum of 46 hours is favoured)

- The recycle time shall be no more than 2 hours

- The energy dissipated per cycle due to cooler operation shall be no more than 860 J. the 4K/1.7K thermal architecture calls for an additional contribution to the average power dissipated - this contribution (not part of the 860 J) shall not be in excess of 0.6 mW

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-COOL-R01 | Temperature at the detectors | Nominal 300 mK |  |
| IRD-COOL-R02 | Operating temperature control | Desirable to be able to vary the temperature of the detectors up to 320 mK and below 300 mK if this is permitted by the temperature drop across the thermal link. The evaporator cold tip temperature can be varied by heating the sorption cooler. Electronic control shall be provided to do this in the flight electronics. |  |
| IRD-COOL-R03 | Temperature drop across thermal link between detectors and evaporator cold tip | Maximum of 25 mK |  |
| IRD-COOL-R04 | Temperature drift | The temperature of the evaporator cold tip should not drift by more than 10 mK/h |  |
| IRD-COOL-R05 | Temperature fluctuations at the evaporator cold tip | No more than 150 nK Hz-1/2 in a frequency band from 0.1-100 Hz. |  |
| IRD-COOL-R06 | System low frequency temperature stability with active temperature control | 150 nK at 0.015 Hz at a maximum power dissipation of 1 W |  |
| IRD-COOL-R07 | Heat lift at evaporator cold tip | Minimum of 10 W at 300 mK |  |
| IRD-COOL-R08 | Hold time | Minimum 46 hours |  |
| IRD-COOL-R09 | Recycle time | Maximum 2 hours |  |

| Requirement ID | Description | Value | Source |
| --- | --- | --- | --- |
| IRD-COOL-R10 | Mechanical interface | Preferred interface is with the instrument 4-K structure |  |
| IRD-COOL-R11 | Thermal Interface with FIRST cryostat | Pumped liquid helium tank at 1.8 K for both sorption pump and evaporator |  |
| IRD-COOL-R12 | Parasitic thermal load onto He bath during cold operation | Maximum 1.2 mW is allowed for the conduction from 4 K to 2 K through the support structure and heat switches. |  |
| IRD-COOL-R13 | Time averaged thermal load onto He bath for 48 hour cycle | Maximum 3 mW (includes 20% margin) |  |
| IRD-COOL-R15 | Maximum envelope | 200x100x100 mm |  |
| IRD-COOL-R17 | Sorption pump heater | The baseline design has a heater resistance of 400  implying a current of up to 20 mA for recycling. It is desirable that this heater resistance is increased so that the allowable resistance of the cryo-harness wiring can, in turn, be increased. The maximum resistance of the heater that can be driven by 28 V is about 5 k |  |
| IRD-COOL-R18 | Thermometers | Thermometers shall be provided on the cooler as necessary to monitor its behaviour and operation. The absolute temperature measurement on the evaporator cold tip shall be 0.5% (<1.5 mK) with a resolution of TBD mK. Thermometers of the same specification shall also be provided on each detector array (q.v.). |  |
| IRD-COOL-R19 | Gas gap heat switches | It is noted that these are a potential single point failure in the instrument operation. Provision of some redundancy (i.e. doubling them up) is desirable but not at the expense of severe limitations on the cooler performance. |  |
| IRD-COOL-R20 | Ground Operation | The cooler must be capable of full operation on the ground, including recycling, when the instrument is in its normal orientation in the test facility. This will be arranged so that the evaporator is below the pump. The cooler must be capable of operating with the instrument rotated to up to 90o about either the S/C Y or Z axes. |  |
| IRD-COOL-R21 | Warm electronics power dissipation | Less than TBD W during cold operation Less than TBD W during recycling |  |

10W heat load to 290 mK

The 3He cooler is required to maintain the temperature of the BDA below the nominal Two chambers, the evaporator and the charcoal sorption pump, are joined by thin-walled stainless steel tubes via the condenser, which is in good thermal contact with the cold heat sink (we will assume a temperature of 2K for the purpose of this description). The refrigerator is charged with 3He gas, and permanently sealed by a crimped tube.

6 litre

48 hours recycle

hard mounted to 4K FPU Optical Bench

Separate straps from Pump and Evaporator to L0 Helium Tank, both via heat switches.

Separate straps necessary to prevent warming of evaporator during cooler recycling.

Evaporator strap has high conductance requirement (>100mW/K @ 2K) in order to limit temperature at end of recycling.

Leak before burst

titanium alloy (Ta6V Eli

3He cooler : 264 mK ultimate T, 20 µW @ 300 mK

Doug to write from Spec doc, presentations, LD papers

## Thermal staps

### 2-K to 4-K

### 300 mK

MJG/AW/PCH

## Harness

There are five sets of harnesses for the SPIRE instrument.

| Name | Prefix | Connection One | Connection Two | Description |
| --- | --- | --- | --- | --- |
| FPU Harnesses | F |  |  |  |
| Cryo-harnesses | C |  |  |  |
| Interconnection Harness | I |  |  |  |
| Warm Interconnect Harness | W |  |  |  |
| Test | T | EGSE | HSDPU, HSFPU | Test harnesses only used during testing of integrated electrical systems |

## Shutter

During the period between final integration of the SPIRE instrument into the Herschel spacecraft and launch (a period of approximately two years), it is desirable to be able to verify that the photometer and spectrometer detectors and the cold and warm read out electronics are operating correctly. To test the optical readout system inside the spacecraft cryostat prior to launch, the background flux of light falling onto the detectors must be similar to the flux seen by the detector when the instrument is imaging astronomical objects during operation. In orbit, this background arises primarily from the self emission of the primary and secondary telescope mirrors which are assumed here to be at 80 K with an overall emissivity of 4%.

Prior to launch, the cryostat must cool the all the Herschel instruments in the focal plane to the nominal cryogenic flight temperature. To do this, a lid must seal the opening at the top of the cryostat vessel. This lid will however be at around 200 K (TBC) due to the relatively inefficient thermal insulation and will provide a very large background flux of photons onto the detectors (see Figure TBD). To prevent this energy from flooding the detectors, a shutter mechanism is employed to prevent the thermal background from the cryostat lid from entering the SPIRE instrument.

Figure 4-29 - Background radiation on the SPIRE detectors in space and on-ground.

When the shutter is closed, the detectors must see the same background flux in the spectrally sensitive band of the instrument as that used to test the detectors in the ground calibration facility (and expected from the telescope in flight). To achieve this, the inside surface of the shutter has a high emissivity coating, and can be heated electrically. Figure 4-28 shows the power absorbed by the photometer detectors as a function of shutter temperature (based on the same assumptions used in Section TBD for the SPRIE sensitivity estimation).

Figure 4-30 - Detector power vs. shutter temperature for the three photometer bands.

The telescope power levels that must be reproduced are nominally 3.9, 3.2 and 2.4 pW at 250, 350 and 500 (m, respectively (see Section TBD). These correspond to shutter temperatures of approximately 17, 14.5, and 12 K. The effective emissivity of the telescope in-orbit is highly uncertain, and the instrument will be tested in the AIV facility over a range of backgrounds. Allowing for a factor of at least two in the power levels to be reproduced, the shutter is therefore designed to operate at up to 25 K. Note that the shutter is not required to reproduce the correct background levels on all arrays at the same time.

With the shutter thus providing a representative background radiation flux, the photometer and spectrometer calibration sources can be used to stimulate the detectors and the resulting signal read out from the read out electronics. These tests can be carried out in the AIV facility and again in the Herschel cryostat, to ensure that the system is operating nominally.

The baseline location of the shutter is shown in Figure TBD. As this shutter is only used in the pre-launch testing of the detectors there is a launch lock device that will prevent the shutter from closing during flight. This launch lock mechanism is a passive device, in that it requires no power to keep the shutter in the locked open position.

The other components of the shutter are described below.

Vane: A metal vane blackened on one side is used to block off the input radiation.

Heater: An electrical heater mounted on the vane of the shutter provides power to heat the vane up to the design temperature.

Temperature sensor: A single Lakeshore Cernox thermistor is mounted on the shutter vane to monitor its temperature.

Actuator: A single rotary stepper motor is used to actuate the shutter vane. It can be driven by a primary set of two windings or alternatively by a separate set of two windings.

Launch latch position indicator: This is a micro switch that indicates the state of the launch lock. The contact on the switch can only be closed when the Shutter has returned to the open position and the launch lock successfully deployed.

Figure 4-31 - Physical layout of shutter within the inlet of the FPU enclosure.

The powering, thermal control and the monitoring of the shutter sub-system are all carried out with the EGSE. When the EGSE is disconnected, the shutter is unable to be operated or controlled. The state of the launch lock indicator must be read out by the spacecraft housekeeping electronics.

# Instrument Operating Modes

The operational modes of the SPIRE instrument are documented in detail in the OMD. These operating modes can be summarised as being the sequence adopted to turn the instrument on, preparing it for scientific observations, making the observations, the recycling of the cryogenic cooler and the abnormal and normal shut down. The implementation of the astronomical observing modes will be briefly described here.

## Spacecraft pointing

It is assumed that SPIRE pointing will be defined with respect to the telescope boresight by two offset positions, one for the centre of the photometer arrays and one for the centre of the spectrometer arrays, as shown below. When the telescope is pointed at a source at the request of SPIRE, it shall be aligned on one of the two positions defined in Fig. TBD below. Any offsetting by the SPIRE BSM or the AOCS shall then be defined with respect to this position.

Figure 5-1 - Definition of SPIRE pointing offsets with respect to the telescope boresight.

## Spacecraft movements during SPIRE observing

The Herschel spacecraft is capable of executing various controlled movements to allow observations of various kinds by the three instruments. These are described in detail in the IID-A. SPIRE will mainly use the following spacecraft movements.

### Nod

The NOD function of the telescope is an operation in which the target source is periodically moved from one instrument chop position to the other chop position by re-pointing the satellite. The pointing direction will change in the direction of the chopper throw - see Fig. TBD

Figure 5-2 - Pointing positions for a telescope NOD function. The circles represent the size of the telescope Airy pattern projected onto the sky. The two nod positions have been offset left and right for clarity.(A1 and B2 are co-aligned).

The purpose of nodding is to subtract out the signal offset that will be present at some level due to the fact that the detector will receive different amounts of ambient background power in the two chop positions. This offset is likely to be larger than low-level signals to be detected, and may also vary significantly with time over the course of an observation if the primary mirror has a time-varying temperature gradient in the chopping direction (Griffin, 1998). Nodding takes out the offset in the following way:

Let 		S = Source signal 	OA = Offset signal for beam A		OB = Offset signal for beam B

Then if we subtract the total signals measured in the two nod positions, we get

(A1 - B1) - (A2 - B2) = [(S + OA) - OB] - [OA - (S + OB)] = 2S.

### Raster

The RASTER Spacecraft Function is a series of fine pointing operations of separated by slews such that the pointing of the telescope axis moves in a raster pattern. Figure TBD shows how the raster pattern will be constructed.

Figure 5-3 - Pointing positions for a telescope RASTER function. The observation is specified in terms of M pointings per line separated by d1 by N lines separated by d2 arcsec.

### Line Scan

In the LINE SCAN Spacecraft Function the satellite is slewed at a constant angular velocity along short parallel lines on the sky. Figure 2-4 shows how the operation is carried out.

Figure 5-4 - The LINE SCAN function of the telescope consists of a number of short slews of length D1 separated by a distance d2. The slews will be carried out in the order shown here

SPIRE observations are implemented by combinations of the following:

Spacecraft Functions: These are the operations that can be carried by the spacecraft to point the telescope, such as line scan; raster; staring etc. They are fully described in the IID-A. Spacecraft Functions also include operations by the CDMS to switch power to the instrument; send commands; collect data etc.

Instrument Functions: These are the operations to be carried out with the instrument such as chopping , jiggling, scanning the FTS mirror, operation of the internal calibrators, etc. Combined with the Spacecraft Functions they fully define how an observation is to be carried out.

Instrument Data Configurations: In addition to specifying how the instrument is to be operated for a given operation, the on-board data processing needs to be specified along with the data to be sampled and the manner in which the detector data is sampled. This will be done be choosing from a number of Data Configurations such as Photometer Full Field Chop, Spectrometer Single Pixel, etc.

Observatory Functions: Combinations of Spacecraft and Instrument Functions and Instrument Data Configurations which, with the appropriate input parameters, allow any Observation to be carried out. 
There are two categories of Observatory Functions: Photometer Observatory Functions (POFs) and Spectrometer Observatory Functions (SOFs).

## Photometer Observatory Functions

| Observation | Observatory Function | Name | Comments |
| --- | --- | --- | --- |
| Point source Photometry | POF1 | Chop without jiggling | Accurate pointing and source position |
|  | POF2 | Seven-point jiggle map | Inaccurate pointing or source position |
| Jiggle Mapping | POF3 | n-point jiggle map | Field mapping |
|  | POF4 | n-point jiggle map with raster | Extended field mapping |
| Scan Mapping | POF5 | Scan map (no chopping) | Large-area mapping |
|  | POF6 | Scan map with chopping | Large area mapping (with 1/f noise) |
| Peak-up | POF7 | Photometer peak-up (TBD) | Determination of pointing offsets |
| Calibrate | POF8 | Photometer calibrate | Responsivity tracking |
| Engineering Modes | POF9 | Special engineering/ commissioning modes |  |

Table TBD: Photometer Observatory Functions

### Point Source Photometry (POF1 and POF2)

For point source photometry, pixel-pixel chopping is used as described in Section 2.2.1.3. The required chop throw is 4F( at 500 (m, corresponding to a 10-mm motion of the beam on the array and to a chop angle on the sky of (chop = 126 arcsec.

Simple pixel-pixel chopping (POF1) requires pointing accuracy sufficiently good that the loss of signal due to the pointing error is acceptable. The signal loss factors for the photometer beams are shown in Fig. TBDa.

Figure 5-5 - (a) Signal loss vs. pointing error for the three SPIRE photometer bands; (b) beam positions for 7-point map with angular offset (.

The required Absolute Position Error (APE) for Herschel (3.7") corresponds to 11%, 6%, and 3% signal loss at 250, 350, and 500 (m respectively. The corresponding figures for the goal AME value (1.5") are 2%, 1%, and 0.5%. For most observations, 11% is not acceptable, but 2% is. Therefore, the required APE is not good enough to allow accurate photometry at 250 (m without peaking up, but the goal is sufficient to allow this.

In order to avoid calibration errors due to satellite pointing errors or inaccurate knowledge of the source co-ordinates, it is envisaged that many point source observations will be carried out by making a 7-point map (POF2). The BSM is used to do to do small map 7-point hexagonal jiggle map with spacing ( arcsec., as shown in Fig. TBD. A suitable value for ( is ~ 6”: this spacing is 1/3 of the beam at 250 (m (so consistent with full sampling), and is almost twice the APE. From such a 7-point map, the total flux of the source can be computed, and the source position can be recovered with an accuracy that depends on the S/N. The central position is made to coincide with one of sets of three overlapping detectors to allow simultaneous optimised observations in the three bands. The chop throw can be set at any desired value within the available range. A value of 126” improves the overall efficiency by allowing the source to be observed all the time in all bands.

### Jiggle Mapping (POF3 and POF4)

Jiggle-map mode is for mapping objects or regions which are extended with respect to the SPIRE beam but smaller than a few arcminutes in size.

The BSM is used to make an n-point jiggle map while chopping with a throw greater than the size of the object to be mapped. The maximum throw is 4 arcminutes (( 2 arcminutes) as illustrated below.

Figure 5-6 Field mapping with maximum chop throw of 4 arcminutes. The detectors in the central square 2 x 2 arcminute portion of a photometer array are deflected by ( 2 arcminutes by the BSM, so that they are alternately chopped from one side of the available field of view to the other.

For full sampling at all wavelengths, n = 64. Considering the simplified case of square-packed horns, the step size must be 0.5(/D = 9” at 250 (m and the number of steps must accommodate the need to cover the distance between two beams at 500 (m: 2(/D = 72”. Eight steps in each orthogonal direction are thus required. The geometry of the juggle pattern is hexagonal for the hexagonally packed feedhorns, but the number of steps required is still 64. To allow flexibility in the use of this mode, permitted values of n shall be 16, 32 and 64. The jiggle positions shall be defined by angles ((Y and ((Z in the Y and Z directions (with respect to the BSM rest position). The sequence in which the jiggle positions are visited is TBD. It is envisaged that a typical time per 64-point jiggle cycle will be 64 seconds, allowing two chop cycles per jiggle position at the maximum chop frequency of 2 Hz.

POF4 involves a sequence of repetitions of POF3 with the telescope performing the steps of a raster scan between the individual jiggle cycles in order to build up a map of a larger area.

### Scan mapping (POF 4 and POF5)

This mode is for mapping a large region of sky by scanning the telescope to provide modulation of the signal. It does not require the beam steering mechanism and is the preferred observing mode for deep extragalactic surveys because it is more efficient and requires simpler data analysis. It also avoids increasing confusion noise as happens with chopped observations if the individual signals in the two chop half-cycles are not both transmitted to the ground (e.g., Oliver, 2001). This will be done in the case of SPIRE, so chopping can be performed if necessary without degrading the S/N in principle.

All SPIRE mechanisms are inactive. The spacecraft operates in line scan mode with the scan direction across the sky at an angle of 14.5o to one of the array axes so that the individual strips of the map are fully sampled (1/2 beam spacing) in the cross-scan direction. The length of each line-scan should be such that the turn-around time of the telescope (here assumed to be 10 seconds) does not constitute a large overhead. Each line should therefore take at least 60 seconds. The scan rate will depend on the 1/f stability of the whole system, but with a maximum limit of 60 arcsec./second, determined by the spacecraft AOCS. The fastest beam-crossing time (for a 250-(m detector) at this scan rate is 18/60 = 300 ms, which is TBD times the required time constant of the detectors. It will be desirable to scan as slowly as possible to minimise telescope turn-around overheads.

In the event of some 1/f noise component resulting in noise fluctuations on the timescale of a scan, the BSM can perform chopping.

### Photometer peak-up (POF7)

This mode is designed to allow SPIRE to peak up the pointing on a sufficiently strong point-like source. As far as the observations are concerned, it is the same as a 7-Point Jiggle Map). It may not be necessary to use this morde in flight. It is not needed for very bright objects (carrying out a small map is quick compared to overheads from slewing etc.). Nor is it practical for very faint objects (poor S/N of the 7-point data would lead to inaccurate offset calculation). It is therefore only likely to be useful for a particular band of source strengths. POF7 may therefore be deleted for simplicity. For the moment, it is included to account for the possibility that peaking up involving autonomous computation of the required pointing corrections by the SPIRE DPU may be needed in some cases.

SPIRE does a standard seven-point jiggle observation. The offset of the source with respect to the commanded pointing is computed by the DPU using the recorded data. The calculated pointing offsets are 
(a) implemented by the BSM (baseline) or (b) transmitted to the spacecraft AOCS. If (b), the AOCS checks that the required telescope movement is within the acceptable limits and executes it. SPIRE waits for a standard period of time to elapse before flagging the data as valid.

### Operation of photometer internal calibrator (POF8)

It is envisaged that the photometer calibrator will be operated at intervals of an hour or more. Its function is to present a repeatable signal to the detectors. This will allow characterisation of (i) any responsivity or system gain drifts; (ii) any variation of the detectors’ responsivity with radiant loading (e.g., non-linear response when	viewing very bright sources).

In operation, it will be powered by a pre-selected waveform, and the corresponding signals will be recorded. A typical duration of the whole sequence will be ~ 10 sec. The BSM must be fixed at its rest position, and the telescope pointing must also be fixed so that the only signal modulation is from the calibrator. To allow calibration to be performed flexibly, it may be advantageous to incorporate this function within some of the other POFS (e.g., to enable calibrator flashes to be interspersed between the rows of line scanning observations).

## Spectrometer Observatory Functions

There are two different Spectrometer Observatory Functions:

SOF1:	Point Source Spectrum

SOF2: 	Fully sampled spectral map

### Point source spectrum (SOF1)

This mode is for measuring the spectrum of a point source that is well centred on the central detectors of the FTS arrays. The telescope is pointed at a known position, with the source lying on the chosen detector pair - nominally the central detectors of the two arrays - pointing offset ((YS, (ZS).

The FTS mirror mechanism is scanned over the required range with the velocity controlled by the drive electronics. Each scan will take between 7 and TBD seconds to complete. The detectors and the position sensor are read out asynchronously whilst the mechanism is moving – i.e. time-sampling of the FTS mechanism position. Each interferogram for each detector is stored in the DPU memory. The scan is repeated until the desired integration time has been reached.

### Fully-sampled spectral map (SOF2)

This mode allows imaging spectroscopy of a region of sky or an extended source that is within the FOV of the spectrometer – i.e. less than 2.6 arcmin circular. This is achieved by using the beam steering mirror to perform a low-frequency jiggle and taking one or more interferograms at each point of the jiggle pattern. The sequence of operations for an individual pointing is as for SOF1, and the BSM is used to make an n-point jiggle map as in POF3. The mirror is held at each position while the desired number of FTS scans is carried out. For full sampling at all wavelengths, n = 25. Considering the simplified case of square-packed horns, the step size must be 0.5(/D = 9” at 250 (m and the number of steps must accommodate the need to cover the distance between two beams at 350 (m: 2(/D = 45”. Five steps in each orthogonal direction are thus required. The geometry of the juggle pattern is hexagonal for the hexagonally packed feedhorns, but the number of steps required is still 25.

The number of interferograms required per position is selectable (a minimum of three would seem sensible, and could be built up either by repeating the jiggle three times or doing three scans at each position of a single jiggle). With three scans per position, it will take at least 525 seconds to take a fully sampled R = 0.4 cm-1 2.6 arcmin circular map - with overheads this will be around ten minutes.

# Spire Sensitivity Estimation

The sensitivity of SPIRE has been estimated under the assumptions listed below. Pessimistic overall optical efficiencies of 30% for the photometer and 15% for the FTS are assumed, taking into account all losses including filter transmission, mirror reflectivity, diffraction within the instrument and pupil alignment errors.

| Telescope temperature (K) | 80 |
| --- | --- |
| Telescope emissivity | 0.04 |
| Telescope used diameter (m) (1) | 3.29 |
| No. of observable hours per 24-hr period | 21 |
| Photometer |
| Bands ((m) | 250 | 350 | 500 |
| Numbers of detectors | 139 | 88 | 43 |
| Beam FWHM (arcsec.) | 17 | 24 | 35 |
| Bolometer DQE (2) | 0.6 | 0.7 | 0.7 |
| Throughput | (2 |
| Bolometer yield | 0.8 |
| Feed-horn/cavity efficiency (3) | 0.7 |
| Field of view (arcmin.) Scan mapping Field mapping | 4 x 8 4 x 4 |
| Overall instrument transmission | 0.3 |
| Filter widths ((/(() | 3.3 |
| Observing efficiency (slewing, setting up, etc.) | 0.9 |
| Chopping efficiency factor | 0.45 |
| Reduction in telescope background by cold stop (4) | 0.8 |
| FTS spectrometer |
| Bands ((m) | 200-300 | 300-670 |
| Numbers of detectors | 37 | 19 |
| Bolometer DQE | 0.6 | 0.7 |
| Feed-horn/cavity efficiency | 0.70 |
| Field of view diameter (arcmin.) | 2.6 |
| Max. spectral resolution (cm-1) | 0.04 |
| Overall instrument transmission | 0.15 |
| Signal modulation efficiency | 0.5 |
| Observing efficiency | 0.8 |
| Electrical filter efficiency | 0.8 |

Table TBD: Assumptions for SPIRE sensitivity estimation

Notes:

- The telescope secondary mirror is the pupil stop for the system, so that the outer edges of the primary mirror are not seen by the detectors. This is important to make sure that radiation from highly emissive elements beyond the primary reflector does not contribute stray light.

- The bolometer DQE (Detective Quantum Efficiency) is defined as the where NEPph is the photon noise NEP due to the absorbed radiant power and NEPTotal is the overall NEP including the contribution from the bolometer noise.

- This is the overall absorption efficiency of the combination of feed-horn, cavity and bolometer element.

- A fraction of the feedhorn throughput falls outside the solid angle defined by the photometer 2-K cold stop and is thus terminated on a cold (non-emitting) surface rather than on the 4% emissive 80-K telescope. This reduces the background power on the detector.

The detailed calculations for the photometer and spectrometer are presented in Annex TBD and Annex TBD. The main results are summarised below.

Background power and photon noise levels: The background power levels (which are dominated by the telescope emission), and the corresponding photon noise limited NEP values are given in Table TBD.

|  |  | Photometer band ((m) | FTS band ((m) |
| --- | --- | --- | --- |
|  |  | 250 (m | 350 (m | 500 (m | 200-300 (m | 300-670 |
| Background power/detector | pW | 3.9 | 3.2 | 2.0 | 6.0 | 11 |
| Background-limited NEP | W Hz-1/2 x 10-17 | 8.1 | 6.1 | 4.5 | 10 | 11 |
| Overall NEP (inc. detector) | W Hz-1/2 x 10-17 | 10 | 7.3 | 5.4 | 12 | 14 |

Table TBD:	Background power and photon noise-limited NEPs for SPIRE.

Instrument sensitivity: The estimated sensitivity levels for SPIRE are summarised in Table TBD. The figures quoted are the nominal values, with an overall uncertainty of around 50% to take into account uncertainties in instrument parameters, particularly feedhorn efficiency, detector DQE, and overall transmission efficiency. The pixel size will be increasingly mis-matched to the diffraction spot size. The trade-off between wavelength coverage and sensitivity of the long-wavelength FTS band must be studied in detail. At the moment, we estimate an effective loss of efficiency of a factor of two at 670 m, and scale linearly for wavelengths between 400 and 670 m. Performance beyond 400 (m may have to be compromised to maintain the desired sensitivity below 400 (m.

| Photometry |
| --- |
| ( (m |  | 250 | 350 | 500 |
| (S(5-(; 1-hr) mJy | Point source (7-point mode) | 2.5 | 2.6 | 2.9 |
|  | 4’ x 4’ jiggle map | 8.8 | 8.7 | 9.1 |
|  | 4’ x 8’ scan map | 7.3 | 7.2 | 7.5 |
| Time (days) to map 1 deg.2 to 3 mJy 1-( | 1o x 1o scan map | 1.8 | 1.7 | 1.9 |

| Line spectroscopy (( = 0.04 cm-1 |
| --- |
| ( (m |  | 200 | 400 | 670 |
| (S (5-(; 1-hr) W m-2 x 10-17 | Point source | 3.4 | 3.9 | 7.8 |
|  | 2.6’ map | 9.0 | 10 | 21 |

| Low-resolution spectrophotometry (( = 1 cm-1 |
| --- |
| ( (m |  | 200 | 400 | 670 |
| (S (1-(; 1-hr) mJy | Point source | 110 | 130 | 260 |
|  | 2.6’ map | 300 | 350 | 700 |

Table TBD: Estimated sensitivity of SPIRE for broad-band photometry, line spectroscopy and low-resolution spectrophotometry. 

For the FTS, limiting flux density is inversely proportional to spectral resolution (). Limiting line flux is independent of spectral resolution (for an unresolved line).

# References

Ade, P A R., P A Hamilton, & D A Naylor, In Fourier transform spectroscopy: new methods and applications, OSA, 90, 1999

Bock, J J, H G LeDuc, A E Lange, & J Zmuidzinas, “A monolothic bolometer array suitable for FIRST”, Proceedings of ESA Symposium on The Far Infrared and Submillimetre Universe, Grenoble, 15-17 April 1977, ESA SP-401, 349, 1997.

Caldwell, M E, B M Swinyard, & A Richards, Proc. SPIE 4013, TBD, 2000.

Dohlen, K, A Orignea, D Pouliquen, & B Swinyard, Proc. SPIE 4013, 119, 2000.

Duband, L, “Spaceborne helium adsorption coolers”, Proceedings of ESA Symposium on The Far Infrared and Submillimetre Universe, Grenoble, 15-17 April 1977, ESA SP-401, 357, 1997.

Duband, L, “A thermal switch for use at liquid helium temperature in spaceborne cryogenic systems”, Proceedings of the 8th International Cryocooler Conference, Vail, Colorado, p731, 1995.

Griffin, M J, & W S Holland, Int. J. Infrared mm Waves 9, 861, 1988.

Griffin, M J. “The design of a bolometer instrument for FIRST”, Proceedings of ESA Symposium on The Far Infrared and Submillimetre Universe, Grenoble, 15-17 April 1997, ESA SP-401, 31, 1997.

Griffin et al. "SPIRE: A bolometer instrument for FIRST", proposal submitted in response to ESA's Announcement of Opportunity for FIRST payload instruments, 1998a.

Griffin, M, L Vigroux, B Swinyard, and C Cunningham, "SPIRE - a bolometer instrument for FIRST", Proc. SPIE 3357, 404, 1998b.

Griffin, M, "Observing speed for SPIRE feedhorn and filled array options", SPIRE Doc Ref, Jan. 24 2000

Holland, W S, E I Robson, W K Gear, C R Cunningham, J F Lightfoot, T Jenness, R J Ivison , J A Stevens, Ade, P A R, M J Griffin, W D Duncan, J A Murphy, & D A Naylor, Mon. Not. R. Astron. Soc. 303, 659, 1999.

Lee, C, P A R Ade, and C V Haynes, "Self-supporting filters for compact focal plane designs", in Proc. Submillimeter and far-Infrared Space Instrumentation, 30th ESLAB Symposium, Noordwijk, ESA S-P 388, 1996.

Martin, D H, “Polarising interferometric spectrometers for near- and sub-millimetre spectra”, Int. J. Infrared mm Waves, 6, 65, 1982.

Mather, J, Appl. Opt. 21, 1125, 1982.

Mauskopf, P D, J J Bock, , H del Castillo, W L Holzapfl, & A E Lange, “Composite infrared bolometers with Si3N4 micromesh absorbers”, Appl. Opt. Vol. 36, No. 4, 765, 1997.

Murphy et al. TBD

Oliver, S. Toledo paper.

Richards, P L, J. Appl. Phys. 76, 1, 1994.

Swinyard, B M, P A R Ade, M J Griffin, P A Hamilton, K Dohlen, J-P Baluteau, D Pouliquen, D Ferand, P Dargent, G Michel, J Martignac, L Rodriguez, D E Jennings, M E Caldwell, & A G Richards, Proc. SPIE 4013, 196, 2000.

# List of Annexes

SPIRE Photometer Sensitivity Estimation

SPIRE FTS Sensitivity Estimation

# Comments

Is this figure correct? Shouldn’t the cold tip be at the Evaporator end?

Do we want to put in something about the thermal stability requirements of the detectors?

What is this symbol?

How does this time constant differ from the others

This is including the contingency

Peter Hargrave to check this whole section

Peter Hargrave to check this whole section

Matt to spec. the temp

# Text boxes

250 m

Pointing offset (arcsec.)

20

16

12

8

4

0

5

4.5

4

3.5

3

2.5

2

1.5

1

0.5

0

Line N

etc.

D1

Line N-1

Line 3

Line 2

Line 1

d2

N lines

d2

M positions per line

d2

Offset for clarity

A1

B1

B2

A2

Nod position

1

Chop

throw

Source

Centre of FIRST 
focal plane

Centre of Spectrometer fov
Coordinates: ((YS, (ZS)

(Z

Centre of photometer fov
Coordinates: ((YP, (ZP)

(Y

500 m

350 m

(

(a)

(b)

4 x 8 arcminute field of view of the SPIRE array determined by the SPIRE optics

Central 4 x 4

arcminute portion

BSM at rest 
position

BSM at offset = +2 arcmin

BSM at offset = -2 arcmin