| SPIRE & PACS Sorption Coolers INTERFACE CONTROL DOCUMENT |
| --- |
| SBT internal ref : SBT/CT/2001-24 |
|  | Name & Function | Date | Signature |
|  | Name & Function | Date | Signature |
| Prepared | L. Duband - Cooler project manager | April 10rd 2001 |  |
| Verified |  |  |  |
| QA Check | P. Dupont – PA manager | April 10th 2001 |  |
| Project Approval | L. Duband - Cooler project manager | April 10th 2001 |  |

Service des Basses Températures (SBT) Département de Recherche Fondamentale sur la Matière Condensée (DRFMC) COMMISSARIAT A L'ENERGIE ATOMIQUE - GRENOBLE (CEA-Grenoble) 17, rue des Martyrs 38054 GRENOBLE Cédex 9, France.

# Document Status

| Issue | Revision | Date | Nb of pages | Modifications |
| --- | --- | --- | --- | --- |
| 0 | 0 | 10/04/2001 | 15 | First draft |

# Table of Contents

1 Scope of the document	1

2 Documents	2

2.1 Applicable documents	2

3 Functional description and block diagram	3

4 Mechanical interface	4

4.1 Cooler structure – SPIRE	5

4.2 Cooler structure - PACS	6

4.3 Heat switches	6

4.4 Evaporator cold end	7

4.5 Electrical connector	8

5 Thermal interface	9

5.1 Cooler structure	9

5.2 Heat switches	9

5.2.1 Sorption pump heat switch	9

5.2.2 Evaporator heat switch	9

5.3 Evaporator cold tip	9

6 Electrical interface	10

7 Appendix – General interface drawing	11

List of Acronyms

| AD | Applicable Document |  |  |
| --- | --- | --- | --- |
| CEA | Commissariat à l' Energie Atomique |  |  |
| CDR | Critical Design Review | Revue de conception détaillée | RCD |
| CQM | Cryogenic Qualification Model |  |  |
| EV | Evaporator |  |  |
| FIRST | Far Infrared and SubmillimetreTelescope |  |  |
| FS | Flight spare |  |  |
| HSE | Heat Switch (on evaporator) |  |  |
| HSP | Heat Switch (on sorption pump) |  |  |
| N/A | Not Applicable |  |  |
| PACS | Photoconductor. Array Camera and Spectrometer |  |  |
| PFM | ProtoFlight Model |  |  |
| RD | Reference Document |  |  |
| SAp | Service d’Astrophysique |  |  |
| SBT | Service des Basses Températures |  |  |
| SCO | Sorption Cooler (full unit) |  |  |
| SP | Sorption pump |  |  |
| SPIRE | Spectral & Photometric Imaging Receiver |  |  |
| SST | Support Structure |  |  |
| TS | Thermal Shunt |  |  |
| TSES | Thermal Strap to Evaporator Switch |  |  |
| TSPS | Thermal Strap to Pump Switch |  |  |

# 1 Scope of the document

This document defines the various interfaces between the 3He sorption cooler and the SPIRE and PACS instruments. It comprises the mechanical, thermal and electrical interfaces. Both intruments share the same cooler design and consequently all interfaces are designed to be compatible.

# 2 Documents

## 2.1 Applicable documents

All Applicable Documents are listed in the AD chapter of the CIDL (HSO-SBT-LI-010).

# 3 Functional description and block diagram

The cooling of the SPIRE and PACS detectors down to 300 mK will be effected by  helium three sorption coolers. Both instruments share the same cooler design and thermal architecture. For each the cooler is mounted off a 4 K structure with two thermal straps connected to the superfluid helium cryostat. The temperature of the structure and the cryostat are referred as respectively Level 0 and Level 1. This thermal architecture is represented on the figure below.

Thermal architecture

Block diagram

# 4 Mechanical interface

The cooler required volume is a rectangle parallelepiped of dimensions 228.5 mm x 100 mm x 100 mm. Reference drawing : 2000-14 B 000. Throughout this documents, unless otherwise specified, all dimensions are in millimeters.

Cooler overall dimensions

For the purpose of this discussion the top, bottom and sides of the cooler are defined as represented on the figure below.

The total overal mass of the cooler is about 1680 grams.

## 4.1 Cooler structure – SPIRE

Within the SPIRE instrument the cooler is mounted “vertically”. The cooler interfaces with the structure with 8 through holes 4.3 mm in diameter located on the 100 mm x 100 mm bottom base plate, as defined on the following drawing.

3D view and interface drawing for cooler mounting

In addition the top cover features four holes as shown on the figure below. Part (or all ) of these holes will be used as snubber(s): within SPIRE some extension(s) is (are) foreseen which insert in this (these) holes after the cooler is mounted, to prevent any excessive lateral motion.

3D view and interface drawing for top cover (sorption pump)

## 4.2 Cooler structure - PACS

Within the PACS instrument the cooler is mounted along one of its side. The cooler interfaces with the structure with 5 through holes 5.5 mm in diameter located on one of the side wall as defined on the following drawing.

3D view and interface drawing for cooler mounting

## 4.3 Heat switches

The cooler comprises two gas gap heat switches. Each heat switch interfaces with a thermal strap (connected to the superfluid cryostat). The mechanical interface for both switches is similar and is a copper plate gold plated 16 mm x 16 mm featuring two M3 holes.

This interface also features a tool designed to prevent any excessive torque on the gas gap heat switch when mounting the thermal strap. This tool is intended to be used only while screwing the strap and must be removed before cooler operation. The recommended torque for the screws is TBD Nm.

It is important to note that any excessive mechanical load on this interface must be avoided and shall in any case never exceeds 80 Nw in any directions.

3D view and interface drawing for heat switch base / thermal strap

## 4.4 Evaporator cold end

The evaporator cold end interfaces with a 300 mK thermal strap which is then splitted and connected to the detectors. This interface is a copper piece gold plated 14 mm x 12.5 mm, 3 mm thick, featuring two through holes 3.2 mm in diameter. Each side of the copper piece is available as a contacting area.

The quality of this interface is critical to the performance of the cooler.

Any excessive troque or bending force on this interface must be avoided. Always use two wrenches (one on each side) when tightening screws on the cold head.

The mounting of the 300 mK thermal strap to the cold end must be performed by trained people. A specific procedure will be established.

3D view and interface drawing for evaporator cold end / 300 mK  thermal strap

## 4.5 Electrical connector

Two electrical connectors type MDM 37 pins (main and redundant) are provided for the cooler operation. These connectors are supported by a bracket included in the structural box and located on the side of the cooler as shown on the following drawing.

3D view and interface drawing for connector / connector

# 5 Thermal interface

Informations of the mechanics of the thermal interfaces are given in the previous paragraphs. In these paragraphs, we define only the interfaces from the thermal point of view.

The maximum storage temperature is 353 K (80 °C). There is no minimum storage temperature. The thermal interfaces are defined in functional conditions.

## 5.1 Cooler structure

The structure of the cooler is mounted on a 4 K structure.

## 5.2 Heat switches

The cooler comprises two identical gas gap heat switches. They both interface with a strap (one each) thermally connected to the 1.8 K superfluid tank. The two straps are different due to the difference in power flowing to the cryostat from one strap to the other (see hereafter). The interface with the switch to the sorption pump (TSPS) is less critical as it is possible to have this interface temperature raise to 10 K (TBD) during recycling. On the contrary the interface to the evaporator switch has a significant impact on the performance. During the recycling process this interface must be kept as close as possible to the cryostat temperature to guarantee good condensation efficiency, minimal cooldown losses, etc… for the cooler.

### 5.2.1 Sorption pump heat switch

(related to TSPS)

| Temperature range (K) | 1.8 | 10 (TBD) |
| --- | --- | --- |
| Heat flow range (mW) | ≤ 2 mW* | ≤ 1 W |

(* : this number is a function of the detectors dissipation)

### 5.2.2 Evaporator heat switch

(related to TSES)

| Temperature range (K) | 1.8 | 1.8 +  |
| --- | --- | --- |
| Heat flow range (mW) | ≤ 1 mW | ≤ 50 mW |

## 5.3 Evaporator cold tip

| Temperature range (K) | 0.29 | 1.8 |
| --- | --- | --- |
| Heat flow range (µW) | 20 | - |

# 6 Electrical interface

All thermometers used are Cernox thermometers type 1030 with SD package (supplier : LakeShore) measured in four wires mode.

All heaters used are high reliability metal film resistors 402 Ohms resistance type RLR/2H3 (TBC) (supplier : Vishay/Sfernice) driven by four wires.

Two electrical connectors type MDM 37 pins (supplier : Canon) are used, one prime and one redondant. The wiring of both connectors is identical.

All wiring is made using manganin wires.

| Signal | Pin n° | Wire diam. | Max. current | Remarks |
| --- | --- | --- | --- | --- |
| SP temperature I+ | 1 | TBD (0.2 ?) | 1 µA |  |
| SP temperature V+ | 2 | “ | N/A |  |
| SP temperature V- | 3 | “ | N/A |  |
| SP temperature I- | 4 | “ | 1 µA |  |
| EV temperature I+ | 5 | 0.1 mm | 250 nA |  |
| EV temperature V+ | 6 | “ | N/A |  |
| EV temperature V- | 7 | “ | N/A |  |
| EV temperature I- | 8 | “ | 250 nA |  |
| HSP pump temperature I+ | 9 | TBD (0.2 ?) | 1 µA |  |
| HSP pump temperature V+ | 10 | “ | N/A |  |
| HSP pump temperature V- | 11 | “ | N/A |  |
| HSP pump temperature I- | 12 | “ | 1 µA |  |
| HSE pump temperature I+ | 13 | TBD (0.2 ?) | 1 µA |  |
| HSE pump temperature V+ | 14 | “ | N/A |  |
| HSE pump temperature V- | 15 | “ | N/A |  |
| HSE pump temperature I- | 16 | “ | 1 µA |  |
| Thermal shunt temperature I+ | 17 | TBD (0.2 ?) | 1 µA |  |
| Thermal shunt temperature V+ | 18 | “ | N/A |  |
| Thermal shunt temperature V- | 19 | “ | N/A |  |
| Thermal shunt temperature I- | 20 | “ | 1 µA |  |
| SP heater I+ | 21 | TBD (0.2 ?) | 25 mA |  |
| SP heater I+ | 22 | “ | “ |  |
| SP heater I- | 23 | “ | “ |  |
| SP heater I- | 24 | “ | “ |  |
| EV heater I+ | 25 | 0.1 mm | 1 mA |  |
| EV heater I+ | 26 | “ | “ |  |
| EV heater I- | 27 | “ | “ |  |
| EV heater I- | 28 | “ | “ |  |
| HSP heater I+ | 29 | TBD (0.2 ?) | 1.5 mA |  |
| HSP heater I+ | 30 | “ | “ |  |
| HSP heater I- | 31 | “ | “ |  |
| HSP heater I- | 32 | “ | “ |  |
| HSE heater I+ | 33 | TBD (0.2 ?) | 1.5 mA |  |
| HSE heater I+ | 34 | “ | “ |  |
| HSE heater I- | 35 | “ | “ |  |
| HSE heater I- | 36 | “ | “ |  |
| Not used | 37 |  |  |  |

# 7 Appendix – General interface drawing