| SUBJECT: | SPIRE Data ICDSPIRE Data ICD |
| --- | --- |
| PREPARED BY: | K.J. KingK.J. King |
| DOCUMENT No: | SPIRE-RAL-PRJ-001078SPIRE-RAL-PRJ-001078 |
| ISSUE: | Issue 1.0 (Draft 2)Issue 1.0 (Draft 2) | Date: | 15th January 200315th January 2003 |
| APPROVED BY: |  |  |  |

| Checking Engineer (IFSI) R. Cerulli |  |
| --- | --- |
| OBS Engineer (IFSI) S. Molinari |  |
| Project Manager (IFSI) R. Orfei |  |
| Operations Scientist S. Sidher |  |
| Instrument Scientist B.M. Swinyard |  |
| Instrument Development Manager E.C. Sawyer |  |

Distribution

| R. Orfei | IFSI |
| --- | --- |
| R Cerulli | IFSI |
| S. Molinari | IFSI |
| J-L Augerres | CEA, Saclay |
| C. Cara | CEA, Saclay |
| S. D. Sidher | RAL |
| B. M. Swinyard | RAL |
| K. J. King | RAL |
| J.A. Long | RAL |

# Change Record

| ISSUE | DATE |  |
| --- | --- | --- |
| Draft 1 | 15th Jan 2002 | Original Draft |
| Draft 2 | 15th January 2003 | Updated Draft for DPU AVM Acceptance Test Section 1.1: Updated Scope description Section 1.3.2: Updated document references and added RD08 Section 2.1.3: Added note about TIME contents Section 2.2: Renamed APID5 Section 2.3: included details of which parts of the 1553 interface protocol are used Section 3.1: Updated Telecommand service usage Section 3.2.3: Defined both Housekeeping and Diagnostic reports. There are now 4 of each with 2 housekeeping and 1 diagnostic report pre-defined in the OBS Section 3.2.3.1: Updated definition of service (3,1) to send Table_ID as a parameter. This allows different SIDs for the same Table_ID Section 3.2.3.2: Service (3,2) set to Not Used Section 3.2.3.3: Service (3,3) now defined Section 3.2.3.4: Service (3,4) set to Not Used Section 3.2.3.5:Service (3,9) now defined Section 3.2.3.6: Service (3,11) set to Not Used Section 3.2.6.1: Memory addressing changed to use 8 bit Memory ID and 24 bit address Section 3.2.6.2: Memory addressing changed to use 8 bit Memory ID and 24 bit address Section 3.2.6.3: Memory addressing changed to use 8 bit Memory ID and 24 bit address Section 3.2.8.1.2-4: Deleted Cooler, SCAL and 300mk control commands – they are now implemented in software-driven VMs Section 3.2.8.2.2-4: Deleted Cooler, SCAL and 300mk control commands – they are now implemented in software-driven VMs Section 3.2.8.3.1: Changed Load_Table command to Set_Table, with consequent changes to parameters Section 3.2.8.3.2: Changed parameters Section 3.2.8.3.3: Changed parameters Section 3.2.8.3.6: Inserted commands to start and stop software-driven VMs Section 3.2.8.3.20: Deleted data packing command – data frames are packed into science packets until they are full (unless a flush command is received) Section 3.2.9.1: Change to packet definition to all Ack field to be changed Section 3.2.14.1: Service (14,1) now defined Section 3.2.14.2: Service (14,2) now defined Section 3.2.14.3: Service (14,3) now defined Section 3.2.17.1: Updated command description Section 3.2.20.4: This telecommand is no longer used – OTF and TSF provided by ground processing Section 3.2.22: Context Service not used Section 4.1: table updated to reflect changes to services Section 4.2.1.6: Updates Error code table Section 4.2.3.2: Added Critical Housekeeping Report definition Section 4.2.3.3: Updated Nominal Housekeeping Report definition Section 4.2.3.4: Added Detector Housekeeping Report definition Section 4.2.5.1: Removed OTF and TSF reports and added definition of report identified so far Section 4.2.5.2: Added definition of reports Section 4.3.5.3: Removed Error/Alarm report definition – they are not used as the ground does not remain in contact with the spacecraft Section 4.2.6: Defined Memory Dump reports Section 4.2.8: defined Function Activity Report Section 4.2.14: Defined Packet Generation Status Report Section 4.2.17: Defined Link connection Report Section 4.2.21: Science Data packet definition updated Section 4.2.22: Context Saving Service not used Section 5.1.1: Updated TC Parameter List Section 5.1.3: Updated TC Parameter Constraints Added commands to select science data Added Appendix D - description of science data processing |

# Table of Contents

1.	Introduction	10

1.1	Scope	10

1.2	Structure of the Document	10

1.3	Documents	11

1.3.1	Applicable Documents	11

1.3.2	Reference Documents	11

2.	The Packet Interface	12

2.1	Packet Structures	12

2.1.1	Fields	12

2.1.2	Telecommand Packets	12

2.1.3	Telemetry Packets	13

2.2	APIDs	13

2.3	Packet Transfer Protocol	13

3.	SPIRE Telecommand Packets	15

3.1	Telecommand Packet Types	15

3.2	Telecommand Packet Definition	17

3.2.1	Telecommand Verification Service	17

3.2.2	Device Command Distribution	17

3.2.3	Housekeeping and Diagnostic Data Reporting	17

3.2.3.1	Define New Housekeeping Parameter Report (Service 3,1)	17

3.2.3.2	Define New Diagnostic Parameter Report (Service 3,2)	18

3.2.3.3	Clear Housekeeping Parameter Report Definitions (Service 3,3)	18

3.2.3.4	Clear Diagnostic Parameter Report Definitions (Service 3,4)	19

3.2.3.5	Report Housekeeping Parameter Report Definitions (Service 3,9)	19

3.2.3.6	Report Diagnostic Parameter Report Definitions (Service 3,11)	19

3.2.4	19

3.2.5	Event Reporting	19

3.2.6	Memory Management	20

3.2.6.1	Load Memory Using Absolute Addresses (Service 6,2)	20

3.2.6.2	Dump Memory Using Absolute Addresses (Service 6,5)	21

3.2.6.3	Check Memory Using Absolute Addresses (Service 6,9)	21

3.2.7	21

3.2.8	Function Management	22

3.2.8.1	Start Function (Service 8,1)	22

3.2.8.2	Stop Function (Service 8,2)	23

3.2.8.3	Perform Activity of a Function (Service 8,4)	24

3.2.8.4	Report Function Status (Service 8,5)	40

3.2.9	Time Management	41

3.2.9.1	Enable Time Verification (Service 9,7)	41

3.2.10	42

3.2.11	On-Board Scheduling	42

3.2.12	On-Board Monitoring	42

3.2.13	42

3.2.14	Packet Transmission Control	43

3.2.14.1	Enable Generation of Telemetry Packets (Service 14,1)	43

3.2.14.2	Disable Generation of Telemetry Packets (Service 14,2)	43

3.2.14.3	Report Enabled Telemetry Packets (Service 14,3)	43

3.2.15	On-Board Storage and Retrieval	45

3.2.16	On-Board Traffic Management	45

3.2.17	Test Service	46

3.2.17.1	Perform Connection Test (Service 17,1)	46

3.2.18	On-Board Control Procedures	46

3.2.19	Action/Event Service	46

3.2.20	Information Distribution Service	47

3.2.20.1	Enable Distribution of Information TC Packets (Service 20,1)	47

3.2.20.2	Disable Distribution of Information TC Packets (Service 20,2)	47

3.2.20.3	Report Distributed Information Packets (Service 20,3)	47

3.2.20.4	Information Telecommand (Service 20,4)	48

3.2.21	Science Data	48

3.2.22	Context Saving Service	48

4.	SPIRE Telemetry Packets	49

4.1	Telemetry Packet Types	49

4.2	Telemetry Packet definitions	50

4.2.1	TC Verification Service	50

4.2.1.1	Telecommand Acceptance Report - Success (1,1)	50

4.2.1.2	Telecommand Acceptance Report - Failure (1,2)	50

4.2.1.3	Telecommand Execution Report - Started (1,3)	52

4.2.1.4	Telecommand Execution Report - Progress (1,5)	52

4.2.1.5	Telecommand Execution Report - Completion (1,7)	52

4.2.1.6	Telecommand Execution Report - Failure (1,8)	53

4.2.2	Device Command Distribution	54

4.2.3	Housekeeping and Diagnostic Data Reporting	55

4.2.3.1	Housekeeping Parameter Report  (Service 3,25)	55

4.2.3.2	Critical Housekeeping Report (SID=0x0300)	56

4.2.3.3	Nominal Housekeeping Report (SID=0x0301)	58

4.2.3.4	Detector Housekeeping Report (SID=0x0302)	69

4.2.4	77

4.2.5	Event Reporting	77

4.2.5.1	Event Report (5,1)	77

4.2.5.2	Exception Report (5,2)	84

4.2.5.3	Error/Alarm Report (5,4)	86

4.2.6	Memory Management	87

4.2.6.1	Memory Dump, Absolute Addresses	87

4.2.6.2	Memory Check Report, Absolute Addresses	87

4.2.7	87

4.2.8	Function Management	88

4.2.9	Time Management	88

4.2.9.1	Time Verification Report	88

4.2.10	88

4.2.11	On Board Scheduling	88

4.2.12	On Board Monitoring	88

4.2.13	88

4.2.14	Packet Transmission Control	89

4.2.15	On Board Storage and Retrieval	89

4.2.16	On Board Traffic Management	89

4.2.17	Test Service	89

4.2.18	On Board Control Procedures	89

4.2.19	Action/Event Service	89

4.2.20	Information Distribution Service	90

4.2.20.1	Distributed Information Packets Report (Service 20,5)	90

4.2.21	Science Data	91

4.2.21.1	Nominal Science Data Report (21,1)	91

4.2.21.2	Science Type B Data Report (21,2)	98

4.2.21.3	Diagnostic Science Report (21,3)	102

4.2.21.4	Auxiliary Science Data Report (21,4)	106

4.2.22	Context Saving Service	109

5.	Parameters	110

5.1	TC Parameters	110

5.1.1	Parameter Definition	110

5.1.2	Conversion Curves	113

5.1.3	Constraints	113

5.2	TM  Parameters	114

5.2.1	Parameter Definition	114

5.2.2	Conversion Curves	124

5.2.3	Constraints	126

6.	Tables	127

7.	Control Loops	128

7.1	Control Loop Parameters	128

8.	Command List	129

9.	Science Data	130

9.1	Assumptions	130

9.2	Science Processing	130

Figures

Tables

Glossary

| APID | Application ID |
| --- | --- |
| CDMS | Command and Data Management System |
| DPU | Digital Processing Unit |
| OBS | On-Board Software |
| SID | Structure ID |
| SPIRE | Spectral and Photometric Imaging REceiver |
| SVM | Service Module |
| VM | Virtual Machine |

# Introduction

Control of the SPIRE instrument is handled by the Digital Processing Unit (DPU) part of the instrument electronics, which contains the On-Board Software (OBS). This unit is mounted on the spacecraft Service Module (SVM) and interfaces directly to the CDMS bus of the Herschel spacecraft for the purpose of transferring commands to, and collection of telemetry data from the instrument.

All data passing between the instrument and the spacecraft is transferred in the form of telemetry and telecommand packets conforming to the ESA packet standards (RD01 and RD02).  The ESA Packet Utilisation Standard (RD03) defines the types of service that may be provided by units on board ESA spacecraft, but the set of supported packet types within the Herschel project is restricted. This restricted set is defined in the Herschel Packet Structure ICD (AD01) and is the minimum set necessary to meet the satellite operational requirements given in AD02.

The current issue of this document is based on Issue 0.7 of the DRCU/DPU ICD (RD06). Until that document is at Issue 1.0 the contents of this document are subject to change. Items likely to be updated in the future are marked in blue text.

## Scope

This document defines the packet types and their contents that will be accepted and generated by the SPIRE instrument during all operations.  These packets conform to the formats given in the Packet Structure ICD (AD01) and the Ground Segment to Instruments ICD (AD03). They also provide for the instrument functionality described  in the instrument Operating Modes document (RD04) and elaborated in the instrument operating document (RD04).

The document does not describe the use of the commands (to be given in the OBS User Manual) or the way in which the commands are handled by the OBS (to be found in the OBS Architectural Design document).

## Structure of the Document

Section 2 describes the packet interface between the instrument and the spacecraft. This includes the general format of the packets used by the SPIRE instrument for telecommanding and telemetry  (from AD01), the allocation of Application IDs used by the instrument and the functionality of the packet  transfer protocol of the instrument/spacecraft interface (from AD01, appendix 9) that is used by the instrument.

Section 3 defines the format and content of each of the telecommand packets accepted by the instrument. Section 4 defines the corresponding information for the telemetry packets generated by the instrument. A description of how these packets are handled by the instrument is given in RD05.

Section 5 defines, in detail, the parameters used in the telecommand and telemetry packets.

## Documents

### Applicable Documents

| AD01 | Herschel/Planck Packet Structure Interface Control Document. (SCI-PT-ICD-07527) Issue 2.0 (draft2) |
| --- | --- |
| AD02 | Herschel/Planck Operations Interface Requirements Document (SCI-PT-RS-07360), Issue 2.0 (draft3) |
| AD03 | Herschel Science Ground Segment to Instruments Interface Control Document (FIRST-FSC-DOC-0200), Issue 1.0 |

### Reference Documents

| RD01 | Packet Telemetry Standard (ESA PSS-04-106), Issue 1, 1998 |
| --- | --- |
| RD02 | Packet Telecommand Standard (ESA PSS-04-107), Issue 2, 1992 |
| RD03 | Telemetry and Telecommand Packet Utilisation Standard (ECSS-E-70/41) Draft 04, April 1999 |
| RD04 | Operating Modes of the SPIRE Instrument (SPIRE-RAL-PRJ-000320), Issue 3.0, 4th January 2002 |
| RD05 | Operating the SPIRE Instrument (SPIRE-RAL-DOC-000768), Issue 0.4, 29th January 2002 |
| RD06 | DRCU/DPU Interface Control Document (SAp-SPIRE-CCa-076-02) Issue 0.7, 26th June 2002 |
| RD07 | MCU/DPU Command List Interface Control Document and User Manual (LAM/ELE/SPI/011011), Issue 2.0, 11th April 2002 |
| RD08 | SPIRE On-Board Software User Manual (SPIRE-IFS-PRJ-001391) |
| RD09 | Herschel SPIRE DPU Hardware User Manual (SPIRE-IFS-PRJ-001390), Issue 1, 7th October 2002 |

# The Packet Interface

## Packet Structures

The following packet structures are shown as a set of 16 bit words, contained in two consecutive octets in the packet structure (all packets are composed of an even number of octets). The most significant octet of each word comes before the least significant in the packet. The least significant bit of each word is on right of each field:

|  | MS Octet | LS Octet |  |
| --- | --- | --- | --- |
| Most Significant Bit | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Least Significant Bit |

### Fields

Within a field (of any length) the most significant bit is designated bit (0), the least significant bit is bit (length-1).

### Telecommand Packets

The following figure gives the general structure of a SPIRE TC Packet (after AD01)

| Packet Header | Packet ID | 0 | 0 | 0 | 1 | 1 | APID |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Sequence Control | 1 | 1 | Src | Count |
|  | Length | Length |
| Data Field | Data Field Header | 0 | 0 | 0 | 0 | Ack | 1 | Type |
|  |  | Sub-Type | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
|  | Source Data |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | D | A | T | A |  |  |  |  |  |  |
|  | Error Control | Checksum |

Src,Count, Length, TIME and Checksum are defined in AD01

Ack is the most significant 3 bits of Ack as defined in AD01

Type and Sub-Type define the packet type and are also defined in AD01

### Telemetry Packets

The following figure gives the general structure of a SPIRE TM Packet (after AD01)

| Packet Header | Packet ID | 0 | 0 | 0 | 0 | 1 | APID |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Sequence Control | 1 | 1 | Count |
|  | Length | Length |
| Data Field | Data Field Header | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Type |
|  |  | Sub-Type | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
|  |  | TIME |
|  | Source Data |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | D | A | T | A |  |  |  |  |  |  |
|  | Error Control | Checksum |

Count, Length, TIME and Checksum are defined in AD01

Type and Sub-Type define the packet type and are also defined in AD01

Note: the TIME field contains the DPU time of the construction of the packet, not the time of collection of the data.

## APIDs

The Application ID is used to identify the source or destination of a telemetry packet. Herschel uses different APIDs for different types of packet (see AD1) as well as for different instruments. The APIDS to be used by SPIRE are given in the following table:

| ID | Telemetry types | APID (hex) |
| --- | --- | --- |
| APID1 | SPIRE Telecommands, Telecommand Verification and Events | 500 |
| APID2 | SPIRE Periodic Housekeeping | 502 |
| APID3 | SPIRE Photometer Science Data | 504 |
| APID4 | SPIRE Spectrometer Science Data | 505 |
| APID5 | SPIRE Subsystem Science Data | 506 |

## Packet Transfer Protocol

At the low-level the Packet Transfer Protocol provides a series of sub-address messages to control and transfer data between subsystems. Some of these implement the packet transfer itself, others provide alternative methods of transferring data and controlling the transfer. This section identifies the sub-addresses used/accepted by the SPIRE instrument.

| Description | Sub-address(es) | Comments |
| --- | --- | --- |
| Mode Command | SA 0R | Used to identify the RT addressed in this subframe. The instrument responds only to its own address. |
|  | SA 0T, 31T, 31R | Not Used |
| Unit Control | SA 1R | Not Used |
|  | SA 1T | Contains Subframe counter and BIT word of 1553 i/f |
| Data Send | SA 2T, 3T, 4T, 7T, 9T, 29T | Not Used |
| Data Receive | SA 2R, 7R, 9R, 15R-26R, 29R | Not Used |
| Asynchronous Short Commands | SA 3R, 4R | Not used |
| Event Messages | SA 5R, 5T, 6R, 6T | Not Used |
| Time Messages | SA 8T | Not used |
|  | SA 8R | Used to transfer spacecraft time from CDMU to instrument (in subframe 33) rather than use packet service (9,5) |
| Packet Transfer TM Transfer Request | SA 10T | Used by instrument to indicate to CDMU that a TM packet is ready to transfer |
| TM Transfer Confirmation | SA 10R | Used to implement retry at packet level |
| TC Transfer Confirmation | SA 27T | Instrument sends a copy of TC Transfer Descriptor to indicate reception of TC |
| TC Transfer Descriptor | SA 27R | Used by the instrument to prepare for TC transfer |
| TM Data Send | SA 11T-26T | Used to transfer TM packets from instrument to CDMU |
| TC Data Receive | SA 11R-14R | Used to transfer TC packets from CDMU to instrument |
| Low-level Commands | SA 28T, 28R | Not Used |
| Data Wrap | SA 30T, 30R | Not Used by instrument, but made available for future use |

# SPIRE Telecommand Packets

This section defines all the telecommand packets accepted by the SPIRE instrument

## Telecommand Packet Types

The Packet Structure ICD (AD1) defines many types of service that can be provided by an Application. The following table shows the telecommand packet types that will be accepted by the SPIRE instrument.

| Description | Service Type | Service Sub-Type | Comments |
| --- | --- | --- | --- |
| Telecommand Verification Service | 1 |  | N/A |
| Device Command Distribution | 2 |  | Not Used |
| Housekeeping and Diagnostic Data Reporting |  |  |  |
| Define New HK Parameter Report | 3 | 1 |  |
| Define New Diagnostic Parameter Report | 3 | 2 | Not Used |
| Clear HK Parameter Report Definitions | 3 | 3 |  |
| Clear Diagnostic Parameter Report Definitions | 3 | 4 | Not Used |
| Report HK Parameter Report Definitions | 3 | 9 |  |
| Report Diagnostic Parameter Report Definitions | 3 | 11 | Not Used |
| Event Reporting | 5 |  | N/A |
| Memory Management |  |  |  |
| Load Memory Using Absolute Addresses | 6 | 2 |  |
| Dump Memory Using Absolute Addresses | 6 | 5 |  |
| Check Memory Using Absolute Addresses | 6 | 9 |  |
| Function Management |  |  |  |
| Start Function | 8 | 1 |  |
| Stop Function | 8 | 2 |  |
| Perform Activity of Function | 8 | 4 |  |
| Report Function Status | 8 | 5 |  |
| Time Management |  |  |  |
| Synchronise User | 9 | 3 | Not Used |
| Enable Time Synchronisation | 9 | 4 | Not Used |
| Time Code | 9 | 5 | Not Used |
| Verify User Time | 9 | 6 | Not Used |
| Enable Time Verification | 9 | 7 |  |
| Synchronise Central Time Reference | 9 | 10 | Not Used |
| On-Board Scheduling | 11 |  | Not Used |
| On-Board Monitoring |  |  |  |
| Enable Monitoring of Parameters | 12 | 1 | TBD |
| Disable Monitoring of Parameters | 12 | 2 | TBD |
| Clear Monitoring List | 12 | 4 | TBD |
| Modify Monitoring List | 12 | 5 | TBD |
| Delete Parameters from Monitoring List | 12 | 6 | TBD |
| Report Current Monitoring List | 12 | 8 | TBD |
| Packet Transmission Control |  |  |  |
| Enable Generation of Telemetry Packets | 14 | 1 |  |
| Disable Generation of Telemetry Packets | 14 | 2 |  |
| Report Enabled Telemetry Packets | 14 | 3 |  |
| On-Board Storage and Retrieval | 15 |  | Not Used |
| On-Board Traffic Management | 16 |  | Not Used |
| Test Service |  |  |  |
| Perform Connection Test | 17 | 1 |  |
| On-Board Control procedures | 18 |  | Not Used |
| Action/Event Service | 19 |  | Not Used |
| Information Distribution Service |  |  |  |
| Enable Distribution of Information TC Packets | 20 | 1 |  |
| Disable Distribution of Information TC Packets | 20 | 2 |  |
| Report Distributed Information Packets | 20 | 3 |  |
| Information Telecommand | 20 | 4 |  |
| Science Data | 21 |  | N/A |
| Context Saving Service |  |  | TBD |
| Report Context | 22 | 1 (TBC) |  |

## Telecommand Packet Definition

### Telecommand Verification Service

Not Applicable

### Device Command Distribution

Not Used

### Housekeeping and Diagnostic Data Reporting

The periodic Housekeeping and Diagnostic Data is reported, at fixed time intervals, in packets with a fixed format.  A maximum of 4 different reports are available for use at the same time, each allocated a different HK_Packet_ID, in the range 0 to 3. Associated with each HK_Packet_ID is a sampling interval and an on-board table, which contains the definition of the contents of the report.

The 4 reports are defined in the following way in the OBS by default:

| HK_Packet_ID | Report | Table_ID | Sampling Interval | HKSID |
| --- | --- | --- | --- | --- |
| 0 | Critical Housekeeping Report | 0 | 2000m sec | 0x0300 |
| 1 | Nominal Housekeeping Report | 1 | 1000 msec | 0x0301 |
| 2 | Detector Housekeeping Report | 2 | ---- Not sampled ---- | 0x0302 |
| 3 | --- Undefined ---- | 3 | ---- Not sampled ---- | --- undefined--- |

The contents of each report are described in the Housekeeping Parameter Report (Services 3,25).

#### Define New Housekeeping Parameter Report (Service 3,1)

This command is used to define/redefine housekeeping reports

In AD01 the maximum length of a report defined in this way is restricted by the maximum length of a telecommand packet to 115 parameters and it is currently impossible to define a new HK report longer than this!  Therefore a different mechanism is used in the OBS to define the reports:

- An on-board table is loaded with the definition of the report using one or more 'Load Table' commands (see section 3.2.8.3). The format of the table contents is given in RD08.

- A 'Define New Housekeeping Parameter Report' is used to associate the table containing the report definition with the HK_Packet_ID, the sampling interval and a structure ID (SID). In this case:

- The Sampling Interval is set in msec

- A single parameter is defined, which contains the Table_ID of the on-board table containing the report definition

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 15 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| HKPCKTID |
| HKSID |
| HKINTERVAL |
| TABLEID |
| MONITOR_TABLEID |
| Checksum |

Note:  In order to redefine a report, the report must first be ‘cleared’ using service (3,3), which has the effect of stopping report generation in order that the table containing the definition of the report contents may be modified. Report generation starts with the execution of this service.

This definition assumes that only one housekeeping report (frame) will be sent in each packet. In the event that we need more than one (e.g. to increase the efficiency of reporting critical housekeeping) we will need to add a second parameter to this command to specify the number of hsk frames per packet.

An on board table will contain in TBD format the information about parameters to be monitored; since monitoring is performed on parameters sampled as part of the nominal HK packet, a monitoring table_ID must be communicated to the OBS each time the HK table ID is updated. The details are TBD.

##### Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_HK_Packet_ID | (1,8) | 0x0301 | HK_Packet_ID Out of Range |
| Illegal_HK_SID | (1,8) | 0x0302 | HK_SID MSB not 0x03 |
| Illegal_Table_ID | (1,8) | 0x0303 | TABLEID is out of rangeof possible table numbers |
| Illegal_HK_Sampling Interval | (1,8) | 0x0304 | HK_Sampling Interval is less than the absolute limit (10, TBC) |
| Undefined_HK_Table | (1,8) | 0x0306 | Table given by TABLEID or MONITOR_TABLEID is not defined |
| Report_in_Use | (1,8) | 0x0307 | An attempt is being made to modify a report without first ‘clearing’ it. |

#### Define New Diagnostic Parameter Report (Service 3,2)

Not Used

#### Clear Housekeeping Parameter Report Definitions (Service 3,3)

This command is used to clear the definition of a housekeeping report. It is usually used prior to redefining the report with service (3,1) and has the affect of stopping report generation within the OBS.

This command should not be used for disabling reporting (use service (14,2) for this) as the definition of the report contents and sampling interval is lost.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 9 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| HKPCKTID |
| Checksum |

##### Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_HK_Packet_ID | (1,8) | 0x0301 | HK_Packet_ID Out of Range |

#### Clear Diagnostic Parameter Report Definitions (Service 3,4)

Not Used

#### Report Housekeeping Parameter Report Definitions (Service 3,9)

This command is used to report the definition of a housekeeping report.

The information requested by this command is reported in TM service (3,10).  This contains only the information provided in the report definition telecommand (service (3,1)). The actual definition of the report contents may be obtained by dumping the appropriate table.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 9 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 |
| 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| HKPCKTID |
| Checksum |

##### Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_HK_Packet_ID | (1,8) | 0x0301 | HK_Packet_ID Out of Range |

#### Report Diagnostic Parameter Report Definitions (Service 3,11)

Not Used

Not Available

### Event Reporting

Not Applicable

### Memory Management

These commands allow access to the contents of memory locations.  Four areas of memory have been identified, with the following MEMORYID:

- Program Memory (RAM)

- Data memory  (RAM)

- 1553 I/F DPRAM

- EEPROM

When accessing memory the address is specified in terms of the number of Single Addressable Units (SAUs)  from the start of the memory area. The SPIRE SAU is always 16bit words.

#### Load Memory Using Absolute Addresses (Service 6,2)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| MEMORYID |  |
| STARTADDR |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | NSAU |
| Data |
| CRC |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Memory_ID | (1,8) | 0x0601 | Not a valid memory ID |
| Illegal_Start_Address | (1,8) | 0x0602 | Start address outside valid range |
| Illegal_Nsau | (1,8) | 0x0603 | Length will place data outside valid address range |
| Bad_Nsau | (1,8) | 0x0604 | Length does not agree with size of data field |
| Bad_CRC | (1,8) | 0x0605 | CRC in command does not agree with calculated checksum over data in command |
| Bad_Load | (1,8) | 0x0606 | CRC in command does not agree with calculated checksum over data in memory |

#### Dump Memory Using Absolute Addresses (Service 6,5)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 11 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| MEMORYID |  |
| STARTADDR |
| NSAU |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Memory_ID | (1,8) | 0x0601 | Not a valid memory ID |
| Illegal_Start_Address | (1,8) | 0x0602 | Start address outside valid range |
| Illegal_Nsau | (1,8) | 0x0603 | Length will read data outside valid address range |

#### Check Memory Using Absolute Addresses (Service 6,9)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 11 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| MEMORYID |  |
| STARTADDR |
| NSAU |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Memory_ID | (1,8) | 0x0601 | Not a valid memory ID |
| Illegal_Start_Address | (1,8) | 0x0602 | Start address outside valid range |
| Illegal_Nsau | (1,8) | 0x0603 | Length will read data outside valid address range |

Not Available

### Function Management

All command packets of this type may give rise to the following Errors:

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Function_ID | (1,2) | 0x0801 | Function_ID not known |

#### Start Function (Service 8,1)

These commands allow functions that are not started automatically when the OBS runs to be started.

##### Function 0x01: Table Management

Always available - cannot be started

##### Function 0x02: Interrupt-driven VM

Always Available - cannot be started

##### Function 0x03: Software-driven VM1

Always Available - cannot be started

##### Function 0x04: Software-driven VM2

Always Available - cannot be started

##### Function 0x05: Software-driven VM3

Always Available - cannot be started

##### Function 0xC0: Operations

Always Available - cannot be started

##### Function 0xC1: Observations

Always Available - cannot be started

##### Function 0xCA: DPU

Always Available - cannot be started

#### Stop Function (Service 8,2)

##### Function 0x01: Table Management

Always available - cannot be stopped

##### Function 0x02: Interrupt-driven VM

Always Available - cannot be stopped

##### Function 0x03: Software-driven VM1

Always Available - cannot be stopped

##### Function 0x04: Software-driven VM2

Always Available - cannot be stopped

##### Function 0x05: Software-driven VM3

Always Available - cannot be stopped

##### Function 0xC0: Operations

Always Available - cannot be stopped

##### Function 0xC1: Observations

Always Available - cannot be stopped

##### Function 0xCA: DPU

Always Available - cannot be stopped

#### Perform Activity of a Function (Service 8,4)

All command packets of this type and subtype may give rise to the following Errors:

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Function_ID | (1,2) | 0x0801 | Function_ID not known |
| Illegal_Activity_ID | (1,2) | 0x0802 | Activity_ID not known |

##### Function 0x01 Table Management, Activity 0x01: Set Table

This telecommand allocates space for a table. The previous space allocated to the table is lost.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length =11 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| TABLEID |
| TABLESIZE |
| Checksum |

Possible Errors

This telecommand shall be expected to issue a Telecommand Execution Report  - Completed, TM(1,7) on successful completion. In the event of being unable to complete the command the following reports shall be issued:

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Table_Space_Full | (1,8) | 0x0809 | Not enough space available in DPU memory to create new table |

##### Function 0x01 Table Management, Activity 0x02: Report Table

Requests telemetry packets reporting the content s of a table starting with the word at offset INDEX (32 bit words) from the start of the table and reporting COUNT words. The contents are reported in as many TM(21,4) packets as necessary to complete the request.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 13 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | AC TIVITYID |
| TABLEID |
| INDEX |
| COUNT |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Illegal_Table_Index | (1,8) | 0x0806 | Index outside valid range |
| Table_Bounds_Error | (1,8) | 0x080d | Not enough data available in the table to report COUNT words |
| Undefined_Table_Error | (1,8) | 0x0811 | Table is not defined |

##### Function 0x01 Table Management, Activity 0x03: Update Table

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| TABLEID |
| INDEX |
| N |
| Data |
| Checksum |

Possible Errors

This telecommand shall be expected to issue a Telecommand Execution Report  - Completed, TM(1,7) on successful completion. In the event of being unable to complete the command the following reports shall be issued:

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Illegal_Table_Index | (1,8) | 0x0806 | Index outside valid range |
| Table_Bounds_error | (1,8) | 0x080d | Not enough space available in table to write new data |
| Bad_Data | (1,8) | 0x0808 | N does not agree with length of data field |
| Undefined_Table_Error | (1,8) | 0x0811 | Table is not defined |

##### Function 0x02 Interrupt-driven VM, Activity 0x01: Execute Command List

Execute the command list held in the telecommand packet.

The Command list being executed is responsible for issuing all the necessary telecommand execution reports

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| N |
| Data |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Bad_Data | (1,8) | 0x0808 | N does not agree with length of data field |
| VM_Active | (1,8) | 0x080c | The VM is already executing a command list |

##### Function 0x02 Interrupt-driven VM, Activity 0x02: Run VM

Execute a command list held in a table by the Interrupt-driven VM. If the VM is already active the telecommand is rejected with an error packet (see below).

The Command list being executed is responsible for issuing all the necessary telecommand execution reports

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| TABLEID |
| INDEX |
| N |
| Data |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Illegal_Table_Index | (1,8) | 0x0806 | Index outside valid range |
| Bad_Data | (1,8) | 0x0808 | N does not agree with length of data field |
| VM_Active | (1,8) | 0x080c | The VM is already executing a command list |
| VM_Undefined_Table_Error | (1,8) | 0x0810 | Table is not defined |

##### Function 0x02 Interrupt-driven VM, Activity 0x03: Halt VM

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| VM_Inactive | (1,8) | 0x080a | Command List Not Running |

##### Function 0x03 Software-driven VM1, Activity 0x02: Run VM1

Start execution of a command list held in a table by the software-driven VM1. If the VM is already active the telecommand is rejected with an error packet (see below).

The Command list being executed is responsible for issuing all the necessary telecommand execution reports

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| TABLEID |
| INDEX |
| N |
| Data |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Illegal_Table_Index | (1,8) | 0x0806 | Index outside valid range |
| Bad_Data | (1,8) | 0x0808 | N does not agree with length of data field |
| VM_Active | (1,8) | 0x080b | The VM is already executing a command list |
| Undefined_Table_Error | (1,8) | 0x080d | Table is not defined |

##### Function 0x03 Software-driven VM1, Activity 0x03: Halt VM1

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| VM_Inactive | (1,8) | 0x080a | VM Not Running |

##### Function 0x04 Software-driven VM2, Activity 0x02: Run VM2

Start execution of a command list held in a table by the software-driven VM2. If the VM is already active the telecommand is rejected with an error packet (see below).

The Command list being executed is responsible for issuing all the necessary telecommand execution reports

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| TABLEID |
| INDEX |
| N |
| Data |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Illegal_Table_Index | (1,8) | 0x0806 | Index outside valid range |
| Bad_Data | (1,8) | 0x0808 | N does not agree with length of data field |
| VM_Active | (1,8) | 0x080b | The VM is already executing a command list |
| Undefined_Table_Error | (1,8) | 0x080d | Table is not defined |

##### Function 0x04 Software-driven VM2, Activity 0x03: Halt VM2

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| VM_Inactive | (1,8) | 0x080a | VM Not Running |

##### Function 0x05 Software-driven VM3, Activity 0x02: Run VM3

Start execution of a command list held in a table by the software-driven VM3. If the VM is already active the telecommand is rejected with an error packet (see below).

The Command list being executed is responsible for issuing all the necessary telecommand execution reports

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| TABLEID |
| INDEX |
| N |
| Data |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |
| Illegal_Table_Index | (1,8) | 0x0806 | Index outside valid range |
| Bad_Data | (1,8) | 0x0808 | N does not agree with length of data field |
| VM_Active | (1,8) | 0x080b | The VM is already executing a command list |
| Undefined_Table_Error | (1,8) | 0x080d | Table is not defined |

##### Function 0x05 Software-driven VM3, Activity 0x03: Halt VM3

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| VM_Inactive | (1,8) | 0x080a | VM Not Running |

##### Function 0xC0 Operations, Activity 0x40: Peak Up

This is a place holder for a possible command, which is still TBD

Possible parameters to this command are TBD but could include definition of the Photometer array to use to find the peak and the pixel defined to be the centre of the array.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Parameters |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| No_Peak | (1,8) | 0x080c | Unable to find Peak |

##### Function 0xC1 Observations, Activity 0x01: Set Observation ID

Set the observation ID, and set the BBID to zero. The value of these two parameters is reported in the instrument housekeeping packets and all science data packets.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 11 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| OBSID |
| Checksum |

##### Function 0xC1 Observations, Activity 0x02: Set Building Block ID

Set the observation ID. This is reported in the instrument housekeeping packets and all science data packets.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 11 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| BBID |
| Checksum |

##### Function 0xC1 Observations, Activity 0x03: Set Observing Mode

This value may be set by the command system in order to notify processing software of the observing mode. It is reported in the housekeeping packets

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 19 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| MODE |
| Checksum |

##### Function 0xC1 Observations, Activity 0x04: Set Observation Step

This value may be set by the command system in order to notify processing software of the current step in the  observing mode. It is reported in the housekeeping packets

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 19 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| STEP |
| Checksum |

##### Function 0xCA DPU, Activity 0x01: Synchronise DRCU Counters

This telecommand sends a broadcast command to the DRCU to reset the time counters in the DCU, SCU and MCU interfaces

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

This telecommand shall be expected to issue a Telecommand Execution Report  - Completed, TM(1,7) on successful completion.

##### Function 0xCA DPU, Activity 0x02: Flush

This telecommand forces the DPU to read any remaining  science data from the FIFOs place them into science data packets and transmit the packets. I.e. it completely  empties the DPU of science data.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 9 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| FIFOFLAGS |
| Checksum |

This telecommand shall be expected to issue a Telecommand Execution Report  - Completed, TM(1,7) on successful completion.

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Invalid_FIFOFLAGS | (1,8) | 0x080f | Incorrect FIFOFLAGS (TBD) |

##### Function 0xCA DPU, Activity 0x03: Set_TM_Nominal_Mode

This telecommand sets the telemetry interface to the CDMS to operate in Nominal Mode.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

This telecommand shall be expected to issue a Telecommand Execution Report  - Completed, TM(1,7) on successful completion.

##### Function 0xCA DPU, Activity 0x04: Set_TM_Burst_Mode

This telecommand sets the telemetry interface to the CDMS to operate in Burst Mode.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

This telecommand shall be expected to issue a Telecommand Execution Report  - Completed, TM(1,7) on successful completion.

##### Function 0xCA DPU, Activity 0x05: Send_DRCU_Command

This telecommand allows sending a single command to the DRCU.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 11 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| DRCU Command MSB |
| DRCU Command LSB |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| LS_Transmit_Error | (1,8) | 0x080e | Command was not correctly received by DRCU |

##### Function 0xCA DPU, Activity 0x06: Reset_FIFOs

This telecommand is used to reset all FIFOs. All data currently in the FIFOs is lost.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

##### Function 0xCA DPU, Activity 0x07: Write2EEPROM

This telecommand is used to copy PM into EEPROM

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 15 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Start_Address MSB |
| Start_Address LSB |
| End_Address MSB |
| End_Address LSB |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| EEPROM_Failed | (1,8) | 0x0812 | PM to EEPROM copy failed |

##### Function 0xCA DPU, Activity 0x08: Force_Boot

This telecommand is used  to call the DPU Boot procedure

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

##### Function 0xCA DPU, Activity 0x09: DPU_Reset

This telecommand  performs a warm reset of the DPU

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| Checksum |

These telecommands are defined here as a place holder for functionality to select individual data words from science data frames in order to reduce data rate. Their format is still TBC

##### Function 0xCA DPU, Activity 0x10: Enable_Selection

Enable data selection from a science data frame. The DPU will use the contents of a table to define which words are extracted from the data frame into the telemetry packet.

It shall be possible to enable extraction from any or all science frames in order to reduce the instrument TM data rate

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 19 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| FRAMEID |
| TABLEID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Frame_ID | (1,8) | 0x0810 | Not a valid frame ID |
| Illegal_Table_ID | (1,8) | 0x0805 | Not a valid table ID |

##### Function 0xCA DPU, Activity 0x11: Disable_Selection

Disable selection for a science data frame. The complete data frame will now be placed into the science telemetry packets

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 19 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | ACTIVITYID |
| FRAMEID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Frame_ID | (1,8) | 0x0810 | Not a valid frame ID |

#### Report Function Status (Service 8,5)

This telecommand requests a report on the status of a function. The instrument should respond with a TM(8,6) report.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 7 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| FUNCTIONID | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Checksum |

### Time Management

#### Enable Time Verification (Service 9,7)

On receipt of this telecommand the instrument should issue a Time Verification Report, TM(9,9).

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 0 | 0 | 0 | Count |
| Length = 5 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Checksum |

Not Available

### On-Board Scheduling

Not Used

### On-Board Monitoring

TBW

Not Available

### Packet Transmission Control

#### Enable Generation of Telemetry Packets (Service 14,1)

This telecommand enables the transmission of instrument telemetry packets. By default all packets are enabled for transmission when the OBS starts. This telecommand is therefore only needed to restart transmission after transmission has been disabled.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| NPCKTS |
| TYPE | SUBTYPE |
| SID |
| Repeat Type, Subtype and SID NPCKTS-1 times |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Type | (1,8) | 0x0e01 | Not a valid packet type for this APID |
| Illegal_Subtype | (1,8) | 0x0e02 | Not a valid subtype for packet with type TYPE |
| Illegal_SID | (1,8) | 0x0e03 | Not at valid SID for packet with type TYPE and subtype SUBTYPE |
| Bad_NPCKTS | (1.8) | 0x0e04 | Number of packets to be enabled is not correct for telecommand length |

#### Disable Generation of Telemetry Packets (Service 14,2)

This telecommand disables the transmission of instrument telemetry packets. By default all packets are enabled for transmission when the OBS starts.

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| NPCKTS |
| TYPE | SUBTYPE |
| SID |
| Repeat Type, Subtype and SID NPCKTS-1 times |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Type | (1,8) | 0x0E01 | Not a valid packet type for this APID |
| Illegal_Subtype | (1,8) | 0x0E02 | Not a valid subtype for packet with type TYPE |
| Illegal_SID | (1,8) | 0x0E03 | Not at valid SID for packet with type TYPE and subtype SUBTYPE |
| Bad_NPCKTS | (1.8) | 0x0E04 | Number of packets to be enabled is not correct for telecommand length |

#### Report Enabled Telemetry Packets (Service 14,3)

This telecommand requests a list of telemetry packet types and subtypes that are enabled for transmission

The list is reported in TM service  (14,4)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Src | Count |
| Length = 5 |
| 0 | 0 | 0 | 0 | Ack | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Checksum |

### On-Board Storage and Retrieval

Not Used

### On-Board Traffic Management

Not Used

### Test Service

#### Perform Connection Test (Service 17,1)

This test is used to check the End-to-End connection to the instrument. The instrument should respond with a Successful Command Acceptance Report, TM(1,1),  followed by a Link Connection Report, TM(17,2).

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 0 | 0 | 0 | Count |
| Length = 5 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Checksum |

### On-Board Control Procedures

Not Used

### Action/Event Service

Not Used

### Information Distribution Service

#### Enable Distribution of Information TC Packets (Service 20,1)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 0 | 0 | 0 | Count |
| Length = 9 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| APID |
| SID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Packet | (1,8) | 0x1401 | Not a valid packet for this APID and SID |

#### Disable Distribution of Information TC Packets (Service 20,2)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 0 | 0 | 0 | Count |
| Length = 9 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| APID |
| SID |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Packet | (1,8) | 0x1401 | Not a valid packet for this APID and SID |

#### Report Distributed Information Packets (Service 20,3)

| 0 | 0 | 0 | 1 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 0 | 0 | 0 | Count |
| Length = 5 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Checksum |

Possible Errors

| Error | TM Service | Error Code | Description |
| --- | --- | --- | --- |
| Illegal_Packet | (1,8) | 0x1401 | Not a valid packet for this APID and SID |

#### Information Telecommand (Service 20,4)

Not used

### Science Data

Not Applicable

### Context Saving Service

Not Used

# SPIRE Telemetry Packets

This section defines all the TM packets that will be produced by SPIRE.

## Telemetry Packet Types

The Packet Structure ICD (AD1) defines many types of service that can be provided by an Application. The following table shows the telemetry packet types that will be produced by the SPIRE.

| Description | Service Type | Service Sub-Type | Comments |
| --- | --- | --- | --- |
| Telecommand Verification Service |  |  |  |
| Telecommand Acceptance Report - Success | 1 | 1 |  |
| Telecommand Acceptance Report - Failure | 1 | 2 |  |
| Telecommand Execution Report - Started | 1 | 3 |  |
| Telecommand Execution Report - Progress | 1 | 5 |  |
| Telecommand Execution Report - Completed | 1 | 7 |  |
| Telecommand Execution Report - Failure | 1 | 8 |  |
| Telecommand Contents Report | 1 | 9 | Not Used |
| Device Command Distribution | 2 |  | N/A |
| Housekeeping and Diagnostic Data Reporting |  |  |  |
| HK Parameter Report Definitions Report | 3 | 10 |  |
| Diagnostic Parameter Definitions Report | 3 | 12 |  |
| Housekeeping Parameter Report | 3 | 25 |  |
| Diagnostic Parameter Report | 3 | 26 |  |
| Event Reporting |  |  |  |
| Event Report | 5 | 1 |  |
| Exception Report | 5 | 2 |  |
| Error/Alarm Report | 5 | 4 | Not Used |
| Memory Management |  |  |  |
| Memory Dump, Absolute Addresses | 6 | 6 |  |
| Memory Check Report, Absolute Addresses | 6 | 10 |  |
| Function Management |  |  |  |
| Function Status Report | 8 | 6 |  |
| Time Management |  |  |  |
| Central Time Reference | 9 | 8 | Not Used |
| Time Verification Report | 9 | 9 |  |
| On-Board Scheduling | 11 |  | Not Used |
| On-Board Monitoring |  |  |  |
| Current Monitoring List Report | 12 | 9 | TBD |
| Packet Transmission Control |  |  |  |
| Enabled Telemetry Packets Report | 14 | 4 |  |
| On-Board Storage and Retrieval | 15 |  | Not Used |
| Test Service |  |  |  |
| Connection Test Report | 17 | 2 |  |
| On-Board Control procedures | 18 |  | Not Used |
| Action/Event Service | 19 |  | Not Used |
| Information Distribution Service |  |  |  |
| Distributed Information Packets Report | 20 | 5 |  |
| Science Data |  |  |  |
| Nominal Science Data Report | 21 | 1 |  |
| Science Type B Data Report | 21 | 2 |  |
| Diagnostic Science Data Report | 21 | 3 |  |
| Auxiliary Science Data Report | 21 | 4 |  |
| Context Saving Service | 22 |  | Not Used |

## Telemetry Packet definitions

### TC Verification Service

#### Telecommand Acceptance Report - Success (1,1)

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 15 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| TC_Packet_Sequence_Control |
| Checksum |

#### Telecommand Acceptance Report - Failure (1,2)

The structure of this packet depends on the type of error found.

##### Packet Control Errors

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 19 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| TC_Packet_Sequence_Control |
| Failure Code |
| Parameter |
| Checksum |

##### Packet Content Error

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length=57 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| Packet_Sequence_Control |
| Failure Code |
| Parameters |
| Checksum |

#### Telecommand Execution Report - Started (1,3)

This report is generated at the start of execution of a telecommand. It is only transmitted in telemetry if the corresponding bit of the Ack field is set.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 15 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| TC_Packet_Sequence_Control |
| Checksum |

#### Telecommand Execution Report - Progress (1,5)

These reports may be generated by an executing telecommand. They are only transmitted in telemetry if the corresponding bit of the Ack filed is set.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 17 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| TC_Packet_Sequence_Control |
| Step_Number |
| Checksum |

#### Telecommand Execution Report - Completion (1,7)

This report is generated at the successful completion of execution of a telecommand. It is only transmitted in telemetry if the corresponding bit of the Ack field is set.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 15 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| TC_Packet_Sequence_Control |
| Checksum |

#### Telecommand Execution Report - Failure (1,8)

This report is generated at the unsuccessful completion of execution of a telecommand

The structure of this packet depends on the type of error found., but follows the general form:

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TC_Packet_ID |
| TC_Packet_Sequence_Control |
| Failure Code |
| Parameters |
| Checksum |

The following error Failure Codes have been defined with the parameters to be provided in the TM packet:

| Error | Failure Code | Parameters |
| --- | --- | --- |
| Illegal_HK_Packet_ID | 0x0301 | HK_Packet_ID |
| Illegal_HK_SID | 0x0302 | HK_SID |
| Illegal_Table_ID | 0x0303 | HK_SID |
| Illegal_HK_Sampling Interval | 0x0304 | HK_Sampling_Interval |
| Bad_Ndata | 0x0305 | Length Ndata |
| Undefined_HK_Table | 0x0306 | HK_SID |
| Report_in_Use | 0x0307 | HK_Packet_ID |
| Illegal_Monitoring_TableID | 0x0308 | Monitoring_TableID |
| Undefined_Monitoring_TableID | 0x0309 | Monitoring_TableID |
| Illegal_Memory_ID | 0x0601 | Memory_ID |
| Illegal_Start_Address | 0x0602 | Memory_ID Start Adress |
| Illegal_Nsau | 0x0603 | Memory_ID Start_Address Nsau |
| Bad_Nsau | 0x0604 | Length Nsau |
| Bad_CRC | 0x0605 | CRC Calculated Checksum |
| Bad_Load | 0x0606 | CRC Calculated Checksum |
| Illegal_Function_ID | 0x0801 | FunctionID |
| Illegal_Activity_ID | 0x0802 | ActivityID |
| Illegal_Table_ID | 0x0805 | Table_ID |
| Illegal_Table_Index | 0x0806 | Table Index |
| Tabel_Bounds_Error | 0x080d | TableID INDEX COUNT TABLESIZE |
| Bad_Data | 0x0808 | Length N |
| Table_Space_Full | 0x0809 | Table_ID Ndata |
| VM_Inactive | 0x080a | ------- |
| VM_Active | 0x080c | ------- |
| No_Peak | 0x080b | TBD |
| Undefined_Table_ID | 0x0811 | Table_ID |
| LS_Transmission_Error | 0x080e | ------- |
| Illegal_FIFOFlags | 0x080f | FIFOFLAGS |
| VM_Undefined_Table_ID | 0x0810 | Table_ID |
| EEPROM_Failed | 0x0812 | ------- |
| Illegal_Type | 0x0e01 | APID Tyep Subtype SID |
| Illegal_Subtype | 0x0e02 | APID Tyep Subtype SID |
| Illegal_SID | 0x0e03 | APID Tyep Subtype SID |
| Bad_NPCKTS | 0x0e04 | Length NPCKTS |
| Illegal_Packet | 0x1401 | APID SID |

### Device Command Distribution

Not Used

### Housekeeping and Diagnostic Data Reporting

#### Housekeeping Parameter Report  (Service 3,25)

The general packet structure is shown below. The Structure ID identifies the housekeeping packet type.

| 0 | 0 | 0 | 0 | 1 | APID2 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 |
| 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID |
| Parameters |
| Checksum |

#### Critical Housekeeping Report (SID=0x0300)

| Location (msb) | Field Length (bits) | Parameter Name | Ref. | Description |
| --- | --- | --- | --- | --- |
| 18 | 0 | 32 | OBSID | AD03 | Observation ID |
| 22 | 0 | 32 | BBID | AD03 | Building Block ID |
| 26 | 0 | 16 | MODE |  | Observing Mode – set by command |
| 28 | 0 | 16 | STEP |  | Observation Step – set by command |
| 30 | 0 | 16 | TCRECV |  | Count of Telecommands Received |
| 32 | 0 | 16 | TCEXEC |  | Count of Telecommands Executed |
| 34 | 0 | 48 | MEMSTAT |  | Memory Check Flags – updated by background memory checking task |
| 40 | 0 | 16 | MONSTAT |  | Status of subsystems monitoring |
| 42 | 0 | 16 | EVENTSTAT |  | Status of events |
| 44 | 0 0 1 2 3 4 | 16 1 1 1 1 2 | DRCUIFSTAT IFCMDOLAPERR IFBCASTERR IFREADERR IFTIMEOUTERR IFCMDSTAT | RD06 RD06 RD06 RD06 RD06 RD06 | CmdIfStat Command Overlapped Forbidden Broadcast Forbidden Read SubSystem Timeout LastCmdStatus |
| 46 |  |  | DCUSTAT |  | TBC |
| 48 | 0 0 1 2 3 4 5 | 16 1 1 1 1 1 1 | PSWJFETSTAT PSWJFETPWR1 PSWJFETPWR2 PSWJFETPWR3 PSWJFETPWR4 PSWJFETPWR5 PSWJFETPWR6 | RD06 RD06 RD06 RD06 RD06 RD06 | P250_JFET_1 P250_JFET_2 P250_JFET_3 P250_JFET_4 P250_JFET_5 P250_JFET_6 |
| 50 | 0 0 1 2 3 4 5 | 16 1 1 1 1 1 1 | PMLWJFETSTAT PMWJFETPWR1 PMWJFETPWR2 PMWJFETPWR3 PMWJFETPWR4 PLWJFETPWR1 PLWJFETPWR2 | RD06 RD06 RD06 RD06 RD06 RD06 | P350_JFET_1 P350_JFET_2 P350_JFET_3 P350_JFET_4 P500_JFET_1 P500_JFET_2 |
| 52 | 0 0 1 2 | 16 1 1 1 | SPECJFETSTAT SLWJFETPWR1 SSWJFETPWR1 SSWJFETPWR2 | RD06 | SLW_JFET1 SSW_JFET1 SSW_JFET2 |
| 54 | 0 0 1 2 3 4 5 6 7 8 9 10 11 | 16 1 1 1 1 1 1 1 1 1 1 1 1 | LIASTAT LIA01STAT LIA02STAT LIA03STAT LIA04STAT LIA05STAT LIA06STAT LIA07STAT LIA08STAT LIA09STAT LIA10STAT LIA11STAT LIA12STAT | RD06 | PWR_STATUS, LIA 1 to LIA 12 +5V/+9V/-9V status LIA 1 +5V/+9V/-9V status LIA 2 +5V/+9V/-9V status LIA 3 +5V/+9V/-9V status LIA 4 +5V/+9V/-9V status LIA 5 +5V/+9V/-9V status LIA 6 +5V/+9V/-9V status LIA 7 +5V/+9V/-9V status LIA 8 +5V/+9V/-9V status LIA 9 +5V/+9V/-9V status LIA 10 +5V/+9V/-9V status LIA 11 +5V/+9V/-9V status LIA 12 +5V/+9V/-9V status |
| 56 | 0 | 16 | MCUERR | RD06 | ErrorCode |
| 58 | 0 | 16 | SMECSTAT | RD06 | SMECStatus |
| 60 | 0 | 16 | CHOPSTAT | RD06 | ChopStatus |
| 62 | 0 | 16 | JIGGSTAT | RD06 | JigStatus |
| 64 | 0 | 16 | SCUSTAT | RD06 | ScuStatus |
| 66 | 0 | 16 | SUBKTEMP | RD06 | SubKTempP |

#### Nominal Housekeeping Report (SID=0x0301)

The following table lists the field to be found in this report

| Location (msb) | Field Length (bits) | Parameter Name | Ref. | Description |
| --- | --- | --- | --- | --- |
| octet | bit |  |  |  |  |
| 18 | 0 | 32 | OBSID | AD03 | Observation ID |
| 22 | 0 | 32 | BBID | AD03 | Building Block ID |
| 26 | 0 | 16 | MODE |  | Observing Mode – set by command |
| 28 | 0 | 16 | STEP |  | Observation Step – set by command |
| 30 | 0 | 48 | THSK |  | Time of start of housekeeping collection |
| 36 | 0 | 48 | TSYNC |  | Time of last sync cmd to DRCU |
| 42 | 0 | 16 | TCRECV |  | Count of Telecommands Received |
| 44 | 0 | 16 | TCRECN | AD01 | Packet Sequence Control, Sequence count of latest received telecommand |
| 46 | 0 | 16 | TCEXEC |  | Count of Telecommands Executed |
| 48 | 0 | 16 | TCEXEN | AD01 | Packet Sequence Control, Sequence count of latest executed telecommand |
| 50 | 0 | 16 | TM1N | AD01 | Packet Sequence Control, Sequence count of TM Packets for APID1 |
| 52 | 0 | 16 | TM2N | AD01 | Packet Sequence Control, Sequence count of TM Packets for APID2 |
| 54 | 0 | 16 | TM3N | AD01 | Packet Sequence Control, Sequence count of TM Packets for APID3 |
| 56 | 0 | 16 | TM4N | AD01 | Packet Sequence Control, Sequence count of TM Packets for APID4 |
| 58 | 0 | 16 | TM5N | AD01 | Packet Sequence Control, Sequence count of TM Packets for APID5 |
| 60 | 0 | 16 |  |  |  |
| 62 | 0 | 16 |  |  |  |
| 64 | 0 | 16 |  |  |  |
| 66 | 0 | 16 |  |  |  |
| 68 | 0 | 32 | OBSLEN |  | Seconds since last set_OBSID command |
| 72 | 0 | 32 | RATE |  | Average number of bits per second in current observation |
| 76 | 0 | 48 | MEMSTAT |  | Memory Check Flags – updated by background memory checking task |
| 82 | 0 | 16 | MONSTAT |  | Status of subsystems monitoring |
| 84 | 0 | 16 | EVENTSTAT |  | Status of events |
| 86 | 0 | 16 |  |  |  |
| 88 | 0 | 16 |  |  |  |
| 90 | 0 | 16 |  |  |  |
| 92 | 0 | 16 |  |  |  |
| 94 | 0 | 16 |  | RD09 | DPU Spare |
| 96 | 0 | 16 | DPU+5V | RD09 | DPU monitored +5V voltage line |
| 98 | 0 | 16 | DPU+15V | RD09 | DPU monitored +15V voltage line |
| 100 | 0 | 16 | DPU-15V | RD09 | DPU monitored –15V voltageline |
| 102 | 0 | 16 | DPUTEMP | RD09 | DPU monitored temperature |
| 104 |  |  |  | RD09 | DPU Spare |
| 106 |  |  |  | RD09 | DPU Spare |
| 108 | 0 | 16 | DPU+2.5V | RD09 | DPU monitored +2.5V Reference voltage line |
| 110 | 0 | 16 |  |  |  |
| 112 | 0 | 16 |  |  |  |
| 114 | 0 0 1 2 3 4 | 16 1 1 1 1 2 | DRCUIFSTAT IFCMDOLAPERR IFBCASTERR IFREADERR IFTIMEOUTERR IFCMDSTAT | RD06 RD06 RD06 RD06 RD06 RD06 | CmdIfStat Command Overlapped Forbidden Broadcast Forbidden Read SubSystem Timeout LastCmdStatus |
| 116 | 0 0 1 2 | 16 1 1 1 | DRCUIFCTRL IFDATARESET IFSSRESET IFSTATRESET | RD06 RD06 RD06 RD06 | CmdIfCtrl DataIfReset SubSystemRst St RD06atusRst |
| 118 | 0 | 16 | DRCUSSDEL | RD06 | SubSDelay |
| 120 | 0 | 16 |  |  |  |
| 122 | 0 | 16 | DCUSTAT |  | TBC |
| 124 | 0 | 16 | DCUDATAMODE | RD06 | DataMode |
| 126 | 0 | 16 | DCUDATAFRMS | RD06 | FrameCount |
| 128 | 0 | 16 | DCUDATASTAT | RD06 | StartFrame |
| 122 | 0 | 16 |  |  |  |
| 122 | 0 | 16 |  |  |  |
| 124 | 0 | 16 | PHOTSAMPFREQ | RD06 | PhotoBiasDiv |
| 126 | 0 | 16 | PHOTBIASMODE | RD06 | PhotoBiasMode |
| 128 | 0 | 16 | PHOTBIASFREQ | RD06 | PhotoMClkDiv |
| 130 | 0 | 16 | PSWBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 0 |
| 132 | 0 | 16 | PMWBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 1 |
| 134 | 0 | 16 | PLWBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 2 |
| 136 | 0 | 16 | TCBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 3 |
| 138 | 0 | 16 | PSWPHASE | RD06 | PhaseShift for Photo Bias Channel 0 |
| 140 | 0 | 16 | PMWPHASE | RD06 | PhaseShift for Photo Bias Channel 1 |
| 142 | 0 | 16 | PLWPHASE | RD06 | PhaseShift for Photo Bias Channel 2 |
| 144 | 0 | 16 | TCPHASE | RD06 | PhaseShift for Photo Bias Channel 3 |
| 146 | 0 0 1 2 3 4 5 | 16 1 1 1 1 1 1 | PSWJFETSTAT PSWJFETPWR1 PSWJFETPWR2 PSWJFETPWR3 PSWJFETPWR4 PSWJFETPWR5 PSWJFETPWR6 | RD06 RD06 RD06 RD06 RD06 RD06 | P250_JFET_1 P250_JFET_2 P250_JFET_3 P250_JFET_4 P250_JFET_5 P250_JFET_6 |
| 148 | 0 0 1 2 3 4 5 | 16 1 1 1 1 1 1 | PMLWJFETSTAT PMWJFETPWR1 PMWJFETPWR2 PMWJFETPWR3 PMWJFETPWR4 PLWJFETPWR1 PLWJFETPWR2 | RD06 RD06 RD06 RD06 RD06 RD06 | P350_JFET_1 P350_JFET_2 P350_JFET_3 P350_JFET_4 P500_JFET_1 P500_JFET_2 |
| 150 | 0 | 16 | PSWVSS1V | RD06 | PSW_VSS for VSS Channel ID 0 |
| 152 | 0 | 16 | PSWVSS2V | RD06 | PSW_VSS for VSS Channel ID 1 |
| 154 | 0 | 16 | PSWVSS3V | RD06 | PSW_VSS for VSS Channel ID 2 |
| 156 | 0 | 16 | PSWVSS4V | RD06 | PSW_VSS for VSS Channel ID 3 |
| 158 | 0 | 16 | PSWVSS5V | RD06 | PSW_VSS for VSS Channel ID 4 |
| 160 | 0 | 16 | PSWVSS6V | RD06 | PSW_VSS for VSS Channel ID 5 |
| 162 | 0 | 16 | PMWVSS1V | RD06 | PMW_VSS for VSS Channel ID 0 |
| 164 | 0 | 16 | PMWVSS2V | RD06 | PMW_VSS for VSS Channel ID 1 |
| 166 | 0 | 16 | PMWVSS3V | RD06 | PMW_VSS for VSS Channel ID 2 |
| 168 | 0 | 16 | PMWVSS4V | RD06 | PMW_VSS for VSS Channel ID 3 |
| 170 | 0 | 16 | PLWVSS1V | RD06 | PLW_VSS for VSS Channel ID 0 |
| 172 | 0 | 16 | PLWVSS2V | RD06 | PLW_VSS for VSS Channel ID 1 |
| 174 | 0 | 16 | PHOTHTRV | RD06 | PhotoHeaterBias |
| 176 | 0 | 16 |  |  |  |
| 178 | 0 | 16 |  |  |  |
| 180 | 0 | 16 |  |  |  |
| 182 | 0 | 16 |  |  |  |
| 184 | 0 | 16 | SPECSAMPFREQ | RD06 | SpectroBiasDiv |
| 186 | 0 | 16 | SPECBIASMODE | RD06 | SpectroBiasMode |
| 188 | 0 | 16 | SPECBIASFREQ | RD06 | SpectroMClkDiv |
| 190 | 0 | 16 | SSWBIAS | RD06 | SpectroBiasAmpl for Spectro Bias Channel 0 |
| 192 | 0 | 16 | SLWBIAS | RD06 | SpectroBiasAmpl for Spectro Bias Channel 0 |
| 194 | 0 | 16 | SSWPHASE | RD06 | SpectroPhaseShift for Spectro Bias Channel 0 |
| 196 | 0 | 16 | SLWPHASE | RD06 | SpectroPhaseShift for Spectro Bias Channel 01 |
| 198 | 0 0 1 2 | 16 1 1 1 | SPECJFETSTAT SLWJFETPWR1 SSWJFETPWR1 SSWJFETPWR2 | RD06 | SLW_JFET1 SSW_JFET1 SSW_JFET2 |
| 200 | 0 | 16 | SSWJFET1V | RD06 | SSW_VSS for VSS Channel ID 0 |
| 202 | 0 | 16 | SSWJFET2V | RD06 | SSW_VSS for VSS Channel ID 1 |
| 204 | 0 | 16 | SLWJFET1V | RD06 | SLW_VSS for VSS Channel ID 0 |
| 206 | 0 | 16 | SPECHTRV | RD06 | Spectro Heater Bias |
| 208 | 0 | 16 |  |  |  |
| 210 | 0 | 16 |  |  |  |
| 212 | 0 | 16 | TC1TEMP | RD06 | T/C 1 |
| 214 | 0 | 16 | TC2TEMP | RD06 | T/C 2 |
| 216 | 0 | 16 | TC3TEMP | RD06 | T/C 3 |
| 218 | 0 | 16 | PSWBIAS-B | RD06 | BIAS_BDA_1, Ph. Bolo. Bias 1 ampl. setting |
| 220 | 0 | 16 | PMWBIAS-B | RD06 | BIAS_BDA_2, Ph. Bolo. Bias 2 ampl. setting |
| 222 | 0 | 16 | PLWBIAS-B | RD06 | BIAS_BDA_3, Ph. Bolo. Bias 3 amplitude setting |
| 224 | 0 | 16 | SSWBIAS-B | RD06 | BIAS_BDA_4, Sp. Bolo. Bias 1 amplitude setting |
| 226 | 0 | 16 | SLWBIAS-B | RD06 | BIAS_BDA_5, Sp. Bolo. Bias 2 amplitude setting |
| 228 | 0 | 16 | TCBIAS-B | RD06 | BIAS_TEMP_5, T/C. Bias 1 ampl. Setting |
| 230 | 0 | 16 | PHOTBIASFREQ-B | RD06 | F_BIAS_PH/TC, Ph. & T/C Bias Freq div. Setting |
| 232 | 0 | 16 | SPECBIASFREQ-B | RD06 | F_BIAS_SP, Sp. Bias Freq div. setting |
| 234 | 0 | 16 | PSWVSS1V-B | RD06 | JFET_VSS_PSW_1, PSW_1 JFET VSS setting |
| 236 | 0 | 16 | PSWVSS2V-B | RD06 | JFET_VSS_PSW_2, PSW_2 JFET VSS setting |
| 238 | 0 | 16 | PSWVSS3V-B | RD06 | JFET_VSS_PSW_3, PSW_3 JFET VSS setting |
| 240 | 0 | 16 | PSWVSS4V-B | RD06 | JFET_VSS_PSW_4, PSW_4 JFET VSS setting |
| 242 | 0 | 16 | PSWVSS5V-B | RD06 | JFET_VSS_PSW_5, PSW_5 JFET VSS setting |
| 244 | 0 | 16 | PSWVSS6V-B | RD06 | JFET_VSS_PSW_6, PSW_6 JFET VSS setting |
| 246 | 0 | 16 | PMWVSS1V-B | RD06 | JFET_VSS_PMW_1, PMW_1 JFET VSS setting |
| 248 | 0 | 16 | PMWVSS2V-B | RD06 | JFET_VSS_PMW_2, PMW_2 JFET VSS setting |
| 250 | 0 | 16 | PMWVSS3V-B | RD06 | JFET_VSS_PMW_3, PMW_3 JFET VSS setting |
| 252 | 0 | 16 | PMWVSS4V-B | RD06 | JFET_VSS_PMW_4, PMW_4 JFET VSS setting |
| 254 | 0 | 16 | PLWVSS1V-B | RD06 | JFET_VSS_PLW_1, PLW_1 JFET VSS setting |
| 256 | 0 | 16 | PLWVSS2V-B | RD06 | JFET_VSS_PLW_2, PLW_2 JFET VSS setting |
| 258 | 0 | 16 | LIA01TEMP | RD06 | LIA_B1_TEMP, LIA board 1 temperature |
| 260 | 0 | 16 | LIA02TEMP | RD06 | LIA_B2_TEMP, LIA board 2 temperature |
| 262 | 0 | 16 | LIA03TEMP | RD06 | LIA_B3_TEMP, LIA board 3 temperature |
| 264 | 0 | 16 | LIA04TEMP | RD06 | LIA_B4_TEMP, LIA board 4 temperature |
| 266 | 0 | 16 | LIA05TEMP | RD06 | LIA_B5_TEMP, LIA board 5 temperature |
| 268 | 0 | 16 | LIA06TEMP | RD06 | LIA_B6_TEMP, LIA board 6 temperature |
| 270 | 0 | 16 | LIA07TEMP | RD06 | LIA_B7_TEMP, LIA board 7 temperature |
| 272 | 0 | 16 | LIA08TEMP | RD06 | LIA_B8_TEMP, LIA board 8 temperature |
| 274 | 0 | 16 | LIA09TEMP | RD06 | LIA_B9_TEMP, LIA board 9 temperature |
| 276 | 0 | 16 | LIA010TEMP | RD06 | LIA_B10_TEMP, LIA board 01 temperature |
| 278 | 0 | 16 | LIA11TEMP | RD06 | LIA_B11_TEMP, LIA board 11 temperature |
| 280 | 0 | 16 | LIA12TEMP | RD06 | LIA_B12_TEMP, LIA board 21 temperature |
| 282 | 0 | 16 | BIASTEMP | RD06 | BIAS_TEMP, Bias board temperature |
| 284 | 0 | 16 | DAQTEMP | RD06 | DAQ_IF_TEMP, DAQ IF board temperature |
| 286 | 0 0 1 2 3 4 5 6 7 8 9 10 11 | 16 1 1 1 1 1 1 1 1 1 1 1 1 | LIASTAT LIA01STAT LIA02STAT LIA03STAT LIA04STAT LIA05STAT LIA06STAT LIA07STAT LIA08STAT LIA09STAT LIA10STAT LIA11STAT LIA12STAT | RD06 | PWR_STATUS, LIA 1 to LIA 12 +5V/+9V/-9V status LIA 1 +5V/+9V/-9V status LIA 2 +5V/+9V/-9V status LIA 3 +5V/+9V/-9V status LIA 4 +5V/+9V/-9V status LIA 5 +5V/+9V/-9V status LIA 6 +5V/+9V/-9V status LIA 7 +5V/+9V/-9V status LIA 8 +5V/+9V/-9V status LIA 9 +5V/+9V/-9V status LIA 10 +5V/+9V/-9V status LIA 11 +5V/+9V/-9V status LIA 12 +5V/+9V/-9V status |
| 288 |  |  |  |  |  |
| 290 |  |  |  |  |  |
| 292 |  |  |  |  |  |
| 294 | 0 | 16 | MCUMACTEMP | RD06 | MACTemp |
| 296 | 0 | 16 | MCUSMECTEMP | RD06 | SMECTemp |
| 298 | 0 | 16 | MCUBSMTEMP | RD06 | BSMTemp |
| 300 | 0 | 16 | MCU+15V | RD06 | P15V |
| 302 | 0 | 16 | MCU-15V | RD06 | M15V |
| 304 | 0 | 16 | MCU+13V | RD06 | P13V |
| 306 | 0 | 16 | MCU-13V | RD06 | M13V |
| 308 | 0 | 16 | MCU+5V | RD06 | P5V |
| 310 | 0 | 16 | MCUPOLLINGTIME | RD06 RD07 | DPUPolling DPUPollingTime |
| 312 | 0 | 16 | MCUERR | RD06 | ErrorCode |
| 314 |  |  |  |  |  |
| 316 |  |  |  |  |  |
| 318 | 0 | 16 | MCUTM10TSAMPLE | RD06 RD07 | TP10SampFreq TelemetryPacket10Sampling |
| 320 | 0 | 16 | MCUTM11TSAMPLE | RD06 RD07 | TP11SampFreq TelemetryPacket11Sampling |
| 322 | 0 | 16 | MCUTM12TSAMPLE | RD06 RD07 | TP12SampFreq TelemetryPacket12Sampling |
| 324 | 0 | 16 | MCUTM13TSAMPLE | RD06 RD07 | TP13SampFreq TelemetryPacket13Sampling |
| 326 | 0 | 16 | MCUTM14TSAMPLE | RD06 RD07 | TP14SampFreq TelemetryPacket14Sampling |
| 328 | 0 | 16 | MCUTM15TSAMPLE | RD06 RD07 | TP15SampFreq TelemetryPacket15Sampling |
| 330 |  |  |  |  |  |
| 332 |  |  |  |  |  |
| 334 |  | 16 | MCUBOOTSTAT | RD06 | Boot Status Register |
| 336 |  | 16 | MCUDLOADCONF | RD06 | DownloadParam |
| 338 |  | 16 | MCUDLOADWORD | RD06 | DownloadWord |
| 340 |  | 16 | MCUDLOADPC | RD06 | DownloadCounter |
| 342 |  | 16 | MCUULOADCONF | RD06 | UploadParam |
| 344 |  | 16 | MCUULOADWORD | RD06 | UploadWord |
| 346 |  | 16 | MCUULOADPC | RD06 | UploadCounter |
| 348 |  |  |  |  |  |
| 350 |  |  |  |  |  |
| 352 | 0 | 16 | SMECENCPWR | RD06 | SEncoderPwr |
| 354 | 0 | 16 | SMECLVDTPWR | RD06 | SLVDTPwr |
| 356 | 0 | 16 | SMECMOTORMODE | RD06 | MotorMode |
| 358 | 0 | 16 | SMECLATCHSTAT | RD06 | SLaunchLatch |
| 360 | 0 | 16 | SMECLOOPMODE | RD06 | SLoopMode |
| 362 |  |  |  |  |  |
| 364 |  |  |  |  |  |
| 366 | 0 | 16 | SCANSTART | RD06 | STrajStartPos |
| 368 | 0 | 16 | SCANEND | RD06 | STrajEndPos |
| 370 | 0 | 16 | SCANSPEED | RD06 | SScanSpeed |
| 372 | 0 | 16 | SCANS | RD06 RD07 | SscanN SScanNumber |
| 374 | 0 | 16 | SCANMODE | RD06 | STrajMode |
| 376 |  |  |  |  |  |
| 378 |  |  |  |  |  |
| 380 | 0 | 16 | SMECKP | RD06 | SKp |
| 382 | 0 | 16 | SMECKD | RD06 | SKd |
| 384 | 0 | 16 | SMECDFILT | RD06 RD07 | SderivFilt SDerivFilter |
| 386 | 0 | 16 | SMECKI | RD06 | SKi |
| 388 | 0 | 16 | SMECINTLIMIT | RD06 RD07 | SintLimit SIntegrationLimit |
| 390 | 0 | 16 | SMECLVDTFILTER | RD06 | SLVDTFilteringFrequency |
| 392 | 0 | 16 | SMECNOTCHZERO | RD06 | NotchZeroFreq |
| 394 | 0 | 16 | SMECNOTCHPOLE | RD06 | NotchPoleFreq |
| 396 | 0 | 16 | SMECNOTCHZFREQ | RD06 | NotchZeroFreq |
| 398 | 0 | 16 | SMECNOTCHPFREQ | RD06 | NotchPoleFreq |
| 400 | 0 | 16 | SMECNOTCHZDAMP | RD06 | NotchZeroDamp |
| 402 | 0 | 16 | SMECNOTCHPDAMP | RD06 | NotchPoleDamp |
| 404 | 0 | 16 | SMECFEEDGAIN | RD06 RD07 | SFeedFwdGain SFeedForwardGain |
| 406 | 0 | 16 | SMECMAXOUT | RD06 RD07 | SMaxOut SMaxOutput |
| 408 | 0 | 16 | SMECMAXSPEED | RD06 | SMaxSpeed |
| 410 | 0 | 16 | SMECMAXACCEL | RD06 RD07 | SMaxAccel SMaxAcceleration |
| 412 | 0 | 16 | SMECMAXPOSNERR | RD06 RD07 | MaxPosError Max tolerated position error |
| 414 | 0 | 16 | SMECMAXSPEEDERR | RD06 RD07 | MaxSpeedErr Max tolerated speed error |
| 416 | 0 | 16 | SMECMINENCAMPL | RD06 | MinEncodSAmpl Min Encoder signal tolerated amplitude |
| 418 | 0 | 16 | SMECMAXENCBIAS | RD06 | MaxEncodSBias Max Encoder signal tolerated bias |
| 420 |  |  |  |  |  |
| 422 |  |  |  |  |  |
| 424 | 0 | 16 | SMECSTAT | RD06 | SMECStatus |
| 426 | 0 | 16 | SMECENCPOSN | RD06 RD07 | SEncodeIncrPos SEncoder IncrementPosition |
| 428 | 0 | 16 | SMECSINE000SIG | RD06 RD07 | SEncodS1 SencoderSignal1 |
| 430 | 0 | 16 | SMECSINE120SIG | RD06 RD07 | SEncodS2 SencoderSignal2 |
| 432 | 0 | 16 | SMECSINE240SIG | RD06 RD07 | SEncodS3 SencoderSignal3 |
| 434 | 0 | 16 | SMECLVDTPOSN | RD07 | SLVDT position |
| 436 | 0 | 16 | SMECLVDTACSIG | RD06 | SLVDTAC |
| 438 | 0 | 16 | SMEVLVDTDCSIG | RD06 | SLVDTAC |
| 440 | 0 | 16 | SMECINSTPOSNERR | RD06 RD07 | MeanPosError Mean position error |
| 442 | 0 | 16 | SMECOLOOPPOSN | RD06 RD07 | EstOLPos Estimated open loop position |
| 444 | 0 | 16 | SMECHOMEPOSNDIFF | RD06 RD07 | PosDiff Position difference encoder/lvdt home |
| 446 | 0 | 16 | SMECENCDELTAT | RD06 RD07 | EncodDT EncoderDeltaTime |
| 448 | 0 | 16 | SMECENCSHIFT | RD06 RD07 | EncodShift EmcoderShift |
| 450 | 0 | 16 | SMECINSTVEL | RD06 RD07 | ActualVel ActualVelocity |
| 452 | 0 | 16 | SMECMEANSPEED | RD06 | MeanSpeed |
| 454 | 0 | 16 | SMECSCANPOSNERR | RD06 RD07 | MeanPosError MeanPositionError |
| 456 | 0 | 16 | SMECMOTORCURR | RD06 | SMotorCurrent |
| 458 | 0 | 16 | SMECMOTORBEMF | RD06 | SBEMF |
| 460 |  |  |  |  |  |
| 462 |  |  |  |  |  |
| 464 |  |  |  |  |  |
| 466 |  |  |  |  |  |
| 468 | 0 | 16 | CHOPSENSPWR | RD06 | CSensorPwr |
| 470 | 0 | 16 | CHOPLATCHSTAT | RD06 | BSMLaunchLatch1 |
| 472 | 0 | 16 | CHOPLOOPMODE | RD06 | ChopLoopMode |
| 474 | 0 | 16 | CHOPPOSN0 | RD06 | CPosition0 |
| 476 | 0 | 16 | CHOPPOSN1 | RD06 | CPosition1 |
| 478 | 0 | 16 | CHOPPERIOD | RD06 | ChopPeriod |
| 480 | 0 | 16 | CHOPMODE | RD06 | ChopMode |
| 482 | 0 | 16 | CHOPS | RD06 | ChopNumber |
| 484 | 0 | 16 | CHOPKP | RD06 | CKp |
| 486 | 0 | 16 | CHOPKD | RD06 | CKd |
| 488 | 0 | 16 | CHOPKI | RD06 | CKi |
| 490 | 0 | 16 | CHOPINTTHRESH | RD06 RD07 | IntegTh CIntegrationThreshold |
| 492 | 0 | 16 | CHOPPOSNERRLIMIT | RD06 RD07 | CPosErrorLimit CPositionErrorLimit |
| 494 | 0 | 16 | CHOPMOTORTORQUE | RD06 RD07 | CMotorTorqueC CMotorTorqueCst |
| 496 | 0 | 16 | CHOPRATELOOPSF | RD06 | CRateLoopsf |
| 498 | 0 | 16 | CHOPACCELLOOPSF | RD06 | CAccelLoopsf |
| 500 | 0 | 16 | CHOPRATELIMIT | RD06 | CRateLimit |
| 502 | 0 | 16 | CHOPACCELLIMIT | RD06 | CAccelLimi |
| 504 | 0 | 16 | CHOPSTATECOEFF1 | RD06 RD07 | CStateCoeffn CStateCoeff1 |
| 506 | 0 | 16 | CHOPSTATECOEFF2 | RD06 RD07 | CstateCoeffn CStateCoeff2 |
| 508 | 0 | 16 | CHOPSTATECOEFF3 | RD06 RD07 | CstateCoeffn CStateCoeff3 |
| 510 | 0 | 16 | CHOPSTATECOEFF4 | RD06 RD07 | CstateCoeffn CStateCoeff4 |
| 512 | 0 | 16 | CHOPICNCOEFF1 | RD06 RD07 | ICnCoeffn ICnCoeff1 |
| 514 | 0 | 16 | CHOPICNCOEFF2 | RD06 RD07 | ICnCoeffn ICnCoeff2 |
| 516 | 0 | 16 | CHOPICNCOEFF3 | RD06 RD07 | ICnCoeffn ICnCoeff3 |
| 518 | 0 | 16 | CHOPICNCOEFF4 | RD06 RD07 | ICnCoeffn ICnCoeff4 |
| 520 | 0 | 16 | CHOPOUTCOEFF1 | RD06 RD07 | COutCoeffn COutCoeff1 |
| 522 | 0 | 16 | CHOPOUTCOEFF2 | RD06 RD07 | COutCoeffn CoutCoeff2 |
| 524 | 0 | 16 | CHOPOUTCOEFF3 | RD06 RD07 | COutCoeffn CoutCoeff3 |
| 526 | 0 | 16 | CHOPOUTCOEFF4 | RD06 RD07 | COutCoeffn CoutCoeff4 |
| 528 |  |  |  |  | Reserved |
| 530 |  |  |  |  | Reserved |
| 532 |  |  |  |  | Reserved |
| 534 |  |  |  |  | Reserved |
| 536 |  |  |  |  | Reserved |
| 538 |  |  |  |  | Reserved |
| 540 |  |  |  |  | Reserved |
| 542 |  |  |  |  | Reserved |
| 544 |  |  |  |  | Reserved |
| 546 |  |  |  |  | Reserved |
| 548 |  |  |  |  | Reserved |
| 550 |  |  |  |  | Reserved |
| 552 |  |  |  |  | Reserved |
| 554 |  |  |  |  | Reserved |
| 556 |  |  |  |  | Reserved |
| 558 |  |  |  |  | Reserved |
| 560 |  |  |  |  | Reserved |
| 562 |  |  |  |  | Reserved |
| 564 | 0 | 16 | CHOPPROFILEVAL01 | RD06 RD07 | CprofileValuen CprofileValue1 |
| 566 | 0 | 16 | CHOPPROFILEVAL02 | RD06 RD07 | CprofileValuen CprofileValue2 |
| 568 | 0 | 16 | CHOPPROFILEVAL03 | RD06 RD07 | CprofileValuen CprofileValue3 |
| 570 | 0 | 16 | CHOPPROFILEVAL04 | RD06 RD07 | CprofileValuen CprofileValue4 |
| 572 | 0 | 16 | CHOPPROFILEVAL05 | RD06 RD07 | CprofileValuen CprofileValue5 |
| 574 | 0 | 16 | CHOPPROFILEVAL06 | RD06 RD07 | CprofileValuen CprofileValue6 |
| 576 | 0 | 16 | CHOPPROFILEVAL07 | RD06 RD07 | CprofileValuen CprofileValue7 |
| 578 | 0 | 16 | CHOPPROFILEVAL08 | RD06 RD07 | CprofileValuen CprofileValue8 |
| 580 | 0 | 16 | CHOPPROFILEVAL09 | RD06 RD07 | CprofileValuen CprofileValue9 |
| 582 | 0 | 16 | CHOPPROFILEVAL10 | RD06 RD07 | CprofileValuen CprofileValue10 |
| 584 | 0 | 16 | CHOPPROFILEVAL11 | RD06 RD07 | CprofileValuen CprofileValue11 |
| 586 | 0 | 16 | CHOPPROFILEVAL12 | RD06 RD07 | CprofileValuen CprofileValue12 |
| 588 | 0 | 16 | CHOPPROFILEVAL13 | RD06 RD07 | CprofileValuen CprofileValue13 |
| 590 | 0 | 16 | CHOPPROFILEVAL14 | RD06 RD07 | CprofileValuen CprofileValue14 |
| 592 | 0 | 16 | CHOPPROFILEVAL15 | RD06 RD07 | CprofileValuen CprofileValue15 |
| 594 |  |  |  |  | Reserved |
| 596 | 0 | 16 | CHOPSTAT | RD06 | ChopStatus |
| 598 | 0 | 16 | CHOPPOSN | RD06 | CPosition |
| 600 | 0 | 16 | CHOPPOSNERR | RD06 RD07 | CMeanPosError CMeanPositionError |
| 602 | 0 | 16 | CHOPSENSPOSN1 | RD06 RD07 | CMagnetoResPos1 CMagnetoresistive pos #1 |
| 604 | 0 | 16 | CHOPSENSPOSN2 | RD06 RD07 | CMagnetoResPos2 CMagnetoresistive pos #2 |
| 606 | 0 | 16 | CHOPMOTORCURR | RD06 | CMotorCurrent |
| 608 | 0 | 16 | CHOPBEMF | RD06 | CBEMF |
| 610 |  |  |  |  |  |
| 612 |  |  |  |  |  |
| 614 |  |  |  |  |  |
| 616 |  |  |  |  |  |
| 618 | 0 | 16 | JIGGSENSPWR | RD06 | JSensorPwr |
| 620 | 0 | 16 | JIGGLATCHSTAT | RD06 | BSMLaunchLatch2 |
| 622 | 0 | 16 | JIGGLOOPMODE | RD06 | JigLoopMode |
| 624 | 0 | 16 | JIGGPOSN0 | RD06 | JPosition0 |
| 626 | 0 | 16 | JIGGPOSN1 | RD06 | JPosition1 |
| 628 | 0 | 16 | JIGGPERIOD | RD06 | JigPeriod |
| 630 | 0 | 16 | JIGGMODE | RD06 | JigMode |
| 632 | 0 | 16 | JIGGLES | RD06 | JigNumber |
| 634 | 0 | 16 | JIGGKP | RD06 | JKp |
| 636 | 0 | 16 | JIGGKD | RD06 | JKd |
| 638 | 0 | 16 | JIGGKI | RD06 | JKi |
| 640 | 0 | 16 | JIGGINTTHRESH | RD06 | JIntegTh |
| 642 | 0 | 16 | JIGGPOSNERRLIMIT | RD06 | JPosErrorLimit |
| 644 | 0 | 16 | JIGGMOTORTORQUE | RD06 | JMotorTorqueC |
| 646 | 0 | 16 | JIGGRATELOOPSF | RD06 | JRateLoopsf |
| 648 | 0 | 16 | JIGGACCELLOOPSF | RD06 | JAccelLoopsf |
| 650 | 0 | 16 | JIGGRATELIMIT | RD06 | JRateLimit |
| 652 | 0 | 16 | JIGGACCELLIMIT | RD06 | JAccelLimi |
| 654 | 0 | 16 | JIGGSTATECOEFF1 | RD06 | JStateCoeffn |
| 656 | 0 | 16 | JIGGSTATECOEFF2 | RD06 | JStateCoeffn |
| 658 | 0 | 16 | JIGGSTATECOEFF3 | RD06 | JStateCoeffn |
| 660 | 0 | 16 | JIGGSTATECOEFF4 | RD06 | JStateCoeffn |
| 662 | 0 | 16 | JIGGICNCOEFF1 | RD06 | ICnCoeffn |
| 664 | 0 | 16 | JIGGICNCOEFF2 | RD06 | ICnCoeffn |
| 666 | 0 | 16 | JIGGICNCOEFF3 | RD06 | ICnCoeffn |
| 668 | 0 | 16 | JIGGICNCOEFF4 | RD06 | ICnCoeffn |
| 670 | 0 | 16 | JIGGOUTCOEFF1 | RD06 | COutCoeffn |
| 672 | 0 | 16 | JIGGOUTCOEFF2 | RD06 | COutCoeffn |
| 674 | 0 | 16 | JIGGOUTCOEFF3 | RD06 | COutCoeffn |
| 676 | 0 | 16 | JIGGOUTCOEFF4 | RD06 | COutCoeffn |
| 678 |  |  |  |  | Reserved |
| 680 |  |  |  |  | Reserved |
| 682 |  |  |  |  | Reserved |
| 684 |  |  |  |  | Reserved |
| 686 |  |  |  |  | Reserved |
| 688 |  |  |  |  | Reserved |
| 690 |  |  |  |  | Reserved |
| 692 |  |  |  |  | Reserved |
| 694 |  |  |  |  | Reserved |
| 696 |  |  |  |  | Reserved |
| 698 |  |  |  |  | Reserved |
| 700 |  |  |  |  | Reserved |
| 702 |  |  |  |  | Reserved |
| 704 |  |  |  |  | Reserved |
| 706 |  |  |  |  | Reserved |
| 708 |  |  |  |  | Reserved |
| 710 |  |  |  |  | Reserved |
| 712 |  |  |  |  | Reserved |
| 714 | 0 | 16 | JIGGPROFILEVAL01 | RD06 | CprofileValuen |
| 716 | 0 | 16 | JIGGPROFILEVAL02 | RD06 | CprofileValuen |
| 718 | 0 | 16 | JIGGPROFILEVAL03 | RD06 | CprofileValuen |
| 720 | 0 | 16 | JIGGPROFILEVAL04 | RD06 | CprofileValuen |
| 722 | 0 | 16 | JIGGPROFILEVAL05 | RD06 | CprofileValuen |
| 724 | 0 | 16 | JIGGPROFILEVAL06 | RD06 | CprofileValuen |
| 726 | 0 | 16 | JIGGPROFILEVAL07 | RD06 | CprofileValuen |
| 728 | 0 | 16 | JIGGPROFILEVAL08 | RD06 | CprofileValuen |
| 730 | 0 | 16 | JIGGPROFILEVAL09 | RD06 | CprofileValuen |
| 732 | 0 | 16 | JIGGPROFILEVAL10 | RD06 | CprofileValuen |
| 734 | 0 | 16 | JIGGPROFILEVAL11 | RD06 | CprofileValuen |
| 736 | 0 | 16 | JIGGPROFILEVAL12 | RD06 | CprofileValuen |
| 738 | 0 | 16 | JIGGPROFILEVAL13 | RD06 | CprofileValuen |
| 740 | 0 | 16 | JIGGPROFILEVAL14 | RD06 | CprofileValuen |
| 742 | 0 | 16 | JIGGPROFILEVAL15 | RD06 | CprofileValuen |
| 744 |  |  |  |  | Reserved |
| 746 | 0 | 16 | JIGGSTAT | RD06 | JigStatus |
| 748 | 0 | 16 | JIGGPOSN | RD06 | JPosition |
| 750 | 0 | 16 | JIGGPOSNERR | RD06 | JMeanPosError |
| 752 | 0 | 16 | JIGGSENSPOSN1 | RD06 | JMagnetoResPos1 |
| 754 | 0 | 16 | JIGGSENSPOSN2 | RD06 | JMagnetoResPos2 |
| 756 | 0 | 16 | JIGGMOTORCURR | RD06 | JMotorCurrent |
| 758 | 0 | 16 | JIGGBEMF | RD06 | JBEMF |
| 760 |  |  |  |  |  |
| 762 |  |  |  |  |  |
| 764 |  |  |  |  |  |
| 766 |  |  |  |  |  |
| 768 | 0 | 16 | MCUTRACETSAMPLE | RD06 RD07 | TraceSampFreq TraceeSampling |
| 770 | 0 | 16 | MCUTRACEBUFLEN | RD06 RD07 | TraceBufLength TraceBuffer |
| 772 | 0 | 16 | MCUTRACEPARM01 | RD06 RD07 | Trace Param 1 TraceParam#1 |
| 774 | 0 | 16 | MCUTRACEPARM02 | RD06 RD07 | Trace Param 2 TraceParam#2 |
| 776 | 0 | 16 | MCUTRACEPARM03 | RD06 RD07 | Trace Param 3 TraceParam#3 |
| 778 | 0 | 16 | MCUTRACEPARM04 | RD06 RD07 | Trace Param 4 TraceParam#4 |
| 780 | 0 | 16 | MCUTRACEPARM05 | RD06 RD07 | Trace Param 5 TraceParam#5 |
| 782 | 0 | 16 | MCUTRACEPARM06 | RD06 RD07 | Trace Param 6 TraceParam#6 |
| 784 | 0 | 16 | MCUTRACEPARM07 | RD10 | TraceParam#7 |
| 786 | 0 | 16 | MCUTRACEPARM08 | RD10 | TraceParam#8 |
| 788 | 0 | 16 | MCUTRACEPARM09 | RD10 | TraceParam#9 |
| 790 | 0 | 16 | MCUTRACEPARM10 | RD10 | TraceParam#10 |
| 792 | 0 | 16 | MCUTRACEPARM11 | RD10 | TraceParam#11 |
| 794 | 0 | 16 | MCUTRACEPARM12 | RD10 | TraceParam#12 |
| 796 | 0 | 16 | MCUTRACEPARM13 | RD10 | TraceParam#13 |
| 798 | 0 | 16 | MCUTRACEPARM14 | RD10 | TraceParam#14 |
| 800 | 0 | 16 | MCUTRACEPARM15 | RD10 | TraceParam#15 |
| 802 | 0 | 16 | MCUTRACEPARM16 | RD10 | TraceParam#16 |
| 804 | 0 | 16 | MCUTRACEPARM17 | RD10 | TraceParam#17 |
| 806 | 0 | 16 | MCUTRACEPARM18 | RD10 | TraceParam#18 |
| 808 | 0 | 16 | MCUTRACEPARM19 | RD10 | TraceParam#19 |
| 810 |  |  |  |  |  |
| 812 |  |  |  |  |  |
| 814 |  |  |  |  |  |
| 816 |  |  |  |  |  |
| 818 | 0 | 16 | SCUSTAT | RD06 | ScuStatus |
| 820 | 0 | 16 | SCUTEMPSTAT | RD06 | TempOnOff |
| 822 | 0 | 16 | SCUDCDCSTAT | RD06 | DRelOnOff |
| 824 | 0 | 16 | SCU+5V | RD06 | ScuCHTp05 |
| 826 | 0 | 16 | SCU+9V | RD06 | ScuCHTp09 |
| 828 | 0 | 16 | SCU-9V | RD06 | ScuCHTn09 |
| 830 | 0 | 16 | SPHSCURR | RD06 | SPHSHeatB |
| 832 | 0 | 16 | EVHSCURR | RD06 | EVHSHeatB |
| 834 | 0 | 16 | TCHTRCURR | RD06 | TCHeaterB |
| 836 | 0 | 16 | SPHTRCURR | RD06 | SPHeater |
| 838 | 0 | 16 | CCUTEMP | RD06 | CcuTempRd |
| 840 | 0 | 16 | TCUTEMP | RD06 | TcuTempRd |
| 842 | 0 | 16 | PSUTEMP | RD06 | PsuTempRd |
| 844 | 0 0 15 | 16 8 1 | SCUFRAMECONF SCUFRAMERATE SCUFRAMETYPE | RD06 RD06 | FrameConf FrameRAte |
| 846 | 0 | 16 | SCUFRAMES | RD06 | SeqLength |
| 848 |  |  |  |  |  |
| 850 |  |  |  |  |  |
| 852 | 0 | 16 | PCALV | RD06 | PhCalVolt |
| 854 | 0 | 16 | SCAL2V | RD06 | SCal2Volt |
| 856 | 0 | 16 | SCAL4V | RD06 | SCal4Volt |
| 858 | 0 | 16 | PCALCURR | RD06 | PhCalBias |
| 860 | 0 | 16 | SCAL2CURR | RD06 | SCal2Bias |
| 862 | 0 | 16 | SCAL4CURR | RD06 | SCal4Bias |
| 864 |  |  |  |  |  |
| 866 | 0 | 16 | SUBKSTAT | RD06 | SubKpOnOff |
| 868 | 0 | 16 | SUBKTEMP | RD06 | SubKTempP |
| 870 |  |  |  |  |  |
| 872 | 0 | 16 | FPUTEMP01 | RD06 | FpuTemp01 |
| 874 | 0 | 16 | FPUTEMP02 | RD06 | FpuTemp02 |
| 876 | 0 | 16 | FPUTEMP03 | RD06 | FpuTemp03 |
| 878 | 0 | 16 | FPUTEMP04 | RD06 | FpuTemp04 |
| 880 | 0 | 16 | FPUTEMP05 | RD06 | FpuTemp05 |
| 882 | 0 | 16 | FPUTEMP06 | RD06 | FpuTemp06 |
| 884 | 0 | 16 | FPUTEMP07 | RD06 | FpuTemp07 |
| 886 | 0 | 16 | FPUTEMP08 | RD06 | FpuTemp08 |
| 888 | 0 | 16 | FPUTEMP09 | RD06 | FpuTemp09 |
| 890 | 0 | 16 | FPUTEMP10 | RD06 | FpuTemp10 |
| 892 | 0 | 16 | FPUTEMP11 | RD06 | FpuTemp11 |
| 894 | 0 | 16 | FPUTEMP12 | RD06 | FpuTemp12 |
| 896 | 0 | 16 | FPUTEMP13 | RD06 | FpuTemp13 |
| 898 | 0 | 16 | FPUTEMP14 | RD06 | FpuTemp14 |
| 900 | 0 | 16 | FPUTEMP15 | RD06 | FpuTemp15 |
| 902 | 0 | 16 | FPUTEMP16 | RD06 | FpuTemp16 |

#### Detector Housekeeping Report (SID=0x0302)

The following table lists the field to be found in this report

| Location (msb) | Field Length (bits) | Parameter Name | Reference | Description |
| --- | --- | --- | --- | --- |
| octet | bit |  |  |  |  |
| 18 | 0 | 32 | OBSID | AD03 | Observation ID |
| 22 | 0 | 32 | BBID | AD03 | Building Block ID |
| 26 | 0 | 16 | PSWBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 0 |
| 28 | 0 | 16 | PMWBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 1 |
| 30 | 0 | 16 | PLWBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 2 |
| 32 | 0 | 16 | TCBIAS | RD06 | PhotoBiasAmpl for Photo Bias Channel 3 |
| 34 | 0 | 16 | PSWPHASE | RD06 | PhaseShift for Photo Bias Channel 0 |
| 36 | 0 | 16 | PMWPHASE | RD06 | PhaseShift for Photo Bias Channel 1 |
| 38 | 0 | 16 | PLWPHASE | RD06 | PhaseShift for Photo Bias Channel 2 |
| 40 | 0 | 16 | TCPHASE | RD06 | PhaseShift for Photo Bias Channel 3 |
| 42 | 0 0 1 2 3 4 5 | 16 1 1 1 1 1 1 | PSWJFETSTAT PSWJFETPWR1 PSWJFETPWR2 PSWJFETPWR3 PSWJFETPWR4 PSWJFETPWR5 PSWJFETPWR6 | RD06 RD06 RD06 RD06 RD06 RD06 | P250_JFET_1 P250_JFET_2 P250_JFET_3 P250_JFET_4 P250_JFET_5 P250_JFET_6 |
| 44 | 0 0 1 2 3 4 5 | 16 1 1 1 1 1 1 | PMLWJFETSTAT PMWJFETPWR1 PMWJFETPWR2 PMWJFETPWR3 PMWJFETPWR4 PLWJFETPWR1 PLWJFETPWR2 | RD06 RD06 RD06 RD06 RD06 RD06 | P350_JFET_1 P350_JFET_2 P350_JFET_3 P350_JFET_4 P500_JFET_1 P500_JFET_2 |
| 46 | 0 | 16 | PSWVSS1V | RD06 | PSW_VSS for VSS Channel ID 0 |
| 48 | 0 | 16 | PSWVSS2V | RD06 | PSW_VSS for VSS Channel ID 1 |
| 50 | 0 | 16 | PSWVSS3V | RD06 | PSW_VSS for VSS Channel ID 2 |
| 52 | 0 | 16 | PSWVSS4V | RD06 | PSW_VSS for VSS Channel ID 3 |
| 54 | 0 | 16 | PSWVSS5V | RD06 | PSW_VSS for VSS Channel ID 4 |
| 56 | 0 | 16 | PSWVSS6V | RD06 | PSW_VSS for VSS Channel ID 5 |
| 58 | 0 | 16 | PMWVSS1V | RD06 | PMW_VSS for VSS Channel ID 0 |
| 60 | 0 | 16 | PMWVSS2V | RD06 | PMW_VSS for VSS Channel ID 1 |
| 62 | 0 | 16 | PMWVSS3V | RD06 | PMW_VSS for VSS Channel ID 2 |
| 64 | 0 | 16 | PMWVSS4V | RD06 | PMW_VSS for VSS Channel ID 3 |
| 66 | 0 | 16 | PLWVSS1V | RD06 | PLW_VSS for VSS Channel ID 0 |
| 68 | 0 | 16 | PLWVSS2V | RD06 | PLW_VSS for VSS Channel ID 1 |
| 70 | 0 | 16 | PHOTHTRV | RD06 | PhotoHeaterBias |
| 72 | 0 | 16 | LIA1OFFSET1 | RD06 | Offset_P1 for Channel_P1 0 |
| 74 | 0 | 16 | LIA1OFFSET2 | RD06 | Offset_P1 for Channel_P1 1 |
| 76 | 0 | 16 | LIA1OFFSET3 | RD06 | Offset_P1 for Channel_P1 2 |
| 78 | 0 | 16 | LIA1OFFSET4 | RD06 | Offset_P1 for Channel_P1 3 |
| 80 | 0 | 16 | LIA1OFFSET5 | RD06 | Offset_P1 for Channel_P1 4 |
| 82 | 0 | 16 | LIA1OFFSET6 | RD06 | Offset_P1 for Channel_P1 5 |
| 84 | 0 | 16 | LIA1OFFSET7 | RD06 | Offset_P1 for Channel_P1 6 |
| 86 | 0 | 16 | LIA1OFFSET8 | RD06 | Offset_P1 for Channel_P1 7 |
| 88 | 0 | 16 | LIA1OFFSET9 | RD06 | Offset_P1 for Channel_P1 8 |
| 90 | 0 | 16 | LIA1OFFSET10 | RD06 | Offset_P1 for Channel_P1 9 |
| 92 | 0 | 16 | LIA1OFFSET11 | RD06 | Offset_P1 for Channel_P1 10 |
| 94 | 0 | 16 | LIA1OFFSET12 | RD06 | Offset_P1 for Channel_P1 11 |
| 96 | 0 | 16 | LIA1OFFSET13 | RD06 | Offset_P1 for Channel_P1 12 |
| 98 | 0 | 16 | LIA1OFFSET14 | RD06 | Offset_P1 for Channel_P1 13 |
| 100 | 0 | 16 | LIA1OFFSET15 | RD06 | Offset_P1 for Channel_P1 14 |
| 102 | 0 | 16 | LIA1OFFSET16 | RD06 | Offset_P1 for Channel_P1 15 |
| 104 | 0 | 16 | LIA1OFFSET17 | RD06 | Offset_P1 for Channel_P1 16 |
| 106 | 0 | 16 | LIA1OFFSET18 | RD06 | Offset_P1 for Channel_P1 17 |
| 108 | 0 | 16 | LIA1OFFSET19 | RD06 | Offset_P1 for Channel_P1 18 |
| 110 | 0 | 16 | LIA1OFFSET20 | RD06 | Offset_P1 for Channel_P1 19 |
| 112 | 0 | 16 | LIA1OFFSET21 | RD06 | Offset_P1 for Channel_P1 20 |
| 114 | 0 | 16 | LIA1OFFSET22 | RD06 | Offset_P1 for Channel_P1 21 |
| 116 | 0 | 16 | LIA1OFFSET23 | RD06 | Offset_P1 for Channel_P1 22 |
| 118 | 0 | 16 | LIA1OFFSET24 | RD06 | Offset_P1 for Channel_P1 23 |
| 120 | 0 | 16 | LIA1OFFSET25 | RD06 | Offset_P1 for Channel_P1 24 |
| 122 | 0 | 16 | LIA1OFFSET26 | RD06 | Offset_P1 for Channel_P1 25 |
| 124 | 0 | 16 | LIA1OFFSET27 | RD06 | Offset_P1 for Channel_P1 26 |
| 126 | 0 | 16 | LIA1OFFSET28 | RD06 | Offset_P1 for Channel_P1 27 |
| 128 | 0 | 16 | LIA1OFFSET29 | RD06 | Offset_P1 for Channel_P1 28 |
| 130 | 0 | 16 | LIA1OFFSET30 | RD06 | Offset_P1 for Channel_P1 29 |
| 132 | 0 | 16 | LIA1OFFSET31 | RD06 | Offset_P1 for Channel_P1 30 |
| 134 | 0 | 16 | LIA1OFFSET32 | RD06 | Offset_P1 for Channel_P1 31 |
| 136 | 0 | 16 | LIA2OFFSET1 | RD06 | Offset_P2 for Channel_P2 0 |
| 138 | 0 | 16 | LIA2OFFSET2 | RD06 | Offset_P2 for Channel_P2 1 |
| 140 | 0 | 16 | LIA2OFFSET3 | RD06 | Offset_P2 for Channel_P2 2 |
| 142 | 0 | 16 | LIA2OFFSET4 | RD06 | Offset_P2 for Channel_P2 3 |
| 144 | 0 | 16 | LIA2OFFSET5 | RD06 | Offset_P2 for Channel_P2 4 |
| 146 | 0 | 16 | LIA2OFFSET6 | RD06 | Offset_P2 for Channel_P2 5 |
| 148 | 0 | 16 | LIA2OFFSET7 | RD06 | Offset_P2 for Channel_P2 6 |
| 150 | 0 | 16 | LIA2OFFSET8 | RD06 | Offset_P2 for Channel_P2 7 |
| 152 | 0 | 16 | LIA2OFFSET9 | RD06 | Offset_P2 for Channel_P2 8 |
| 154 | 0 | 16 | LIA2OFFSET10 | RD06 | Offset_P2 for Channel_P2 9 |
| 156 | 0 | 16 | LIA2OFFSET11 | RD06 | Offset_P2 for Channel_P2 10 |
| 158 | 0 | 16 | LIA2OFFSET12 | RD06 | Offset_P2 for Channel_P2 11 |
| 160 | 0 | 16 | LIA2OFFSET13 | RD06 | Offset_P2 for Channel_P2 12 |
| 162 | 0 | 16 | LIA2OFFSET14 | RD06 | Offset_P2 for Channel_P2 13 |
| 164 | 0 | 16 | LIA2OFFSET15 | RD06 | Offset_P2 for Channel_P2 14 |
| 166 | 0 | 16 | LIA2OFFSET16 | RD06 | Offset_P2 for Channel_P2 15 |
| 168 | 0 | 16 | LIA2OFFSET17 | RD06 | Offset_P2 for Channel_P2 16 |
| 170 | 0 | 16 | LIA2OFFSET18 | RD06 | Offset_P2 for Channel_P2 17 |
| 172 | 0 | 16 | LIA2OFFSET19 | RD06 | Offset_P2 for Channel_P2 18 |
| 174 | 0 | 16 | LIA2OFFSET20 | RD06 | Offset_P2 for Channel_P2 19 |
| 176 | 0 | 16 | LIA2OFFSET21 | RD06 | Offset_P2 for Channel_P2 20 |
| 178 | 0 | 16 | LIA2OFFSET22 | RD06 | Offset_P2 for Channel_P2 21 |
| 180 | 0 | 16 | LIA2OFFSET23 | RD06 | Offset_P2 for Channel_P2 22 |
| 182 | 0 | 16 | LIA2OFFSET24 | RD06 | Offset_P2 for Channel_P2 23 |
| 184 | 0 | 16 | LIA2OFFSET25 | RD06 | Offset_P2 for Channel_P2 24 |
| 186 | 0 | 16 | LIA2OFFSET26 | RD06 | Offset_P2 for Channel_P2 25 |
| 188 | 0 | 16 | LIA2OFFSET27 | RD06 | Offset_P2 for Channel_P2 26 |
| 190 | 0 | 16 | LIA2OFFSET28 | RD06 | Offset_P2 for Channel_P2 27 |
| 192 | 0 | 16 | LIA2OFFSET29 | RD06 | Offset_P2 for Channel_P2 28 |
| 194 | 0 | 16 | LIA2OFFSET30 | RD06 | Offset_P2 for Channel_P2 29 |
| 196 | 0 | 16 | LIA2OFFSET31 | RD06 | Offset_P2 for Channel_P2 30 |
| 198 | 0 | 16 | LIA2OFFSET32 | RD06 | Offset_P2 for Channel_P2 31 |
| 200 | 0 | 16 | LIA3OFFSET1 | RD06 | Offset_P3 for Channel_P3 0 |
| 202 | 0 | 16 | LIA3OFFSET2 | RD06 | Offset_P3 for Channel_P3 1 |
| 204 | 0 | 16 | LIA3OFFSET3 | RD06 | Offset_P3 for Channel_P3 2 |
| 206 | 0 | 16 | LIA3OFFSET4 | RD06 | Offset_P3 for Channel_P3 3 |
| 208 | 0 | 16 | LIA3OFFSET5 | RD06 | Offset_P3 for Channel_P3 4 |
| 210 | 0 | 16 | LIA3OFFSET6 | RD06 | Offset_P3 for Channel_P3 5 |
| 212 | 0 | 16 | LIA3OFFSET7 | RD06 | Offset_P3 for Channel_P3 6 |
| 214 | 0 | 16 | LIA3OFFSET8 | RD06 | Offset_P3 for Channel_P3 7 |
| 216 | 0 | 16 | LIA3OFFSET9 | RD06 | Offset_P3 for Channel_P3 8 |
| 218 | 0 | 16 | LIA3OFFSET10 | RD06 | Offset_P3 for Channel_P3 9 |
| 220 | 0 | 16 | LIA3OFFSET11 | RD06 | Offset_P3 for Channel_P3 10 |
| 222 | 0 | 16 | LIA3OFFSET12 | RD06 | Offset_P3 for Channel_P3 11 |
| 224 | 0 | 16 | LIA3OFFSET13 | RD06 | Offset_P3 for Channel_P3 12 |
| 226 | 0 | 16 | LIA3OFFSET14 | RD06 | Offset_P3 for Channel_P3 13 |
| 228 | 0 | 16 | LIA3OFFSET15 | RD06 | Offset_P3 for Channel_P3 14 |
| 230 | 0 | 16 | LIA3OFFSET16 | RD06 | Offset_P3 for Channel_P3 15 |
| 232 | 0 | 16 | LIA3OFFSET17 | RD06 | Offset_P3 for Channel_P3 16 |
| 234 | 0 | 16 | LIA3OFFSET18 | RD06 | Offset_P3 for Channel_P3 17 |
| 236 | 0 | 16 | LIA3OFFSET19 | RD06 | Offset_P3 for Channel_P3 18 |
| 238 | 0 | 16 | LIA3OFFSET20 | RD06 | Offset_P3 for Channel_P3 19 |
| 240 | 0 | 16 | LIA3OFFSET21 | RD06 | Offset_P3 for Channel_P3 20 |
| 242 | 0 | 16 | LIA3OFFSET22 | RD06 | Offset_P3 for Channel_P3 21 |
| 244 | 0 | 16 | LIA3OFFSET23 | RD06 | Offset_P3 for Channel_P3 22 |
| 246 | 0 | 16 | LIA3OFFSET24 | RD06 | Offset_P3 for Channel_P3 23 |
| 248 | 0 | 16 | LIA3OFFSET25 | RD06 | Offset_P3 for Channel_P3 24 |
| 250 | 0 | 16 | LIA3OFFSET26 | RD06 | Offset_P3 for Channel_P3 25 |
| 252 | 0 | 16 | LIA3OFFSET27 | RD06 | Offset_P3 for Channel_P3 26 |
| 254 | 0 | 16 | LIA3OFFSET28 | RD06 | Offset_P3 for Channel_P3 27 |
| 256 | 0 | 16 | LIA3OFFSET29 | RD06 | Offset_P3 for Channel_P3 28 |
| 258 | 0 | 16 | LIA3OFFSET30 | RD06 | Offset_P3 for Channel_P3 29 |
| 260 | 0 | 16 | LIA3OFFSET31 | RD06 | Offset_P3 for Channel_P3 30 |
| 262 | 0 | 16 | LIA3OFFSET32 | RD06 | Offset_P3 for Channel_P3 31 |
| 264 | 0 | 16 | LIA4OFFSET1 | RD06 | Offset_P4 for Channel_P4 0 |
| 266 | 0 | 16 | LIA4OFFSET2 | RD06 | Offset_P4 for Channel_P4 1 |
| 268 | 0 | 16 | LIA4OFFSET3 | RD06 | Offset_P4 for Channel_P4 2 |
| 270 | 0 | 16 | LIA4OFFSET4 | RD06 | Offset_P4 for Channel_P4 3 |
| 272 | 0 | 16 | LIA4OFFSET5 | RD06 | Offset_P4 for Channel_P4 4 |
| 274 | 0 | 16 | LIA4OFFSET6 | RD06 | Offset_P4 for Channel_P4 5 |
| 276 | 0 | 16 | LIA4OFFSET7 | RD06 | Offset_P4 for Channel_P4 6 |
| 278 | 0 | 16 | LIA4OFFSET8 | RD06 | Offset_P4 for Channel_P4 7 |
| 280 | 0 | 16 | LIA4OFFSET9 | RD06 | Offset_P4 for Channel_P4 8 |
| 282 | 0 | 16 | LIA4OFFSET10 | RD06 | Offset_P4 for Channel_P4 9 |
| 284 | 0 | 16 | LIA4OFFSET11 | RD06 | Offset_P4 for Channel_P4 10 |
| 286 | 0 | 16 | LIA4OFFSET12 | RD06 | Offset_P4 for Channel_P4 11 |
| 288 | 0 | 16 | LIA4OFFSET13 | RD06 | Offset_P4 for Channel_P4 12 |
| 290 | 0 | 16 | LIA4OFFSET14 | RD06 | Offset_P4 for Channel_P4 13 |
| 292 | 0 | 16 | LIA4OFFSET15 | RD06 | Offset_P4 for Channel_P4 14 |
| 294 | 0 | 16 | LIA4OFFSET16 | RD06 | Offset_P4 for Channel_P4 15 |
| 296 | 0 | 16 | LIA4OFFSET17 | RD06 | Offset_P4 for Channel_P4 16 |
| 298 | 0 | 16 | LIA4OFFSET18 | RD06 | Offset_P4 for Channel_P4 17 |
| 300 | 0 | 16 | LIA4OFFSET19 | RD06 | Offset_P4 for Channel_P4 18 |
| 302 | 0 | 16 | LIA4OFFSET20 | RD06 | Offset_P4 for Channel_P4 19 |
| 304 | 0 | 16 | LIA4OFFSET21 | RD06 | Offset_P4 for Channel_P4 20 |
| 306 | 0 | 16 | LIA4OFFSET22 | RD06 | Offset_P4 for Channel_P4 21 |
| 308 | 0 | 16 | LIA4OFFSET23 | RD06 | Offset_P4 for Channel_P4 22 |
| 310 | 0 | 16 | LIA4OFFSET24 | RD06 | Offset_P4 for Channel_P4 23 |
| 312 | 0 | 16 | LIA4OFFSET25 | RD06 | Offset_P4 for Channel_P4 24 |
| 314 | 0 | 16 | LIA4OFFSET26 | RD06 | Offset_P4 for Channel_P4 25 |
| 316 | 0 | 16 | LIA4OFFSET27 | RD06 | Offset_P4 for Channel_P4 26 |
| 318 | 0 | 16 | LIA4OFFSET28 | RD06 | Offset_P4 for Channel_P4 27 |
| 320 | 0 | 16 | LIA4OFFSET29 | RD06 | Offset_P4 for Channel_P4 28 |
| 322 | 0 | 16 | LIA4OFFSET30 | RD06 | Offset_P4 for Channel_P4 29 |
| 324 | 0 | 16 | LIA4OFFSET31 | RD06 | Offset_P4 for Channel_P4 30 |
| 326 | 0 | 16 | LIA4OFFSET32 | RD06 | Offset_P1 for Channel_P4 31 |
| 328 | 0 | 16 | LIA5OFFSET1 | RD06 | Offset_P5 for Channel_P5 0 |
| 330 | 0 | 16 | LIA5OFFSET2 | RD06 | Offset_P5 for Channel_P5 1 |
| 332 | 0 | 16 | LIA5OFFSET3 | RD06 | Offset_P5 for Channel_P5 2 |
| 334 | 0 | 16 | LIA5OFFSET4 | RD06 | Offset_P5 for Channel_P5 3 |
| 336 | 0 | 16 | LIA5OFFSET5 | RD06 | Offset_P5 for Channel_P5 4 |
| 338 | 0 | 16 | LIA5OFFSET6 | RD06 | Offset_P5 for Channel_P5 5 |
| 340 | 0 | 16 | LIA5OFFSET7 | RD06 | Offset_P5 for Channel_P5 6 |
| 342 | 0 | 16 | LIA5OFFSET8 | RD06 | Offset_P5 for Channel_P5 7 |
| 344 | 0 | 16 | LIA5OFFSET9 | RD06 | Offset_P5 for Channel_P5 8 |
| 346 | 0 | 16 | LIA5OFFSET10 | RD06 | Offset_P5 for Channel_P5 9 |
| 348 | 0 | 16 | LIA5OFFSET11 | RD06 | Offset_P5 for Channel_P5 10 |
| 350 | 0 | 16 | LIA5OFFSET12 | RD06 | Offset_P5 for Channel_P5 11 |
| 352 | 0 | 16 | LIA5OFFSET13 | RD06 | Offset_P5 for Channel_P5 12 |
| 354 | 0 | 16 | LIA5OFFSET14 | RD06 | Offset_P5 for Channel_P5 13 |
| 356 | 0 | 16 | LIA5OFFSET15 | RD06 | Offset_P5 for Channel_P5 14 |
| 358 | 0 | 16 | LIA5OFFSET16 | RD06 | Offset_P5 for Channel_P5 15 |
| 360 | 0 | 16 | LIA5OFFSET17 | RD06 | Offset_P5 for Channel_P5 16 |
| 362 | 0 | 16 | LIA5OFFSET18 | RD06 | Offset_P5 for Channel_P5 17 |
| 364 | 0 | 16 | LIA5OFFSET19 | RD06 | Offset_P5 for Channel_P5 18 |
| 366 | 0 | 16 | LIA5OFFSET20 | RD06 | Offset_P5 for Channel_P5 19 |
| 368 | 0 | 16 | LIA5OFFSET21 | RD06 | Offset_P5 for Channel_P5 20 |
| 370 | 0 | 16 | LIA5OFFSET22 | RD06 | Offset_P5 for Channel_P5 21 |
| 372 | 0 | 16 | LIA5OFFSET23 | RD06 | Offset_P5 for Channel_P5 22 |
| 374 | 0 | 16 | LIA5OFFSET24 | RD06 | Offset_P5 for Channel_P5 23 |
| 376 | 0 | 16 | LIA5OFFSET25 | RD06 | Offset_P5 for Channel_P5 24 |
| 378 | 0 | 16 | LIA5OFFSET26 | RD06 | Offset_P5 for Channel_P5 25 |
| 380 | 0 | 16 | LIA5OFFSET27 | RD06 | Offset_P5 for Channel_P5 26 |
| 382 | 0 | 16 | LIA5OFFSET28 | RD06 | Offset_P5 for Channel_P5 27 |
| 384 | 0 | 16 | LIA5OFFSET29 | RD06 | Offset_P5 for Channel_P5 28 |
| 386 | 0 | 16 | LIA5OFFSET30 | RD06 | Offset_P5 for Channel_P5 29 |
| 388 | 0 | 16 | LIA5OFFSET31 | RD06 | Offset_P5 for Channel_P5 30 |
| 390 | 0 | 16 | LIA5OFFSET32 | RD06 | Offset_P5 for Channel_P5 31 |
| 392 | 0 | 16 | LIA6OFFSET1 | RD06 | Offset_P6 for Channel_P6 0 |
| 394 | 0 | 16 | LIA6OFFSET2 | RD06 | Offset_P6 for Channel_P6 1 |
| 396 | 0 | 16 | LIA6OFFSET3 | RD06 | Offset_P6 for Channel_P6 2 |
| 398 | 0 | 16 | LIA6OFFSET4 | RD06 | Offset_P6 for Channel_P6 3 |
| 400 | 0 | 16 | LIA6OFFSET5 | RD06 | Offset_P6 for Channel_P6 4 |
| 402 | 0 | 16 | LIA6OFFSET6 | RD06 | Offset_P6 for Channel_P6 5 |
| 404 | 0 | 16 | LIA6OFFSET7 | RD06 | Offset_P6 for Channel_P6 6 |
| 406 | 0 | 16 | LIA6OFFSET8 | RD06 | Offset_P6 for Channel_P6 7 |
| 408 | 0 | 16 | LIA6OFFSET9 | RD06 | Offset_P6 for Channel_P6 8 |
| 410 | 0 | 16 | LIA6OFFSET10 | RD06 | Offset_P6 for Channel_P6 9 |
| 412 | 0 | 16 | LIA6OFFSET11 | RD06 | Offset_P6 for Channel_P6 10 |
| 414 | 0 | 16 | LIA6OFFSET12 | RD06 | Offset_P6 for Channel_P6 11 |
| 416 | 0 | 16 | LIA6OFFSET13 | RD06 | Offset_P6 for Channel_P6 12 |
| 418 | 0 | 16 | LIA6OFFSET14 | RD06 | Offset_P6 for Channel_P6 13 |
| 420 | 0 | 16 | LIA6OFFSET15 | RD06 | Offset_P6 for Channel_P6 14 |
| 422 | 0 | 16 | LIA6OFFSET16 | RD06 | Offset_P6 for Channel_P6 15 |
| 424 | 0 | 16 | LIA6OFFSET17 | RD06 | Offset_P6 for Channel_P6 16 |
| 426 | 0 | 16 | LIA6OFFSET18 | RD06 | Offset_P6 for Channel_P6 17 |
| 428 | 0 | 16 | LIA6OFFSET19 | RD06 | Offset_P6 for Channel_P6 18 |
| 430 | 0 | 16 | LIA6OFFSET20 | RD06 | Offset_P6 for Channel_P6 19 |
| 432 | 0 | 16 | LIA6OFFSET21 | RD06 | Offset_P6 for Channel_P6 20 |
| 434 | 0 | 16 | LIA6OFFSET22 | RD06 | Offset_P6 for Channel_P6 21 |
| 436 | 0 | 16 | LIA6OFFSET23 | RD06 | Offset_P6 for Channel_P6 22 |
| 438 | 0 | 16 | LIA6OFFSET24 | RD06 | Offset_P6 for Channel_P6 23 |
| 440 | 0 | 16 | LIA6OFFSET25 | RD06 | Offset_P6 for Channel_P6 24 |
| 442 | 0 | 16 | LIA6OFFSET26 | RD06 | Offset_P6 for Channel_P6 25 |
| 444 | 0 | 16 | LIA6OFFSET27 | RD06 | Offset_P6 for Channel_P6 26 |
| 446 | 0 | 16 | LIA6OFFSET28 | RD06 | Offset_P6 for Channel_P6 27 |
| 448 | 0 | 16 | LIA6OFFSET29 | RD06 | Offset_P6 for Channel_P6 28 |
| 450 | 0 | 16 | LIA6OFFSET30 | RD06 | Offset_P6 for Channel_P6 29 |
| 452 | 0 | 16 | LIA6OFFSET31 | RD06 | Offset_P6 for Channel_P6 30 |
| 454 | 0 | 16 | LIA6OFFSET32 | RD06 | Offset_P6 for Channel_P6 31 |
| 456 | 0 | 16 | LIA7OFFSET1 | RD06 | Offset_P7 for Channel_P7 0 |
| 458 | 0 | 16 | LIA7OFFSET2 | RD06 | Offset_P7 for Channel_P7 1 |
| 460 | 0 | 16 | LIA7OFFSET3 | RD06 | Offset_P7 for Channel_P7 2 |
| 462 | 0 | 16 | LIA7OFFSET4 | RD06 | Offset_P7 for Channel_P7 3 |
| 464 | 0 | 16 | LIA7OFFSET5 | RD06 | Offset_P7 for Channel_P7 4 |
| 466 | 0 | 16 | LIA7OFFSET6 | RD06 | Offset_P7 for Channel_P7 5 |
| 468 | 0 | 16 | LIA7OFFSET7 | RD06 | Offset_P7 for Channel_P7 6 |
| 470 | 0 | 16 | LIA7OFFSET8 | RD06 | Offset_P7 for Channel_P7 7 |
| 472 | 0 | 16 | LIA7OFFSET9 | RD06 | Offset_P7 for Channel_P7 8 |
| 474 | 0 | 16 | LIA7OFFSET10 | RD06 | Offset_P7 for Channel_P7 9 |
| 476 | 0 | 16 | LIA7OFFSET11 | RD06 | Offset_P7 for Channel_P7 10 |
| 478 | 0 | 16 | LIA7OFFSET12 | RD06 | Offset_P7 for Channel_P7 11 |
| 480 | 0 | 16 | LIA7OFFSET13 | RD06 | Offset_P7 for Channel_P7 12 |
| 482 | 0 | 16 | LIA7OFFSET14 | RD06 | Offset_P7 for Channel_P7 13 |
| 484 | 0 | 16 | LIA7OFFSET15 | RD06 | Offset_P7 for Channel_P7 14 |
| 486 | 0 | 16 | LIA7OFFSET16 | RD06 | Offset_P7 for Channel_P7 15 |
| 488 | 0 | 16 | LIA7OFFSET17 | RD06 | Offset_P7 for Channel_P7 16 |
| 490 | 0 | 16 | LIA7OFFSET18 | RD06 | Offset_P7 for Channel_P7 17 |
| 492 | 0 | 16 | LIA7OFFSET19 | RD06 | Offset_P7 for Channel_P7 18 |
| 494 | 0 | 16 | LIA7OFFSET20 | RD06 | Offset_P7 for Channel_P7 19 |
| 496 | 0 | 16 | LIA7OFFSET21 | RD06 | Offset_P7 for Channel_P7 20 |
| 498 | 0 | 16 | LIA7OFFSET22 | RD06 | Offset_P7 for Channel_P7 21 |
| 500 | 0 | 16 | LIA7OFFSET23 | RD06 | Offset_P7 for Channel_P722 |
| 502 | 0 | 16 | LIA7OFFSET24 | RD06 | Offset_P7 for Channel_P7 23 |
| 504 | 0 | 16 | LIA7OFFSET25 | RD06 | Offset_P7 for Channel_P7 24 |
| 506 | 0 | 16 | LIA7OFFSET26 | RD06 | Offset_P7 for Channel_P7 25 |
| 508 | 0 | 16 | LIA7OFFSET27 | RD06 | Offset_P7 for Channel_P7 26 |
| 510 | 0 | 16 | LIA7OFFSET28 | RD06 | Offset_P7 for Channel_P7 27 |
| 512 | 0 | 16 | LIA7OFFSET29 | RD06 | Offset_P7 for Channel_P7 28 |
| 514 | 0 | 16 | LIA7OFFSET30 | RD06 | Offset_P7 for Channel_P7 29 |
| 516 | 0 | 16 | LIA7OFFSET31 | RD06 | Offset_P7 for Channel_P7 30 |
| 518 | 0 | 16 | LIA7OFFSET32 | RD06 | Offset_P7 for Channel_P7 31 |
| 520 | 0 | 16 | LIA8OFFSET1 | RD06 | Offset_P8 for Channel_P8 0 |
| 522 | 0 | 16 | LIA8OFFSET2 | RD06 | Offset_P8 for Channel_P8 1 |
| 524 | 0 | 16 | LIA8OFFSET3 | RD06 | Offset_P8 for Channel_P8 2 |
| 526 | 0 | 16 | LIA8OFFSET4 | RD06 | Offset_P8 for Channel_P8 3 |
| 528 | 0 | 16 | LIA8OFFSET5 | RD06 | Offset_P8 for Channel_P8 4 |
| 530 | 0 | 16 | LIA8OFFSET6 | RD06 | Offset_P8 for Channel_P8 5 |
| 532 | 0 | 16 | LIA8OFFSET7 | RD06 | Offset_P8 for Channel_P8 6 |
| 534 | 0 | 16 | LIA8OFFSET8 | RD06 | Offset_P8 for Channel_P8 7 |
| 536 | 0 | 16 | LIA8OFFSET9 | RD06 | Offset_P8 for Channel_P8 8 |
| 538 | 0 | 16 | LIA8OFFSET10 | RD06 | Offset_P8 for Channel_P8 9 |
| 540 | 0 | 16 | LIA8OFFSET11 | RD06 | Offset_P8 for Channel_P8 10 |
| 542 | 0 | 16 | LIA8OFFSET12 | RD06 | Offset_P8 for Channel_P8 11 |
| 544 | 0 | 16 | LIA8OFFSET13 | RD06 | Offset_P8 for Channel_P8 12 |
| 546 | 0 | 16 | LIA8OFFSET14 | RD06 | Offset_P8 for Channel_P8 13 |
| 548 | 0 | 16 | LIA8OFFSET15 | RD06 | Offset_P8 for Channel_P8 14 |
| 550 | 0 | 16 | LIA8OFFSET16 | RD06 | Offset_P8 for Channel_P8 15 |
| 552 | 0 | 16 | LIA8OFFSET17 | RD06 | Offset_P8 for Channel_P8 16 |
| 554 | 0 | 16 | LIA8OFFSET18 | RD06 | Offset_P8 for Channel_P8 17 |
| 556 | 0 | 16 | LIA8OFFSET19 | RD06 | Offset_P8 for Channel_P8 18 |
| 558 | 0 | 16 | LIA8OFFSET20 | RD06 | Offset_P8 for Channel_P8 19 |
| 560 | 0 | 16 | LIA8OFFSET21 | RD06 | Offset_P8 for Channel_P8 20 |
| 562 | 0 | 16 | LIA8OFFSET22 | RD06 | Offset_P8 for Channel_P8 21 |
| 564 | 0 | 16 | LIA8OFFSET23 | RD06 | Offset_P8 for Channel_P8 22 |
| 566 | 0 | 16 | LIA8OFFSET24 | RD06 | Offset_P8 for Channel_P8 23 |
| 568 | 0 | 16 | LIA8OFFSET25 | RD06 | Offset_P8 for Channel_P8 24 |
| 570 | 0 | 16 | LIA8OFFSET26 | RD06 | Offset_P8 for Channel_P8 25 |
| 572 | 0 | 16 | LIA8OFFSET27 | RD06 | Offset_P8 for Channel_P8 26 |
| 574 | 0 | 16 | LIA8OFFSET28 | RD06 | Offset_P8 for Channel_P8 27 |
| 576 | 0 | 16 | LIA8OFFSET29 | RD06 | Offset_P8 for Channel_P8 28 |
| 578 | 0 | 16 | LIA8OFFSET30 | RD06 | Offset_P8 for Channel_P8 29 |
| 580 | 0 | 16 | LIA8OFFSET31 | RD06 | Offset_P8 for Channel_P8 30 |
| 582 | 0 | 16 | LIA8OFFSET32 | RD06 | Offset_P8 for Channel_P8 31 |
| 584 | 0 | 16 | LIA9OFFSET1 | RD06 | Offset_P9 for Channel_P9 0 |
| 586 | 0 | 16 | LIA9OFFSET2 | RD06 | Offset_P9 for Channel_P9 1 |
| 588 | 0 | 16 | LIA9OFFSET3 | RD06 | Offset_P9 for Channel_P9 2 |
| 590 | 0 | 16 | LIA9OFFSET4 | RD06 | Offset_P9 for Channel_P9 3 |
| 592 | 0 | 16 | LIA9OFFSET5 | RD06 | Offset_P9 for Channel_P9 4 |
| 594 | 0 | 16 | LIA9OFFSET6 | RD06 | Offset_P9 for Channel_P9 5 |
| 596 | 0 | 16 | LIA9OFFSET7 | RD06 | Offset_P9 for Channel_P9 6 |
| 598 | 0 | 16 | LIA9OFFSET8 | RD06 | Offset_P9 for Channel_P9 7 |
| 600 | 0 | 16 | LIA9OFFSET9 | RD06 | Offset_P9 for Channel_P9 8 |
| 602 | 0 | 16 | LIA9OFFSET10 | RD06 | Offset_P9 for Channel_P9 9 |
| 604 | 0 | 16 | LIA9OFFSET11 | RD06 | Offset_P9 for Channel_P9 10 |
| 606 | 0 | 16 | LIA9OFFSET12 | RD06 | Offset_P9 for Channel_P9 11 |
| 608 | 0 | 16 | LIA9OFFSET13 | RD06 | Offset_P9 for Channel_P9 12 |
| 610 | 0 | 16 | LIA9OFFSET14 | RD06 | Offset_P9 for Channel_P9 13 |
| 612 | 0 | 16 | LIA9OFFSET15 | RD06 | Offset_P9 for Channel_P9 14 |
| 614 | 0 | 16 | LIA9OFFSET16 | RD06 | Offset_P9 for Channel_P9 15 |
| 616 | 0 | 16 | LIA9OFFSET17 | RD06 | Offset_P9 for Channel_P9 16 |
| 618 | 0 | 16 | LIA9OFFSET18 | RD06 | Offset_P9 for Channel_P9 17 |
| 620 | 0 | 16 | LIA9OFFSET19 | RD06 | Offset_P9 for Channel_P9 18 |
| 622 | 0 | 16 | LIA9OFFSET20 | RD06 | Offset_P9 for Channel_P9 19 |
| 624 | 0 | 16 | LIA9OFFSET21 | RD06 | Offset_P9 for Channel_P9 20 |
| 626 | 0 | 16 | LIA9OFFSET22 | RD06 | Offset_P9 for Channel_P9 21 |
| 628 | 0 | 16 | LIA9OFFSET23 | RD06 | Offset_P9 for Channel_P9 22 |
| 630 | 0 | 16 | LIA9OFFSET24 | RD06 | Offset_P9 for Channel_P9 23 |
| 632 | 0 | 16 | LIA9OFFSET25 | RD06 | Offset_P9 for Channel_P9 24 |
| 634 | 0 | 16 | LIA9OFFSET26 | RD06 | Offset_P9 for Channel_P9 25 |
| 636 | 0 | 16 | LIA9OFFSET27 | RD06 | Offset_P9 for Channel_P9 26 |
| 638 | 0 | 16 | LIA9OFFSET28 | RD06 | Offset_P9 for Channel_P9 27 |
| 640 | 0 | 16 | LIA9OFFSET29 | RD06 | Offset_P9 for Channel_P9 28 |
| 642 | 0 | 16 | LIA9OFFSET30 | RD06 | Offset_P9 for Channel_P9 29 |
| 644 | 0 | 16 | LIA9OFFSET31 | RD06 | Offset_P9 for Channel_P9 30 |
| 646 | 0 | 16 | LIA9OFFSET32 | RD06 | Offset_P9 for Channel_P9 31 |
| 648 | 0 | 16 | SSWBIAS | RD06 | SpectroBiasAmpl for Spectro Bias Channel 0 |
| 650 | 0 | 16 | SLWBIAS | RD06 | SpectroBiasAmpl for Spectro Bias Channel 0 |
| 652 | 0 | 16 | SSWPHASE | RD06 | SpectroPhaseShift for Spectro Bias Channel 0 |
| 654 | 0 | 16 | SLWPHASE | RD06 | SpectroPhaseShift for Spectro Bias Channel 01 |
| 656 | 0 0 1 2 | 16 1 1 1 | SPECJFETSTAT SLWJFETPWR1 SSWJFETPWR1 SSWJFETPWR2 | RD06 | SLW_JFET1 SSW_JFET1 SSW_JFET2 |
| 658 | 0 | 16 | SSWJFET1V | RD06 | SSW_VSS for VSS Channel ID 0 |
| 660 | 0 | 16 | SSWJFET2V | RD06 | SSW_VSS for VSS Channel ID 1 |
| 662 | 0 | 16 | SLWJFET1V | RD06 | SLW_VSS for VSS Channel ID 0 |
| 664 | 0 | 16 | SPECHTRV | RD06 | Spectro Heater Bias |
| 666 | 0 | 16 | SLIA1OFFSET1 | RD06 | Offset_S1 for Channel_S1 0 |
| 668 | 0 | 16 | SLIA1OFFSET2 | RD06 | Offset_S1 for Channel_S1 1 |
| 670 | 0 | 16 | SLIA1OFFSET3 | RD06 | Offset_S1 for Channel_S1 2 |
| 672 | 0 | 16 | SLIA1OFFSET4 | RD06 | Offset_S1 for Channel_S1 3 |
| 674 | 0 | 16 | SLIA1OFFSET5 | RD06 | Offset_S1 for Channel_S1 4 |
| 676 | 0 | 16 | SLIA1OFFSET6 | RD06 | Offset_S1 for Channel_S1 5 |
| 678 | 0 | 16 | SLIA1OFFSET7 | RD06 | Offset_S1 for Channel_S1 6 |
| 680 | 0 | 16 | SLIA1OFFSET8 | RD06 | Offset_S1 for Channel_S1 7 |
| 682 | 0 | 16 | SLIA1OFFSET9 | RD06 | Offset_S1 for Channel_S1 8 |
| 684 | 0 | 16 | SLIA1OFFSET10 | RD06 | Offset_S1 for Channel_S1 9 |
| 686 | 0 | 16 | SLIA1OFFSET11 | RD06 | Offset_S1 for Channel_S1 10 |
| 688 | 0 | 16 | SLIA1OFFSET12 | RD06 | Offset_S1 for Channel_S1 11 |
| 690 | 0 | 16 | SLIA1OFFSET13 | RD06 | Offset_S1 for Channel_S1 12 |
| 692 | 0 | 16 | SLIA1OFFSET14 | RD06 | Offset_S1 for Channel_S1 13 |
| 694 | 0 | 16 | SLIA1OFFSET15 | RD06 | Offset_S1 for Channel_S1 14 |
| 696 | 0 | 16 | SLIA1OFFSET16 | RD06 | Offset_S1 for Channel_S1 15 |
| 698 | 0 | 16 | SLIA1OFFSET17 | RD06 | Offset_S1 for Channel_S1 16 |
| 700 | 0 | 16 | SLIA1OFFSET18 | RD06 | Offset_S1 for Channel_S1 17 |
| 702 | 0 | 16 | SLIA1OFFSET19 | RD06 | Offset_S1 for Channel_S1 18 |
| 704 | 0 | 16 | SLIA1OFFSET20 | RD06 | Offset_S1 for Channel_S1 19 |
| 706 | 0 | 16 | SLIA1OFFSET21 | RD06 | Offset_S1 for Channel_S1 20 |
| 708 | 0 | 16 | SLIA1OFFSET22 | RD06 | Offset_S1 for Channel_S1 21 |
| 710 | 0 | 16 | SLIA1OFFSET23 | RD06 | Offset_S1 for Channel_S1 22 |
| 712 | 0 | 16 | SLIA1OFFSET24 | RD06 | Offset_S1 for Channel_S1 23 |
| 714 | 0 | 16 | SLIA2OFFSET1 | RD06 | Offset_S2 for Channel_S2 0 |
| 716 | 0 | 16 | SLIA2OFFSET2 | RD06 | Offset_S2 for Channel_S2 1 |
| 718 | 0 | 16 | SLIA2OFFSET3 | RD06 | Offset_S2 for Channel_S2 2 |
| 720 | 0 | 16 | SLIA2OFFSET4 | RD06 | Offset_S2 for Channel_S2 3 |
| 722 | 0 | 16 | SLIA2OFFSET5 | RD06 | Offset_S2 for Channel_S2 4 |
| 724 | 0 | 16 | SLIA2OFFSET6 | RD06 | Offset_S2 for Channel_S2 5 |
| 726 | 0 | 16 | SLIA2OFFSET7 | RD06 | Offset_S2 for Channel_S2 6 |
| 728 | 0 | 16 | SLIA2OFFSET8 | RD06 | Offset_S2 for Channel_S2 7 |
| 730 | 0 | 16 | SLIA2OFFSET9 | RD06 | Offset_S2 for Channel_S2 8 |
| 732 | 0 | 16 | SLIA2OFFSET10 | RD06 | Offset_S2 for Channel_S2 9 |
| 734 | 0 | 16 | SLIA2OFFSET11 | RD06 | Offset_S2 for Channel_S2 10 |
| 736 | 0 | 16 | SLIA2OFFSET12 | RD06 | Offset_S2 for Channel_S2 11 |
| 738 | 0 | 16 | SLIA2OFFSET13 | RD06 | Offset_S2 for Channel_S2 12 |
| 740 | 0 | 16 | SLIA2OFFSET14 | RD06 | Offset_S2 for Channel_S2 13 |
| 742 | 0 | 16 | SLIA2OFFSET15 | RD06 | Offset_S2 for Channel_S2 14 |
| 744 | 0 | 16 | SLIA2OFFSET16 | RD06 | Offset_S2 for Channel_S2 15 |
| 746 | 0 | 16 | SLIA2OFFSET17 | RD06 | Offset_S2 for Channel_S2 16 |
| 748 | 0 | 16 | SLIA2OFFSET18 | RD06 | Offset_S2 for Channel_S2 17 |
| 750 | 0 | 16 | SLIA2OFFSET19 | RD06 | Offset_S2 for Channel_S2 18 |
| 752 | 0 | 16 | SLIA2OFFSET20 | RD06 | Offset_S2 for Channel_S2 19 |
| 754 | 0 | 16 | SLIA2OFFSET21 | RD06 | Offset_S2 for Channel_S2 20 |
| 756 | 0 | 16 | SLIA2OFFSET22 | RD06 | Offset_S2 for Channel_S2 21 |
| 758 | 0 | 16 | SLIA2OFFSET23 | RD06 | Offset_S2 for Channel_S2 22 |
| 760 | 0 | 16 | SLIA2OFFSET24 | RD06 | Offset_S2 for Channel_S2 23 |
| 762 | 0 | 16 | SLIA3OFFSET1 | RD06 | Offset_S3 for Channel_S3 0 |
| 764 | 0 | 16 | SLIA3OFFSET2 | RD06 | Offset_S3 for Channel_S3 1 |
| 766 | 0 | 16 | SLIA3OFFSET3 | RD06 | Offset_S3 for Channel_S3 2 |
| 768 | 0 | 16 | SLIA3OFFSET4 | RD06 | Offset_S3 for Channel_S3 3 |
| 770 | 0 | 16 | SLIA3OFFSET5 | RD06 | Offset_S3 for Channel_S3 4 |
| 772 | 0 | 16 | SLIA3OFFSET6 | RD06 | Offset_S3 for Channel_S3 5 |
| 774 | 0 | 16 | SLIA3OFFSET7 | RD06 | Offset_S3 for Channel_S3 6 |
| 776 | 0 | 16 | SLIA3OFFSET8 | RD06 | Offset_S3 for Channel_S3 7 |
| 778 | 0 | 16 | SLIA3OFFSET9 | RD06 | Offset_S3 for Channel_S3 8 |
| 780 | 0 | 16 | SLIA3OFFSET10 | RD06 | Offset_S3 for Channel_S3 9 |
| 782 | 0 | 16 | SLIA3OFFSET11 | RD06 | Offset_S3 for Channel_S3 10 |
| 784 | 0 | 16 | SLIA3OFFSET12 | RD06 | Offset_S3 for Channel_S3 11 |
| 786 | 0 | 16 | SLIA3OFFSET13 | RD06 | Offset_S3 for Channel_S3 12 |
| 788 | 0 | 16 | SLIA3OFFSET14 | RD06 | Offset_S3 for Channel_S3 13 |
| 790 | 0 | 16 | SLIA3OFFSET15 | RD06 | Offset_S3 for Channel_S3 14 |
| 792 | 0 | 16 | SLIA3OFFSET16 | RD06 | Offset_S3 for Channel_S3 15 |
| 794 | 0 | 16 | SLIA3OFFSET17 | RD06 | Offset_S3 for Channel_S3 16 |
| 796 | 0 | 16 | SLIA3OFFSET18 | RD06 | Offset_S3 for Channel_S3 17 |
| 798 | 0 | 16 | SLIA3OFFSET19 | RD06 | Offset_S3 for Channel_S3 18 |
| 800 | 0 | 16 | SLIA3OFFSET20 | RD06 | Offset_S3 for Channel_S3 19 |
| 802 | 0 | 16 | SLIA3OFFSET21 | RD06 | Offset_S3 for Channel_S3 20 |
| 804 | 0 | 16 | SLIA3OFFSET22 | RD06 | Offset_S3 for Channel_S3 21 |
| 806 | 0 | 16 | SLIA3OFFSET23 | RD06 | Offset_S3 for Channel_S3 22 |
| 808 | 0 | 16 | SLIA3OFFSET24 | RD06 | Offset_S3 for Channel_S3 23 |

Not Available

### Event Reporting

#### Event Report (5,1)

These reports provide information on events in the instrument which may impact on the scientific results, but which should have no impact on the health and safety of the instrument. These will include the reporting of anomalies which have been dealt with by the DPU

##### New Step Report

Indicates a new step in the current operation Mode

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 25 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0501 |
| OBSID |
| BBID |
| MODE |
| STEP |
| Checksum |

##### Peak Up Report

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0504 |
| OBSID |
| BBID |
| Parameters |
| Checksum |

##### Block Not Allocated

The OBS failed to allocate a memory block from the VIRTUOSO Memory Pools

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 29 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0505 |
| OBSID |
| BBID |
| POOLID |
| SIZE |
| ALLOC_RESULT |
| TASKID |
| Checksum |

##### Unknown DRCU Command

The command sent to the DRCU is unknown

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 27 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0509 |
| OBSID |
| BBID |
| COMMAND |
| ACK |
| Checksum |

##### Forbidden DRCU Command

The command sent to the DRCU is forbidden

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 27 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x050A |
| OBSID |
| BBID |
| COMMAND |
| ACK |
| Checksum |

##### Frame ID Error

When copying data frames from the FIFO to the internal science data buffer, the frame ID  is not correct.

(The DPU should also execute a procedure to resynchronise with the science data in the FIFO).

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 25 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0511 |
| OBSID |
| BBID |
| FRAMEID |
| FIFOID |
| Checksum |

##### Frame Length Error

When copying data frames from the FIFO to the internal science data buffer, the frame length is not correct (RD06).

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 27 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0507 |
| OBSID |
| BBID |
| FRAMEID |
| LENGTHRD |
| LENGTHEX |
| Checksum |

##### Frame Checksum Error

When copying data frames from the FIFO to the internal science data buffer, the Checksum word at the end of the frame is not correct.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 27 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0508 |
| OBSID |
| BBID |
| FRAMEID |
| CHKWORDEX |
| CHKWORDRD |
| Checksum |

##### Subsystem Timeout Error

This is issued when the DRCU responds with a timeout to a sent command.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 25 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x050B |
| OBSID |
| BBID |
| CID |
| ACK |
| Checksum |

##### Subsystem Response Error

This is issued when there is an error in the transmission of a GET command to the DRCU, i.e. when the CID returned by the DRCU is not identical to the CID sent.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 25 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x050C |
| OBSID |
| BBID |
| CID |
| RESPONSE |
| Checksum |

##### DPU Pool Full Error

This (a warning more than an error) is issued when the DPU internal memory pools (where the TM packets are stored before being sent to the 1553 I/F, it will be detailed in the ADD some time) reach a 80% occupation.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 21 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x050D |
| OBSID |
| BBID |
| Checksum |

##### DPU FIFO Full Error

This (a warning more than an error) is issued when the DPU internal VIRTUOSO FIFOs (used to exchange messages between OBS tasks, it will be detailed in the ADD some time) reach a 80% occupation.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 21 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x050E |
| OBSID |
| BBID |
| Checksum |

#### Exception Report (5,2)

These reports indicate an anomaly with the instrument that cannot be dealt with by the DPU and that action should be taken by the spacecraft.

##### DRCU Anomaly

OBS has detected an unrecoverable anomaly in the DRCU  - expected action is for the CDMS to switch off the DRCU

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0520 |
| OBSID |
| BBID |
| ANOMALYID |
| Checksum |

##### DPU Anomaly

OBS has detected an unrecoverable anomaly in the DPU  - expected action is for the CDMS to switch off the DPU

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0521 |
| OBSID |
| BBID |
| ANOMALYID |
| Checksum |

##### Operations  Anomaly

OBS has detected a problem during an observation  - expected action is for the CDMS to suspend SPIRE commanding until the instrument is ready to resume.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0530 |
| OBSID |
| BBID |
| ANOMALYID |
| Checksum |

##### Resume

Resume SPIRE Commanding - with next subschedule

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID = 0x0531 |
| OBSID |
| BBID |
| Checksum |

#### Error/Alarm Report (5,4)

##### Memory Check Error

Generated when the Boot Software memory check fails

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| EVENTID |
| SID = 0x0540 |
| MEMID | NPAGES |
| Page IDs |
| Checksum |

### Memory Management

#### Memory Dump, Absolute Addresses

This report is issued in response to TC (6,5). Multiple packets may be generated in order to dump the requested address range.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| MEMORYID |  |
| STARTADDR |
| NSAU |
| Data |
| CHK |
| Checksum |

#### Memory Check Report, Absolute Addresses

This report is issued in response to TC (6,9). Multiple packets may be generated in order to dump the requested address range.

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length =17 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| MEMORYID |  |
| STARTADDR |
| NSAU |
| CHK |
| Checksum |

Not Available

### Function Management

This report is issued in response to TC (8,5).

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| FUNCTIONID | ACTIVITYID |
| SID |
| Parameters |
| Checksum |

### Time Management

#### Time Verification Report

This report is generated in response to TC(9,7)

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 17 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| TDPU |
| Checksum |

Not Available

### On Board Scheduling

Not Used

### On Board Monitoring

TBW

Not Available

### Packet Transmission Control

This report is issued in response to TC (14,3).

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| NDEF |
| Type Subtype SID |
| Type Subtype SID |
| Checksum |

### On Board Storage and Retrieval

Not Used

### On Board Traffic Management

Not Used

### Test Service

This report is generated in response to TC (17,1)

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length = 17 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Checksum |

### On Board Control Procedures

Not Used

### Action/Event Service

Not Used

### Information Distribution Service

#### Distributed Information Packets Report (Service 20,5)

| 0 | 0 | 0 | 0 | 1 | APID1 |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 |
| 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| NPKT |
| APID1 |
| SID1 |
| : : : |
| APIDNPKT |
| SIDNPKT |
| Checksum |

### Science Data

#### Nominal Science Data Report (21,1)

| 0 | 0 | 0 | 0 | 1 | APID |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID |
| OBSID |
| BBID |
| Parameters |
| Checksum |

The parameters field contains one or more blocks of science data all of the same type.  These blocks are the same as the science data frames issued by the DRCU, with the possibility of the frame header and trailer having been removed.

Normally the DPU fills the TM packet with the maximum number of blocks before the packet is made available for transmission. In the event of a flush telecommand the packet is transmitted immediately and will not be filled. The length field, with the APID and SID, allows the number of blocks in the packet to be calculated.

The possible block types are as follows:

##### Photometer Full Array Block

This block of data is inserted in the Parameter field once.

Block Length=294

FrameID=00

288 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### Spectrometer Full Array Block

Block Length=78

FrameID=01

72 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### SMEC Block

In order to meet the telemetry bit rate requirements, this block of data is derived from the science data frame (FrameID=11) received from the MCU by stripping it of its header and trailer.

Acquisition Time MSW

Acquisition Time LSW

Optical encoder Coarse Posn

Optical encoder Fine Posn

LVDT Fine Position

Commanded Current

Actuator Back EMF

This structure is TBC

##### BSM Block

Block Length=13

FrameID=12

Acquisition Time MSW

Acquisition Time LSW

Chop Position

Chop Commanded Current

Chop Back EMF

Jiggle Position

Jiggle Commanded Current

Jiggle Back EMF

Frame_Time MSW

Frame_Time LSW

Checkword

This structure is TBC

##### SCU Block

Block Length=30

FrameID=20

HSK Data

ADC _Flags

Frame_Time MSW

Frame_Time LSW

Checkword

This structure is TBC

The content of the HSK Data is TBD

##### DPU Structure

This block of data is dependant on the observational mode executing.

This structure is TBC

#### Science Type B Data Report (21,2)

| 0 | 0 | 0 | 0 | 1 | APID |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID |
| OBSID |
| BBID |
| Parameters |
| Checksum |

The parameters field contains one or more blocks of science data all of the same type.  These blocks are the same as the science data frames issued by the DRCU, with the possibility of the frame header and trailer having been removed.

Normally the DPU fills the TM packet with the maximum number of blocks before the packet is made available for transmission. In the event of a flush telecommand the packet is transmitted immediately and will not be filled. The length field, with the APID and SID, allows the number of blocks in the packet to be calculated.

The possible block types are as follows:

##### PSW Array Block

Block Length=150

FrameID=02

144 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### PMW Array Block

Block Length=102

FrameID=03

96 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### PLW Array Block

Block Length=54

FrameID=04

48 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### SSW Array Block

Block Length=54

FrameID=05

48 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### SLW Array Block

Block Length=30

FrameID=06

24 Detector Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detectors in the block is TBD

##### SMEC Step Block

It is TBD whether a separate block type is needed for this mode

#### Diagnostic Science Report (21,3)

| 0 | 0 | 0 | 0 | 1 | APID |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID |
| OBSID |
| BBID |
| Parameters |
| Checksum |

The parameters field contains one or more blocks of science data all of the same type.  These blocks are the same as the science data frames issued by the DRCU, with the possibility of the frame header and trailer having been removed.

Normally the DPU fills the TM packet with the maximum number of blocks before the packet is made available for transmission. In the event of a flush telecommand the packet is transmitted immediately and will not be filled. The length field, with the APID and SID, allows the number of blocks in the packet to be calculated.

The possible block types are as follows:

##### Photometer Test Pattern Block

This block of data is inserted in the Parameter field once.

Block Length=294

FrameID=07

288 Test Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The content of the Test Values is TBD

##### Spectrometer Test Pattern Block

Block Length=78

FrameID=0A

72 Test Values

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The content of the Test Values is TBD

##### MCU Trace Block

Block Length=24

FrameID=14

19 Trace Values

Frame_Time MSW

Frame_Time LSW

Checkword

The content of the Trace values is set by command

It is TBD if other trace block structures are possible

##### MCU Test Pattern Block

TBD

##### SCU Test Pattern Block

Block Length=30

FrameID=21

Test Data

Frame_Time MSW

Frame_Time LSW

Checkword

This structure is TBC

The content of the Test Data is TBD

#### Auxiliary Science Data Report (21,4)

| 0 | 0 | 0 | 0 | 1 | APID |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | Count |
| Length |
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 |
| 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| TIME |
| SID |
| OBSID |
| BBID |
| Parameters |
| Checksum |

The parameters field contains one or more blocks of science data all of the same type.  These blocks are the same as the science data frames issued by the DRCU, with the possibility of the frame header and trailer having been removed.

Normally the DPU fills the TM packet with the maximum number of blocks before the packet is made available for transmission. In the event of a flush telecommand the packet is transmitted immediately and will not be filled. The length field, with the APID and SID, allows the number of blocks in the packet to be calculated.

The possible block types are as follows:

##### Photometer Offsets Block

This block of data is inserted in the Parameter field once.

Block Length=294

FrameID=08

288 Offsets

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detector offsets in the block is TBD

##### Spectrometer Offsets Block

Block Length=78

FrameID=09

72 Offsets

ADC_Flags

Frame_Time MSW

Frame_Time LSW

Checkword

The order of the detector offsets in the block is TBD

### Context Saving Service

Not Used

# Parameters

## TC Parameters

### Parameter Definition

| Parameter Name | Service Reference | Type | Size (bits) | Conversion Curve | Constraint Table | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| Activity_ID | (8,4) | Integer | 8 | None | None |  |
| APID | (20,1) (20,2) | Integer | 11 | None | None | Application ID |
| BBID BBINTR BBTYPE BBCOUNT | (8,4) | Constant Integer Integer | 32 2 14 16 | None None None | None None None | Field is split into 3 parts: Location: Bits 0-1 Value: 2 Location: Bits 2-15 Location: Bits 16-31 |
| COUNT | (8,4) | Integer | 16 | None | None | Number of words to be reported |
| CRC | (6,2) | Unsigned Integer | 16 | None | None | Cyclic Redundancy Check - algorithm TBD |
| Data | (6,2) (8,4) |  |  |  |  | Variable length field of any data type. Must be an integer number of 16 bits in length |
| FIFOFLAGS FIFODPU FIFODCU FIFOSCU FIFOMCU | (8,4) | Integer Integer Integer Integer | 16 1 1 1 1 | Flush Flush Flush Flush | None None None None | Flags indicating the FIFO(s) to flush Location: bit 15: DPU science buffer Location: bit 14: DCU FIFO Location: bit 13: SCU FIFO Location: bit 12: MCU FIFO |
| FRAMEID | (8,4) | Integer | 16 | None | Frame |  |
| FUNCTIONID | (8,1) (8,2) (8,4) (8,5) | Integer | 8 | Func_ID | None |  |
| HKPCKTID | (3,1) (3,3) (3,9) | Integer | 16 | None | Hsk_ID | Identifies housekeeping packet definitions |
| HKSID | (3,1) | Unsigned Integer | 16 | None | None | Unique identifier for housekeeping parameter list Value = 0x03nn where nn = a unique number (starting from 0) defining the report structure |
| HKINTERVAL | (3,1) | Unsigned Integer | 16 | None | Period | Time between each packet generated, in milliseconds |
| INDEX | (8,4) | Unsigned Integer | 16 | None | None | Offset into a table( in octets) |
| Length | (6,2) (6,5) (6,9) | Unsigned Integer | 16 | None | None |  |
| MEMORYID | (6,2) (6,5) (6,9) | Integer | 16 | Mem_ID | None | ID of Memory Area to be addressed |
| MODE | (8,4) | Unsigned Integer | 16 | None | None | Observing Mode |
| N | (8,4) | Unsigned Integer | 16 | None | N32 |  |
| NCHOP | (8,4) | Integer | 16 | None | TBD | Number of Chop science data frames to combine into one Chop Science TM Packet |
| NDATA | (8,4) | Unsigned Integer | 16 | None | N32 |  |
| NDCU | (8,4) | Integer | 16 | None | TBD | Number of DCU science data frames to combine into one Chop Science TM Packet |
| NDPU | (8,4) | Integer | 16 | None | TBD | Number of DPU science data frames to combine into one Chop Science TM Packet |
| NJIGGLE | (8,4) | Integer | 16 | None | TBD | Number of Jiggle science data frames to combine into one Chop Science TM Packet |
| NPCKTS | (14,1) (14,2) | Integer | 16 | None | None | Number of packet types to follow |
| NSAU | (6,2) (6,5) (6,9) | Integer | 16 | None | None | Number of SAUs to transfer |
| NSCU | (8,4) | Integer | 16 | None | TBD | Number of SCU science data frames to combine into one Chop Science TM Packet |
| NSMEC | (8,4) | Integer | 16 | None | TBD | Number of SMEC science data frames to combine into one Chop Science TM Packet |
| OBSID | (8,4) | Unsigned Integer | 32 | None | None | Observation ID |
| Parameters | (8,4) |  |  |  |  | Variable length field of 32 data words. |
| SID | (8,1) (8,4) (14,1) (14,2) (20,1) (20,2) | Unsigned Integer | 16 | None | None | Unique identifier for parameter list |
| STARTADDR | (6,2) (6,5) (6,9) | Unsigned Integer | 16 | None | None | Start address of memory area to be addressed |
| STEP | (8,4) | Unsigned Integer | 16 | None | None | Observation Step |
| SUBTYPE | (14,1) (14,2) | Integer | 8 | None | None | Packet Subtype |
| TABLEID | (8,4) | Integer | 16 | None | Tab_ID | Number of an OBS table |
| TABLESIZE | (8,4) | Integer | 16 | None | TBD | Table size in 32 bit words |
| TYPE | (14,1) (14,2) | Integer | 8 | None | None | Packet Type |

### Conversion Curves

| Name | Type | Raw Value | Converted Value | Comments |
| --- | --- | --- | --- | --- |
| Data_Type | Enumerated | 0 1 2 3 | 8Bits 16Bits 24Bits 32Bits |  |
| Func_ID | Enumerated | 0x01 0x02 0x03 0x04 0xC0 0xC1 0xCA | Table Cooler SCAL 300mK Ops Obs DPU | TableManagement CoolerControl SCALControl 300mKControl Operations Observations DPU |
| Info_Type | Enumerated | 0 1 2 3 | Off Target On Target Start Scan End Scan |  |
| Mem_ID | Enumerated | TBD TBD TBD | PROG DATA EEPROM |  |
| Flush | Enumerated | 0 1 | Do not flush FIFO Flush FIFO |  |

### Constraints

| Name | Type | Minimum Value | Maximum Value | Comments |
| --- | --- | --- | --- | --- |
| Bit2 | Range | 0 | 3 | Value held in lowest 2 bits of field |
| Frame | Set |  |  | Set of possible frame IDs, TBD |
| Hsk_ID | Range | 0 | 3 | Allows up to 4 housekeeping packets to be defined |
| Hsk_Value | Range | 0x00000000 | 0x000FFFFF | The value is held in lowest 20 bits of the field |
| N16 | Range | 1 | 112 | Number of 16 bit data words held in a command |
| N32 | Range | 1 | 56 | Number of 32 bit data words held in a command |
| Nsamples | Range | 1 | 32 |  |
| Period | Range | 10 | 60000 | Time period expressed as milliseconds - range is 0.01 to 60 seconds |
| Tab_ID | Range | 0 | 127 |  |

## TM  Parameters

### Parameter Definition

| Name | DRCU Name | Length (bits) | Conversion | Limits | Description |
| --- | --- | --- | --- | --- | --- |
| BBID BBINSTR BBTYPE BBCNT |  | 32 2 14 16 |  |  | Building Block ID Instrument (Bits 0,1) BB type (Bits 2-15) BB sequence count (Bits 16-31) |
| BIASTEMP | BIAS_TEMP | 8 | TBD |  | BIAS Board Temperature |
| CCUTEMP | CcuTempRd | TBD | TBD | TBD | Temperature from sensor on CCHK Board |
| CHOPACCELLIMIT | CAccelLimi | 16, TBC |  |  |  |
| CHOPACCELLOOPSF | CAccelLoopsf | 16, TBC |  |  |  |
| CHOPBEMF | CBEMF | 16, TBC | TBD |  |  |
| CHOPINTTHRESH | IntegTh | 16 |  |  | Integration threshold for the Chop PID Controller |
| CHOPKD | CKd | 16 |  |  | Derivative Gain of Chop Axis PID Controller |
| CHOPKI | CKi | 16 |  |  | Integral Gain of Chop Axis PID Controller |
| CHOPKP | CKp | 16 |  |  | Proportional Gain of Chop Axis PID Controller |
| CHOPLATCHSTAT | BSMLaunchLatch1 | 1 | LATCH |  | Chopper Axis Launch Latch Status |
| CHOPLOOPMODE | ChopLoopMode | 2 | CLOOPMODE |  | Chopper Axis Loop Mode |
| CHOPMODE | ChopMode | 2 | CHOPMODE |  | Chopping Mode |
| CHOPMOTORCURR | CMotorCurrent | 16, TBC | TBD |  |  |
| CHOPMOTORTORQUE | CMotorTorqueC | 16, TBC | TBD |  |  |
| CHOPPERIOD | ChopPeriod | 16 | PERIOD |  | Chop Cycle Period |
| CHOPICNCOEFF1 | ICnCoeff1 | 16, TBC | TBD |  |  |
| CHOPICNCOEFF2 | ICnCoeff2 | 16, TBC | TBD |  |  |
| CHOPICNCOEFF3 | ICnCoeff3 | 16, TBC | TBD |  |  |
| CHOPICNCOEFF4 | ICnCoeff4 | 16, TBC | TBD |  |  |
| CHOPOUTCOEFF1 | CoutCoeff1 | 16, TBC | TBD |  |  |
| CHOPOUTCOEFF2 | CoutCoeff2 | 16, TBC | TBD |  |  |
| CHOPOUTCOEFF3 | CoutCoeff3 | 16, TBC | TBD |  |  |
| CHOPOUTCOEFF4 | CoutCoeff4 | 16, TBC | TBD |  |  |
| CHOPPOSN | ActualPosition | 16 |  |  | Last Absolute position of the Chopper |
| CHOPPOSNERR | CMeanPositionError | 16 |  |  | Mean Position Error over a Chopper movement |
| CHOPPOSNERRLIMIT | CPositionErrorLimit | 16 |  |  | Minimum Value of Chop Position Error that causes a PID error to be reported |
| CHOPPOSN0 | Position0 | 16 |  |  | Chopper Axis Position 0 |
| CHOPPOSN1 | Position1 | 16 |  |  | Chopper Axis Position 1 |
| CHOPRATELIMIT | CRateLimit | 16, TBC |  |  |  |
| CHOPRATELOOPSF | CRateLoopsf | 16, TBC |  |  |  |
| CHOPPROFILEVAL01 | CprofileValue1 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL02 | CprofileValue2 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL03 | CprofileValue3 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL04 | CprofileValue4 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL05 | CprofileValue5 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL06 | CprofileValue6 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL07 | CprofileValue7 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL08 | CprofileValue8 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL09 | CprofileValue9 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL10 | CprofileValue10 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL11 | CprofileValue11 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL12 | CprofileValue12 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL13 | CprofileValue13 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL14 | CprofileValue14 | 16, TBC | TBD |  |  |
| CHOPPROFILEVAL15 | CprofileValue15 | 16, TBC | TBD |  |  |
| CHOPS | ChopNumber | 16 |  |  | Number of Chop Cycles to perform |
| CHOPSENSPOSN1 | CMagnetoResPos1 | 16, TBC | TBD |  |  |
| CHOPSENSPOSN2 | CMagnetoResPos2 | 16, TBC | TBD |  |  |
| CHOPSENSPWR | SensorPwr | 1 | CHOPSENS |  | Chopper axis sensor power status |
| CHOPSTAT | ChopStatus | 16, TBC | TBD |  | Contents of Chop Axis Activity Status Register |
| CHOPSTATECOEFF1 | CStateCoeff1 | 16, TBC | TBD |  |  |
| CHOPSTATECOEFF2 | CStateCoeff2 | 16, TBC | TBD |  |  |
| CHOPSTATECOEFF3 | CStateCoeff3 | 16, TBC | TBD |  |  |
| CHOPSTATECOEFF4 | CStateCoeff4 | 16, TBC | TBD |  |  |
| DAQBIAS5V | DcuDAQp05 | TBD | TBD | TBD | Voltage to DAQ Board on +5V line |
| DAQ9V | DcuDAQp09 | TBD | TBD | TBD | Voltage to DAQ Board on +9V line |
| DAQ9VN | DcuDAQn09 | TBD | TBD | TBD | Voltage to DAQ Board on -9V line |
| DAQTEMP | DAQ_I/F_TEMP | 8 | TBD |  | DAQ I/F Board Temperature |
| DCUDATAFRMS | frame | 8 | DCUFRMS |  | Number of DCU data frames to generate |
| DCUDATAMODE | mode | 5 | DCUMODE |  | DCU data collection mode |
| DCUDATASTAT | start | 1 | DCUSTAT |  | Status of DCU data frame generation |
| DCUFLAGS |  | 16 |  |  | DCU Status -TBC |
| DCUTEMP | DcuTempRd | TBD | TBD | TBD | Temperature from sensor on Distribution Board |
| DPU5V |  | 12 | TBD |  | DPU +5V line voltage |
| DPU15V |  | 12 | TBD |  | DPU +15V line voltage |
| DPU15VN |  | 12 | TBD |  | DPU -15V line voltage |
| DPUI |  | 12 | TBD |  | DPU measured input current |
| DPU_Spare1 |  | 12 | TBD |  |  |
| DPU_Spare2 |  | 12 | TBD |  |  |
| DPU_Spare3 |  | 12 | TBD |  |  |
| DPUTEMP |  | 12 | TBD |  | DPU DC-DC Converter Temperature |
| JIGGKDHIGH | KdHigh | 16, TBC |  |  | Maximum Derivative Gain of Chop Axis PID Controller? |
| JIGGKDLOW | KdLow | 16, TBC |  |  | Minimum Derivative Gain of Chop Axis PID Controller? |
| JIGGFILTHIGH | DerivFilterHigh | 16, TBC |  |  | Filtering Time Constant to calculate the derivative term of the Chop PID Controller |
| JIGGFILTLOW | DerivFilterLow | 16, TBC |  |  | Filtering Time Constant to calculate the derivative term of the Chop PID Controller |
| JIGGKIHIGH | KiHigh | 16, TBC |  |  | Maximum Integral Gain of Chop Axis PID Controller? |
| JIGGKILOW | KiLow | 16, TBC |  |  | Minimum Integral Gain of Chop Axis PID Controller? |
| JIGGKPHIGH | KpHigh | 16, TBC |  |  | Maximum Proportional Gain of Chop Axis PID Controller? |
| JIGGKPLOW | KpLow | 16, TBC |  |  | Minimum Proportional Gain of Chop Axis PID Controller? |
| JIGGLATSTAT | JiggleLaunchLatch | 1 | LATCH |  | Jiggle Axis Launch Latch Status |
| JIGGLOOPMODE | JiggleLoopMode | 2 | JLOOPMODE |  | Jiggle Axis Loop Mode |
| JIGGMEANPERR | MeanPositionError | 16, TBC |  |  | Mean Position Error over a Jiggle movement? |
| JIGGMODE | JiggleMode | 2 | JIGGMODE |  | Jiggling Mode |
| JIGGPERIOD | JigglePeriod | 16, TBC | PERIOD |  | Jiggle Cycle Period – TBD if applicable |
| JIGGPERRLIM | PositionErrorLimit | 16, TBC |  |  | Minimum Value of Chop Position Error that causes a PID error to be reported? |
| JIGGPOSN | ActualPosition | 16, TBC |  |  | Last Absolute position of the Jiggle Axis |
| JIGGPOSN0 | Position0 | 16, TBC |  |  | Jiggle Axis Position 0 |
| JIGGPOSN1 | Position1 | 16, TBC |  |  | Jiggle Axis Position 1 – TBD if applicable |
| JIGGLES | JiggleNumber | 16, TBC |  |  | Number of Jiggle Cycles to perform – TBD if applicable |
| JIGGSATNLIM | IntegrationLimit | 16, TBC |  |  | Integration Saturation Limit for the Chop PID Controller? |
| JIGGSENSPWR | SensorPwr | 1 | JIGGSENS |  | Jiggle Axis sensor power status |
| JIGGSTAT | JiggleStatus | 16, TBC | TBD |  | Contents of Jiggle Axis Activity Status Register |
| LIA01TEMP | LIA_B1TEMP | 8 | TBD |  | LIA Board 1 Temperature |
| LIA02TEMP | LIA_B2_TEMP | 8 | TBD |  | LIA Board 2 Temperature |
| LIA03TEMP | LIA_B3_TEMP | 8 | TBD |  | LIA Board 3 Temperature |
| LIA04TEMP | LIA_B4_TEMP | 8 | TBD |  | LIA Board 4 Temperature |
| LIA05TEMP | LIA_B5_TEMP | 8 | TBD |  | LIA Board 5 Temperature |
| LIA06TEMP | LIA_B6_TEMP | 8 | TBD |  | LIA Board 6 Temperature |
| LIA07TEMP | LIA_B7_TEMP | 8 | TBD |  | LIA Board 7 Temperature |
| LIA08TEMP | LIA_B8_TEMP | 8 | TBD |  | LIA Board 8 Temperature |
| LIA09TEMP | LIA_B9_TEMP | 8 | TBD |  | LIA Board 9 Temperature |
| LIA10TEMP | LIA_B10_TEMP | 8 | TBD |  | LIA Board 10 Temperature |
| LIA11TEMP | LIA_B11_TEMP | 8 | TBD |  | LIA Board 11Temperature |
| LIA12TEMP | LIA_B12_TEMP | 8 | TBD |  | LIA Board 12 Temperature |
| MCUFLAGS |  | 16 |  |  | MCU Status -TBC |
| MCUTIMEOUT | DPU_PollingTime |  |  |  | Maximum Time between commands before an IO error is reported |
| MCUTELSTAT | Telemetry | TBD | TBD |  | Define TM packets to be sent |
| MCUTELSAMP | TelemetrySampling | TBD | TBD |  | Sampling Rate of Telemetry Packets |
| MEMCHK |  | 48 | TBD |  | Memory Check Flags |
| MODE |  | 16 | TBD |  | Instrument Operating Mode |
| MONFLAGS |  | 16 |  |  | Subsystem Monitoring Status -TBC |
| OBSID |  | 32 |  |  | Observation ID |
| OBSLEN |  | 32 |  |  | Time since start of Observation (secs) |
| OTFERROR |  | 16, TBC | TBD |  | On Target Flag Error |
| OTFTIME |  | 48 |  |  | On Target Flag Time |
| OTF |  | 1 | OTF |  | On Target Flag |
| PHOTBIAS9V | DcuBPHp09 | TBD | TBD | TBD | Voltage to Photometer Bias Board on +9V line |
| PHOTBIAS9VN | DcuBPHn09 | TBD | TBD | TBD | Voltage to Photometer Bias Board on -9V line |
| PHOTBIASFREQ | Div_photo_bias | 9 | BIASFREQ |  | Photometry Bias frequency |
| PHOTBIASMODE | Mode_photo_bias | 8 | BIASMODE |  | Photometry bias generator mode |
| PHOTLIA5V | Dc uLPHp05 | TBD | TBD | TBD | Voltage to Photometer LIA Board on +5V line |
| PHOTLIA9V | Dc uLPHp09 | TBD | TBD | TBD | Voltage to Photometer LIA Board on +9V line |
| PHOTLIA9VN | Dc uLPHn09 | TBD | TBD | TBD | Voltage to Photometer LIA Board on -9V line |
| PHOTSAMPFREQ | Div_photo_sampl | 8 | PSAMPFREQ |  | Photometry sampling frequency |
| PHTRV | Ampl_photo_heater | 8 | HTR |  | Voltage applied to photometer heaters |
| PLWBIAS | Ampl_P500 | 8 | BIAS |  | Amplitude of bias for photometry long wave channels |
| PLWJFET1V | Ampl_P500_VSS1 | 8 | JFETV |  | Voltage applied to JFET for photometry long wave channels (group1) |
| PLWJFET2V | Ampl_P500_VSS2 | 8 | JFETV |  | Voltage applied to JFET for photometry long wave channels (group2) |
| PLWPHSE | Phase_shift_P500 | 8 | PHASE |  | Phase shift for long wave photometry channels demodulation |
| PMLWJFETSTAT PMWJFET1STAT PMWJFET2STAT PMWJFET3STAT PMWJFET4STAT PLWJFET1STAT PLWJFET2STAT | P350_JFET_1 P350_JFET_2 P350_JFET_3 P350_JFET_4 P500_JFET_1 P500_JFET_2 | 6 1 1 1 1 1 1 | JFETSTAT JFETSTAT JFETSTAT JFETSTAT JFETSTAT JFETSTAT |  | Status of JFET power for photometry medium wave channels (group1) (Bit 7) Status of JFET power for photometry medium wave channels (group2) (Bit 6) Status of JFET power for photometry medium wave channels (group3) (Bit 5) Status of JFET power for photometry medium wave channels (group4) (Bit 4) Status of JFET power for photometry long wave channels (group1) (Bit 3) Status of JFET power for photometry long wave channels (group2) (Bit 2) |
| PMWBIAS | Ampl_P350 | 8 | BIAS |  | Amplitude of bias for photometry medium wave channels |
| PMWJFET1V | Ampl_P350_VSS1 | 8 | JFETV |  | Voltage applied to JFET for photometry medium wave channels (group1) |
| PMWJFET2V | Ampl_P350_VSS2 | 8 | JFETV |  | Voltage applied to JFET for photometry medium wave channels (group2) |
| PMWJFET3V | Ampl_P350_VSS3 | 8 | JFETV |  | Voltage applied to JFET for photometry medium wave channels (group3) |
| PMWJFET4V | Ampl_P350_VSS4 | 8 | JFETV |  | Voltage applied to JFET for photometry medium wave channels (group4) |
| PMWPHSE | Phase_shift_P350 | 8 | PHASE |  | Phase shift for medium wave photometry channels demodulation |
| PSUTEMP | PsuTempRd | TBD | TBD | TBD | PSU Temperature Sensor reading |
| PSWBIAS | Ampl_P250 | 8 | BIAS |  | Amplitude of bias for photometry short wave channels |
| PSWJFET1V | Ampl_P250_VSS1 | 8 | JFETV |  | Voltage applied to JFET for photometry short wave channels (group1) |
| PSWJFET2V | Ampl_P250_VSS2 | 8 | JFETV |  | Voltage applied to JFET for photometry short wave channels (group2) |
| PSWJFET3V | Ampl_P250_VSS3 | 8 | JFETV |  | Voltage applied to JFET for photometry short wave channels (group3) |
| PSWJFET4V | Ampl_P250_VSS4 | 8 | JFETV |  | Voltage applied to JFET for photometry short wave channels (group4) |
| PSWJFET5V | Ampl_P250_VSS5 | 8 | JFETV |  | Voltage applied to JFET for photometry short wave channels (group5) |
| PSWJFET6V | Ampl_P250_VSS6 | 8 | JFETV |  | Voltage applied to JFET for photometry short wave channels (group6) |
| PSWJFETSTAT PSWJFET1STAT PSWJFET2STAT PSWJFET3STAT PSWJFET4STAT PSWJFET5STAT PSWJFET6STAT | P250_JFET_1 P250_JFET_2 P250_JFET_3 P250_JFET_4 P250_JFET_5 P250_JFET_6 | 6 1 1 1 1 1 1 | JFETSTAT JFETSTAT JFETSTAT JFETSTAT JFETSTAT JFETSTAT |  | Status of JFET power for photometry short wave channels (group1) (Bit 7) Status of JFET power for photometry short wave channels (group2) (Bit 6) Status of JFET power for photometry short wave channels (group3) (Bit 5) Status of JFET power for photometry short wave channels (group4) (Bit 4) Status of JFET power for photometry short wave channels (group5) (Bit 3) Status of JFET power for photometry short wave channels (group6) (Bit 2) |
| PSWPHSE | Phase_shift_P250 | 8 | PHASE |  | Phase shift for short wave photometry channels demodulation |
| RATE |  | 32 |  |  | Data Rate (average bits/sec over observation) |
| SCANMODE | ScanMode | 2 | SCANMODE |  | SMEC Scan Mode |
| SCANS | ScanNumber | 16 |  |  | Number of SMEC Scans to perform TBC |
| SCANSPEED | ScanSpeed | 16 |  |  | SMEC Scan Speed |
| SCANSTART | ScanStart | 16 |  |  | SMEC Scan Start Position |
| SCANSTAT | StartScan | 1 | SCANSTAT |  | SMEC Scan Status |
| SCU5V | ScuCHTp05 | 12, TBC | TBD |  | Voltage on SCU 5Vline |
| SCU9V | ScuCHTp09 | 12, TBC | TBD |  | Voltage on SCU+9Vline |
| SCU9VN | ScuCHTn09 | 12, TBC | TBD |  | Voltage on SCU -9Vline |
| SCUCAL1I | CalibraI1 | 12, TBC | TBD |  | Current in Calibrator 1 |
| SCUCAL1V | CalibraV1 | 12, TBC | TBD |  | Voltage Applied to Calibrator 1 |
| SCUCAL2I | CalibraI2 | 12, TBC | TBD |  | Current in Calibrator 2 |
| SCUCAL2V | CalibraV2 | 12, TBC | TBD |  | Voltage Applied to Calibrator 2 |
| SCUCAL3I | CalibraI3 | 12, TBC | TBD |  | Current in Calibrator 3 |
| SCUCAL3V | CalibraV3 | 12, TBC | TBD |  | Voltage Applied to Calibrator 3 |
| SCUCMDSTAT |  | TBD | TBD |  | Command Interface Control Status Word - TBD |
| SCUDATAFREQ | FrameRate | 16,TBC | TBD |  | Frame Rate |
| SCUDELAY | SubsDelay | 16,TBC | TBD |  | Subsystem Response Time |
| SCUFLAGS |  | 16 |  |  | SCU Status -TBC |
| SCUFRAMES | SeqLength | 16,TBC |  |  | Number of frames |
| SCUHHTRV | HheaterV1 | 12, TBC | TBD |  | Voltage Applied to High Power Heater |
| SCULHTR1V | LHeaterV1 | 12, TBC | TBD |  | Voltage Applied to Low Power Heater1 |
| SCULHTR2V | LheaterV2 | 12, TBC | TBD |  | Voltage Applied to Low Power Heater2 |
| SCULHTR3V | LheaterV3 | 12, TBC | TBD |  | Voltage Applied to Low Power Heater3 |
| SCUSTAT | DRelOnOff | 8 | TBD |  | Status of Distribution Board relays |
| SCUSUBKFREQ | SubKpRate | 16,TBC | TBD |  | Sampling Rate of SubK temperatures |
| SCUSUBSSTAT | SubStatus | TBD | TBD |  | Subsystem Status |
| SCUTEMPSTAT Allocation of bits TBD | TempOnOff | 12 | TBD |  | Power on status of temperature probes |
| SHTRV | Ampl_spectro_heater | 8 | HTR |  | Voltage applied to spectrometer heaters |
| SJFETSTAT SLWJFET1STAT SSWJFET1STAT SSWJFET2STAT | S-LW_JFET_1 S-SW_JFET_1 S-SW_JFET_2 | 6 1 1 1 | JFETSTAT JFETSTAT JFETSTAT |  | Status of JFET power for spectrometry long wave channels (Bit 7) Status of JFET power for spectrometry short wave channels (group1) (Bit 6) Status of JFET power for spectrometry short wave channels (group2) (Bit 5) |
| SLWBIAS | Ampl_S-LW | 8 | BIAS |  | Amplitude of bias for spectrometry long wave channels |
| SLWJFET1V | Ampl_S-LW_VSS1 | 8 | JFETV |  | Voltage applied to JFET for spectrometry long wave channels |
| SLWPHSE | Phase_shift_S-LW | 8 | PHASE |  | Phase shift for long wave spectrometry channels demodulation |
| SMECOENCPWR | EncoderPwr | 3 | TBD |  | SMEC Optical Encoder LED power |
| SMECFILTHIGH | DerivFilterHigh | 16,TBC |  |  | Filtering Time Constant to calculate the derivative term of the SMEC PID Controller |
| SMECFILTLOW | DerivFilterLow | 16,TBC |  |  | Filtering Time Constant to calculate the derivative term of the SMEC PID Controller |
| SMECSATNLIM | IntegrationLimit | 16,TBC |  |  | Integration Saturation Limit for the SMEC PID Controller? |
| SMECINITSTAT |  | TBD |  |  | Initialisation Status of SMEC |
| SMECKDHIGH | KdHigh | 16,TBC |  |  | Maximum Derivative Gain of SMEC PID Controller? |
| SMECKDLOW | KdLow | 16,TBC |  |  | Minimum Derivative Gain of SMEC PID Controller? |
| SMECKIHIGH | KiHigh | 16,TBC |  |  | Maximum Integral Gain of SMEC PID Controller? |
| SMECKILOW | KiLow | 16,TBC |  |  | Minimum Integral Gain of SMEC PID Controller? |
| SMECKPHIGH | KpHigh | 16,TBC |  |  | Maximum Proportional Gain of SMEC PID Controller? |
| SMECKPLOW | KpLow | 16,TBC |  |  | Maximum Proportional Gain of SMEC PID Controller? |
| SMECLAT1STAT | LaunchLatch1 | 1 | LATCH |  | SMEC Launch Latch1 Status |
| SMECLAT2STAT | LaunchLatch2 | 1 | LATCH |  | SMEC Launch Latch2 Status |
| SMECLOOPMODE | LoopMode | 2 | LOOPMODE |  | SMEC Control Loop Mode |
| SMECLVDTPWR | LVDTPwr | 1 | TBD |  | SMEC LVDT Oscillator On/Off |
| SMECMEANPERR | MeanPositionError | 16,TBC | TBD |  | Mean SMEC Position Error over scan |
| SMECMEANVEL | MeanSpeed | 16,TBC | TBD |  | Mean SMEC Speed over scan |
| SMECNOTCHIGH | NotchParamHigh | 16,TBC |  |  | TBW |
| SMECNOTCHLOW | NotchParamLow | 16,TBC |  |  | TBW |
| SMECPOSN | ActualPosition | 16,TBC |  |  | Last Absolute position of the SMEC |
| SMECPERRLIM | PositionErrorLimit | 16,TBC |  |  | Minimum Value of SMEC Position Error that causes a PID error to be reported? |
| SMECSTAT | SmecStatus | 16,TBC | TBD |  | Contents of SMEC Activity Status Register |
| SMECVEL | ActualVelocity | 16,TBC |  |  | Instantaneous Velocity (20Hz filtered) of the SMEC |
| SPECBIAS9V | DcuBSPp09 | TBD | TBD | TBD | Voltage to Spectrometer Bias Board on +9V line |
| SPECBIAS9VN | DcuBSPn09 | TBD | TBD | TBD | Voltage to Spectrometer Bias Board on -9V line |
| SPECBIASFREQ | Div_spectro_bias | 9 | BIASFREQ |  | Spectrometry Bias frequency |
| SPECBIASMODE | Mode_spectro_bias | 8 | BIASMODE |  | Spectrometry bias generator mode |
| SPECLIA5V | DcuLSPp05 | TBD | TBD | TBD | Voltage to Spectrometer LIA Board on +5V line |
| SPECLIA9V | DcuLSPp09 | TBD | TBD | TBD | Voltage to Spectrometer LIA Board on +9V line |
| SPECLIA9VN | DcuLSPn09 | TBD | TBD | TBD | Voltage to Spectrometer LIA Board on -9V line |
| SPECSAMPFREQ | Div_spectro_sampl | 8 | SSAMPFREQ |  | Spectrometry sampling frequency |
| SSWBIAS | Ampl_S-SW | 8 | BIAS |  | Amplitude of bias for spectrometry short wave channels |
| SSWJFET1V | Ampl_S-SW_VSS1 | 8 | JFETV |  | Voltage applied to JFET for spectrometry short wave channels (group1) |
| SSWJFET2V | Ampl_S-SW_VSS2 | 8 | JFETV |  | Voltage applied to JFET for spectrometry short wave channels (group2) |
| SSWPHSE | Phase_shift_S-SW | 8 | PHASE |  | Phase shift for short wave spectrometry channels demodulation |
| STEP |  | 16 |  |  | Number of current step in an observation |
| SUBKTEMP | SubKTempP | TBD | TBD | TBD | Temperature from SubK temperature probe |
| TRACEPERIOD | TraceSampling |  | PERIOD |  | Time period between successive Trace Values |
| TRACESIZE | TraceBuffer | 16, TBC |  |  | Trace Buffer Size |
| TRACEPARM1 | TraceParam#1 | TBD |  |  | Trace Parameter 1 |
| TRACEPARM2 | TraceParam#2 | TBD |  |  | Trace Parameter 2 |
| TRACEPARM3 | TraceParam#3 | TBD |  |  | Trace Parameter 3 |
| TRACEPARM4 | TraceParam#4 | TBD |  |  | Trace Parameter 4 |
| TRACEPARM5 | TraceParam#5 | TBD |  |  | Trace Parameter 5 |
| TRACEPARM6 | TraceParam#6 | TBD |  |  | Trace Parameter 6 |
| TC1TEMP | T/C1 | 20 | TBD |  | Thermal Control Thermistor #1 value |
| TC2TEMP | T/C2 | 20 | TBD |  | Thermal Control Thermistor #2 value |
| TC3TEMP | T/C3 | 20 | TBD |  | Thermal Control Thermistor #3 value |
| TCBIAS | Ampl_T/C | 8 | BIAS |  | Amplitude of bias for thermal control channels |
| TCEXEC |  | 16 |  |  | Number of telecommand packets executed since switch on |
| TCEXEN |  | 16 |  |  | Sequence Control word of last executeded telecommand packet |
| TCPHSE | Phase_shift_T/C | 8 | PHASE |  | Phase shift for thermal control channels demodulation |
| TCRECN |  | 16 |  |  | Sequence Control word of last received telecommand packet |
| TCRECV |  | 16 |  |  | Number of telecommand packets received since switch on |
| TCUTEMP | TcuTempRd | TBD | TBD | TBD | Temperature from sensor on Temp Board |
| TDPU |  | 48 |  |  | Local DPU On Board Time |
| TEMP01 | FpuTemp01 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP02 | FpuTemp02 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP03 | FpuTemp03 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP04 | FpuTemp04 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP05 | FpuTemp05 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP06 | FpuTemp06 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP07 | FpuTemp07 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP08 | FpuTemp08 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP09 | FpuTemp09 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP10 | FpuTemp10 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP11 | FpuTemp11 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP12 | FpuTemp12 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP13 | FpuTemp13 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP14 | FpuTemp14 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP15 | FpuTemp15 | TBD | TBD | TBD | To be replaced by correct name |
| TEMP16 | FpuTemp16 | TBD | TBD | TBD | To be replaced by correct name |
| THSK |  | 48 |  |  | DPU Time of start of last housekeeping data collection period |
| TSFTIME |  | 48 |  |  | Telescope Scan Time |
| TSF |  | 1 | TSF |  | Telscope Scan Flag |
| TSYNC |  | 48 |  |  | DPU Time of last DRCU Synchronisation |
| SCUSPARE | Reserved | TBD |  |  |  |

### Conversion Curves

| Name | Type | Raw Value | Converted Value | Units | Comments |
| --- | --- | --- | --- | --- | --- |
| BIAS | Analogue | 0 255 | 0.0 Vbmax | Volts |  |
| BIASFREQ | Derived Calculated | Raw | Eng | Hz |  |
| BIASMODE | Enumerated | 0c00 0x01 0x02 …… …… 0xFD 0xFE 0xFF | OFF DC001 DC002 …….. …….. DC253 DC252 SINE |  | Values 0x01 to 0xFE are test levels |
| CHOPMODE | Enumerated | 0 1 2 | STOP STEP TOGGLE |  | Chopping Stopped Chopping in step mode Chopping in toggle mode |
| CHOPSENS | Enumerated | 0 1 | ON OFF |  | Values TBC |
| CLOOPMODE | Enumerated | 0 1 2 | OPEN BEMF SENS |  | Loop Open Loop closed using Back EMF Loop Closed using Magnetoresistive Sensor |
| DCUFRMS | Enumerated | 0 1 2 ….. ….. 254 255 | CONT. 1 2 ….. ….. 254 255 |  |  |
| DCUMODE | Enumerated | 0x00 0x01 0x02 0x03 0x04 0x05 0x06 0x08 0x0C 0x10 0x14 0x18 0x1C | PHOT PSW PMW PLW SPEC SSW SLW PTEST STEST POFFSET SOFFSET POFFGET SOFFGET |  |  |
| DCUSTAT | Enumerated | 0 1 | OFF RUN |  |  |
| HTR | Analogue | 0 255 | 0.0 Vhmax | Volts |  |
| JFETSTAT | Enumerated | 0 1 | OFF ON |  |  |
| JFETV | Analogue | 0 255 | 0.0 VSSmax | Volts |  |
| JIGGLEMODE | Enumerated | 0 1 2 | STOP STEP TOGGLE |  | Jiggling Stopped Jiggling in step mode Jiggling in toggle mode - TBD if applicable |
| JIGGLESENS | Enumerated | 0 1 | ON OFF |  | Values TBC |
| JLOOPMODE | Enumerated | 0 1 2 | OPEN BEMF SENS |  | Loop Open Loop closed using Back EMF Loop Closed using Magnetoresistive Sensor |
| LATCH | Enumerated | 0 1 | Engaged Disengaged |  | Values TBC |
| LOOPMODE | Enumerated | 0 1 2 3 | OPEN BEMF LVDT OENC |  | Loop Open Loop closed using Back EMF Loop Closed using LVDT Loop Closed using Optical Encoder |
| OTF | Enumerated | 0 1 | Off Target On Target |  |  |
| PERIOD |  |  |  | mS |  |
| PHASE | Analogue | 0 255 | 0 360 | Degrees |  |
| PSAMPFREQ | Derived Calculated | Raw | Eng | Hz |  |
| SCANMODE | Enumerated | 0 1 2 3 | STOP STEP SAWTOOTH TRIANGLE |  |  |
| SCANSTAT | Enumerated | 0 1 | STOP RUN |  |  |
| SSAMPFREQ | Derived Calculated | Raw | Eng | Hz |  |
| TSF | Enumerated | 2 3 | Scanning Pointing |  |  |

### Constraints

Appendix A

This section describes the use of  the On Board Software (OBS) commands and telemetry in sufficient detail that the interface can be understood.

# Tables

Much of the data used by the OBS for operations will be held in tables. These will be used to hold, for example:

- Housekeeping packet parameter lists

- Jiggle position tables

- Command Lists

The OBS will provide for the definition of up to 128 tables, within a fixed-size memory area and will automatically move data within this memory area in order to allow update of tables, including a change in size.

By default the following table numbers are assigned and populated in the OBS, though they may be changed by command:

| Table ID | Description |
| --- | --- |
| 00 | Critical Housekeeping Report Definition |
| 01 | Nominal Housekeeping Report Definition |
| 02 | Detector Housekeeping Report Definition |
| 03 | undefined |
| 21 | Cooler Recycle |
| 22 | Cooler Temperature Control |
| 23 | 300mK Temperature Control |
| 24 | SCAL Temperature Control |
| 31 | CHOP |
| 32 | SCAN |
| 100 | Functions |

All tables are accessed using a Table_ID and individual locations in the table are accessed by an index, specified as a number of words, offset from the beginning of the table. The word size is 32bits.

Appendix B

# Control Loops

Several control loops are executed by the On-board Software. These are all implemented as PID (Proportional, Integral, Differential) control algorithms, where the control variable (C) is adjusted to maintain the measurement variable (M) at the set point (S). The equation giving this relationship is

Where:

CN = Control variable value for sample period N

MN = Measurement Variable value at sample N

S = Required Set Point of the measurement variable

KP = Proportional Gain

KI = Integral Gain

KD = Differential Gain

n = number of samples in the integral

## Control Loop Parameters

When starting the control loop the following parameters have to be provided

- S = Required Set Point of the measurement variable

- KP = Proportional Gain

- KI = Integral Gain

- KD = Differential Gain

- n = number of samples in the integral (max 32, TBC)

- dt = the time interval between samples

- PM = The housekeeping parameter to use as the measurement variable

- PC =The housekeeping parameter to use as the control variable

It is expected that the software shall maintain a buffer containing the last n samples in order to calculate the integral term for each sample period.

If the accuracy of the measurement sampling interval cannot be guaranteed, the software will also need to maintain a buffer containing the time of each of the n samples in order to make a more accurate estimate of the integral and differential terms (the equation above assumes a fixed sampling period).

The software must ensure that the accuracy of the calculation is adequate (e.g. by use of floating point, or fixed point, math functions)

Appendix C

# Command List

This section (TBW) defines the command list entries that are available.

Appendix D

# Science Data

Science data is extracted in the form of science frames from the three FIFOs filled via the high speed telemetry interface. This appendix indicates how the OBS will handle this data.

## Assumptions

The OBS provides 3 buffers (one per FIFO) to hold the telemetry packet as it is accumulated

## Science Processing

This section describes the processing for the DCU telemetry (the other science telemetry will work in exactly the same way)

Assume that the FIFO is empty (following a Flush_FIFOs command)

When science data is generated by the DCU it is inserted into the FIFO. When the FIFO is half full an interrupt is passed to the OBS and it executes the following steps:

- Check that a complete frame is available (each frame starts with a word indicating the length  of the frame). If not, wait for more data

- Check the content of the frame is correct using the checkword  in the frame

- Check the Frame ID is valid - if not, issue an event packet  (Bad_Science_Data) containing the frame and flush the FIFO

- If the telemetry buffer is empty -

- Create a SID for the telemetry packet from the FRAMEID and put it in the telemetry buffer

- Add the current OBSID and BBID to the telemetry buffer

- Check the FRAMEID is the same as the FRAMEID in the SID of the telemetry buffer. If not, issue an event packet (Bad_Frame_ID) containing the frame

- Move the frame to the telemetry buffer, removing the LENGTH and FRAMEID words and selecting only those words defined in the selection table associated with this FRAMEID. (If no selection table is associated with FRAMEID, then all data is selected)

- If NDCU frames are now in the telemetry buffer, add the appropriate packet header (depends on FRAMEID) and send it to the instrument output telemetry stream

- Clear the telemetry buffer

- Repeat from Step 1

# Text boxes

| Error | Failure Code | Parameter |
| --- | --- | --- |
| Illegal APID | 0 | TC_Packet_APID |
| Incomplete Packet or invalid Length | 1 | TC_Packet_Length |
| Incorrect Checksum | 2 | TC_Packet_Checksum |
| Illegal Packet Type | 3 | TC_Packet_Type |
| Illegal Packet Sub-Type | 4 | TC_Packet_Sub-Type |

Note: The parameter is placed in the least significant bits of the16 bit 'parameter'  field and the most significant bits are padded with zeros

Parameters

| Name | Comments |
| --- | --- |
| HKPCKTID | HK_Packet_ID, in range 0 - 3 |
| HKSID |  |
| HKINTERVAL | msec |
| TABLEID | Table_ID |

Note: The parameter is placed in the least significant bits of the16 bit 'parameter'  field and the most significant bits are padded with zeros.

| Error | Failure Code | Parameters |
| --- | --- | --- |
| Illegal or inconsistent Application Data | 5 | See Note |
| Other TBD errors | 16-255 | See Note |

Note: The parameters field should contain the first 20 words from the 'source data' field of the received telecommand packet, unless this field is less than 20 words in length, in which case all words from the 'source data' field will be included and padded with zeros.

| SID | Packet type | Default Period (msec) |
| --- | --- | --- |
| 0x0300 | Critical Housekeeping Report | 2000 |
| 0x0301 | Nominal Housekeeping Report | 1000 |
| 0x0302 | Detector Diagnostic Report | Undefined |

The reports available are identified by the APID (APID3 = Photometer, APID4 = Spectrometer) and the SID (composed of the Frame Structure (MSByte) and FrameID (LSByte)):

| APID | SID | Report |
| --- | --- | --- |
| APID3 | 0x0208 | Photometer Offsets |
| APID4 | 0x0209 | Spectrometer Offsets |

Parameter

| Name | Comment |
| --- | --- |
| FUNCTIONID | Function to report on |

The reports available are identified by the APID (APID3 = Photometer, APID4 = Spectrometer) and the SID (composed of the Frame Structure (MSByte) and FrameID (LSByte)):

| APID | SID | Report |
| --- | --- | --- |
| APID3 | 0x0307 | Photometer Test Pattern |
| APID3 | 0x020A | Spectrometer Test Pattern |
| APID5 | 0x0814 | MCU Trace |
| APID5 | 0x0915 | MCU Test Pattern |
| APID5 | 0x1121 | SCU Test Pattern |

The reports available are identified by the APID (APID3 = Photometer, APID4 = Spectrometer) and the SID (composed of the Frame Structure (MSByte) and FrameID (LSByte)):

| APID | SID | Report |
| --- | --- | --- |
| APID3 | 0x0102 | PSW Array |
| APID3 | 0x0103 | PMW Array |
| APID3 | 0x0104 | PLW Array |
| APID4 | 0x0105 | SSW Array |
| APID4 | 0x0106 | SLW Array |
| APID5 | 0x0510 | SMEC Step (TBC) |

Parameter

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x01 |
| ACTIVITYID | 0x02 |
| TABLEID |  |
| INDEX | Starting offset to first word to report |
| COUNT | Number of (32 bit) words to report |

The reports available are identified by the APID (APID3 = Photometer, APID4 = Spectrometer) and the SID (composed of the Frame Structure (MSByte) and FrameID (LSByte)):

| APID | SID | Report |
| --- | --- | --- |
| APID3 | 0x0200 | Photometer Full Array |
| APID4 | 0x0201 | Spectrometer Full Array |
| APID5 | 0x0511 | SMEC |
| APID5 | 0x0612 | BSM |
| APID5 | 0x1020 | SCU |
| APID5 | 0x8080 | DPU |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x03 |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x04 |

| Parameter | Comment |
| --- | --- |
| SID | 0x0511 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| FRAMEID | The Frame ID received from the DRCU |
| FIFOID | The ID of the FIFO where the frame has been received |

| Parameter | Comment |
| --- | --- |
| SID | 0x0530 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| ANOMALYID | Anomaly ID |

| Parameter | Comment |
| --- | --- |
| SID | 0x0521 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| ANOMALYID | Anomaly ID |

| Parameter | Comment |
| --- | --- |
| SID | 0x0520 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| ANOMALYID | Anomaly ID |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x01 |
| ACTIVITYID | 0x01 |
| TABLEID |  |
| TABLESIZE | Number of 32 bit words allocated to table |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x01 |
| ACTIVITYID | 0x03 |
| TABLEID |  |
| INDEX | Start location at which to place Data |
| N | Number of 32bit words to place in table |
| Data | N (32 bit) words to be placed in the table |

Contents of the Data is dependant on the function of the table and will be specified by IFSI

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x02 |
| ACTIVITYID | 0x01 |
| N | Number of 32bit words in the command List |
| Data | N * 32 bit data words containing the command list to execute |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x02 |
| ACTIVITYID | 0x02 |
| TABLEID | Table_ID of list to execute |
| INDEX | Index within Table at which to start |
| N | Number of data words |
| Data | N * 32 bit data words. These are the parameters passed to the VM in registers R(00) to R(N-1) |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0x02 |
| ACTIVITYID | 0x03 |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0xC1 |
| ACTIVITYID | 0x01 |
| OBSID | Observation ID (32 bits) |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0xC1 |
| ACTIVITYID | 0x02 |
| BBID | Building Block ID (32 bits) |

Parameters

| Name | Comments |
| --- | --- |
| MEMORYID | Memory Area to be addressed |
| STARTADDR | Start Address in SAUs |
| NSAU | Number of SAUs to be loaded |
| DATA | Data to be loaded |
| CRC | CRC Checksum of data field |

The DATA field will contain the words to be written into DPU memory. 48 or 32bit words will be split into 3 or 2 16bit words and written into the DATA field of the telecommand starting from the MSB.

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x01 |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x02 |
| FIFOFLAGS |  |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x10 |
| FRAMEID | Science frame to extract data from |

Parameters

| Name | Value and Comments |
| --- | --- |
| MEMORYID | Memory Area to be addressed |
| STARTADDR | Start Address in SAUs |
| NSAU | Number of SAUs to be dumped |

Parameters

| Name | Value and Comments |
| --- | --- |
| MEMORYID | Memory Area to be addressed |
| STARTADDR | Start Address in SAUs |
| NSAU | Number of SAUs to be checked |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0x03 |
| ACTIVITYID | 0x03 |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x03 |
| ACTIVITYID | 0x03 |
| TABLEID | Table_ID of list to execute |
| INDEX | Index within Table at which to start |
| N | Number of data words |
| Data | N * 32 bit data words. These are the parameters passed to the VM in registers R(00) to R(N-1) |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0xC1 |
| ACTIVITYID | 0x03 |
| MODE | Observing Mode |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0xC1 |
| ACTIVITYID | 0x04 |
| STEP | Observation Step |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xC0 |
| ACTIVITYID | 0x40 |

| Parameter | Comment |
| --- | --- |
| SID | 0x0509 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| COMMAND | copy of command sent to DRCU (32 bits) |
| ACK | Echo of command word received by the DPU |

| Parameter | Comment |
| --- | --- |
| OTF | Copy of INFOTYPE from command 0 = Off Target 1 = On Target |

| Parameter | Comment |
| --- | --- |
| SID | 0x0504 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| Parameters | TBD, but will contain information for the AOCS to offset to bring the peak on to the required pixel |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x10 |
| FRAMEID | Science frame to extract data from |
| TABLEID | Table containing the selection (0 = no selection) |

Parameters

| Name | Comments |
| --- | --- |
| HKPCKTID | HK_Packet_ID, in range 0 - 3 |

Note: The parameter is placed in the least significant bits of the16 bit 'parameter'  field and the most significant bits are padded with zeros.

Parameters

| Name | Comments |
| --- | --- |
| HKPCKTID | HK_Packet_ID, in range 0 - 3 |

Note: The parameter is placed in the least significant bits of the16 bit 'parameter'  field and the most significant bits are padded with zeros.

Parameters

| Name | Value and Comments |
| --- | --- |
| NPCKTS | Number of packets to enable |
| TYPE | Packet Type to enable |
| SUBTYPE | Packet Subtype to enable |
| SID | Packet SID to enable |

Parameters

| Name | Value and Comments |
| --- | --- |
| NPCKTS | Number of packets to enable |
| TYPE | Packet Type to enable |
| SUBTYPE | Packet Subtype to enable |
| SID | Packet SID to enable |

Parameters

| Name | Value and Comments |
| --- | --- |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x04 |
| ACTIVITYID | 0x02 |
| TABLEID | Table_ID of list to execute |
| INDEX | Index within Table at which to start |
| N | Number of data words |
| Data | N * 32 bit data words. These are the parameters passed to the VM in registers R(00) to R(N-1) |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0x04 |
| ACTIVITYID | 0x03 |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | 0x05 |
| ACTIVITYID | 0x02 |
| TABLEID | Table_ID of list to execute |
| INDEX | Index within Table at which to start |
| N | Number of data words |
| Data | N * 32 bit data words. These are the parameters passed to the VM in registers R(00) to R(N-1) |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0x05 |
| ACTIVITYID | 0x03 |

Parameters

| Name | Value and Comments |
| --- | --- |
| MEMORYID | Memory Area to be addressed |
| STARTADDR | Start Address in SAUs |
| NSAU | Number of SAUs to be dumped |
| Data | Contents of memory in increasing address |
| CHK | Checksum over the data included in this packet |

Parameters

| Name | Value and Comments |
| --- | --- |
| MEMORYID | Memory Area to be addressed |
| STARTADDR | Start Address in SAUs |
| NSAU | Number of SAUs to be checked |
| CHK | Checksum over the data included in the address range (STARTADDR to STARTADDR+NSAU) |

Parameters

| Name | Value and Comments |
| --- | --- |
| FUNCTIONID | ID of function to be reported |
| ACTIVITYID | Current Activity |
| SID | 0x0806 |
| Parameters | Parameters send in corresponding TC packet starting the activity |

Parameters

| Name | Value and Comments |
| --- | --- |
| NDEF | Number of sets of (Type, Subtype, SID) to follow |
| Type | Packet Type |
| Subtype | Packet Subtype |
| SID | Parameters send in corresponding TC packet starting the activity |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x05 |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x06 |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x07 |
| Start_Address | Starting address in PM |
| End_Address | End address in PM |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x08 |

Parameters

| Name | Value |
| --- | --- |
| FUNCTIONID | 0xCA |
| ACTIVITYID | 0x09 |

| Parameter | Comment |
| --- | --- |
| SID | 0x0505 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| POOLID | ID of the VIRTUOSO Pool which failed to allocate block. IDs are defined in RD08 |
| SIZE | Size in Bytes of the requested memory block |
| ALLOC_RESULT | Allocation Result: Allocation Failed Allocation Timed out |
| TASKID | ID of the OBS task where failure occurred. IDs are defined in RD08 |

| Parameter | Comment |
| --- | --- |
| SID | 0x0507 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| FRAMEID | The Frame ID received form the DRCU |
| LENGTHRD | Frame length as read from the DRCU frame |
| LENGTHEX | Expected frame length for that frame ID |

| Parameter | Comment |
| --- | --- |
| SID | 0x0508 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| FRAMEID | The Frame ID received form the DRCU |
| CHKWORDEX | XOR checksum computed on the read frame |
| CHKWORDRD | Checksum word read at the end of teh received frame |

| Parameter | Comment |
| --- | --- |
| SID | 0x050A |
| OBSID | Observation ID |
| BBID | Building Block ID |
| COMMAND | copy of command sent to DRCU (32 bits) |
| ACK | Echo of command word received by the DPU |

| Parameter | Comment |
| --- | --- |
| SID | 0x050B |
| OBSID | Observation ID |
| BBID | Building Block ID |
| CID | 16 MSb of the command sent to the DRCU |
| ACK | MS 16 bits of response from DRCU containing the Ack bits received from the DRCU |

| Parameter | Comment |
| --- | --- |
| SID | 0x050C |
| OBSID | Observation ID |
| BBID | Building Block ID |
| CID | 16 MSb of the command sent to the DRCU |
| RESPONSE | The CID returned by the DRCU |

| Parameter | Comment |
| --- | --- |
| SID | 0x050D |
| OBSID | Observation ID |
| BBID | Building Block ID |

| Parameter | Comment |
| --- | --- |
| SID | 0x050E |
| OBSID | Observation ID |
| BBID | Building Block ID |

| Parameter | Comment |
| --- | --- |
| SID | 0x0501 |
| OBSID | Observation ID |
| BBID | Building Block ID |
| MODE | Current Mode |
| STEP | Step Number |

| Parameter | Comment |
| --- | --- |
| SID | 0x0531 |
| OBSID | Observation ID |
| BBID | Building Block ID |

| Parameter | Comment |
| --- | --- |
| EVENTID | TBD |
| SID | 0x0540 |
| MEMID | Memory ID |
| NPAGES | Number of failed memory pages |
| Page IDs | ID of each failed page (8 bits) Up to 40 pages may be reported in one event packet |

Note: the TIME field will have the MSb set to 1 to indicate that time is not yet synchronised