Spectrometer Calibrator

Failure Modes Effects & Criticality Analysis (FMECA) Report

SPIRE Ref.: SPIRE-QMW-PRJ-001102

Cardiff Ref.: HSO-CDF-RP-030

Issue: 1.1

Prepared by: Peter Hargrave

Last Modified on: 15 November 2001

Approved by:

Distribution list

Update history

| Date | Version | Remarks |
| --- | --- | --- |
| 31/08/01 | 1.0 | First Issue for DDR |
| 15/11/01 | 1.1 | Minor corrections |

List of Acronyms

| Term | Meaning | Term | Meaning |
| --- | --- | --- | --- |
| AD | Applicable Document | IR | Infrared |
| ADC | Analogue to Digital Converter | IRD | Instrument Requirements Document |
| AIV | Assembly, Integration and Verification | IRTS | Infrared Telescope in Space |
| AME | Absolute Measurement Error | ISM | Interstellar Medium |
| AOCS | Attitude and Orbit Control System | JFET | Junction Field Effect Transistor |
| APART | Arizona's Program for the Analysis of Radiation Transfer | ISO | Infrared Space Observatory |
| APE | Absolute Pointing Error | LCL | Latching Current Limiter |
| ASAP | Advanced Systems Analysis Program | LIA | Lock-In Amplifier |
| ATC | Astronomy Technology Centre, Edinburgh | LVDT | Linear Variable Differential Transformer |
| AVM | Avionics Model | LWS | Long Wave Spectrometer (an instrument used on ISO) |
| BDA | Bolometer Detector Array | MAC | Multi Axis Controller |
| BFL | Back Focal Length | MAIV | Manufacturing, Assembly, Integration and Verification |
| BRO | Breault Research Organization | MCU | Mechanism Control Unit = HSMCU |
| BSM | Beam Steering Mirror | MGSE | Mechanical Ground Support Equipment |
| CBB | Cryogenic Black Body | M-P | Martin-Puplett |
| CDF | Cardiff, Department of Physics & Astronomy | NEP | Noise Equivalent Power |
| CDMS | Command and Data Management System | NTD | Neutron Transmutation Doped |
| CDMU | Command and Data Management Unit | OBS | On-Board Software |
| CDR | Critical Design Review | OGSE | Optical Ground Support Equipment |
| CEA | Commissariat a l’Energie Atomique | OMD | Observing Modes Document |
| CMOS | Complimentary Metal Oxide Silicon | OPD | Optical Path Difference |
| CoG | Centre of Gravity | PACS | Photodetector Array Camera and Spectrometer |
| CPU | Central Processing Unit | PCAL | Photometer Calibration source |
| CQM | Cryogenic Qualification Model | PFM | Proto-Flight Model |
| CVV | Cryostat Vacuum Vessel | PID | Proportional, Integral and Differential (used in the context of feedback control loop architecture) |
| DAC | Digital to Analogue Converter | PLW | Photometer, Long Wavelength |
| DAQ | Data Acquisition | PMW | Photometer, Medium Wavelength |
| DCU | Detector Control Unit = HSDCU | POF | Photometer Observatory Function |
| DDR | Detailed Design Review | PROM | Programmable Read Only Memory |
| DM | Development Model | PSW | Photometer, Short Wavelength |
| DPU | Digital Processing Unit = HSDPU | PUS | Packet Utilisation Standard |
| DSP | Digital Signal Processor | RAL | Rutherford Appleton Laboratory, |
| DQE | Detective Quantum Efficiency | RD | Reference Document |
| EDAC | Error Detection and Correction | RMS | Root Mean Squared |
| EGSE | Electrical Ground Support Equipment | SCAL | Spectrometer Calibration Source |
| EM | Engineering Model | SCUBA | Submillimetre Common User Bolometer Array |
| EMC | Electro-magnetic Compatibility | SED | Spectral Energy Distribution |
| EMI | Electro-magnetic Interference | SMEC | Spectrometer Mechanics |
| ESA | European Space Agency | SMPS | Switch Mode Power Supply |
| FCU | FCU Control Unit = HSFCU | SOB | SPIRE Optical Bench |
| FIR | Far Infrared | SOF | Spectrometer Observatory Function |
| FIRST | Far Infra-Red and Submillimetre Telescope | SPIRE | Spectral and Photometric Imaging Receiver |
| FOV | Field of View | SRAM | Static Random Access Memory |
| F-P | Fabry-Perot | SSSD | SubSystem Specification Document |
| FPGA | Field Programmable Gate Array | STP | Standard Temperature and Pressure |
| FPU | Focal Plane Unit | SVM | Service Module |
| FS | Flight Spare | TBC | To Be Confirmed |
| FTS | Fourier Transform Spectrometer | TBD | To Be Determined |
| FWHM | Full Width Half maximum | TC | Telecommand |
| GSFC | Goddard Space Flight Center | URD | User Requirements Document |
| HK | House Keeping | UV | Ultra Violet |
| HOB | Herschel Optical Bench | WE | Warm Electronics |
| HPDU | Herschel Power Distribution Unit | ZPD | Zero Path Difference |
| HSDCU | Herschel-SPIRE Detector Control Unit |  |  |
| HSDPU | Herschel-SPIRE Digital Processing Unit |  |  |
| HSFCU | Herschel-SPIRE FPU Control Unit |  |  |
| HSO | Herschel Space Observatory |  |  |
| IF | Interface |  |  |
| IID-A | Instrument Interface Document - Part A |  |  |
| IID-B | Instrument Interface Document - Part B |  |  |
| IMF | Initial Mass Function |  |  |

Table of Contents

1.	Scope	5

2.	Documents	5

2.1.	Applicable documents	5

2.2.	Reference documents	5

3.	Details of the analysis	5

# Scope

This document presents the results of the FMECA carried out on the SCAL design.

# Documents

## Applicable documents

All applicable documents are listed in the AD chapter of the CIDL (HSO-CDF-LI-029).

## Reference documents

# Details of the analysis

A failure modes effects and criticality analysis has been performed on all functional elements of SCAL (excluding structural elements whose integrity has been assessed with stress analysis and fracture mechanics analysis as necessary) which can cause failure effects within the experiment or cause damage to or interfere with, the proper functioning of the SPIRE instrument or Herschel spacecraft.

Each failure effect identified has been given a criticality category according to the definition below:

- Category 1:	The failure effect is not confined to the subsystem.  When this failure results also in loss or degradation of the instruments function this shall be stated.

- Category 2:	The failure results in loss or degradation of the subsystems function but the effect is confined to the subsystem.

- Category 3: 	Minor internal subsystem failures.

The following attributes have been added to the criticality category as appropriate:

-	"R", if the design contains a redundant item which can perform the same function

-	"SH", if the failure effect causes a safety hazard

-	"SPF" if the failure is caused by a single point failure.

The following failure modes have been considered: -

Premature operation

Failure to operate (at the prescribed time)

Failure to cease operation (at the prescribed time)

Failure during operation

Degradation or out of tolerance operation

For failure at component level e.g. hardware interface

-	short circuit

-	open circuit

-	incorrect function e.g. from single event upset  - ex: latch-ups.

Incorrect commands or sequence of commands

Incorrect software functions

Mechanical failure

Design specifications, descriptions functional diagrams etc. used in the preparation of the FMECA shall be attached or referenced.

Table 1 Results of FMECA of SCAL subsystem.

| FAILURE MODES EFFECTS AND CRITICALITY ANALYSIS (FMECA) Product: SPIRE Instrument Project/Phase: Herschel System/Subsystem/Equipment: SCAL Mission phase/Operational Mode: Space Flight Prepared by: P.Hargrave Approved by: Date: 25/08/01 Document reference: Issue: |
| --- |
| Id number | Item/ block | Function | Failure mode | Failure cause | Failure effects a. Local effects b. End effects | Severity | Failure detection method/ observable symptoms | Compensation provisions | Correction actions | Remarks |
| 000.001 | SCAL Assembly |  | Input power short to ground (source A or B) | Connector Failure | a. Loss of SCAL_A or SCAL_B sources b. Degraded dynamic range for spectrometer detectors | 2R | No source heating for one of the prime sources | Switch to redundant side | Can run second prime source – impaired telescope nulling | There are two prime sources with full redundancy. Running the other prime source will still allow some degree of telescope nulling with perhaps increased power dissipation |
| 000.002 | SCAL assembly |  | Input power open circuit (source A or B) | Connector Failure | a. Loss of SCAL_A or SCAL_B sources b. Reduced dynamic range for spectrometer detectors | 2R | No source heating for one of the prime sources | Switch to redundant side | Can run second prime source – impaired telescope nulling | There are two prime sources with full redundancy. Running the other prime source will still allow some degree of telescope nulling with perhaps increased power dissipation |
| 000.003 | SCAL assembly |  | Thermometry short (source A or B) | Connector Failure | a. Inability to monitor SCAL_A or SCAL_B temperatures b. Loss of thermometry data for source | 3R | Low impedance reading on source thermometer | Switch to redundant side | Can still run source “blind” based on experimental data until we maximise the central peak in the interferogram |  |
| 000.004 | SCAL assembly |  | Thermometry open circuit (source A or B) | Connector Failure | a. Inability to monitor SCAL_A or SCAL_B temperatures b. Loss of thermometry data for source | 3R | Open circuit reading for thermometer | Switch to redundant side | Can still run source “blind” based on experimental data until we maximise the central peak in the interferogram |  |
| 100.001 | SCAL_A(P) – prime 4% source assembly | Radiant source for telescope nulling | Heater body breaks off strut | Manufacturing error. Part failure. | a. Loss of SCAL_A(P) source b. Degraded dynamic range for spectrometer detectors. Possible damage to other subsystem and instrument components (e.g. SM8B, beam splitter) | 1R | Open circuit reading for thermometer and heater | a. Switch to redundant side b. None | Can run second prime source – impaired telescope nulling |  |
| 101.001 | SCAL_A(P) – heater body |  | Loss of black coating | Manufacturing error. | a. Reduced emissivity of source – increased power dissipation to achieve required level of nulling. Impaired nulling. b. Possible damage to other subsystem and instrument components (e.g. SM8B, beam splitter) from coating fragments | 1R |  | a. Switch to redundant side b. None | Can run second prime source – impaired telescope nulling | May not be possible to match spectrum within requirements |
| 102.001 | SCAL_A(P) – strut | Supports heater body – provides thermal isolation | Breaks | Material failure | a. Loss of source b. Reduced dynamic range for spectrometer detectors. Possible damage to other subsystem and instrument components (e.g. SM8B, beam splitter) | 1R |  | a. Switch to redundant side b. None | Can run second prime source – impaired telescope nulling |  |
| 103.001 | SCAL_A(P) – sapphire plate | Provides electrical isolation for heater resitor & good thermal path to heater body | Cracks | Manufacturing error. Differential thermal contraction | a. Impaired thermal path to heater body – increased warm-up time b. Longer warm-up | 3R |  | Switch to redundant side |  | May be detectable by performing cross-calibration with redundant side sources. |
| 104.001 | SCAL_A(P) – thermometer | Monitors source temperature | Short | Manufacturing error. Part failure. | a. Inability to monitor SCAL_A temperature b. Loss of thermometry data for source | 3R | Low impedance reading on source thermometer | Switch to redundant side | Can still run source “blind” based on experimental data until we maximise the central peak in the interferogram |  |
| 104.002 |  |  | Open | Manufacturing error. Part failure. | a. Inability to monitor SCAL_A temperature b. Loss of thermometry data for source | 3R | Open circuit reading for thermometer | Switch to redundant side | Can still run source “blind” based on experimental data until we maximise the central peak in the interferogram |  |
| 104.003 |  |  | Erratic output | Manufacturing error. Part failure. | a. Inability to monitor SCAL_A temperature b. Loss of thermometry data for source | 3R |  | Switch to redundant side | Can still run source “blind” based on experimental data until we maximise the central peak in the interferogram |  |
| 104.004 |  |  | No output | Manufacturing error. Part failure. | a. Inability to monitor SCAL_A temperature b. Loss of thermometry data for source | 3R |  | Switch to redundant side | Can still run source “blind” based on experimental data until we maximise the central peak in the interferogram |  |
| 105.001 | SCAL_A(P) – Heater resistor | Heater for source | Open | Manufacturing error. Part failure. | a. Loss of source b. Reduced dynamic range for spectrometer detectors | 2R | No source heating | Switch to redundant side | Can run second prime source – impaired telescope nulling | There are two prime sources with full redundancy. Running the other prime source will still allow some degree of telescope nulling with perhaps increased power dissipation |
| 105.002 |  |  | Short | Manufacturing error. Part failure. | a. Loss of source b. Reduced dynamic range for spectrometer detectors | 2R | No source heating | Switch to redundant side | Can run second prime source – impaired telescope nulling | There are two prime sources with full redundancy. Running the other prime source will still allow some degree of telescope nulling with perhaps increased power dissipation |
| 105.003 |  |  | Changed value | Manufacturing error. Part failure. | a. Change in drive requirements b. None | 3R | Monitor I/V values for heater (4 wire configuration) | Change drive current Switch to redundant side? | Change drive current |  |