|  | HERSCHEL SPIRE | Ref: SPIRE-RAL-DOC-001652 Author: S. Ronayette | Page: 1 Issue: 0.1 Date: 08 April 2003 |
| --- | --- | --- | --- |
| SPIRE Functional Test Specifications |

| Prepared by: |  |  |
| --- | --- | --- |
|  | S. Ronayette (U of L) | Date |
| Checked: |  |  |
|  | D.L. Smith (RAL) | Date |
| Approved: |  |  |
|  | B.M. Swinyard (RAL) | Date |

Distribution

| RAL | Dave Smith |
| --- | --- |
|  | Ken King |
|  | Eric Sawyer |
|  | Bruce Swinyard |
|  | Doug Griffin |
|  | Tanya Lim |
|  | Sunil Sidher |

| Host system | Windows 2000 SP2 |
| --- | --- |
| Word Processor | Microsoft Word 2000 SR1 |
| File | Cryostat Test Plan 0.3.Doc |

# Document Change Record

| Date | Index | Affected Pages | Changes |
| --- | --- | --- | --- |
|  | 0.1 | All | First Draft |

Contents

1	Scope of Document	6

2	Documents	6

2.1	Applicable Documents	6

2.2	Reference Documents	6

3	Acronyms	7

4	Test Summaries	8

5	Test Specifications	10

5.0	SCU-00, SCU Science packet generation check	10

5.1	SCU-01, SCU Science Data Check	11

5.2	SCU-02, SCU DC Thermometry Check	12

5.3	SCU-03, SCU PCAL	13

5.4	SCU-04, SCU SCAL Check	14

5.5	SCU-05, SCU AC thermometry check	15

5.6	SCU-06, SCU Cooler Heaters check	16

5.7	SCU-07, SCU Test pattern test	17

5.8	MCU-01, MCU Power On	18

5.9	MCU-02_0, MCU Science Packet generation check	19

5.10	MCU-02, MCU Science Data check	20

5.11	MCU-03, MCU Test Pattern test	21

5.12	MCU-04, MCU launch latch check	22

5.13	SMEC-01, SMEC Switch On	23

5.14	SMEC-02, SMEC Initialisation	24

5.15	SMEC-03, SMEC Position test	25

5.16	SMEC-04, SMEC Multiple position test	26

5.17	SMEC-05, SMEC Saw tooth scan test	27

5.18	SMEC-06, SMEC triangular test	28

5.19	BSM-01, BSM Power on motors and sensors	29

5.20	BSM-02, BSM Move to position	30

5.21	BSM-03, BSM Scan test	31

5.22	BSM-04, BSM Operating mode test	32

5.23	PCAL-01, PCAL Characterisation test	33

5.24	SCAL-01, SCAL Characterisation test	34

5.25	SCAL-02 SCAL PID test	35

# Scope of Document

This document describes in detail how the SPIRE AIV facility will be integrated and tested in preparation for the STM and CQM testing, in particular the test cryostat and harness. It describes the step-by-step sequence of activities, high-level procedures, organisation, resources and facilities required. It does not describe the general facility build and telescope simulator integration and testing which will have been completed as separate tasks.

# Documents

## Applicable Documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD 1 | SPIRE Data ICD | K.J. King | SPIRE-RAL-PRJ-001078, issue 1.0 | 15/01/2003 |
| AD2 | MCU/DCU Command list ICD an user manual | Didier Ferrand | LAM/ELE/SPI/011011, issue 3.0 | 15/01/2003 |
| AD3 | DRCU/DPU ICD | Pinsard, Ferrand, Mur | Sap-SPIRE-Cca-076-02, issue 1.0 | 14/02/2003 |

## Reference Documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| RD 1 | EGSE specification |  |  |  |
| RD 2 |  |  |  |  |
| RD 3 |  |  |  |  |

# Acronyms

| AD | Applicable Document |  |  |
| --- | --- | --- | --- |
| AIV | Assembly Integration and Verification |  |  |
| AVM | Avionics Model |  |  |
| BSM | Beam Steering Mechanism |  |  |
| CQM | Cryogenic Qualification Model |  |  |
| DCU | Detector Control Unit |  |  |
| DPU | Digital Processing Unit |  |  |
| DRCU | Detector Readout and Control Unit |  |  |
| EGSE | Electrical Ground Support Equipment |  |  |
| FTS | Fourier Transform Spectrometer |  |  |
| HOB | Herschel Optical Bench |  |  |
| ILT | Instrument Level Test |  |  |
| MCU | Mechanism Control Unit |  |  |
| MGSE | Mechanical Ground Support Equipment |  |  |
| N/A | Not Applicable |  |  |
| PCAL | Photometer Calibrator source |  |  |
| PFM | Proto-Flight Model |  |  |
| RAL | Rutherford Appleton Laboratory |  |  |
| RD | Reference Document |  |  |
| SCAL | Spectrometer Calibrator source |  |  |
| SCU | Subsystem Control Unit |  |  |
| SMEC | Spectrometer Mechanism |  |  |
| SPIRE | Spectral and Photometric Imaging Receiver |  |  |
| STM | Structural Test Model |  |  |
| TBC | To Be Confirmed |  |  |
| TBD | To Be Defined |  |  |
| TFCS | Test Facility Control System |  |  |

# Test Summaries

(Characterisation test are in italic)

## SCU-00

Check the integrity of the SCU science data. Also check that the number of frames generated matches the number commanded.

## SCU-01

Check the integrity of the SCU science data. Check that content of science frames matches that in HK telemetry.

## SCU-02

Switch on thermometers and check values – could have different short and full functional tests.

## SCU-03

Switch on PCAL, set a current and check voltage and current, then switch off.

## SCU-04

Switch on SCAL, set a current and check voltage, current and temperature changes, then switch off.

## SCU-05

Switch on AC thermometers and check values.

## SCU-06

Switch on heaters, check voltage and temperature.

## SCU-07

Load test pattern and check content of packets. Perform test for each type of packet (SMEC, BSM, ENG).

## MCU-01

Power on MCU from SCU, check all voltage, check threshold flags, check if DSP booted. Check HK.

## MCU-02_0

Check that the number of frames generated matches the number commanded. Perform test for each type of packet. (SMEC, BSM, ENG).

## MCU-02

Check of integrity of MCU science data. Perform test for each type of packet. (SMEC, BSM, ENG).

## MCU-03

Load test pattern and check contents of packets. Perform test for each type of packet (SMEC, BSM, ENG).

## MCU-04

Open launch latch, move SMEC short distance and check response.

## SMEC-01

Switch on LVDT, fringe encoder and motor. Check encoder signal, status bits, and current.

## SMEC-02

Initialise SMEC. Check LVDT, fringe encoder and current.

## SMEC-03

Move to set positions and measure read back position, currents, back emf.

## SMEC-04

Move to number of set positions and measure positions, current, back emf – repeat for different scan speed.

## SMEC-05

Scan over range (saw tooth) and measure positions, current, back-emf – repeat for different scan speed.

## SMEC-06

Scan over range (triangular) and measure positions, current, back-emf – repeat for different scan speed

## BSM-01

Check telemetry responses, consumed power.

## BSM-02

Move BSM to set positions. Check telemetry responses, consumed power.

## BSM-03

Move BSM to multiple set positions. Check telemetry response, consumed power.

## BSM-04

Set BSM to each mode (Chop, jiggle…). Check telemetry response, consumed power.

## PCAL-01

Step through each setting and measure telemetry response.

## SCAL-01

Set current and record temperatures.

## SCAL-02

Set temperatures and record temperatures.

# Test Specifications

## SCU-00 (SCU Science packet generation check)

| ID: | SCU-00 |
| --- | --- |
| Test name: | SCU Science Packet Generation Check |
| Description of test: | Check the integrity of the SCU science data. Also check that the number of frames generated matches the number commanded. |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | Prime and Redundant Units |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required |
| Test Conditions: | Warm or Cold |
| Constraints: |  |
| Outline procedure and analysis: | In SCOS, select the appropriate display (DPU and OBS PARAMETER) Send request for N SCU science packets Wait for packets to be generated Check that N SCU science packets were generated, i.e. that parameter TM5N (sequence count of TM packets for SCU Science) increases until N |
| Success/Failure Criteria: | Test passed if number of packets generated = number requested If not? |

## SCU-01 (SCU Science Data Check)

| ID: | SCU-01 |
| --- | --- |
| Test name: | SCU Science Data Check |
| Description of test: | Check the integrity of the SCU science data. Check that content of science frames matches that in HK telemetry |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | Prime and Redundant Units |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required QLA may be required |
| Test Conditions: | Warm or Cold |
| Constraints: | SCU thermometry needs to be switched on – see test SCU-02 |
| Outline procedure and analysis: | Send a request for N SCU science packets. Wait for packets to be generated Use either QLA or “Packet display” with SCOS to see the content of the frames Compare it with the HK parameters Example of parameters to compare: TempOnOff ScuCHTp05 ScuCHTp09 PCALV SCAL2V SCAL2Temp … |
| Success/Failure Criteria: | Test passed if values in the science frame are identical to those in the HK nominal report. If not? |

## SCU-02 (SCU DC Thermometry Check)

| ID: | SCU-02 |
| --- | --- |
| Test name: | SCU DC Thermometry Check |
| Description of test: | Switch on thermometers in-turn and check values |
| Test Type: | Integrity Check |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and Redundant Units |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required |
| Test Conditions: | Warm or Cold |
| Constraints: |  |
| Outline procedure and analysis: | Select the appropriate display (SCU PARAMETER) Send command to switch on SCU thermometer channel TBD Wait 5 seconds (allow time for sensor to activate) Check values of the parameters if Warm, T>290K if Cold, T<4.2K Repeat steps 1 to 4 for other thermometer channels Parameters to check: CPHPtemp, CPHStemp, CEHStemp, CSHTtemp, SOBtemp, SL0temp, PL0temp, SUBtemp, BAFtemp, BSMStemp, SCL2temp, SCL4temp, SCSTtemp, FTSStemp, FTSMtemp, BSMMtemp |
| Success/Failure Criteria: | Test passed if all thermometer channels are reading expected values. If not – raise a NCR and proceed to next test. |

## SCU-03 (SCU PCAL)

| ID: | SCU-03 |
| --- | --- |
| Test name: | SCU PCAL |
| Description of test: | Switch on PCAL, set a current and check voltage and current, then switch off |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Switch on PCAL Select the appropriate display (SCU PARAMETER) Send a command to set a current value Check values of the parameters (PCALV and PCALCURR). Switch off PCAL. |
| Success/Failure Criteria: | Test passed if value of current read = current set and voltage has value expected |

## SCU-04 (SCU SCAL Check)

| ID: | SCU-04 |
| --- | --- |
| Test name: | SCAL check |
| Description of test: | Switch on SCAL, set a current and check voltage, current and temperature changes, then switch off |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Switch on SCAL Send a command to set a current value Use the appropriate display (SCU PARAMETER) and check values of the parameters for voltage, current and temperature (SCAL2V, SCAL4V, SCAL2CURR, SCAL4CURR, SCAL2TEMP, SCAL4TEMP). Switch off SCAL |
| Success/Failure Criteria: | Test passed if value of current read = current set and voltage and temperature have values expected |

## SCU-05 (SCU AC thermometry check)

| ID: | SCU-05 |
| --- | --- |
| Test name: | AC thermometry check |
| Description of test: | Switch on AC thermometers and check values |
| Test Type: | integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Select the appropriate display (SCU PARAMETER) Send command to switch on SCU thermometer channel TBD Wait 5 seconds (allow time for sensor to activate) Check values of the parameters if Warm, T > 290K if Cold, T < 1K Parameter to check: CEVTemp |
| Success/Failure Criteria: | Test passed if thermometer channel is reading expected value If not – raise a NCR and proceed to next test. |

## SCU-06 (SCU Cooler Heaters check)

| ID: | SCU-06 |
| --- | --- |
| Test name: | SCU Cooler Heater Check |
| Description of test: | Switch on heaters, check voltage and temperature |
| Test Type: | integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | Instrument Level Tests (ILT) SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: | SCU-05 should work first |
| Outline procedure and analysis: | Select the appropriate display (SCU PARAMETER) Execute command list to switch on the heaters Send a command to set a value of current Check values of parameters (SPHSCURR, EVHSCURR, TCHTRCURR, SPHTRCURR…) …what about parameters for voltage? (in SCU science but not in HK nominal report?) parameters for temperatures? |
| Success/Failure Criteria: | Test passed if voltage and temperature are read and have the values expected. |

## SCU-07 (SCU Test pattern test)

| ID: | SCU-07 |
| --- | --- |
| Test name: | Test pattern test |
| Description of test: | Load test pattern and check content of packets. Perform test for each type of packet (SMEC, BSM, ENG) |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at SCU test pattern packet Request N test pattern packet. Check content. Repeat for each type of packet (SMEC, BSM, ENG) |
| Success/Failure Criteria: | test passed if the test pattern is properly generated |

## MCU-01 (MCU Power On)

| ID: | MCU-01 |
| --- | --- |
| Test name: | MCU Power On |
| Description of test: | Power on MCU from SCU, check all voltage, check threshold flags, check if DSP booted. Check HK |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Select the appropriate display (MCU PARAMETER) Send command to power MCU from SCU Wait 10 seconds Check that all voltage come up and are in range (parameters MCUP15V, MCUM15V, MCUP13V, MCUM13V, MCUP5V) and check threshold flag Check boot status register (MCUBOOT STAT) to see if DSP booted Switch off MCU |
| Success/Failure Criteria: | if DSP not booted, then OBS should cope with this |

## MCU-02_0 (MCU Science Packet generation check)

| ID: | MCU-02_0 |
| --- | --- |
| Test name: | MCU Science packet generation check |
| Description of test: | Check that the number of frames generated matches the number commanded Perform test for each type of packet. (SMEC, BSM, ENG) |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required QLA may be required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | In SCOS, select the appropriate display (DPU and OBS PARAMETER) Send request for N SMEC science packets Wait for packets to be generated Check that N science packets were generated, i.e. that parameter TM5N (sequence count of TM packets for SCU Science) increases until N Repeat steps 2 to 4 for BMS and ENG science packet |
| Success/Failure Criteria: | Test passed if number of packets generated = number requested. If not? |

## MCU-02 (MCU Science Data check)

| ID: | MCU-02 |
| --- | --- |
| Test name: | MCU Science Data Check |
| Description of test: | Check of integrity of MCU science data. Perform test for each type of packet. (SMEC, BSM, ENG) |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required QLA may be required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Send a request for N SMEC science packets. Wait for packets to be generated Use either QLA or “Packet display” with SCOS to see the content of the frames Compare it with the HK parameter. Repeat steps 1 to 4 for BSM and ENG science packet parameters to compare: SMEC: Optical encoder coarse position, fine position, LVDT DC signal, Motor BEMF BSM: Chop position, Chop motor current, chop BEMF, jiggle position, Jiggle motor current, Jiggle BEMF ENG: SMEC encoder signal 1/2/3, SMEC LVDT AC/DC signal, SMEC motor current/voltage, chop/jiggle magneto resistor, chop/jiggle motor current, chop/jiggle servo error. |
| Success/Failure Criteria: | Test passed if values in the science frames are identical to those in the HK nominal report If not? |

## MCU-03 (MCU Test Pattern test)

| ID: | MCU-03 |
| --- | --- |
| Test name: | MCU Test Pattern test |
| Description of test: | Load test pattern and check contents of packets. Perform test for each type of packet (SMEC, BSM, ENG) |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at MCU test pattern packet Request N test pattern packet. Check content. Repeat for each type of packet (SMEC, BSM, ENG) |
| Success/Failure Criteria: | test passed if the test pattern is properly generated |

## MCU-04 (MCU launch latch check)

| ID: | MCU-04 |
| --- | --- |
| Test name: | MCU launch latch check |
| Description of test: | Open launch latch, move SMEC short distance and check response |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT LAM EGSE only |
| Test Conditions: | Warm and cold |
| Constraints: | Can only be performed in cryostat configuration |
| Outline procedure and analysis: | Select appropriate display (SMEC PARAMETERS) Send command to open launch latch Send command to move SMEC by a short distance Check response (parameter SMECLVDTPOSN ?) |
| Success/Failure Criteria: | Test passed if SMEC moves. |

## SMEC-01 (SMEC Switch On)

| ID: | SMEC-01 |
| --- | --- |
| Test name: | SMEC Switch On |
| Description of test: | Switch on LVDT, fringe encoder and motor. Check encoder signal, status bits, current |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Select appropriate display (SMEC PARAMETERS) Send command to switch on LVDT Send command to switch on fringe encoder Send a command to switch on motors Check LVDT position (SMECLVDTPOSN) Check encoder signals (SMECSINE000SIG SMECSINE120SIG, SMECSINE240SIG) Check status bit (SMECSTAT) and current (SMECMOTORCURR) |
| Success/Failure Criteria: | Test passed if values read are values expected |

## SMEC-02 (SMEC Initialisation)

| ID: | SMEC-02 |
| --- | --- |
| Test name: | SMEC initialisation |
| Description of test: | Initialise SMEC. Check LVDT, fringe encoder and current |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Send command to initialise SMEC Check LVDT position (SMECLVDTPOSN) Check encoder signals (SMECENCPOSN, SMECSINE000SIG SMECSINE120SIG, SMECSINE240SIG) Check current (SMECMOTORCURR) |
| Success/Failure Criteria: | Test passed if values read are values expected |

## SMEC-03 (SMEC Position test)

| ID: | SMEC-03 |
| --- | --- |
| Test name: | SMEC position test |
| Description of test: | Move to set positions and measure read back position, currents, back emf |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Select appropriate display (SMEC PARAMETERS) Send command to move SMEC to a chosen position Read position (SMECENCPOSN, SMECLVDTPOSN) Check current (SMECMOTORCURR) and back emf (SMECMOTORBEMF) |
| Success/Failure Criteria: | Test passed if SMEC moves to the set position and current and back emf have expected values |

## SMEC-04 (SMEC Multiple position test)

| ID: | SMEC-04 |
| --- | --- |
| Test name: | SMEC Multiple position test |
| Description of test: | Move to number of set positions and measure positions, current, back emf – repeat for different scan speed |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for SMEC position Record position sensor signal, motor current, back emf Send command to move SMEC to a chosen position Check SMEC movement (length, velocity and stability of the scan) Set a new scan speed, a new position and repeat steps 1 to 4 Parameters present in an “Engineering” frame: Servo error, Encoder signal 1/2/3, LVDT AC/DC signal, motor current, motor voltage Parameters present in a “SMEC” frame: opt. enc. coarse position, opt. enc. fine position, LVDT DC signal, Back EMF |
| Success/Failure Criteria: | Test passed if SMEC moves to the set position with the requested scan speed, and current and back emf have expected values |

## SMEC-05 (SMEC Saw tooth scan test)

| ID: | SMEC-05 |
| --- | --- |
| Test name: | SMEC Saw tooth scan test |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scan speed |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for SMEC position Record position sensor signal, motor current, back emf Send command for SMEC to do a saw tooth scan Check SMEC movement (length, velocity and stability of the scan) Set a new scan speed, a new position and repeat steps 1 to 4 Parameters present in an “Engineering” frame: Servo error, Encoder signal 1/2/3, LVDT AC/DC signal, motor current, motor voltage Parameters present in a “SMEC” frame: opt. enc. coarse position, opt. enc. fine position, LVDT DC signal, Back EMF |
| Success/Failure Criteria: | Test passed if SMEC does the requested scan and current and back emf have expected values |

## SMEC-06 (SMEC triangular test)

| ID: | SMEC-06 |
| --- | --- |
| Test name: | SMEC triangular scan test |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scan speed |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | Prime and redundant |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for SMEC position Record position sensor signal, motor current, back emf Send command for SMEC to do a triangular scan Check SMEC movement (length, velocity and stability of the scan) Set a new scan speed, a new position and repeat steps 1 to 4 Parameters present in an “Engineering” frame: Servo error, Encoder signal 1/2/3, LVDT AC/DC signal, motor current, motor voltage Parameters present in a “SMEC” frame: opt. enc. coarse position, opt. enc. fine position, LVDT DC signal, Back EMF |
| Success/Failure Criteria: | Test passed if SMEC does the requested scan and current and back emf have expected values |

## BSM-01 (BSM Power on motors and sensors)

| ID: | BSM-01 |
| --- | --- |
| Test name: | BSM Power on motors and sensors |
| Description of test: | Check telemetry responses, consumed power |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Select appropriate displays (CHOP PARAMETERS and JIGGLE PARAMETERS) Send command to switch on BSM motors Check parameters: MCUBSMTEMP, CHOPSENSPOSN0, CHOPSENSPOSN1, CHOPMOTORCURR, CHOPBEMF, JIGGSENSPOSN0, JIGGSENSPOSN1, JIGGMOTORCURR, JIGGBEMF |
| Success/Failure Criteria: | test passed if BSM is switched on and parameters have the values expected |

## BSM-02 (BSM Move to position)

| ID: | BSM-02 |
| --- | --- |
| Test name: | BSM Move to Position |
| Description of test: | Move BSM to set positions. Check telemetry responses, consumed power |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | ILT SCOS-2000 required |
| Test Conditions: | Warm and cold |
| Constraints: |  |
| Outline procedure and analysis: | Select appropriate displays (CHOP PARAMETERS and JIGGLE PARAMETERS, TBC) Send command to move BSM to set positions Check parameters: MCUBSMTEMP, CHOPPOSN, CHOPPOSNERR, CHOPSENSPOSN0, CHOPSENSPOSN1, CHOPMOTORCURR, CHOPBEMF, JIGGPOSN, JIGGPOSNERR, JIGGSENSPOSN0, JIGGSENSPOSN1, JIGGMOTORCURR, JIGGBEMF |
| Success/Failure Criteria: | test passed if BSM moves to the requested position and parameters have the values expected |

## BSM-03 (BSM Scan test)

| ID: | BSM-03 |
| --- | --- |
| Test name: | BSM Scan test |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for BSM Record motor current, position sensor signal Send command to move BSM to a chosen position Repeat for a TBD set of position Derive characteristics from Position vs Time (scan velocity and stability, consumed power…) Parameters present in an “Engineering” frame: Magneto resistor signal, Motor current, Servo Error (for chopper and jiggle) Parameters present in a “BSM” frame: magneto resistor signal, Motor setting current, motor measured voltage (for chopper and jiggle) |
| Success/Failure Criteria: | test passed if BSM moves to the set positions, if characteristics are as expected |

## BSM-04 (BSM Operating mode test)

| ID: | BSM-04 |
| --- | --- |
| Test name: | BSM Operating mode test |
| Description of test: | Set BSM to each mode (Chop, jiggle…). Check telemetry response, consumed power |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for BSM Record motor current, position sensor signal Send command for BSM to chop Derive characteristics from Position vs Time (period, amplitude, velocity, stability…), check consumed power. Repeat for a TBD set of chop throws, TBD set of chop periods Repeat for the jiggle mode Parameters present in an “Engineering” frame: Magneto resistor signal, Motor current, Servo Error (for chopper and jiggle) Parameters present in a “BSM” frame: Magneto resistor signal, Motor setting current, motor measured voltage (for chopper and jiggle) |
| Success/Failure Criteria: | test passed if BSM behave as requested, if characteristics are as expected |

## PCAL-01 (PCAL Characterisation test)

| ID: | PCAL-01 |
| --- | --- |
| Test name: | PCAL characterisation test |
| Description of test: | Step through each setting and measure telemetry response |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for PCAL Set a current Record temperatures, current and voltage Repeat for a TBD set of currents Derive PCAL characteristics (stability, consumed power) Parameters present in a SCU science packet: PhCalCur, PhCalVolt …PhCal temp ?? |
| Success/Failure Criteria: | Test passed if PCAL characteristics are as expected |

## SCAL-01 (SCAL Characterisation test)

| ID: | SCAL-01 |
| --- | --- |
| Test name: | SCAL Characterisation tests |
| Description of test: | Set current and record temperatures |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for SCAL Record temperatures and voltage Set a current (Heaviside step) Derive time constant from Tp vs Time Repeat for a TBD set of currents Set current to 0 Derive cooling time constant from Tp vs Time Derive other characteristics? (Consumed power…) Parameters present in a SCU science packet: Scal2Temp, Scal4Temp, Scal2curr, Scal4curr, Sca2V, Scal4V |
| Success/Failure Criteria: | Test passed if SCAL characteristics are as expected |

## SCAL-02 (SCAL PID test)

| ID: | SCAL-02 |
| --- | --- |
| Test name: | SCAL PID tests |
| Description of test: | Set temperatures and record temperatures |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: |  |
| Instrument Configuration: | SPIRE in READY mode |
| EGSE Configuration: | QLA required |
| Test Conditions: | Cold |
| Constraints: |  |
| Outline procedure and analysis: | Set-up QLA to look at the parameters for SCAL Set temperature Record temperatures and voltages Repeat for a TBD set of temperature Derive stability from Tp vs Time Parameters present in a SCU science packet: Scal2Temp, Scal4Temp, Scal2curr, Scal4curr, Sca2V, Scal4V |
| Success/Failure Criteria: | Test passed if SCAL characteristics are as expected, if the stability is within requirement |