| SUBJECT: | SPIRE Functional test specificationsSPIRE Functional test specifications |
| --- | --- |
| PREPARED BY: | Samuel RonayetteSamuel Ronayette (U o L/RAL) |
| DOCUMENT No: | unknownSPIRE-RAL-DOC-001652 |
| ISSUE: | draftissue 1.2 | Date: | 28th April 200305/07/2004 |
| CHECKED BY: | Asier Abreu (U o L/ RAL) | Date: |  |
| APPROVED BY: |  | Date: |  |

Distribution

| RAL | Dave Smith |
| --- | --- |
|  | Ken King |
|  | Bruce Swinyard |
|  | Tanya Lim |
|  | Sunil Sidher |
|  | Asier Abreu |

Change Record

| ISSUE | DATE | Changes |
| --- | --- | --- |
| 0.1 | 08th April 2003 | 1st Draft for comment |
| 0.2 | 28th April 2003 | Updated Draft incorporating comments |
| 0.3 | 23rd May 2003 | test FUNC-MCU-05 changed into test FUNC-SMEC-02 10 DCU test specification added |
| 0.3_1 | 2nd June 2003 | Page Headers Corrected |
| 0.4 | 20th June 2003 | test FUNC-DCU-06 updated |
| 0.5 | 25th July 2003 | LEDs test added (FUNC-SMEC-04) DCU tests modified, new numbering |
| 0.6 |  | Characterisation tests described in more details FUNC-SMEC-01 includes initialisation and SMEC initialisation removed Definition of a switch on detector “test” Test condition summary added Test flow added |
| 0.7 | 2nd Sept. 2003 | New format Command names added Values of parameters added |
| 1.0 (draft 1) | 22/09/2003 | BSM-04 modified FUNC-SCAL-02 modified FUNC-SCU-07 modified FUNC-DCU-04 to FUNC-DCU-11 split in photometer and spectrometer tests FUNC-DCU-08 split in short and full test “switch off” procedure added in each test concerned Use of the command SetDRelOnOff corrected |
| 1.0 (draft 2) | 05/12/2003 | New definition of BSM tests (open loop) |
| 1.0 | 23/12/2003 | New BSM test (BSM-05c and j) Other BSM tests revised Spectro sampling frequency corrected (in DCU tests) Flow chart updated SMEC launch latch test modified Other SMEC test revised (cf use of open loop command) |
| 1.1 | ? | ? |
| 1.2 | 05/07/2004 | Correct current values for SCU-07 Cooler Heater Check DCU Frequency test: new frequencies BBID and STEPS added (Asier) |

Table of Contents

1.	Introduction	6

1.1	Scope	6

1.2	Structure of Document	6

1.3	Documents	6

2.	Test Summaries	7

2.1	SCU tests	7

2.2	MCU tests	7

2.3	SMEC tests	8

2.4	BSM tests	8

2.5	PCAL test	9

2.6	SCAL tests	9

2.7	DCU tests	9

3.	Summary of test condition and test flow	11

3.1	Summary of test condition	11

3.2	Test flow	12

4.	Test Specifications	18

4.1	FUNC-SCU-01, SCU science packet generation check	18

4.2	FUNC-SCU-02, SCU Science data check	20

4.3	FUNC-SCU-03, SCU DC thermometry check	22

4.4	FUNC-SCU-04, SCU PCAL check	24

4.5	FUNC-SCU-05, SCU SCAL check	26

4.6	FUNC-SCU-06, SCU AC thermometry check	28

4.7	FUNC-SCU-07, SCU cooler heaters check	30

4.8	FUNC-SCU-08, SCU Test pattern test	32

4.9	FUNC-MCU-01, MCU power on	34

4.10	FUNC-MCU-02, MCU Science Packet Generation Test	37

4.11	FUNC-MCU-03, MCU Science data check	40

4.12	FUNC-MCU-04, MCU test pattern test	43

4.13	FUNC-SMEC-01, SMEC switch on and initialisation	45

4.14	FUNC-SMEC-02, SMEC launch latch check	49

4.15	FUNC-SMEC-03, SMEC LEDs test	51

4.16	FUNC-SMEC-04, SMEC position test	53

4.17	FUNC-SMEC-05, SMEC step and look scan test	55

4.18	FUNC-SMEC-06, SMEC saw tooth scan test	58

4.19	FUNC-SMEC-07, SMEC triangular scan test	61

4.20	FUNC-SMEC-08, SMEC open loop position test	64

4.21	FUNC-SMEC-09, SMEC open loop position test	66

4.22	FUNC-BSM-01c, BSM power on chop motor and sensors	69

4.23	FUNC-BSM-01j, BSM power on jiggle motor and sensors	71

4.24	FUNC-BSM-02c, BSM chop axis in open loop	73

4.25	FUNC-BSM-02j, BSM jiggle axis in open loop	75

4.26	FUNC-BSM-03c, BSM position test, loop closed on MRS, chop axis	77

4.27	FUNC-BSM-03j, BSM position test, loop closed on MRS, jiggle axis	79

4.28	FUNC-BSM-04c, BSM chop close loop scan test	81

4.29	FUNC-BSM-04j, BSM jiggle close loop scan test	83

4.30	FUNC-BSM-05c, BSM chop open loop scan test	85

4.31	FUNC-BSM-05j, BSM jiggle open loop scan test	87

4.32	FUNC-BSM-06, BSM operating mode test	89

4.33	FUNC-PCAL-01, PCAL characterisation test	92

4.34	FUNC-SCAL-01, SCAL characterisation test	95

4.35	FUNC-SCAL-02, SCAL PID test	98

4.36	FUNC-DCU-01, DCU Science Packet generation check	100

4.37	FUNC-DCU-02, DCU Science data check	102

4.38	FUNC-DCU-03, DCU Test pattern test	105

4.39	FUNC-DCU-04P, DCU Photometer LIAs switch on	108

4.40	FUNC-DCU-04S, DCU Spectrometer LIAs switch on	110

4.41	FUNC-DCU-05P, DCU Photometer Offset test	112

4.42	FUNC-DCU-05S, DCU Spectrometer Offset test	115

4.43	FUNC-DCU-06P, DCU Photometer JFET heaters	118

4.44	FUNC-DCU-06S, DCU Spectrometer JFET heaters	121

4.45	FUNC-DCU-07P, DCU Photometer JFET test	124

4.46	FUNC-DCU-07S, DCU Spectrometer JFET test	128

4.47	FUNC-DCU-08P_full, DCU Full photometer phase shift test	132

4.48	FUNC-DCU-08P_short, DCU Short Photometer phase shift test	135

4.49	FUNC-DCU-08S_full, DCU Full spectrometer Phase shift test	137

4.50	FUNC-DCU-08S_short, DCU Short spectrometer Phase shift test	140

4.51	FUNC-DCU-09P, DCU Photometer bias frequency test	142

4.52	FUNC-DCU-09S, DCU Spectrometer bias frequency test	144

4.53	FUNC-DCU-10P, DCU Photometer bias amplitude test	146

4.54	FUNC-DCU-10S, DCU Spectrometer bias amplitude test	150

4.55	FUNC-DCU-11P, DCU Photometer detector switch on	154

4.56	FUNC-DCU-11S, DCU Spectrometer detector switch on	157

4.57	FUNC-DCU-12P, DCU Photometer detector settings	160

4.58	FUNC-DCU-12S, DCU Detector settings	162

Figures

Test flow	12

Tables

Summary of test condition	11

Glossary

| AD | Applicable Document |
| --- | --- |
| AIV | Assembly Integration and Verification |
| AVM | Avionics Model |
| BSM | Beam Steering Mechanism |
| CQM | Cryogenic Qualification Model |
| DCU | Detector Control Unit |
| DPU | Digital Processing Unit |
| DRCU | Detector Readout and Control Unit |
| EGSE | Electrical Ground Support Equipment |
| FLT | Flight Level Test |
| FTS | Fourier Transform Spectrometer |
| HOB | Herschel Optical Bench |
| ILT | Instrument Level Test |
| IST | Integrated System Test |
| MCU | Mechanism Control Unit |
| MGSE | Mechanical Ground Support Equipment |
| N/A | Not Applicable |
| PCAL | Photometer Calibrator source |
| PFM | Proto-Flight Model |
| PLT | Payload Level Test |
| RAL | Rutherford Appleton Laboratory |
| RD | Reference Document |
| SCAL | Spectrometer Calibrator source |
| SCU | Subsystem Control Unit |
| SMEC | Spectrometer Mechanism |
| SPIRE | Spectral and Photometric Imaging Receiver |
| STM | Structural Test Model |
| TBC | To Be Confirmed |
| TBD | To Be Defined |
| TFCS | Test Facility Control System |
| AD | Applicable Document |
| AIV | Assembly Integration and Verification |
| AVM | Avionics Model |

# Introduction

## Scope

This document contains the functional test specification for SPIRE. These are tests of different subsystems (SCU, MCU, SMEC, BSM, PCAL, SCAL, DCU) that will be performed on different model of the instrument (AVM, CQM, PFM, FS). They consist of integrity check and characterisation tests. These tests are not calibration or performance tests.

## Structure of Document

A summary of each test is given first (name, short description) and in the following section, we give the specification of each test. They are grouped together per subsystem (respectively SCU, MCU, SMEC, BSM, PCAL, SCAL, and DCU)

## Documents

Applicable Documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD 1 | SPIRE Data ICD | K.J. King | SPIRE-RAL-PRJ-001078, issue 1.0 | 15/01/2003 |
| AD2 | MCU/DCU Command list ICD an user manual | Didier Ferrand | LAM/ELE/SPI/011011, issue 3.0 | 15/01/2003 |
| AD3 | DRCU/DPU ICD | Pinsard, Ferrand, Mur | Sap-SPIRE-Cca-076-02, issue 1.0 | 14/02/2003 |

Reference Documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| RD 1 | Subsystem reaction for specification for the instrument simulator | Bruce Swinyard | SPIRE-RAL-NOT-001715, issue 1.0 | 11/06/2003 |
| RD 2 | Operating modes for the SPIRE instrument | Bruce Swinyard, Matt Griffin | SPIRE-RAL-DOC-000320, issue 3.0 | 04/01/2002 |
| RD 3 | MCU/BSM DM Mechanism, Integration and test report | Didier Ferrand | LAM/ELE/SPI/031020, issue 1.0 | 20/10/2003 |

# Test Summaries

## SCU tests

### FUNC-SCU-01, SCU Science packet generation check

To check the integrity of SCU science packet interface and that the SCU generates the nominal science frame data. Request SCU science frame and check that all the frames are received correctly by the DPU

### FUNC-SCU-02, SCU Science data check

Checking the integrity of the SCU science data. Request science frames and housekeeping data and compare content

### FUNC-SCU-03, SCU DC thermometry check

To check the integrity of thermometers. Switch on thermometers all at once and check values

### FUNC-SCU-04, SCU PCAL check

Checking the integrity of PCAL. Switch on PCAL, set a current and check voltage and current, then switch off

### FUNC-SCU-05, SCU SCAL check

Checking the integrity of SCAL. Switch on SCAL, set a current and check voltage, current and temperature changes, then switch off

### FUNC-SCU-06, SCU AC thermometry check

Checking the integrity of the AC thermometer. Switch on AC thermometers and check values

### FUNC-SCU-07, SCU cooler heater check

Checking the integrity of the cooler heaters. Switch on heaters, check voltage

### FUNC-SCU-08, SCU Test pattern test

Checking the integrity of SCU test pattern packet. Load test pattern and check content of packets

## MCU tests

### FUNC-MCU-01, MCU power on

To power on the MCU into a state ready to execute SMEC or BSM commands.

This procedure also tests the low speed interface to the MCU, used for command and housekeeping data transfer. Apply power to MCU from SCU and boot the MCU DSP ROM software, checking voltages and status

### FUNC-MCU-02, MCU Science packet generation check

To check the integrity of MCU science packet interface and that the MCU generates all the nominal science data frames. Generate science data from the SMEC and BSM at typical data rates and check that all the frames are received correctly by the DPU.

### FUNC-MCU-03, MCU Science data check

Checking the integrity of MCU science data. Compare science data with HK. Perform test for each type of packet. (SMEC, BSM, ENG)

### FUNC-MCU-04, MCU test pattern test

Checking the integrity of MCU test pattern packets. Load MCU test pattern and check contents of packets.

## SMEC tests

### FUNC-SMEC-01, SMEC switch on and initialisation

Switching on and initialising SMEC. SMEC ends up at home position and is ready for following tests. Switch on various part of the SMEC system then initialise it. Check parameters

### FUNC-SMEC-02, SMEC launch latch check

Checking that the launch latch works properly. Open launch latch, move SMEC short distance and check response

### FUNC-SMEC-03, SMEC LEDs test

Checking integrity of the LEDs. Switch on LEDs and step up voltages

### FUNC-SMEC-04, SMEC position test

Checking that SMEC moves to the position commanded. Move to set positions and measure read back position, currents, back emf

### FUNC-SMEC-05, SMEC multiple position test

Characterising SMEC movement for a step-and-look scan. Perform a step-and-look scan, measure positions, current, back emf

### FUNC-SMEC-06, SMEC saw tooth scan test

Characterising SMEC movement during a saw tooth scan. Scan over range and measure positions, current, back-emf – repeat for different scans

### FUNC-SMEC-07, SMEC triangular scan test

Characterising SMEC movement during a triangular scan. Scan over range and measure positions, current, back-emf – repeat for different scans

## BSM tests

### FUNC-BSM-01c/j, BSM power on motor and sensor

Checking the integrity of BSM motors and sensors. Switch BSM motors on. Check telemetry responses, consumed power

### FUNC-BSM-02c/j, position test, open loop

Observing BSM behaviour in open loop. Determining sensor sensibility and scaling the back emf

### FUNC-BSM-03c/j, position test, close loop

Checking that BSM moves to the position commanded and that the control loop on MRS works

### FUNC-BSM-04c/j, BSM scan test, close loop

Characterising BSM movement for a single scan, loop closed on MRS. Move BSM to multiple set positions. Check telemetry response, consumed power.

### FUNC-BSM-05c/j, BSM scan test, open loop

Characterising BSM movement for a single scan, loop open (no sensor). Move BSM to multiple set positions. Check telemetry response, consumed power.

### FUNC-BSM-06, BSM operating mode test

Characterising BSM movement on chop axis for various positions on jiggle axis. Set BSM to each mode (Chop, jiggle…). Check telemetry response, consumed power

## PCAL test

### FUNC-PCAL-01, PCAL characterisation test

Characterising PCAL. Go through each setting and measure telemetry response

## SCAL tests

### FUNC-SCAL-01, SCAL characterisation test

Characterising SCAL. Set current and record temperatures

### FUNC-SCAL-02, SCAL PID test

Verifying that the PID controller works properly. Set temperatures and record temperatures

## DCU tests

### FUNC-DCU-01, DCU Science Packet generation check

Checking the integrity of the DCU science packet generation. Request DCU science packet and check that the number of frames generated matches the number commanded

### FUNC-DCU-02, DCU Science data check

Checking the integrity of the DCU science packet. Request DCU science packet and check some values

### FUNC-DCU-03, DCU Test pattern test

Checking the integrity of the DCU test pattern. Load test pattern and check content of packets. Perform for each type of packet (Photometer and Spectrometer)

### FUNC-DCU-04P/S, DCU LIAs switch on

Checking the integrity of the LIAs. Switch on LIAs. Check parameter values

### FUNC-DCU-05P/S, DCU Offset test

Characterising the detector output for various values of offset. Change offsets up and down, record output and analyse off line

### FUNC-DCU-06P/S, DCU JFET heaters

Characterising influence of JFET heaters on detectors output. Switch on JFET heaters. Record detector signal, derive noise

### FUNC-DCU-07P/S, DCU JFET test

Characterising detector signal for various JFET drain voltages. Switch on JFETS with Vdd on. Activate auto-offset control. Measure noise. Vary Vss and measure noise

### FUNC-DCU-08P/S_full, DCU Phase shift test

Adjusting the demodulation phase. Set phase to /2 and 3/2 adjust phase by small amounts to find minimum – set phase back to +d

### FUNC-DCU-08P/S_short, DCU Phase shift test

Set phase at /2+d and verify that signal = 0

### FUNC-DCU-09P/S, DCU Bias frequency test

Characterising detector signal for various bias frequencies. Set up bias frequency. Measure noise

### FUNC-DCU-10P/S, DCU Bias amplitude test

Characterising detector signal for various bias amplitudes. Set bias amplitude. Measure change in signal level, measure noise, read out offsets

### FUNC-DCU-11P/S, DCU detectors switch on

Switching on the all chain for reading detectors signal. This procedure must be performed before test FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10 – see flowchart p17

### FUNC-DCU-12P/S, DCU detectors settings

Setting sampling frequency, bias frequency, bias amplitude and demodulation phase when LIA and JFET are already on. This procedure must be performed after FUNC-DCU-04 and FUNC-DCU-07 before carrying on with tests FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10. – see flowchart p17

# Summary of test condition and test flow

In this section is table that sum up the condition in which each test is performed, i.e. instrument configuration, if the test is performed during short functional test, during warm functional test. It is followed by the test flow

## Summary of test condition

S: performed during Short functional test

F: performed during Full functional test (all?)

W: performed during Warm functional test

C: performed during Cold functional test (all?)

*: test that have a short and a full version (refer to procedure for details)

!: different script according to instrument model (refer  to procedure for details)

?: TBD or TBC

| TEST | TEST CONDITION | Estimated dur. |
| --- | --- | --- |
| FUNC-SCU-01 | S F W C | 1 min |
| FUNC-SCU-02 | F W C | 2 min |
| FUNC-SCU-03 | S F W C | 1 min |
| FUNC-SCU-04 | S F W C | 1 min |
| FUNC-SCU-05 | S F W C | 3 min |
| FUNC-SCU-06 | S F W C | 1 min |
| FUNC-SCU-07 | S F W C | 2 min |
| FUNC-SCU-08 | F W C | 1 min |
| FUNC-MCU-01 | S F W C | 1 min |
| FUNC-MCU-02 | S F W C | 2 min |
| FUNC-MCU-03 | F W C | 2 min |
| FUNC-MCU-04 | F W C | 1 min |
| FUNC-SMEC-01 | S F W C | 3 min |
| FUNC-SMEC-02 | S F W C | 3 min |
| FUNC-SMEC-03 | S F W C | 1 min |
| FUNC-SMEC-04 | S F W C * | 3 min |
| FUNC-SMEC-05 | F W C | 18 min |
| FUNC-SMEC-06 | F W C | 5 min |
| FUNC-SMEC-07 | F W C | 5 min |
| FUNC-BSM-01 | S F W C | 1 min |
| FUNC-BSM-02c | F W C | 1 min |
| FUNC-BSM-02j | F W C | 5 min |
| FUNC-BSM-03c | F W C | 5 min |
| FUNC-BSM-03j | F W C | 5 min |
| FUNC-BSM-04 | F W C | 5 min |
| FUNC-BSM-05 | F W C | 45 min ? |
| FUNC-PCAL-01 | F W? C | 2 min |
| FUNC-SCAL-01 | F W? C | 15 min |
| FUNC-SCAL-02 | F W? C |  |
| FUNC-DCU-01 | S F W C * | 1 ¦ 3 min |
| FUNC-DCU-02 | F W C * | 2 ¦ 5 min |
| FUNC-DCU-03 | F W C * | 2 ¦ 5 min |
| FUNC-DCU-04P/S | S F W C | 1 min |
| FUNC-DCU-05P/S | F C ! | 3 min |
| FUNC-DCU-06P/S | F C ! | 3 min |
| FUNC-DCU-07P/S | F C ! | 4 min |
| FUNC-DCU-08P/S_full | F C ! | 10 min |
| FUNC-DCU-08P/S_short | F C ! | 3 min |
| FUNC-DCU-09P/S | F C ! | 3 min |
| FUNC-DCU-10P/S | F C ! | 3 min |
| FUNC-DCU-11P/S | F C ! | 1 min |
| FUNC-DCU-12P/S |  |  |

##### Total duration:

##### Comment/open issues

-There might be other tests that deserve a * (different short and full), TBD

## Test flow

Warm short functional test: perform test in yellow boxes only

Warm full functional test: perform test in yellow and pink boxes only

Cold short functional test: perform test in yellow and blue boxes only

Cold full functional test: perform all tests

Test into brackets: cannot be performed with the CQM

Comment/open issues:

-Maybe FUNC-SCU-06 (AC thermometry) should be repeated when the detectors are at 300 mK?

# Test Specifications

## FUNC-SCU-01, SCU science packet generation check

| ID: | FUNC-SCU-01 |
| --- | --- |
| Purpose | To check the integrity of SCU science packet interface and that the SCU generates the nominal science frame data |
| Description of test: | Request SCU science frame and check that all the frames are received correctly by the DPU |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration: | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | < 1 min |
| Constraints: | none |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
|  | Select SCOS display (DPU and OBS PARAMETER) |  |  |
|  | Request nominal SCU science frames Check frame counter Set OBS STEP 1 Request 200 SCU TM frames, sampled at 80Hz Flush SCU FIFO Set OBS STEP(0xffff) | Read SCUFRAMECNT Set_OBS_STEP(0x0001) SetFrameConf(0x0000) A0830000 SetSeqLength(0x00C8) A08400C8 SetFrameCtrl(0x0001) A0820001 TM(8,4: CA-02) (0x4000) Set_OBS_STEP(0xffff) |  |
|  | Check generation of packets Check that frame counter has incremented by ~200 Check No Science Packet Events occurred | Read SCUFRAMECNT ? |  |

Success/Failure Criteria: Test passed if correct number of frame generated

Comments/Open issues:

## FUNC-SCU-02, SCU Science data check

| ID: | FUNC-SCU-02 |
| --- | --- |
| Purpose | Checking the integrity of the SCU science data |
| Description of test: | Request science frames and housekeeping data and compare content |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial Configuration: | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 2 min |
| Constraints: | SCU thermometry needs to be switched on – see test FUNC-SCU-03 and FUNC-SCU-06 FUNC-SCU-01 successfully passed |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
|  | Run QLA script FUNC-SCU-02 |  |  |
|  | With SCOS, request nominal SCU science frames Mark beginning of test (change BBId) Request 200 SCU TM frames, sampled at 80Hz Flush SCU FIFO Mark end of data stream (step to 0xffff) Once all packet are received, mark end of test (reset BBId to 0x8000000) | BBTYPE =0x8A00 ,STEP=0x0001 SetFrameConf(0x0000) A0830000 SetSeqLength(0x00C8) A08400C8 SetFrameCtrl(0x0001) A0820001 TM(8,4: CA-02) (0x4000) STEP=0xffff BBTYPE =0x8000 |  |
|  | Display parameters with QLA Display the last science packet received Display the HK parameters received just after BBId changed to 0x8000000 | TCHeaterVolt, PhCalCur, SCal4Cur, SCal2Cur, PhCalVolt, SCal4Volt, SCal2Volt, CPHPtemp, CPHStemp, CEHStemp, CSHTtemp, SOBtemp, SL0temp, PL0temp, SUBtemp, BAFtemp, BSMStemp, SCL2temp, SCL4temp, SCSTtemp, FTSStemp, FTSMtemp, BSMMtemp, CEVTemp |  |
|  | Analyse with QLA Compare science frame with HK Write in log file the result of the comparison | Check values on the ‘clock’ displays on the screen |  |

Success/Failure Criteria: Test passed if values in the science frame are identical to those in the HK nominal report.

Comments/Open issues:

## FUNC-SCU-03, SCU DC thermometry check

| ID: | FUNC-SCU-03 |
| --- | --- |
| Purpose | To check the integrity of thermometers |
| Description of test: | Switch on thermometers all at once and check values |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode - DC thermometers off |
| Final configuration | SPIRE in REDY mode - DC thermometers on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
|  | Select SCOS display (SCU parameters) |  |  |
|  | Switch On thermometers Switch on SCU DC Thermometers, all at once Wait 5 seconds (time for sensor to activate) | TempOnOff(0xFFFF) A085FFFF Wait 5 seconds |  |
|  | Check parameters values | CPHPtemp= 0x08CA CPHStemp= 0x08CA CEHStemp= 0x08CA CSHTtemp= 0x08CA SOBtemp= 0x08CA SL0temp= 0x0DAC PL0temp= 0x0DAC SUBtemp= 0x08CA BAFtemp= 0x08CA BSMStemp= 0x08CA SCL2temp= 0x08CA SCL4temp= 0x08CA SCSTtemp= 0x08CA FTSStemp= 0x08CA FTSMtemp= 0x08CA BSMMtemp= 0x08CA |  |

Switch off procedure:

| Action | Command |
| --- | --- |
| Switch off thermometers | TempOnOff(0x0000) A0850000 |

Success/Failure Criteria: Test passed if all thermometer channels are reading expected values

Comments/Open issues:

## FUNC-SCU-04, SCU PCAL check

| ID: | FUNC-SCU-04 |
| --- | --- |
| Purpose | Checking the integrity of PCAL |
| Description of test: | Switch on PCAL, set a current and check voltage and current, then switch off |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm (TBC) and cold |
| Total duration | < 1 min |
| Constraints: | None |

Procedure and analysis:

| Steps | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU parameters) |  |  |
| 2 | Switch On PCAL Set a current (2.25 mA) Wait 2 seconds | SetPhCalBias(0x050A) A0C8050A Wait 2 seconds |  |
| 3 | Check parameters values | PCALCURR = 0x2366 (2.25 mA) PCALV = 1FDD (567 mV) |  |
| 4 | Switch off PCAL | SetPhCalbias(0x0000) A0C80000 |  |

Success/Failure Criteria: Test passed if value of current read = current set and voltage has value expected

Comments/Open issues:

This test needs a calibration curve. Voltage expected is dependent on instrument model and version of MIB

Is it possible to perform this test warm?

Answer from Peter Hargrave: good question, I don’t know, maybe, has never been done

## FUNC-SCU-05, SCU SCAL check

| ID: | FUNC-SCU-05 |
| --- | --- |
| Purpose | Checking the integrity of SCAL |
| Description of test: | Switch on SCAL, set a current and check voltage, current and temperature changes, then switch off |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 3 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU parameters) |  |  |
| 2 | Switch on SCAL-4 Set a current (2.5 mA) Wait 1 minute (TBC) | SetSCal4Bias(0x072F) A0CC072F Wait 1 minute |  |
| 3 | Check parameters values for SCAL-4 | SCAL4CURR = 0x36F6 (2.5mA) SCAL4V = 0x30DA (1.25V) SCAL4TEMP = 0x |  |
| 4 | Switch on SCAL-2 Set a current (2.5 mA) Wait 1 minute (TBC) | SetSCal2Bias(0x072E) A0CA072E Wait 1 minute |  |
| 5 | Check parameters values for SCAL-2 | SCAL2CURR = 0x36FC (2.5mA) SCAL2V = 0x0x30E6 (1.25V) SCAL2TEMP = 0x |  |
| 6 | Switch off SCAL2 Switch off SCAL4 | SetScal2Bias(0x0000) A0CC0000 SetScal4Bias(0x0000) A0CA0000 |  |

Success/Failure Criteria Test passed if value of current read = current set and voltage and temperature have values expected

Comments/Open issues:

This test needs a calibration curve.

Time constant is long and may be we do not wait enough for the stable temperature to be achieved, but it is not the purpose of this test. We just make sure that the wires are OK (shall we wait even less than 1 min then?)

## FUNC-SCU-06, SCU AC thermometry check

| ID: | FUNC-SCU-06 |
| --- | --- |
| Purpose | Checking the integrity of the AC thermometer |
| Description of test: | Switch on AC thermometers and check values |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode – AC thermometers off |
| Final configuration | SPIRE in REDY mode – AC thermometers on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | < 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU PARAMETERS) |  |  |
| 2 | Switch on thermometer Switch on AC thermometer Wait 5 seconds | SubKOnOff(0x0001) A0860001 Wait 5 seconds |  |
| 3 | Check parameter value | CEVtemp (SUBKTEMP) = 0x05DC or 0x0247 or…(cf comment) |  |

Switch off procedure:

| Action | Command |
| --- | --- |
| Switch off thermometer | SubKOnOff(0x0000) A0850000 |

Success/Failure Criteria: Test passed if thermometer channel is reading expected value

Comments/Open issues:

Temperature is dependent on cooler condition. If SPIRE is in READY mode, temperature expected is > 4K If in STAND-BY mode, temperature expected is 300 mK

## FUNC-SCU-07, SCU cooler heaters check

| ID: | FUNC-SCU-07 |
| --- | --- |
| Purpose | Checking the integrity of the cooler heaters |
| Description of test: | Switch on heaters, check voltage |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | FUNC-SCU-06 should be completed successfully |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU PARAMETERS) |  |  |
| 2 | Heaters switch on Apply 1mA to EVHS heater Apply 1mA to SPHS heater Apply 1mA to SP heater Wait 1 minute | SetEVHSHeatB(0x09F3) A0C409F3 SetSPHSHeatB(0x09F2) A0C509F2 SetSPHeaterB(0x0054) A0C70054 Wait 1 minute |  |
| 3 | Check parameter values | EVHSHeatVolt = 0x3DBA (401.99 mV) SPHSHeatVolt = 0x3DB5 (401.95 mV) SPHeaterVolt = 0x028A (401.33 mV) |  |
| 4 | Switch off heaters Set all currents down to 0 | SetEVHSHeatB(0x0000) A0C40000 SetSPHSHeatB(0x0000) A0C50000 SetSPHeaterB(0x0000) A0C70000 |  |

Success/Failure Criteria: Test passed if voltages are read and have the value expected

Comments/Open issues:

Waiting 1 minute is not enough for the voltage to stabilize but it is not the purpose of this test. It is long enough for the voltage to move. We just check that we can measure something

## FUNC-SCU-08, SCU Test pattern test

| ID: | FUNC-SCU-08 |
| --- | --- |
| Purpose | Checking the integrity of SCU test pattern packet |
| Description of test: | Load test pattern and check content of packets |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final Configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | warm and cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-SCU-08 |  |  |
| 2 | With SCOS, request nominal SCU science frames Mark beginning of test (change BBId) Request 200 SCU TEST frames, sampled at 80Hz Flush SCU FIFO Mark end of data stream (step to 0xffff) Once all packet are received, mark end of test (reset BBId to 0x8000000) Display contents of last test packet | BBTYPE =0x8A07 ,STEP=0x0001 SetFrameConf(0x0000) A0838000 SetSeqLength(0x00C8) A08400C8 SetFrameCtrl(0x0001) A0820001 TM(8,4: CA-02) (0x4000) STEP=0xffff BBTYPE =0x8000 |  |
| 3 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-SCU-08 Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |

Success/Failure criteria: Test passed if test pattern generated is similar to the one generated during previous run of this test.

Comments/Open issues:

## FUNC-MCU-01, MCU power on

| ID: | FUNC-MCU-01 |
| --- | --- |
| Purpose | To power on the MCU into a state ready to execute SMEC or BSM commands. This procedure also tests the low speed interface to the MCU, used for command and housekeeping data transfer |
| Description of test: | Apply power to MCU from SCU and boot the MCU DSP ROM software, checking voltages and status |
| Test Type: | Configuration |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode – MCU off |
| Final configuration | SPIRE in REDY mode – MCU on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (MCU_PARAMETERS) |  |  |
| 2 | Power on MCU from SCU | SetDRelOnOff(0x0004) A0870004 |  |
| 3 | Check MCU power status Check parameters values | SCUDCDCSTAT = 4 MCU+5V = 0xA804 MCU+15V = 0xF810 FUNC-MCU-15V = 0x07ED MCU+13V = 0xE80E FUNC-MCU-13V = 0x17EF MCUMACTEMP = 0x1000 MCUSMECTEMP = 0x1000 MCUBSMTEMP = 0x1000 |  |
| 4 | Check threshold flag |  |  |
| 5 | Reset MCU Subsystem Reset on Wait 5 seconds Reset off Wait 5 seconds Check Boot Status Check Interface Status | SetMCUCmdIfCtrl(0x0005) 90010005 Wait 5 seconds SetMCUCmdIfCtrl(0x0007) 90010007 Wait 5 seconds MCUBOOTSTAT = 0x0001 MCUIFSTAT = 0x0000 |  |
| 6 | Boot MCU Download Code to RAM Wait 5 seconds Check Boot Status Start Code in RAM | SetDownloadConfig(0xC000) 9021C000 Wait 5 seconds MCUBOOTSTAT = 0x0001 SetBootRam(0x0001) 90240001 |  |
| 7 | Check MCU status | MCUBOOTSTAT = 0x0001 MCUIFSTAT = 0x0000 MCUERR = 0x0000 MCUSCHEDCNTLSW = TBD MCUSCHEDCNTMSW = TBD |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Switch off MCU from SCU | SetDRelOnOff(0x0000) A0870000 |

Success/Failure Criteria Procedure completed with no errors

Comment/Open Issue:

How does this change for testing the redundant MCU?

If DSP is not booted correctly, then should the OBS cope with this?

## FUNC-MCU-02, MCU Science Packet Generation Test

| ID: | FUNC-MCU-02 |
| --- | --- |
| Purpose | To check the integrity of MCU science packet interface and that the MCU generates all the nominal science data frames |
| Description of test: | Generate science data from the SMEC and BSM at typical data rates and check that all the frames are received correctly by the DPU. |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode -MCU Powered on (following FUNC-MCU-01) |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display(s) (DPU PARAMETERS, OBS PARAMETERS) |  |  |
| 2 | Request ENG Frames Check Frame Counter Start ENG TM Packet Sampling at 64Hz Wait 10 seconds Stop ENG TM Packet Sampling Flush MCU FIFO Check Frame counter has incremented by ~640 Check No Science Packet Events occurred | Read MCUFRAMECNT SetTP14SampFreq(0x002C) 91C4002C Wait 10 seconds SetTP14SampFreq(0x0000) 91C40000 TM(8,4: CA-02) (0x2000) Read MCUFRAMECNT ???? |  |
| 3 | Request SMEC Frames Check Frame Counter Start SMEC TM Packet Sampling at 240Hz Wait 10 seconds Stop SMEC TM Packet Sampling Flush MCU FIFO Check Frame counter has incremented by ~2400 Check No Science Packet Events occurred | Read MCUFRAMECNT SetTP10SampFreq(0x000B) 91C0000B Wait 10 seconds SetTP10SampFreq(0x0000) 91C00000 TM(8,4: CA-02) (0x2000) Read MCUFRAMECNT ???? |  |
| 4 | Request BSM Frames Check Frame Counter Start BSM TM Packet Sampling at 64Hz Wait 10 seconds Stop BSM TM Packet Sampling Flush MCU FIFO Check Frame counter has incremented by ~640 Check No Science Packet Events occurred | Read MCUFRAMECNT SetTP12SampFreq(0x002C) 91C2002C Wait 10 seconds SetTP12SampFreq(0x0000) 91C20000 TM(8,4: CA-02) (0x2000) Read MCUFRAMECNT ???? |  |
| 5 | Request SMEC+BSM Frames Check Frame Counter Start SMEC TM Packet Sampling at 240Hz Start BSM TM Packet Sampling at 64Hz Wait 10 seconds Stop BSM TM Packet Sampling Stop SMEC TM Packet Sampling Flush MCU FIFO Check Frame counter has incremented by ~3040 Check No Science Packet Events occurred | Read MCUFRAMECNT SetTP10SampFreq(0x000B) 91C0000B SetTP12SampFreq(0x002C) 91C2002C Wait 10 seconds SetTP12SampFreq(0x0000) 91C20000 SetTP10SampFreq(0x0000) 91C00000 TM(8,4: CA-02) (0x2000) Read MCUFRAMECNT ???? |  |

Success/Failure Criteria: Correct number of frames generated for each frame type

Comment/Open Issue:

This procedure is currently only valid for QM0, which implements the start and stop of TM packets using a single SetSampFreq command. For later version sof the MCU there will be two commands: one to set the sample frequency and one to start/stop the sampling, for a number of frames.

## FUNC-MCU-03, MCU Science data check

| ID: | FUNC-MCU-03 |
| --- | --- |
| Purpose of test | Checking the integrity of MCU science data |
| Description of test: | Compare science data with HK. Perform test for each type of packet. (SMEC, BSM, ENG) |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode- MCU Powered on (following FUNC-MCU-01) |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | FUNC-MCU-01 and FUNC-MCU-02 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-MCU-03 |  |  |
| 2 | Request ENG Frames Mark beginning of test (change BBId) Request 256 ENG TM frames, sampled at 64Hz Once all packet are received, mark end of ENG packet transmission (reset BBId to 0x80000000) Flush MCU FIFO Display science parameters from last frame Display HK parameters received just after BBId changed to value3 | ? SetTP14SampFreq(0x002C) 91C4002C Wait 4 sec SetTP14SampFreq(0x0000) 91C40000 ? TM(8,4: CA-02) (0x2000) SMECENCSIG1, SMECENCSIG2, SMECENCSIG3, SMECLVDTACSIG, SMECLVDTDCSIG, SMECMOTORCURR, CHOPSENSSIG, CHOPMOTORCURR, JIGGSENSSIG, JIGGMOTORCURR |  |
| 3 | Request SMEC science Frames Mark beginning of SMEC packet transmission (change BBId ) Request 720 SMEC TM frames, sampled at 240Hz Once all packet are received, mark end of SMEC packet transmission (reset BBId to 0x80000000) Flush MCU FIFO Display science parameters from last frame Display HK parameters received just after BBId changed to value5 | ? SetTP10SampFreq(0x000B) 91C0000B Wait 3 sec SetTP10SampFreq(0x0000) 91C00000 ? TM(8,4: CA-02) (0x2000) SMECENCPOSN, SMECENCFINEPOSN, SMECLVDTDCSIG, SMECMOTORBEMF |  |
| 4 | Request BSM Frames Mark beginning of BSM packet transmission (change BBId) Request 256 BSM TM frames, sampled at 64Hz Once all packet are received, mark end of BSM packet transmission (reset BBId to 0x80000000) Flush MCU FIFO Display science parameters from last frame Display HK parameters received just after BBId changed to value7 | ? SetTP12SampFreq(0x002C) 91C2002C Wait 4 sec SetTP12SampFreq(0x0000) 91C20000 ? TM(8,4: CA-02) (0x2000) CHOPSENSSIG, CHOPMOTORCURR, CHOPBEMF, JIGGSENSSIG, JIGGMOTORCURR, JIGGBEMF |  |
| 5 | Analyse with QLA Compare science frame with HK Write in a file the result of the comparison | Compare the parameters in the two ‘clock displays on the screen for each type of packet (ENG, SMEC and BSM) Write the results of the comparison to a file using the ‘save’ button on the FUNC-MCU-03 GUI |  |

Success/failure criteria: test passed if values in HK are identical to those in the science frames

Comments/Open issues: Is there a way to specify PRECISELY the number of packet desired?

## FUNC-MCU-04, MCU test pattern test

| ID: | FUNC-MCU-04 |
| --- | --- |
| Purpose | Checking the integrity of MCU test pattern packets |
| Description of test: | Load MCU test pattern and check contents of packets |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode- MCU Powered on (following FUNC-MCU-01) |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-MCU-04 |  |  |
| 2 | Request MCU test pattern frames Request 100 MCU test packet at 100Hz Read frame counter to know the exact number of packet received Flush MCU FIFO Display test pattern with QLA | SetTP15SampFreq(0x001C) 91C5001C Wait 1 sec SetTP15SampFreq(0x0000) 91C50000 MCUFRAMECNT ~100 TM(8,4: CA-02) (0x2000 The content of the test data is TBD |  |
| 3 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-MCU-04 Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |

Success/Failure criteria: test passed is test pattern is properly generated

Comments/Open issues: Is there a way to set PRECISELY the number of frames desired?

## FUNC-SMEC-01, SMEC switch on and initialisation

| ID: | FUNC-SMEC-01 |
| --- | --- |
| Purpose | Switching on and initialising SMEC. SMEC ends up at home position and is ready for following tests |
| Description of test: | Switch on various part of the SMEC system then initialise it. Check parameters |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on – SMEC off |
| Final configuration | SPIRE in REDY mode +MCU on + SMEC on and initialised, at home position |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2-3 min |
| Constraints: | FUNC-MCU-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display |  |  |
| 2 | Switch on fringe encoder Set encoder LED to half power Check SMECStatus Check SencoderSignal1 Check SencoderSignal2 Check SencoderSignal3 | SetSEncoderPwr(0x0003) 90400003 SMECSTAT = 0x0080 SMECENCSIG1=0x7FFF SMECENCSIG2=0xC9E8 SMECENCSIG3=0x3616 |  |
| 3 | Switch on LVDT Power on LVDT Check LVDTAC Check LVDTDC Check LVDT position | SetSLVDTPwr(0x0001) 90410001 SMECLVDTACSIG = 0x1DB7 SMECLVDTDCSIG = 0x1DBA SMECLVDTPOSN = 0x0000 |  |
| 4 | Switch on motors | ??? |  |
| 5 | Initialise SMEC Set the mode to initialisation Check SMECStatus Check SencoderSignal1 Check SencoderSignal2 Check SencoderSignal3 Check LVDTAC Check LVDTDC Check LVDT position Check motor current Check back emf Check encoder increment position Check encoder fine position | SetSTrajMode(0x0004) 90490004 SMECSTAT = 0x0188 SMECENCSIG1=0xD527 SMECENCSIG2=0x5A4F SMECENCSIG3=0x5085 SMECLVDTACSIG = 0x315F SMECLVDTDCSIG = 0x3165 SMECLVDTPOSN = 0x0709 SMECMOTORCURR = 0x03E8 SMECBEMF = 0x03E8 SMECENCPOSN = 0x0708 SMECENCFINEPOS = 0x0000 |  |

Success/Failure Criteria Test passed if SMEC ends up at home position and parameters values reflect this

Comment/Open Issue: Need conversion curve between signals and physical position. Command to switch on motors?

Switch off procedure:

| Action | Command |
| --- | --- |
| Send SMEC to 0 Switch off LVDT Switch off encoder Engage launch latch | SetTrajEndPosition(0x0000) 90450000 SetSTrajMode(0x0001) 90490001 SetSLVDTPwr(0x0000) 90410000 SetSEncoderPwr(0x0000) 90400000 SetSLaunchLatch(0x0001) 90430001 |

## FUNC-SMEC-02, SMEC launch latch check

| ID: | FUNC-SMEC-02 |
| --- | --- |
| Purpose | Checking that the launch latch works properly |
| Description of test: | Open launch latch, move SMEC short distance and check response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 |
| Level | ILT/PLT/IST |
| Test Conditions: | Warm and cold |
| Total duration | 2-3 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 2 | Send SMEC to position for launch latch Move SMEC to position Read SMEC position | SetTrajEndPosition(0x0000) 90450000 SetSTrajMode(0x0001) 90490001 SMECENCPOSN=0x???? SMECENCFINEPOSN=0x???? SMECLVDTACSIG=0x???? SMECLVDTPOSN=0x???? SMECPOSNDELTA=0x???? |  |
| 3 | Test launch latch Engage launch latch Check latch status Try and move SMEC by a short distance (500 um) Make sure SMEC hasn’t moved Disengage launch latch Check latch status | SetSLaunchLatch(0x0001) 90430001 SMECLATCHSTAT=0x0001 SetTrajEndPosition(0x01F4) 904501F4 SetSTrajMode(0x0001) 90490001 Check values on SCOS display SetSLaunchLatch(0x0002) 90430002 SMECLATCHSTAT=0x0002 |  |
| 4 | Check Move SMEC by a short distance (500 um) Read SMEC position Make sure that SMEC moved by the distance requested | SetTrajEndPosition(0x01F4) 904501F4 SetSTrajMode(0x0001) 90490001 SMECENCPOSN=0x???? SMECENCFINEPOSN=0x???? SMECLVDTACSIG=0x???? SMECLVDTPOSN=0x???? SMECPOSNDELTA=0x???? Check values on SCOS display |  |
| 5 | Drive SMEC back to its home position | SetSTrajMode(0x0004) 90490004 |  |

Success/Failure Criteria: Test passed if launch latch is successfully engaged/disengaged

Comments/Open issues:

Is the actuator for the latch redundant?

## FUNC-SMEC-03, SMEC LEDs test

| ID: | FUNC-SMEC-03 |
| --- | --- |
| Purpose | Checking integrity of the LEDs |
| Description of test: | Switch on LEDs and step up voltages |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | none |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display |  |  |
| 2 | Test LEDs Apply power to LEDs by stepping up voltages from 0 to 5V Read photodiode signal for each step | SetSEncoderPwr(0x0000) 90400000 SetSEncoderPwr(0x0001) 90400001 SetSEncoderPwr(0x0002) 90400002 SetSEncoderPwr(0x0003) 90400003 SetSEncoderPwr(0x0004) 90400004 SetSEncoderPwr(0x0005) 90400005 SetSEncoderPwr(0x0006) 90400006 SetSEncoderPwr(0x0007) 90400007 SMECENCSIG1=0x???? SMECENCSIG2=0x???? SMECENCSIG3=0x???? |  |
| 3 | Apply normal power to LEDs (half power?) | SetSEncoderPwr(0x0003) 90400003 |  |

Success/Failure Criteria: Test passed if encoder signal increases as expected, i.e. with the power applied to the LEDs

Comment/open issues:

Are they the correct parameter to check???

I need to specify the values expected

## FUNC-SMEC-04, SMEC position test

| ID: | FUNC-SMEC-04 |
| --- | --- |
| Purpose | Checking that SMEC moves to the position commanded |
| Description of test: | Move to set positions and measure read back position, currents, back emf |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 3 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 2 | Move SMEC Close loop on encoder 0th and 1st order set a position to reach (e.g. +1mm from mechanical stop) Move SMEC to position Read SMEC position Read motor current Read back emf | SetSLoopMode(0x0001) 90440001 SetTrajEndPosition(0x03E8) 904503E8 SetSTrajMode(0x0001) 90490001 SMECENCPOSN=0x03E8 SMECENCFINEPOSN=0x???? SMECLVDTACSIG=0x1DB7 SMECLVDTPOSN=0x03E8 SMECPOSNDELTA=0x0000 SMECMOTORCURR = 0x???? SMECBEMF = 0x???? |  |
| 3 | Repeat for a small set of position (for full functional test only?) 2mm from mechanical stop 3mm from mechanical stop 4mm from mechanical stop 5mm from mechanical stop 6mm from mechanical stop 7mm from mechanical stop 8mm from mechanical stop 9mm from mechanical stop | SetTrajEndPosition(0x07D0) 904507D0 SetTrajEndPosition(0x0BB8) 90450BB8 SetTrajEndPosition(0x0FA0) 90450FA0 SetTrajEndPosition(0x1388) 90451388 SetTrajEndPosition(0x1770) 90451770 SetTrajEndPosition(0x1B58) 90451B58 SetTrajEndPosition(0x1F40) 90451F40 SetTrajEndPosition(0x2328) 90452328 |  |
| 4 | Drive SMEC back to its home position Open loop (DAC to 0) | SetSTrajMode(0x0004) 90490004 SetSLoopMode 90440006 |  |

Success/Failure Criteria: Test passed if SMEC moves to the set position and current and back emf have expected values

Comment/Open issue:

Need conversion curve between signal and physical position. I need to specify some of the values expected

I use conversion curves in Bruce’s document

## FUNC-SMEC-05, SMEC step and look scan test

| ID: | FUNC-SMEC-05 |
| --- | --- |
| Purpose | Characterising SMEC movement for a step-and-look scan |
| Description of test: | Perform a step-and-look scan, measure positions, current, back emf |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | ~18 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start MCU ENG and SMEC packet sampling at 240Hz | SetTP10SampFreq(0x000B) 91C0000B |  |
| 2 | Run QLA script FUNC-SMEC-05 in order to: Display SMEC parameters Record a time series of the parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, SMECMOTORVOLT, (+others???) |  |
| 3 | Move SMEC Close loop on encoder 0th and 1st order Set start position of the scan (e.g. -0.5mm from ZPD) Move to position record parameters for 10 sec For i=1..100{ increment position to reach by 10 um move to position record parameters for 10 sec } | SetSLoopMode(0x0001) 90440001 SetSTrajEndPosition(0x1194) 90451194 SetSTrajMode(0x0001) 90490001 Wait 10 sec pos = 0x1194 + i*0x000A SetSTrajEndPosition(0x(pos)) 9045(pos) SetSTrajMode(0x0001) 90490001 Wait 10 sec |  |
| 4 | Write in a file the parameters recorded with QLA |  |  |
| 5 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 91C00000 |  |
| 6 | Drive SMEC back to its home position Open loop (DAC forced to 0) | SetSTrajMode(0x0004) 90490004 SetSLoopMode 90440006 |  |
| 7 | Offline analysis Characterise SMEC scan | Derive Average value of each parameter for each step and plot average value vs position compare LVDT and encoder signal (plot LVDT vs encoder) consumed power |  |

Success/Failure Criteria Test passed if SMEC does the scan required

Comment/Open issue:

This test, how I’ve written it, is long…Shall we do less steps?

Need conversion curve between signals and physical position.

I use conversion curves in Bruce’s document

## FUNC-SMEC-06, SMEC saw tooth scan test

| ID: | FUNC-SMEC-06 |
| --- | --- |
| Purpose | Characterising SMEC movement during a saw tooth scan |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scans |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-SMEC-06 in order to: Display SMEC parameters Record a time series of the parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECENCPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, SMECMOTORVOLT, (+others???) |  |
| 2 | With SCOS, start MCU ENG and SMEC packet sampling at 240Hz | SetTP10SampFreq(0x000B) 91C0000B |  |
| 3 | Move SMEC Close loop on encoder 0th and 1st order Set a scan speed (e.g. 500 um/s) Set number of scan to 2 Set a scan amplitude (e.g. from -1mm to +1 mm) Run the scan | SetSLoopMode(0x0001) 90440001 SetSScanSpeedForward(0x1388) 90471388 SetSScanRevSpeed(0xFFFF) 9056FFFF SetSScanNumber(0x0002) 90480002 SetStrajStartPosition(0x0FA0) 90460FA0 SetSTrajEndPosition(0x1770) 90451770 SetSTrajMode(0x0002) 90490002 |  |
| 4 | Repeat for different scan amplitude Set scan amplitude from –4 to +4 mm Run the scan | SetStrajStartPosition(0x03E8) 904603E8 SetSTrajEndPosition(0x2328) 90452328 SetSTrajMode(0x0002) 90490002 |  |
| 5 | Repeat for different scan speed Set the speed to 1mm/s Run the scan Set the speed to 1.5mm/s Run the scan Set the speed to 2mm/s Run the scan | SetSScanSpeedForward(0x2710) 90472710 SetSScanRevSpeed(0xFFFF) 9056FFFF SetSTrajMode(0x0002) 90490002 SetSScanSpeedForward(0x3A98) 90473A98 SetSScanRevSpeed(0xFFFF) 9056FFFF SetSTrajMode(0x0002) 90490002 SetSScanSpeedForward(0x4E20) 90474E80 SetSScanRevSpeed(0xFFFF) 9056FFFF SetSTrajMode(0x0002) 90490002 |  |
| 6 | Write in a file the parameters recorded with QLA |  |  |
| 7 | Drive SMEC back to its home position Open loop (DAC forced to 0) | SetSTrajMode(0x0004) 90490004 SetSLoopMode 90440006 |  |
| 8 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 91C00000 |  |
| 9 | Offline analysis Characterise SMEC movement: | Derive: scan length, accuracy of position reached velocity, velocity stability consumed power |  |

Success/Failure Criteria: Test passed if SMEC does the scan required

Comment/Open issue:

Is there something to do to perform a saw tooth scan? Or is it like a triangular scan with the reverse scan speed set to its maximum value?

I use conversion curves in Bruce’s document

## FUNC-SMEC-07, SMEC triangular scan test

| ID: | FUNC-SMEC-07 |
| --- | --- |
| Purpose | Characterising SMEC movement during a triangular scan |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scans |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start MCU ENG and SMEC packet sampling at 240Hz | SetTP10SampFreq(0x000B) 91C0000B |  |
| 2 | Run QLA script FUNC-SMEC-07 in order to: Display SMEC parameters Record a time series of the parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECENCPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, SMECMOTORVOLT, (+others???) |  |
| 3 | Move SMEC Close loop on encoder 0th and 1st order Set a scan speed (e.g. 500 um/s) Set number of scan to 2 Set a scan amplitude (e.g. from -1mm to +1 mm) Run the scan | SetSLoopMode(0x0001) 90440001 SetSScanSpeedForward(0x1388) 90471388 SetSScanRevSpeed(0x1388) 90561388 SetSScanNumber(0x0002) 90480002 SetStrajStartPosition(0x0FA0) 90460FA0 SetSTrajEndPosition(0x1770) 90451770 SetSTrajMode(0x0002) 90490002 |  |
| 4 | Repeat for different scan amplitude Set scan amplitude from –4 to +4 mm Run the scan | SetStrajStartPosition(0x03E8) 904603E8 SetSTrajEndPosition(0x2328) 90452328 SetSTrajMode(0x0002) 90490002 |  |
| 5 | Repeat for different scan speed Set the speed to 1mm/s Run the scan Set the speed to 1.5mm/s Run the scan Set the speed to 2mm/s Run the scan | SetSScanSpeedForward(0x2710) 90472710 SetSScanRevSpeed(0x2710) 90562710 SetSTrajMode(0x0002) 90490002 SetSScanSpeedForward(0x3A98) 90473A98 SetSScanRevSpeed(0x3A98) 90563A98 SetSTrajMode(0x0002) 90490002 SetSScanSpeedForward(0x4E20) 90474E80 SetSScanRevSpeed(0x4E20) 90564E20 SetSTrajMode(0x0002) 90490002 |  |
| 6 | Write in a file the parameters recorded with QLA | Use save button on QLA GUI |  |
| 7 | Drive SMEC back to its home position Open loop (DAC forced to 0) | SetSTrajMode(0x0004) 90490004 SetSLoopMode 90440006 |  |
| 8 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 91C00000 |  |
| 9 | Offline analysis Characterise SMEC movement: | Derive: scan length, accuracy of position reached velocity, velocity stability consumed power |  |

Success/Failure Criteria: Test passed if SMEC does the scan required

Comment/Open issue:

I use conversion curves in Bruce’s document

## FUNC-SMEC-08, SMEC open loop position test

| ID: | FUNC-SMEC-08 |
| --- | --- |
| Purpose | Moving SMEC in open loop mode. Scaling back emf |
| Description of test: | Move SMEC short distance and check response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 |
| Level | ILT/PLT/IST |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Start MCU ENG and SMEC packet sampling at 240Hz | SetTP10SampFreq(0x000B) 91C0000B |  |
| 2 | Move SMEC Open loop Set a position to reach (e.g. 5mm for mechanical stop) Start recording parameters with QLA Move SMEC | SetSLoopMode 90440006 SetTrajEndPosition 90451388 SMECLVDTDCSIG, SMECENCPOSN, SMECMOTORBEMF SetTrajMode 90490001 |  |
| 3 | Scale back emf TBD | SetMotorResistance |  |
| 4 | Move SMEC to a new position Open loop Set a position to reach (e.g. 5mm for mechanical stop) Start recording parameters with QLA Move SMEC | SetSLoopMode 90440006 SetTrajEndPosition 90451388 SMECLVDTDCSIG, SMECENCPOSN, SMECMOTORBEMF SetTrajMode 90490001 |  |
| 5 | Drive SMEC back to its home position | SetSTrajMode(0x0004) 90490004 |  |

## FUNC-SMEC-09, SMEC open loop position test

| ID: | FUNC-SMEC-09 |
| --- | --- |
| Purpose | Scanning in open loop mode (feedback on back emf only) |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scans. Test identical to SMEC-07 but with open loop |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 and SMEC-08 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start MCU ENG and SMEC packet sampling at 240Hz | SetTP10SampFreq(0x000B) 91C0000B |  |
| 2 | Run QLA script FUNC-SMEC-07 in order to: Display SMEC parameters Record a time series of the parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECENCPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, SMECMOTORVOLT, (+others???) |  |
| 3 | Move SMEC Open loop (no sensor) Set a scan speed (e.g. 500 um/s) Set number of scan to 2 Set a scan amplitude (e.g. from -1mm to +1 mm) Run the scan | SetSLoopMode(0x0006) 90440006 SetSScanSpeedForward(0x1388) 90471388 SetSScanRevSpeed(0x1388) 90561388 SetSScanNumber(0x0002) 90480002 SetStrajStartPosition(0x0FA0) 90460FA0 SetSTrajEndPosition(0x1770) 90451770 SetSTrajMode(0x0002) 90490002 |  |
| 4 | Repeat for different scan amplitude Set scan amplitude from –4 to +4 mm Run the scan | SetStrajStartPosition(0x03E8) 904603E8 SetSTrajEndPosition(0x2328) 90452328 SetSTrajMode(0x0002) 90490002 |  |
| 5 | Repeat for different scan speed Set the speed to 1mm/s Run the scan Set the speed to 1.5mm/s Run the scan Set the speed to 2mm/s Run the scan | SetSScanSpeedForward(0x2710) 90472710 SetSScanRevSpeed(0x2710) 90562710 SetSTrajMode(0x0002) 90490002 SetSScanSpeedForward(0x3A98) 90473A98 SetSScanRevSpeed(0x3A98) 90563A98 SetSTrajMode(0x0002) 90490002 SetSScanSpeedForward(0x4E20) 90474E80 SetSScanRevSpeed(0x4E20) 90564E20 SetSTrajMode(0x0002) 90490002 |  |
| 6 | Write in a file the parameters recorded with QLA | Use save button on QLA GUI |  |
| 7 | Drive SMEC back to its home position Open loop (DAC forced to 0) | SetSTrajMode(0x0004) 90490004 SetSLoopMode 90440006 |  |
| 8 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 91C00000 |  |
| 9 | Offline analysis Characterise SMEC movement: | Derive: scan length, accuracy of position reached velocity, velocity stability consumed power |  |

Success/Failure Criteria: Test passed if SMEC does the scan required

Comment/Open issue:

I use conversion curves in Bruce’s document

## FUNC-BSM-01c, BSM power on chop motor and sensors

| ID: | FUNC-BSM-01c |
| --- | --- |
| Purpose | Checking the integrity of BSM motors and sensors on chop axis |
| Description of test: | Switch BSM chop motors on. Check telemetry responses, consumed power |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode +MCU on – BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Switch on BSM chop axis Open loop (no sensor) Power on chop sensors Read chop status Read chop magneto resistive signal Read chop motor current Read chop back emf | SetCSensorPwr(0x0003) 90C00003 CHOPTSTAT CHOPSENSSIG CHOPMOTORCURR CHOPBEMF CHOPDACVAL |  |
| 3 | Evaluate consumed power | Proportional to (CHOPMOTORCURR)^2 |  |
| 4 | Switch off Power off Chop sensors Open loop (DAC forced to 0) | SetCSensorPwr(0x0000) 90C00000 SetChopLoopMode(0x0000) 90C20000 |  |

Success/Failure Criteria: Test passed if BSM sensors are switched on and parameters have the values expected

Comment/Open issue:

Need conversion curve

## FUNC-BSM-01j, BSM power on jiggle motor and sensors

| ID: | FUNC-BSM-01j |
| --- | --- |
| Purpose | Checking the integrity of BSM motors and sensors in jiggle axis |
| Description of test: | Switch BSM jigglle motors on. Check telemetry responses, consumed power |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode +MCU on – BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGG PARAMETERS) |  |  |
| 2 | Switch on BSM jiggle axis Open loop (no sensor) Power on jiggle sensors Read jiggle status Read jiggle magneto resistive signal Read jiggle motor current Read jiggle back emf | SetJigLoopMode(0x0003) 91420003 SetJSensorPwr(0x0001) 91400001 JIGGSTAT JIGGSENSSIG JIGGMOTORCURR JIGGBEMF JIGGDACVAL |  |
| 3 | Evaluate consumed power | Proportional to (JIGGMOTORCURR)^2 |  |
| 4 | Switch off Power off jiggle sensors Open loop (DAC forced to 0) | SetJSensorPwr(0x0000) 91400000 SetJigLoopMode(0x0000) 91420000 |  |

Success/Failure Criteria: Test passed if BSM sensors are switched on and parameters have the values expected

Comment/Open issue:

Need conversion curve

## FUNC-BSM-02c, BSM chop axis in open loop

| ID: | BSM-02c |
| --- | --- |
| Purpose | Observing BSM behaviour on chop axis in open loop. Determining the sensor sensibility and scaling the back emf |
| Description of test: | Perform a step of 1 deg, record sensor signal and plot it with QLA. From motor current and voltage, set the chop motor resistance in order to damp the natural axis oscillation. Set a new target value, record and plot sensor signal. |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: |  |
| Initial configuration: | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-200 QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | MCU-01, BSM-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP Parameter) |  |  |
| 2 | Run QLA script FUNC-BSM-02c in order to: Display BSM parameters Record parameters | CHOPSENSIG, CHOPMOTORCURR, CHOPBEMF |  |
| 3 | Move BSM in open loop Power on chop sensors Set the open loop constant between target position and DAC value = 1/32768 Set slew rate limit Set mode to open loop (no sensor) Start BSM packet telemetry Start recording CHOPSENSIG with QLA Wait 500 ms Perform a step to 20000 (–1 degree) Wait 500ms Plot CHOPSENSIG | SetCSensorPwr(0x0001) 90C00001 SetCPositionScaleFactor 90D60BEB SetCRateLimit 90D103E8 SetChopLoopMode 90C20003 SetTelemetryPacket12Sample 91C20008 Qla operation SetChopTargetPosition 90C34E20 Plot should show oscillations during ~500 ms |  |
| 4 | Scaling the back emf (With QLA, plot V/I (i.e. (CHOPBEMF-32768)/(CHOPMOTORCURR-32768) ) Find max of the plot == R) Set Cmotor resistance with the value R*10000 in hex) Set Cmotor resistance with R=2.0224 ohms Apply some gain on the feedback | (R = (CHOPBEMF-32768)/(CHOPMOTORCURR-32768) SetCmotor resistance 90D3(0xValue)) SetCmotor resistance 90D34F00 SetCmotorbackEmfGain 90D201F4 |  |
| 5 | Move BSM Start recording CHOPSENSIG with QLA Wait 500 ms Perform a step to 46000 (+1 degree) Wait 500ms Perform a step to 20000 (-1 degree) Wait 500 ms Perform a step to 0 Wait 500 ms Plot CHOPSENSIG | SetChopTargetPosition 90C3B3B0 SetChopTargetPosition 90C34E20 SetChopTargetPosition 90C38000 Oscillations should be much smaller and during ~100ms |  |
| 6 | Switch off Power off Chop sensors Open loop (DAC forced to 0) | SetCSensorPwr(0x0000) 90C00000 90C20000 |  |

Success/Failure criteria: test passed if BSM moves as required

Comment/open issues:

## FUNC-BSM-02j, BSM jiggle axis in open loop

| ID: | BSM-02j |
| --- | --- |
| Purpose | Observing BSM behaviour on jiggle axis in open loop. Determining the sensor sensibility and scaling the back emf |
| Description of test: | Perform a step of 1 deg, record sensor signal and plot it with QLA. From motor current and voltage, set the jiggle motor resistance in order to damp the natural axis oscillation. Set a new target value, record and plot sensor signal. |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: |  |
| Initial configuration: | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | Idem + |
| EGSE Configuration: | SCOS-200 QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | MCU-01, BSM-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGGLE Parameter) |  |  |
| 2 | Run QLA script FUNC-BSM-02j in order to: Display BSM parameters Record parameters | JIGGSENSIG, JIGGMOTORCURR, JIGGBEMF |  |
| 3 | Move BSM in open loop Power on jiggle sensor Set the open loop constant between target position and DAC value = 1/32768 Set slew rate limit Set mode to open loop (no sensor) Start BSM packet telemetry Start recording JIGGSENSIG with QLA Wait 500 ms Perform a step to 20000 (–0.25 degree) Wait 500ms Plot JIGGSENSIG | SetJSensorPwr(0x0001) 91400001 SetJPositionScaleFactor 91560BEB SetJRateLimit 915103E8 SetJiggLoopMode 91420003 SetTelemetryPacket12Sample 91C20008 Qla operation SetJiggTargetPosition 91434E20 Plot should show oscillations |  |
| 4 | Scaling the back emf (With QLA, plot V/I (i.e. (JIGGBEMF-32768)/(JIGGMOTORCURR-32768) ) Find max of the plot == R Set Jmotor resistance with the value R*10000 in hex) Set Jmotor resistance with the value R = TBD Apply some gain on the feedback | (R = (JIGGBEMF-32768)/(JIGGMOTORCURR-32768) SetJmotor resistance 9153(0xValue)) SetJMotorResistance TBD SetJmotorbackEmfGain 915201F4 |  |
| 5 | Move BSM Start recording JIGGSENSIG with QLA Wait 500 ms Perform a step to 46000 (+0.25 degree) Wait 500ms Perform a step to 20000 (-0.25 degree) Wait 500 ms Perform a step to 0 Wait 500 ms Plot JIGGSENSIG | SetJiggTargetPosition 9143B3B0 SetJiggTargetPosition 91434E20 SetJiggTargetPosition 91438000 Oscillation should be much smaller |  |
| 6 | Switch off Power off jiggle sensors Open loop (DAC forced to 0) | SetJSensorPwr(0x0000) 91400000 91420000 |  |

Success/Failure criteria: test passed if BSM moves as required

Comment/open issues:

## FUNC-BSM-03c, BSM position test, loop closed on MRS, chop axis

| ID: | FUNC-BSM-03c |
| --- | --- |
| Purpose | Checking that BSM moves to the position commanded and that control loop on MRS works, for chop axis |
| Description of test: | Move BSM to a set position. Check telemetry response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required, QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01, FUNC-BSM-01c, FUNC-BSM-02c successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Run QLA script Display parameter | CHOPSENSIG, CHOPMOTORCURR, CHOPBEMF |  |
| 2 | Set the axis to move independently Send command Check status bit | SetBSMMove(0x0000) 90C60000 CHOPSTAT=1 JIGGSTAT=1 |  |
| 3 | Closed loop settings Power on chop sensor Go to 0 position Set PID proportional gain Set PID derivative gain Set PID integral gain Close the loop on the sensor Decrease slew rate | SetCSensorPwr(0x0001) 90C00001 90C38000 90C84DBA 90C91A19 90CA3342 90C20001 90D10320 |  |
| 4 | Move BSM on chop axis Start BSM packet telemetry Start recording CHOPSENSIG with QLA Wait 100 ms Perform a step to 20000 (–1 degree) Wait 100ms Plot CHOPSENSIG | SetTelemetryPacket12Sample 91C20008 Qla operation SetChopTargetPosition 90C34E20 Plot should show a settling time of about 20 ms |  |
| 5 | Evaluate consumed power | Proportional to (CHOPMOTORCURR)^2 |  |
| 6 | On SCOS display, check position error (should be 0) | CmeanPositionError |  |
| 7 | Switch off Send BSM to “off” position (where current =0?) if not already Power off Chop sensors Open loop (DAC forced to 0) | SetChopTargetPosition(0x0000) 90C38000 SetCSensorPwr(0x0000) 90C00000 90C20000 |  |

Success/Failure Criteria: test passed if BSM moves to the requested position

Comment/Open issue:

## FUNC-BSM-03j, BSM position test, loop closed on MRS, jiggle axis

| ID: | FUNC-BSM-03j |
| --- | --- |
| Purpose | Checking that BSM moves to the position commanded and that control loop on MRS works, for jiggle axis |
| Description of test: | Move BSM to a set position. Check telemetry response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01, FUNC-BSM-01j, FUNC-BSM-02j successful |

Procedure and analysis:

| Step | Action | Command/TM | comment |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGG PARAMETERS) |  |  |
| 2 | Run QLA script Display parameter | JIGGSENSIG, JIGGMOTORCURR, JIGGBEMF |  |
| 3 | Closed loop settings Power on jiggle sensor Go to 0 position Set PID proportional gain Set PID derivative gain Set PID integral gain Close the loop on the sensor Decrease slew rate | SetJSensorPwr(0x0001) 91400001 91438000 91484DBA 91491A19 914A3342 91420001 91510320 |  |
| 4 | Move BSM on chop axis Start BSM packet telemetry Start recording JIGGSENSIG with QLA Wait 100 ms Perform a step to 20000 (–0.25 degree) Wait 100ms Plot JIGGSENSIG | SetTelemetryPacket12Sample 91C20008 Qla operation SetJiggTargetPosition 91534E20 Plot should show a settling time of about 20 ms |  |
| 5 | Evaluate consumed power | Proportional to (JIGGMOTORCURR)^2 |  |
| 6 | On SCOS display, check position error (should be 0) | JmeanPositionError |  |
| 4 | Switch off Send BSM to “off” position (where current =0?) if not already Power off jiggle sensors Open loop (DAC forced to 0) | SetJiggTargetPosition(0x0000) 91438000 SetJSensorPwr(0x0000) 91400000 SetJigLoopMode(0x0000) 91420000 |  |

Success/Failure Criteria: test passed if BSM moves to the requested position

Comment/Open issue:

## FUNC-BSM-04c, BSM chop close loop scan test

| ID: | FUNC-BSM-04c |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on chop axis, in close loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01c, 02c, 03c successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start TM packet sampling for BSM parameters at 64Hz | SetTP12SampFreq(0x002C) 91C2002C |  |
| 2 | Run QLA script FUNC-BSM-03 in order to: Display BSM parameters Record a time series of the parameters | CHOPMOTORCURR, CHOPBEMF, CHOPSENSSIG, JIGGMOTORCURR, JIGGBEMF, JIGGSENSSIG |  |
| 3 | Set the axis to move independently Send command Check status bit | SetBSMMove(0x0000) 90C60000 CHOPSTAT=1 JIGGSTAT=1 |  |
| 4 | Move BSM on chop axis Power on chop sensor Close loop Set the position to reach to -120 arcsec (on-sky) Record parameters for 2 sec For i=1..60 { set OSN for QLA to trigger increment position parameter by +4arcsec (on-sky) move to position record parameters for 2 sec set OSN for QLA to trigger } | SetCSensorPwr(0x0001) 90C00001 90C20001 SetChopTargetPosition(0x2AAB) 90C32AAB Wait 2 sec pos = pos + 0x02D8 SetChopTargetPosition(0x(pos)) 90C3(pos) Wait 2 sec |  |
| 5 | Write in a file the parameters recorded with QLA |  |  |
| 8 | Stop TM packet sampling | SetTP12SampFreq(0x0000) 91C20000 |  |
| 9 | Offline analysis: Characterise BSM movement | Derive for both axis and each position: Consumed power. Plot current vs position Find position where current = 0 |  |
| 7 | Switch off Send BSM to “off” position (where current =0?) if not already Power off Chop sensors Open loop (DAC forced to 0) | SetChopTargetPosition(0x0000) 90C38000 SetCSensorPwr(0x0000) 90C00000 90C20000 |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/Open issue: I use conversion curves in Bruce’s document

## FUNC-BSM-04j, BSM jiggle close loop scan test

| ID: | FUNC-BSM-04j |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on jiggle axis, in close loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01j, 02j, 03j successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start TM packet sampling for BSM parameters at 64Hz | SetTP12SampFreq(0x002C) 91C2002C |  |
| 2 | Run QLA script FUNC-BSM-03j in order to: Display BSM parameters Record a time series of the parameters | CHOPMOTORCURR, CHOPBEMF, CHOPSENSSIG, JIGGMOTORCURR, JIGGBEMF, JIGGSENSSIG |  |
| 6 | Move BSM on jiggle axis Power on jigg sensors Close loop Set the position to reach to -30 arcsec (on-sky) Record parameters for 2 sec For i=1..60 { set OSN for QLA to trigger increment position parameter by +1arcsec (on-sky) move to position record parameters for 2 sec set OSN for QLA to trigger } | SetJSensorPwr(0x0000) 91400001 SetJigLoopMode(0x0000) 91420001 SetJiggTargetPosition(0x6AAB) 90C36AAB Wait 2 sec pos = pos + 0x00B6 SetJiggTargetPosition(0x(pos)) 90C3(pos) Wait 2 sec |  |
| 7 | Write in a file the parameters recorded with QLA |  |  |
| 8 | Stop TM packet sampling | SetTP12SampFreq(0x0000) 91C20000 |  |
| 9 | Offline analysis: Characterise BSM movement | Derive for both axis and each position: Consumed power. Plot current vs position Find position where current = 0 |  |
| 4 | Switch off Send BSM to “off” position (where current =0?) if not already Power off jiggle sensors Open loop (DAC forced to 0) | SetJiggTargetPosition(0x0000) 91438000 SetJSensorPwr(0x0000) 91400000 SetJigLoopMode(0x0000) 91420000 |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/Open issue: I use conversion curves in Bruce’s document

## FUNC-BSM-05c, BSM chop open loop scan test

| ID: | FUNC-BSM-05c |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on chop axis, in open loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode +MCU on + BSM off |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01c, 02c successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start TM packet sampling for BSM parameters at 64Hz | SetTP12SampFreq(0x002C) 91C2002C |  |
| 2 | Run QLA script FUNC-BSM-03 in order to: Display BSM parameters Record a time series of the parameters | CHOPMOTORCURR, CHOPBEMF, CHOPSENSSIG, JIGGMOTORCURR, JIGGBEMF, JIGGSENSSIG |  |
| 3 | Set the axis to move independently Send command Check status bit | SetBSMMove(0x0000) 90C60000 CHOPSTAT=1 JIGGSTAT=1 |  |
| 4 | Move BSM on chop axis Power on chop sensor Set mode to open loop (no sensor) Set the position to reach to -120 arcsec (on-sky) Record parameters for 2 sec For i=1..60 { set OSN for QLA to trigger increment position parameter by +4arcsec (on-sky) move to position record parameters for 2 sec set OSN for QLA to trigger } | SetCSensorPwr(0x0001) 90C00001 90C20003 SetChopTargetPosition(0x2AAB) 90C32AAB Wait 2 sec pos = pos + 0x02D8 SetChopTargetPosition(0x(pos)) 90C3(pos) Wait 2 sec |  |
| 5 | Write in a file the parameters recorded with QLA |  |  |
| 8 | Stop TM packet sampling | SetTP12SampFreq(0x0000) 91C20000 |  |
| 9 | Offline analysis: Characterise BSM movement | Derive for both axis and each position: Consumed power. Plot current vs position Find position where current = 0 |  |
| 7 | Switch off Send BSM to “off” position (where current =0?) if not already Power off Chop sensors Open loop (DAC forced to 0) | SetChopTargetPosition(0x0000) 90C38000 SetCSensorPwr(0x0000) 90C00000 90C20000 |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/Open issue:

I use conversion curves in Bruce’s document

## FUNC-BSM-05j, BSM jiggle open loop scan test

| ID: | FUNC-BSM-05j |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on jiggle axis, in open loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode +MCU on + BSM on |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01j, 02j successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start TM packet sampling for BSM parameters at 64Hz | SetTP12SampFreq(0x002C) 91C2002C |  |
| 2 | Run QLA script FUNC-BSM-03j in order to: Display BSM parameters Record a time series of the parameters | CHOPMOTORCURR, CHOPBEMF, CHOPSENSSIG, JIGGMOTORCURR, JIGGBEMF, JIGGSENSSIG |  |
| 6 | Move BSM on jiggle axis Power on jiggle sensors Set mode to open loop (no sensor) Set the position to reach to -30 arcsec (on-sky) Record parameters for 2 sec For i=1..60 { set OSN for QLA to trigger increment position parameter by +1arcsec (on-sky) move to position record parameters for 2 sec set OSN for QLA to trigger } | SetJSensorPwr(0x0000) 91400001 91420003 SetJiggTargetPosition(0x6AAB) 90C36AAB Wait 2 sec pos = pos + 0x00B6 SetJiggTargetPosition(0x(pos)) 90C3(pos) Wait 2 sec |  |
| 7 | Write in a file the parameters recorded with QLA |  |  |
| 8 | Stop TM packet sampling | SetTP12SampFreq(0x0000) 91C20000 |  |
| 9 | Offline analysis: Characterise BSM movement | Derive for both axis and each position: Consumed power. Plot current vs position Find position where current = 0 |  |
| 4 | Switch off Send BSM to “off” position (where current =0?) if not already Power off jiggle sensors Open loop (DAC forced to 0) | SetJiggTargetPosition(0x0000) 91438000 SetJSensorPwr(0x0000) 91400000 SetJigLoopMode(0x0000) 91420000 |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/Open issue: I use conversion curves in Bruce’s document

## FUNC-BSM-06, BSM operating mode test

| ID: | FUNC-BSM-06 |
| --- | --- |
| Purpose | Characterising BSM movement on chop axis for various position on juggle axis |
| Description of test: | Chop. Vary position on jiggle axis and chop again. Check telemetry response, consumed power |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode +MCU on + BSM on |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 45 min |
| Constraints: | FUNC-MCU-01 and BSM-01c/j, 02c/j successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, start TM packet sampling for BSM parameters at 64Hz | SetTP12SampFreq(0x002C) 91C2002C |  |
| 2 | Run QLA script FUNC-BSM-04 in order to: Display BSM parameters Record a time series of the parameters | CHOPMOTORCURR, CHOPBEMF, CHOPSENSSIG, JIGGMOTORCURR, JIGGBEMF, JIGGSENSSIG |  |
| 3 | Set the axis to move independently Send command Check status bit Power on sensors on both axis Close loop on both axis | SetBSMMove(0x0000) 90C60000 CHOPSTAT=1 JIGGSTAT=1 90C00001 91400001 90C20001 91420001 |  |
| 4 | Set position on Jiggle axis | SetJiggTarget Position(0x8000) 91438000 |  |
| 5 | Chop Set a chop frequency Set a chop amplitude (e.g. 126” on-sky) Set a number of chop throw to do start chopping: for i=1..Nchop { move to 1st position Wait 1/2Fchop move to 2nd position Wait 1/2Fchop } | Fchop = 2Hz POS1 = 0x5333, POS2 =0xACCD (POS2 – POS1 = 126”) Nchop = 20 for i=1..Nchop { SetChopTargetPosition(0x”POS1”) 90C3”POS1” Wait 1/2Fchop SetChopTargetPosition(0x”POS2”) 90C3”POS2” Wait 1/2Fchop } |  |
| 6 | Repeat step 4 for a set of chop frequencies | Fchop = 0.2 Hz Fchop = 0.5 Hz Fchop = 1 Hz |  |
| 7 | Repeat steps 4 for a set of amplitudes | POS1 = 0x6E39, POS2 =0x91C7 (POS2 – POS1 = 50”) POS1 = 0x5333, POS2 =0xA38E (POS2 – POS1 = 100”) POS1 = 0x38E4, POS2 =0xC71C (POS2 – POS1 = 200”) |  |
| 8 | Vary position on jiggle axis and repeat steps 4 to 7 30” 15” -15” -30” | SetJiggTarget Position(0x9555) 91439555 SetJiggTarget Position(0x8AAAA) 91438AAAA SetJiggTarget Position(0x7555) 91437555 SetJiggTarget Position(0x6AAA) 91436AAA |  |
| 9 | Write in a file the parameters recorded with QLA |  |  |
| 10 | Stop TM packet sampling | SetTP12SampFreq(0x0000) 91C20000 |  |
| 11 | Switch off Send BSM to “off” position (where current =0?) if not already Power off sensors Open loop (DAC forced to 0) | SetJiggTargetPosition(0x8000) 91438000 SetChopTargetPosition(0x8000) 90C38000 SetJSensorPwr(0x0000) 91400000 SetChopSensorPwr(0x0000) 90C00000 SetJigLoopMode(0x0000) 91420000 SetChopLoopMode(0x0000) 90C20000 |  |
| 12 | Offline analysis: Characterise BSM movement | plot current Vs position on chop axis for each position on the jiggle axis |  |

Success/Failure Criteria: test passed if BSM does the scans required

Comment/Open Issue:

I use conversion curves in Bruce’s document

This test is a bit long defined like this. Shall I remove some scan?

## FUNC-PCAL-01, PCAL characterisation test

| ID: | FUNC-PCAL-01 |
| --- | --- |
| Purpose | Characterising PCAL |
| Description of test: | Step through each setting and measure telemetry response |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in SPIRE in REDY mode |
| Final Configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 2 min |
| Constraints: | FUNC-SCU-04 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-PCAL-01 in order to: Display PCAL parameters Record parameters at each current setting | PCALCURR, PCALV |  |
| 2 | Start SCU science packet sampling at 80 Hz a) mark beginning of PCAL bblock | a) BBTYPE=0x8300 ,STEP =0x0001 SetFrameConf(0x0000) A0830000 SetFrameCtrl(0x0001) A0820001 |  |
| 2 | Apply current to PCAL a) Set STEP 0x0002 Set current to 1 mA (Read voltage) record during 2 sec then Set STEP 0xffff b) Set STEP 0x0003 Set current to 2.5 mA (Read voltage) record during 2 sec then Set STEP 0xffff c) Set STEP 0x0004 Set current to 4 mA (Read voltage) record during 2 sec then Set STEP 0xffff d) Set STEP 0x0005 Set current to 5.5 mA (Read voltage) record during 2 sec then Set STEP 0xffff e) Set STEP 0x0006 Set current to 7 mA (Read voltage) record during 2 sec then Set STEP 0xffff e) Set STEP 0x0007 Set current to 0 mA (Read voltage) record during 2 sec then Set STEP 0xffff e) Set STEP 0x0008 | a) Set STEP= 0x0002 SetPhCalBias(0x0240) A0C80240 PCALV = 0x135B (252mV) Wait 2 sec then STEP =0xffff b) Set STEP=0x0003 SetPhCalBias(0x0599) A0C80599 PCALV = 0x3066 (630mV) Wait 2 sec then STEP =0xffff c) Set STEP=0x0004 SetPhCalBias(0x08F2) A0C808F2 PCALV = 0x4D70 (1.01V) Wait 2 sec then STEP =0xffff d) Set STEP=0x0005 SetPhCalBias(0x0C4B) A0C80C4B PCALV = 0x6A7A (1.39V) Wait 2 sec then STEP =0xffff e) Set STEP=0x0006 SetPhCalBias(0x0FA5) A0C80FA5 PCALV = 0x8784 (1.76V) Wait 2 sec then STEP =0xffff e) Set STEP=0x0007 SetPhCalBias(0x0000) A0C80000 PCALV = 0x0000 Wait 2 sec then STEP =0xffff e) Set STEP=0x0007 |  |
| 3 | Write in a file parameters recorded |  |  |
| 4 | Stop TM packet sampling | SetFrameCtrl(0x0000) A0820000 |  |
| 5 | Offline analysis Characterise PCAL | Derive for each setting: Voltage stablity Time constant (cooling and warming) Consumed power |  |

##### Success/Failure Criteria: Test passed if each bias current is set as required

Comment/Open issue:

## FUNC-SCAL-01, SCAL characterisation test

| ID: | FUNC-SCAL-01 |
| --- | --- |
| Purpose | Characterising SCAL |
| Description of test: | Set current and record temperatures |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in SPIRE in REDY mode |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 15 min |
| Constraints: | FUNC-SCU-05 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-SCAL-01 in order to: Display SCAL parameters a) Mark beginning od SCAL2 bblock | SCAL2CURR, SCAL2V, SCAL2TEMP, SCAL4CURR, SCAL4V, SCAL4TEMP a) BBTYPE =0x8400 ,STEP=0x0001 |  |
| 2 | Apply current to SCAL2 Set current to 1 mA Set STEP 0x0002 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 2 mA Set STEP 0x0003 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 3 mA Set STEP 0x0004 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 4 mA Set STEP 0x0005 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 5 mA Set STEP 0x0006 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 5.5 mA Set STEP 0x0007 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 0 mA Set STEP 0x0008 (Read voltage) Record during 1 min then set STEP 0xffff | SetSCal2Bias(0x02E2) A0CA02E2 STEP =0x0002 SCAL2V = 0x138E (0.5V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x05BF) A0CA05BF STEP=0x0003 SCAL2V = 0x271E (1V) Wait 1 min then STEP =0xffff SetSCal2Bias(0x089D) A0CA089D STEP=0x0004 SCAL2V = 0x3AAE (1.5V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0B7B) A0CA0B7B STEP=0x0005 SCAL2V = 0x4E3D (2V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0E58) A0CA0E58 STEP =0x0006 SCAL2V = 0x61CD (2.5V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0FC7) A0CA0FC7 STEP =0x0007 SCAL2V = 0x6B95 (2.75V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0000) A0CA0000 STEP 0x0008 SCAL2V = 0x0000 Wait 1 min then STEP 0xffff |  |
| 3 | Mark beginning of SCAL4 bblock Repeat test for SCAL4 with the same step layout as SCAL2 Set current to 1 mA (Read voltage) Record during 1 min Set current to 2 mA (Read voltage) Record during 1 min Set current to 3 mA (Read voltage) Record during 1 min Set current to 4 mA (Read voltage) Record during 1 min Set current to 5 mA (Read voltage) Record during 1 min Set current to 5.5 mA (Read voltage) Record during 1 min Set current to 0 mA (Read voltage) Record during 1 min | BBTYPE=0x8401 ,STEP=0x0001 SetSCal2Bias(0x02E2) A0CC02E2 SCAL2V = 0x138A (0.5V) Wait 1 min SetSCal2Bias(0x05C0) A0CC05C0 SCAL2V = 0x2714 (1V) Wait 1 min SetSCal2Bias(0x089E) A0CC089E SCAL2V = 0x3A9F (1.5V) Wait 1 min SetSCal2Bias(0x0B7B) A0CC0B7B SCAL2V = 0x4E2A (2V) Wait 1 min SetSCal2Bias(0x0E59) A0CC0E59 SCAL2V = 0x61B5 (2.5V) Wait 1 min SetSCal2Bias(0x0FC8) A0CC0FC8 SCAL2V = 0x6B7A (2.75V) Wait 1 min SetSCal2Bias(0x0000) A0CC0000 SCAL2V = 0x0000 Wait 1 min |  |
| 4 | Write in a file parameters recorded |  |  |
| 5 | Offline analysis Characterise SCAL | Derive for each setting: Temperature stablity Time constant (cooling and warming) Consumed power |  |

Success/Failure Criteria: Test passed if each bias current is set as required

Comment/Open issue: 1 min< SCAL time constant. Shall we record data during 10 or 15 min instead of 1?

## FUNC-SCAL-02, SCAL PID test

| ID: | FUNC-SCAL-02 |
| --- | --- |
| Purpose | Verifying that the PID controller works properly |
| Description of test: | Set temperatures and record temperatures |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in SPIRE in REDY mode |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 2 hours |
| Constraints: | FUNC-SCU-05 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA Display SCAL parameters (in HK) Record a time series of the parameters | SCAL2CURR, SCAL2V, SCAL2TEMP, SCAL4CURR, SCAL4V, SCAL4TEMP |  |
| 2 | Apply current to SCAL2 Set current to 2.5 mA Start PID Read voltage, temperature Record during 1 hour Switch off SCAL2 | SetSCal2Bias(0x072E) A0CA072E SCAL2V = 0x30E6 (1.25V) SCAL2TEMP Wait 1 hour SetSCal2Bias(0x0000) A0CC0000 |  |
| 3 | Apply current to SCAL4 Set current to 2.5 mA Start PID Read voltage, temperature Record during 1 hour Switch off SCAL4 | SetSCal4Bias(0x072F) A0CC072F SCAL4V = 0x30DA (1.25V) SCAL4TEMP Wait 1 hour SetSCal4Bias(0x0000) A0CC0000 |  |
| 4 | Write in a file parameters recorded |  |  |
| 5 | Offline analysis Characterise SCAL PID | Derive standard deviation around nominal temperature |  |

Success/Failure Criteria Test passed if bias current is set as required

Comment/Open issue:

Can’t this test and the previous one be merged into a single test?

Do we have to wait until the stable temperature is achieved? I guess if the purpose of this test is indeed to see if the PID controller works well, we have to record temperature for a long time (…1 hr?)

## FUNC-DCU-01, DCU Science Packet generation check

| ID: | FUNC-DCU-01 |
| --- | --- |
| Purpose | Checking the integrity of the DCU science packet generation |
| Description of test: | Request DCU science packet and check that the number of frames generated matches the number commanded |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min (short) – 3 min (full) |
| Constraints: | none |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (DPU and OBS PARAMETER) |  |  |
| 2 | Request nominal DCU photometer science frames Set mode to photometer full array Check frame counter Request 200 frames at 15.3Hz Flush DCU FIFO Check that frame counter has incremented by ~200 Check No Science Packet Events occurred | SetDataMode(0x0000) 843C0000 Read DCUFRAMECNT SetFrameCount(0x00C8) 843D00C8 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) Read DCUFRAMECNT ? |  |
| 3 | Request nominal DCU spectrometer science frames Set mode to spectrometer full array Check frame counter Request 200 frames, sampled at 80Hz Flush DCU FIFO Check that frame counter has incremented by ~200 Check No Science Packet Events occurred | SetDataMode(0x0004) 843C0004 Read DCUFRAMECNT SetFrameCount(0x00C8) 843D0008 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) Read DCUFRAMECNT ? |  |
| 4 | For full functional test, repeat step 3 for: PSW PMW&T/C PLW SLW SSW | SetDataMode(0x0001) 843C0001 SetDataMode(0x0002) 843C0002 SetDataMode(0x0003) 843C0003 SetDataMode(0x0005) 843C0005 SetDataMode(0x0006) 843C0006 |  |

Success/Failure Criteria: Correct number of frames generated for each frame type

Comment/Open issue:

## FUNC-DCU-02, DCU Science data check

| ID: | FUNC-DCU-02 |
| --- | --- |
| Purpose | Checking the integrity of the DCU science packet |
| Description of test: | Request DCU science packet and check some values |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 2 min (short) – 5 min( full) |
| Constraints: | FUNC-DCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-DCU-02 |  |  |
| 2 | Request nominal Photometer science frame a) Mark beginning of data bblobk Request 200 frames, sampled at 15.3Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) | a) BBTYPE =0x8800, STEP=0x0001 SetDataMode(0x0000) 843C0000 SetFrameCount(0x00C8) 843D00C8 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff |  |
| 3 | Analyse with QLA Check reception of packets | Not sure what kind of analysis can be done, but for example: Verify regularity of the reception: time difference between two consecutive frame should be constant, or time of each frame vs frame# should be linear |  |
| 4 | Repeat for spectrometer frames a) Mark beginning of data bblobk Request 200 frames, sampled at 80Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) | a) BBTYPE =0x8801, STEP=0x0001 SetDataMode(0x0004) 843C0004 SetFrameCount(0x00C8) 843D00C8 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff |  |
| 5 | Same analysis with QLA |  |  |
| 6 | For full functional test, repeat step 2 for: PSW BBTYPE =0x8802 PMW&T/C BBTYPE =0x8803 PLW BBTYPE =0x8804 SLW BBTYPE =0x8805 SSW BBTYPE =0x8806 | SetDataMode(0x0001) 843C0001 SetDataMode(0x0002) 843C0002 SetDataMode(0x0003) 843C0003 SetDataMode(0x0005) 843C0005 SetDataMode(0x0006) 843C0006 |  |

Success/Failure Criteria: Test passed if frames are properly generated

Comment/Open issue:

Require different scripts for CQM (not all BDAs) and PFM

## FUNC-DCU-03, DCU Test pattern test

| ID: | FUNC-DCU-03 |
| --- | --- |
| Purpose | Checking the integrity of the DCU test pattern |
| Description of test: | Load test pattern and check content of packets. Perform for each type of packet (Photometer and Spectrometer) |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 2 min (short) – 5 min (full) |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-DCU-03 |  |  |
| 2 | Request DCU photometer test frames a) Mark beginning of data bblobk Set mode to photometer test pattern Request 100 frames at 15.3Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) Display test pattern with QLA | a) BBTYPE= 0x8807 , STEP =0x0001 SetDataMode(0x0008) 843C0008 SetFrameCount(0x00C8) 843D00C8 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff The content is TBD |  |
| 3 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-DCU-03-PHOT-FULL Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |
| 4 | Request DCU spectrometer test frames a) Mark beginning of data bblobk Set mode to spectrometer test pattern Request 200 frames at 80Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) Display test pattern with QLA | a) BBTYPE= 0x8808 , STEP =0x0001 SetDataMode(0x000C) 843C000C SetFrameCount(0x00C8) 843D0008 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff The content is TBD |  |
| 5 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-DCU-03-SPEC-FULL Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |
| 6 | For full functional test, repeat step 2 and 3 for PSW BBTYPE =0x880B PMW&T/C BBTYPE =0x880C PLW BBTYPE =0x880D SLW BBTYPE =0x880E SSW BBTYPE =0x880F | SetDataMode(0x0009) 843C0009 SetDataMode(0x000A) 843C000A SetDataMode(0x000B) 843C000B SetDataMode(0x000D) 843C000D SetDataMode(0x000E) 843C000E |  |

Success/failure criteria: Test passed if test pattern generated is similar to the one generated during previous run of this test

Comment/Open issue:

Require different scripts for CQM (not all BDAs) and PFM

## FUNC-DCU-04P, DCU Photometer LIAs switch on

| ID: | FUNC-DCU-04P |
| --- | --- |
| Purpose | Checking the integrity of the Photometer LIAs |
| Description of test: | Switch on photometer LIAs. Check parameter values |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode – Photometer LIAs cards off |
| Final configuration: | SPIRE in REDY mode + Photometer LIAs cards on |
| EGSE Configuration: | SCOS 2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | none |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (DCU PARAMETERS) |  |  |
| 2 | Switch on photometer LIAs Switch on Photometer LIA cards (command sent to SCU) Check status Read voltages Read temperatures | SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A0870005 if MCU is on PWR_STATUS = 0x01FF LIAP_P5 = 0xB554 LIAP_P9 = 0xE00D LIAP_N9 = 0x1FF0 LIA01TEMP = ? LIA02TEMP = ? LIA03TEMP = ? LIA04TEMP = ? LIA05TEMP = ? LIA06TEMP = ? LIA07TEMP = ? LIA08TEMP = ? LIA09TEMP = ? |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Switch off Photometer LIA cards | SetDRelOnOff(0x0000) A0870000 or SetDRelOnOff(0x0004) A0870004 if MCU is on |

Success/Failure Criteria: Test passed if parameters have expected values

Comment/Open issue: By using QLA, we could also look at the detector output and evaluate noise, which should be very low…should it be another test?

## FUNC-DCU-04S, DCU Spectrometer LIAs switch on

| ID: | FUNC-DCU-04S |
| --- | --- |
| Purpose | Checking the integrity of the spectrometer LIAs |
| Description of test: | Switch on spectrometer LIAs. Check parameter values |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode + spectrometer LIAs cards off |
| Final configuration: | SPIRE in REDY mode – spectrometer LIAs cards on |
| EGSE Configuration: | SCOS 2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | none |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (DCU PARAMETERS) |  |  |
| 2 | Switch on spectrometer LIA Switch on spectrometer LIA cards (command sent to SCU) Check status Read voltages Read temperatures | SetDRelOnOff(0x0002) A0870002 or SetDRelOnOff(0x0006) A0870006 if MCU is on PWR_STATUS = 0x0E00 LIAS_P5 = 0xB554 LIAS_P9 = 0xE00D LIAS_N9 = 0x1FF0 LIA09TEMP = ? LIA10TEMP = ? LIA11TEMP = ? |  |

Note: for QM1, the command to switch on LIA cards is SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A0870005 if MCU is on

##### Switch off procedure

| Action | Command |
| --- | --- |
| Switch of spectrometer LIA cards | SetDRelOnOff(0x0000) A0870000 Or SetDRelOnOff(0x0004) A0870004 if MCU is on |

Success/Failure Criteria: Test passed if parameters have expected values

Comment/Open issue: By using QLA, we could also look at the detector output and evaluate noise, which should be very low…should it be another test?

## FUNC-DCU-05P, DCU Photometer Offset test

| ID: | FUNC-DCU-05P |
| --- | --- |
| Purpose | Characterising the detector output for various values of offset, on the photometer side |
| Description of test: | Change offsets up and down, record output and analyse off line. |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode + photometer LIAs cards on |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | FUNC-DCU-04P successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8816, STEP=0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA script FUNC-DCU-05P Look at detector signals Record detector signals |  |  |
| 3 | Set photometer LIAs offsets Set offset to 0V Record output during 20 sec Set offset to 1V Record output during 20 sec Set offset to 2V Record output during 20 sec Set offset to 3V Record output during 20 sec Set offset to 4V Record output during 20 sec Set offset to 5V Record output during 20 sec Reset offset to 0V | STEP=0x0002 Wait 20 sec then STEP=0xffff STEP=0x0003 Wait 20 sec then STEP=0xffff STEP=0x0004 Wait 20 sec then STEP=0xffff STEP=0x0005 Wait 20 sec then STEP=0xffff STEP=0x0006 Wait 20 sec then STEP=0xffff STEP=0x0007 Wait 20 sec then STEP=0xffff STEP=0x0008 Wait 20 sec then STEP=0xffff |  |
| 4 | Write in a file signal recorded |  |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 6 | Offline analysis | Verify that detector signal follows the offsets Derive noise for each setting? Analyse noise Vs offset value? |  |

##### Success/Failure Criteria: Test passed if the analysis shows relevant results

Comment/Open issue:

Require different scripts for CQM and PFM. Is it the way to test it (same voltage on all channel at the same time)? Or should we do it channel par channel?

## FUNC-DCU-05S, DCU Spectrometer Offset test

| ID: | FUNC-DCU-05S |
| --- | --- |
| Purpose | Characterising the detector output for various values of offset, on the spectrometer side |
| Description of test: | Change offsets up and down, record output and analyse off line. |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | SPIRE in REDY mode + spectrometer LIAs cards on |
| Final configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | FUNC-DCU-04S successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8817, STEP=0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA script FUNC-DCU-05S Look at detector signals Record detector signals |  |  |
| 3 | Set spectrometer LIAs offsets Set offset to 0V Record output during 20 sec Set offset to 1V Record output during 20 sec Set offset to 2V Record output during 20 sec Set offset to 3V Record output during 20 sec Set offset to 4V Record output during 20 sec Set offset to 5V Record output during 20 sec Reset offset to 0V | STEP=0x0002 Wait 20 sec then STEP=0xffff STEP=0x0003 Wait 20 sec then STEP=0xffff STEP=0x0004 Wait 20 sec then STEP=0xffff STEP=0x0005 Wait 20 sec then STEP=0xffff STEP=0x0006 Wait 20 sec then STEP=0xffff STEP=0x0007 Wait 20 sec then STEP=0xffff STEP=0x0008 Wait 20 sec then STEP=0xffff |  |
| 4 | Write in a file signal recorded |  |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 6 | Offline analysis | Verify that detector signal follows the offsets Derive noise for each setting? Analyse noise Vs offset value? |  |

##### Success/Failure Criteria: Test passed if the analysis shows relevant results

Comment/Open issue:

Require different scripts for CQM and PFM. Is it the way to test it (same voltage on all channel at the same time)? Or should we do it channel par channel?

## FUNC-DCU-06P, DCU Photometer JFET heaters

| ID: | FUNC-DCU-06P |
| --- | --- |
| Purpose | Characterising influence of JFET heaters on detectors output, for the photometer |
| Description of test: | Switch on photometer JFET heaters. Record detector signal, derive noise |
| Test Type: | Characterisation |
| Instrument model | STM/CQM |
| Redundancy: | no |
| Initial configuration | SPIRE in REDY mode – Photometer JFET heaters off |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 3 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8818, STEP=0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-06P Look at detector signals Record detector signals Display parameter | PHOTHTRV |  |
| 3 | apply power to photometer JFET heaters Set bias voltage to 0V Record output for 20 sec Set bias voltage to -1V Record output for 20 sec Set bias voltage to -2V Record output for 20 sec Set bias voltage to -3V Record output for 20 sec Set bias voltage to -4V Record output for 20 sec Set bias voltage to -5V Record output for 20 sec Set voltage down to 0 | SetPhotoHeaterBias(0x0000) 84110000, then STEP=0x0002 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0033) 84110033, then STEP =0x0003 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0066) 84110066, then STEP=0x0004 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0099) 84110099, then STEP=0x0005 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x00CC) 841100CC, then STEP=0x0006 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x00FF) 841100FF, then STEP=0x0007 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0000) 84110000, then STEP=0x0008 Wait 20 sec then STEP=0xffff |  |
| 4 | Write in a file the signal recorded |  |  |
| 5 | Offline analysis | Derive noise for each setting Analyse noise Vs voltage |  |

##### Success/Failure Criteria: Test passed is bias voltages are set as required

Comment/Open issue:

Require different scripts for CQM and PFM

A Test Facility thermometer might be checked as well (TBC)

There should be thermometers on the instrument to check as well? But which ones?

## FUNC-DCU-06S, DCU Spectrometer JFET heaters

| ID: | FUNC-DCU-06S |
| --- | --- |
| Purpose | Characterising influence of JFET heaters on detectors output, for the spectrometer |
| Description of test: | Switch on Spectrometer JFET heaters. Record detector signal, derive noise |
| Test Type: | Characterisation |
| Instrument model | STM/CQM |
| Redundancy: | no |
| Initial configuration | SPIRE in REDY – Spectrometer JFET heaters off |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 3 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8819, STEP=0x0001 SetDataMode(0x0000) 843C0004 SetSpectroBiasFreq(0x062) 84390062 SetSpectroSampFreq(0x000C) 8438000C SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-06S Look at detector signals Record detector signals Display parameter | SPECHTRV |  |
| 3 | Apply power to spectrometer JFET heaters Set bias voltage to 0V Record output for 20 sec Set bias voltage to -1V Record output for 20 sec Set bias voltage to -2V Record output for 20 sec Set bias voltage to -3V Record output for 20 sec Set bias voltage to -4V Record output for 20 sec Set bias voltage to -5V Record output for 20 sec Set voltage down to 0 | SetSpectroHeaterBias(0x0000) 84330000, then STEP=0x0002 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0033) 84330033, then STEP =0x0003 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0066) 84330066, then STEP=0x0004 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0099) 84330099, then STEP=0x0005 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x00CC)843300CC then STEP=0x0006 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x00FF)843300FF, then STEP=0x0007 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0000) 84330000, then STEP=0x0008 Wait 20 sec then STEP=0xffff |  |
| 4 | Write in a file the signal recorded |  |  |
| 5 | Offline analysis | Derive noise for each setting Analyse noise Vs voltage |  |

##### Success/Failure Criteria: Test passed is bias voltages are set as required

##### Comment/Open issue:

Require different scripts for CQM and PFM

A Test Facility thermometer might be checked as well (TBC)

There should be thermometers on the instrument to check as well? But which ones?

## FUNC-DCU-07P, DCU Photometer JFET test

| ID: | FUNC-DCU-07P |
| --- | --- |
| Purpose | Characterising detector signal for various photometer JFET drain voltages |
| Description of test: | Switch on Photometer JFETs with Vdd on. Activate auto-offset control. Measure noise. Vary Vss and measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode – Photometer JFETs off |
| Final configuration | SPIRE in REDY mode – Photometer JFETs on and warm |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3-4 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE=0x881C , STEP=0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-07P Look at detector signal Record detector signal Display parameters | PSWJFETSTAT, PMLWJFETSTAT, PSWJFET#V (#=1..6), PMWJFET#V (#=1..4), PLWJFET#V (#=1..2), TCJFETV |  |
| 3 | Switch on photometer JFETs Switch drain voltages to high level Switch on JFET heaters to half power during 1 min Switch off heaters Activate auto-offset control | SetPhSWJfetPwr(0x003F) 8412003F SetPhMLWJfetPwr(0x007F) 8413007F SetPhotoHeaterBias(0x007F) 8411007F Wait 1 min SetPhotoHeaterBias(0x0000) 84110000 SetDataMode(0x0010) 843C0010 Wait xx sec |  |
| 4 | Vary Vss and record signal a) Set STEP 0x0002 Set Vss to 0V on all channel Record signal during 20 sec Set STEP 0x0003 Set Vss to -1V on all channel Record signal during 20 sec Set STEP 0x0004 Set Vss to -2V on all channel Record signal during 20 sec Set STEP =0x0005 Set Vss to -3V on all channel Record signal during 20 sec Set STEP =0x0006 Set Vss to -4V on all channel Record signal during 20 sec Set STEP =0x0007 Set Vss to -5V on all channel Record signal during 20 sec Set STEP =0x0008 | a) Set STEP =0x0002 For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 / For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 / SetPhTCJfetVSS1(0x0000) 8414000 Wait 20 send then STEP =0xffff Set STEP =0x0003 For #=1..6, SetPhSWJfetVSS#(0x0033) 8(404+#)0033 / For #=1..4, SetPhMWJfetVSS#(0x0033) 8(40A+#)0033 For #=1..2, SetPhLWJfetVSS#(0x0033) 8(40E+#)0033 / SetPhTCJfetVSS1(0x0033) 84140033 Wait 20 send then STEP =0xffff Set STEP =0x0004 For #=1..6, SetPhSWJfetVSS#(0x0066) 8(404+#)0066 / For #=1..4, SetPhMWJfetVSS#(0x0066) 8(40A+#)0066 For #=1..2, SetPhLWJfetVSS#(0x0066) 8(40E+#)0066 / SetPhTCJfetVSS1(0x0066) 84140066 Wait 20 send then STEP =0xffff Set STEP =0x0005 For #=1..6, SetPhSWJfetVSS#(0x0099) 8(404+#)0099 / For #=1..4, SetPhMWJfetVSS#(0x0099) 8(40A+#)0099 For #=1..2, SetPhLWJfetVSS#(0x0099) 8(40E+#)0099 / SetPhTCJfetVSS1(0x0099) 84140099 Wait 20 send then STEP =0xffff Set STEP =0x0006 For #=1..6, SetPhSWJfetVSS#(0x00CC) 8(404+#)00CC / For #=1..4, SetPhMWJfetVSS#(0x00CC) 8(40A+#)00CC For #=1..2, SetPhLWJfetVSS#(0x00CC) 8(40E+#)00CC / SetPhTCJfetVSS1(0x00CC) 841400CC Wait 20 send then STEP =0xffff Set STEP =0x0007 For #=1..6, SetPhSWJfetVSS#(0x00FF) 8(404+#)00FF / For #=1..4, SetPhMWJfetVSS#(0x00FF) 8(40A+#)00FF For #=1..2, SetPhLWJfetVSS#(0x00FF) 8(40E+#)00FF / SetPhTCJfetVSS1(0x00FF) 841400FF Wait 20 send then STEP =0xffff Set STEP =0x0008 |  |
| 5 | Set Vss back to its operating value (half level) | For #=1..6, SetPhSWJfetVSS#(0x007F) 8(404+#)007F For #=1..4, SetPhMWJfetVSS#(0x007F) 8(40A+#)007F For #=1..2, SetPhLWJfetVSS#(0x007F) 8(40E+#)007F SetPhTCJfetVSS1(0x007F) 8414007F |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each value of Vss Analyse noise vs Vss |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set Vss to 0V on all channel Switch off drain voltages Switch off LIA cards | For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 / For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 / SetPhTCJfetVSS1(0x0000) 8414000 SetPhSWJfetPwr(0x0000) 84120000 SetPhMLWJfetPwr(0x0000) 84130000 SetDRelOnOff(0x0000) A087000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

##### Success/Failure Criteria: Test passed if each voltage is set as required

Comment/Open issue:

All JFETs are tested together

Refer also to performance test ILT-PERF-DAN

Require different scripts for CQM and PFM

Is that the way to turn on JFETs (drain then source)? Does the order matter?

## FUNC-DCU-07S, DCU Spectrometer JFET test

| ID: | FUNC-DCU-07S |
| --- | --- |
| Purpose | Characterising detector signal for various spectrometer JFET drain voltages |
| Description of test: | Switch on spectrometer JFETS with Vdd on. Activate auto-offset control. Measure noise. Vary Vss and measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode – spectrometer JFETs off |
| Final configuration | SPIRE in REDY mode – spectrometer JFETs on and warm |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3-4 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a) BBTYPE=0x881D, STEP=0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-07S Look at detector signal Record detector signal Display parameters | SPECJFETSTAT, SSWJFET#V (#=1..2), SLWJFETV |  |
| 3 | Switch on spectrometer JFETs Switch drain voltages to high level Switch on JFET heaters to half power during 1 min Switch off heaters Activate auto-offset control | SetSpSLWJfetPwr(0x0007) 84370007 SetPhotoHeaterBias(0x007F) 8433007F Wait 1 min SetPhotoHeaterBias(0x0000) 84330000 SetDataMode(0x0014) 843C0014 Wait xx sec |  |
| 4 | Vary Vss and record signal (Same step setting as previous test) Set Vss to 0V on all channel Record signal during 20 sec Set Vss to -1V on all channel Record signal during 20 sec Set Vss to -2V on all channel Record signal during 20 sec Set Vss to -3V on all channel Record signal during 20 sec Set Vss to -4V on all channel Record signal during 20 sec Set Vss to -5V on all channel Record signal during 20 sec | For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x0033) 8(434+#)0033 SetSpLWJfetVSS(0x0033) 84340033 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x0033) 8(434+#)0066 SetSpLWJfetVSS(0x0033) 84340066 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x0099) 8(434+#)0099 SetSpLWJfetVSS(0x0099) 84340099 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x00CC) 8(434+#)00CC SetSpLWJfetVSS(0x00CC) 843400CC Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x00FF) 8(434+#)00FF SetSpLWJfetVSS(0x00FF) 843400FF Wait 20 sec |  |
| 5 | Write in a file the data recorded |  |  |
| 6 | Set Vss back to its operating value (half level) | For #=1..2, SetSpSWJfetVSS#(0x007F) 8(434+#)007F SetSpLWJfetVSS(0x007F) 8434007F |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each value of Vss Analyse noise vs Vss |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set Vss to 0V on all channel Switch off drain voltages Switch off LIA cards | For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 SetSpSLWJfetPwr(0x0000) 84370000 SetDRelOnOff(0x0000) A087000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

##### Success/Failure Criteria: Test passed if each voltage is set as required

Comment/Open issue:

All JFETs are tested together

Refer also to performance test ILT-PERF-DAN

Require different scripts for CQM and PFM

Is that the way to turn on JFETs (drain then source)? Does the order matter?

## FUNC-DCU-08P_full, DCU Full photometer phase shift test

| ID: | FUNC-DCU-08P_full |
| --- | --- |
| Purpose | Adjusting the demodulation phase |
| Description of test: | Set phase to /2 and 3/2 adjust phase by small amounts to find minimum – set phase back to +d |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in PHOT-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in PHOT-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 5 min |
| Constraints: | Photometer detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometers detectors |  |
| 2 | a) Mark beginning of data bblock Set_BBID, Set _OBS_STEP(0x0001) With SCOS, request bolometer array science packet at 15.3Hz | a) BBTYPE =0x881C , STEP =0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08P_full Look at detector signal Record detector signal Display parameter | PSWPHASE, PMWPHASE, PLWPHASE, TCPHASE |  |
| 4 | Set _OBS_STEP(0x0002) Set phase on photometer channel Set demodulation phase to /2 on all channel Set _OBS_STEP(0x0003) Adjust phase by small amount to find minimum signal (=0) at /2+d then set step to 0xffff Repeat part b until minimum signal is found incrementing step by 1. Set phase back to d Check if signal changes | Set_OBS_STEP(0x0002) SetPhotoDemodSW(0x0040) 841A0040 SetPhotoDemodMW(0x0040) 841B0040 SetPhotoDemodLW(0x0040) 841C0040 SetPhotoDemodTC(0x0040) 841D0040 Set_OBS_STEP(0x0003) Set d parameter on SCOS prompt (between 0 and 360) then set step to 0xffff c) Set step to 3+$i until minimum signal is obtained SetPhotoDemodSW(0x00xx) 841A00xx SetPhotoDemodMW(0x00xx) 841B00xx SetPhotoDemodLW(0x00xx) 841C00xx SetPhotoDemodTC(0x00xx) 841D00xx |  |
| 5 | Set phase to 3/2 and repeat previous part with the same step setting , | SetPhotoDemodSW(0x00BF) 841A00BF SetPhotoDemodMW(0x00BF) 841B00BF SetPhotoDemodLW(0x00BF) 841C00BF SetPhotoDemodTC(0x00BF) 841D00BF |  |
| 6 | Write in file data recorded and value of d |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

Require different scripts for CQM and PFM

## FUNC-DCU-08P_short, DCU Short Photometer phase shift test

| ID: | FUNC-DCU-08P_short |
| --- | --- |
| Purpose | Adjusting the demodulation phase? |
| Description of test: | Set phase to /2+d and verify that signal = 0 |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in PHOT-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in PHOT-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | Photometer detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | comment |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometers detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE =0x881C , STEP =0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08P_short Look at detector signal Record detector signal Display parameter | PSWPHASE, PMWPHASE, PLWPHASE, TCPHASE |  |
| 4 | Set STEP 0x0002 Set demodulation phase to /2+d (d found by FUNC-DCU-08P_full) Verify that signal = 0 Set phase back to d | Set STEP=0x0002 SetPhotoDemodSW(0x00xx) 841A00xx SetPhotoDemodMW(0x00xx) 841B00xx SetPhotoDemodLW(0x00xx) 841C00xx SetPhotoDemodTC(0x00xx) 841D00xx |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

The value of dis found by FUNC-DCU-08P_full

Require different scripts for CQM and PFM

## FUNC-DCU-08S_full, DCU Full spectrometer Phase shift test

| ID: | FUNC-DCU-08S_full |
| --- | --- |
| Purpose | Adjusting the demodulation phase? |
| Description of test: | Set phase to /2 and 3/2 adjust phase by small amounts to find minimum – set phase back to +d |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in SPEC-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in SPEC-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 5 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a)BBTYPE=0x881D,STEP=0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08S_full Look at detector signal Record detector signal Display parameter | SSWPHASE, SLWPHASE |  |
| 4 | Set_OBS_STEP(0x0002) Set phase on spectrometer channel Set demodulation phase to /2 on all channel Set_OBS_STEP(0x0003) Adjust phase by small amount to find minimum signal (=0) at /2+dincrementing step in each change Set phase back to d Check if signal changes | SET_OBS_STEP(0x0002) SetSpectroDemodSW(0x0040) 843A0040 SetSpectroDemodLW(0x0040) 843B0040 SET_OBS_STEP(0x0003) Set d parameter on SCOS prompt (between 0 and 360). Set STEP 3+$i c) Set Phase to d SetSpectroDemodSW(0x00xx) 843A00xx SetSpectroDemodMW(0x00xx) 843B00xx |  |
| 5 | Set phase to 3/2 and repeat part 4 with the same step setting | SetSpectroDemodSW(0x00BF) 843A00BF SetSpectroDemodMW(0x00BF) 843B00BF |  |
| 6 | Write in file data recorded and value of d |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

Require different scripts for CQM and PFM

## FUNC-DCU-08S_short, DCU Short spectrometer Phase shift test

| ID: | FUNC-DCU-08S_short |
| --- | --- |
| Purpose | Adjusting the demodulation phase? |
| Description of test: | Set phase at /2+d and verify that signal = 0 |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in SPEC-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in SPEC-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 80Hz a) Mark the beginning of data bblock | a) BBTYPE=0x881C ,STEP =0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08S_short Look at detector signal Record detector signal Display parameter | SSWPHASE, SLWPHASE |  |
| 4 | Set STEP 0x0002 Set demodulation phase to /2+d (d found by FUNC-DCU-08S_full) Verify that signal = 0 Set phase back to d | b) Set STEP 0x0002 SetSpectroDemodSW(0x00xx) 843A00?xx SetSpectroDemodMW(0x00xx) 843B00xx |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

Require different scripts for CQM and PFM

## FUNC-DCU-09P, DCU Photometer bias frequency test

| ID: | FUNC-DCU-09P |
| --- | --- |
| Purpose | Characterising detector signal for various bias frequencies on photometer side |
| Description of test: | Set up bias frequency. Measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in PHOT-STBY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometers detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz a)Mark beginning of data bblock | a)BBTYPE=0x8820,STEP=0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-09P Look at detector signal Record detector signal Display parameters | PHOTOBIASFREQ |  |
| 4 | Vary bias frequency Set bias frequency at 55Hz Set STEP 0x0002 Execute phase shift full procedure. Record signal during 20 sec then set STEP 0xffff Set bias frequency at 130Hz Set STEP 0x0003 Execute phase shift full procedure. Record signal during 20 sec then set STEP 0xffff Set bias frequency at 190Hz Set STEP 0x0004 Execute phase shift full procedure. Record signal during 20 sec then set STEP 0xffff | SetPhotoBiasFreq(0x01FF) 84190163 STEP=0x0002 Wait 20 sec then STEP=0xffff SetPhotoBiasFreq(0x00C3) 84190096 STEP=0x0003 Wait 20 sec then STEP=0xffff SetPhotoBiasFreq(0x007A) 84190066 STEP=0x0004 Wait 20 sec then STEP=0xffff |  |
| 5 | Set Bias frequency back to its operating value (200Hz) | SetPhotoBiasFreq(0x0062) 84190062 |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each frequency Analyse noise vs frequency |  |

Success/Failure Criteria: Test passed if bias frequency is set to each value as required

Comment/Open issue:

Require different scripts for CQM and PFM

## FUNC-DCU-09S, DCU Spectrometer bias frequency test

| ID: | FUNC-DCU-09S |
| --- | --- |
| Purpose | Characterising detector signal for various bias frequencies on spectrometer side |
| Description of test: | Set up bias frequency. Measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in SPEC-STBY |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a)BBTYPE=0x8821,STEP=0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-09S Look at detector signal Record detector signal Display parameters | SPECBIASFREQ |  |
| 4 | Vary bias frequency with same step setting as the photometer one Set bias frequency at 55Hz Execute phase shift test (FUNC-DCU-08S_ full) Record detector signal during 20 sec Set bias frequency at 130Hz Execute phase shift test (FUNC-DCU-08S_ full) Record detector signal during 20 sec Set bias frequency at 190Hz Execute phase shift test (FUNC-DCU-08S_ full) Record detector signal during 20 sec | SetSpectroBiasFreq(0x01FF) 84390163 Wait 20 sec SetSpectroBiasFreq(0x00C3) 84390096 Wait 20 sec SetSpectroBiasFreq(0x007A) 84390066 Wait 20 sec |  |
| 5 | Set Bias frequency back to its operating value (160Hz) | SetSpectroBiasFreq(0x0062) 8439007A |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each frequency Analyse noise vs frequency |  |

Success/Failure Criteria: Test passed if bias frequency is set to each value as required

Comment/Open issue:

Require different scripts for CQM and PFM

## FUNC-DCU-10P, DCU Photometer bias amplitude test

| ID: | FUNC-DCU-10P |
| --- | --- |
| Purpose | Characterising detector signal for various bias amplitude, on photometer side |
| Description of test: | Set bias amplitude. Measure change in signal level, measure noise, read out offsets |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in PHOT-STBY |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA script FUNC-DCU-10P Look at detector signal Record detector signal Display parameters | PSWBIAS, PLWBIAS, PMWBIAS, TCBIAS |  |
| 4 | Vary bias amplitude a) Mark beginning of bias ampl bblock Set bias amplitude to 0mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 40mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 80mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 120mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 160mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 200m Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff | a) BBID=0x88220001,STEP=0x0001 SetPhotoBiasAmplSW(0x0000) 841A0000; SetPhotoBiasAmplMW(0x0000) 841B0000; SetPhotoBiasAmplLW(0x0000) 841C0000; SetPhotoBiasAmplTC(0x0000) 841D0000 BBID=0x881E0001,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040001,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88220002,STEP=0x0001 SetPhotoBiasAmplSW(0x0033) 841A0033; SetPhotoBiasAmplMW(0x0033) 841B0033; SetPhotoBiasAmplLW(0x0033) 841C0033; SetPhotoBiasAmplTC(0x0033) 841D0033 BBID=0x881E0002,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040002,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88220003,STEP=0x0001 SetPhotoBiasAmplSW(0x0066) 841A0066; SetPhotoBiasAmplMW(0x0066) 841B0066 SetPhotoBiasAmplLW(0x0066) 841C0066; SetPhotoBiasAmplTC(0x0066) 841D0066 BBID=0x881E0003,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040003,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88220004,STEP=0x0001 SetPhotoBiasAmplSW(0x0099) 841A0099; SetPhotoBiasAmplMW(0x0099) 841B0099 SetPhotoBiasAmplLW(0x0099) 841C0099; SetPhotoBiasAmplTC(0x0099) 841D0099 BBID=0x881E0004,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040004,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88220005,STEP=0x0001 SetPhotoBiasAmplSW(0x00CC) 841A00CC; SetPhotoBiasAmplMW(0x00CC) 841B00CC SetPhotoBiasAmplLW(0x00CC) 841C00CC; SetPhotoBiasAmplTC(0x00CC) 841D00CC BBID=0x881E0005,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040005,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88220006,STEP=0x0001 SetPhotoBiasAmplSW(0x00FF) 841A00FF; SetPhotoBiasAmplMW(0x00FF) 841B00FF SetPhotoBiasAmplLW(0x00FF) 841C00FF; SetPhotoBiasAmplTC(0x00FF) 841D00FF BBID=0x881E0006,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040006,STEP=0x0001 Wait 20 sec then STEP=0xffff |  |
| 5 | Set Bias amplitude back to its operating value (60mV) | SetPhotoBiasAmplSW(0x004D) 841A004D SetPhotoBiasAmplMW(0x004D) 841B004D SetPhotoBiasAmplLW(0x004D) 841C004D SetPhotoBiasAmplTC(0x004D) 841D004D |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Evaluate change in signal level and derive noise for each value of bias amplitude Analyse signal level and noise Vs bias amplitude |  |

Success/Failure Criteria: Test passed if bias amplitude is set to each value as required and offset procedure works each time it is called.

Comment/Open issue:

How can we check that the offset procedure has done the job properly?

Require different scripts for CQM and PFM

What is the set offset procedure?

## FUNC-DCU-10S, DCU Spectrometer bias amplitude test

| ID: | FUNC-DCU-10S |
| --- | --- |
| Purpose | Characterising detector signal for various bias amplitude on spectrometer side |
| Description of test: | Set bias amplitude. Measure change in signal level, measure noise, read out offsets |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in SPEC-STBY |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz Same layout as previous test but with Spectrometer Bias Amplitud building block type 0x8823 And spec data bblock 0x8801 if Specfull , or 0x8805 if SSW or 0x8806 if SLW | SetDataMode(0x0004) 843C0004 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA script FUNC-DCU-10S Look at detector signal Record detector signal Display parameters | SSWBIAS, SLWBIAS |  |
| 4 | Vary bias amplitude Set bias amplitude to 0mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 40mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 80mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 120mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 160mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 200mV Run set offset procedure Record detector signal during 20 sec | SetSpectroBiasAmplSW(0x0000) 843A0000 SetSpectroBiasAmplLW(0x0000) 843B0000 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x0033) 843A0033 SetSpectroBiasAmplLW(0x0033) 843B0033 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x0066) 843A0066 SetSpectroBiasAmplLW(0x0066) 843B0066 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x0099) 843A0099 SetSpectroBiasAmplLW(0x0099) 843B0099 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x00CC) 843A00CC SetSpectroBiasAmplLW(0x00CC) 843B00CC SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x00FF) 843A00FF SetSpectroBiasAmplLW(0x00FF) 843B00FF SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec |  |
| 5 | Set Bias amplitude back to its operating value (60mV) | SetSpectroBiasAmplSW(0x004D) 843A004D SetSpectroBiasAmplLW(0x004D) 843B004D |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Evaluate change in signal level and derive noise for each value of bias amplitude Analyse signal level and noise Vs bias amplitude |  |

Success/Failure Criteria: Test passed if bias amplitude is set to each value as required and offset procedure works each time it is called.

Comment/Open issue:

How can we check that the offset procedure has done the job properly?

Require different scripts for CQM and PFM

What is the set offset procedure?

## FUNC-DCU-11P, DCU Photometer detector switch on

| ID: | FUNC-DCU-11P |
| --- | --- |
| Purpose | Switching on the all chain for reading detectors signal. This procedure must be performed before test FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10 – see flowchart p17 |
| Description of test: | Switch on photometer LIAs and JFETs, set sampling frequency, bias amplitude, bias frequency, demodulation phase to their operating values |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | detectors off |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | none |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Switch on photometer LIAs card Check HK parameters Switch on JFET source voltages to half-level Switch drain voltages to high level Switch on JFET heaters to half power during 1 min Switch off JFET heaters Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set photo bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A087005 if MCU is on PWR_STATUS = 0x01FF LIAP_P5 = 0xB554 LIAP_P9 = 0xE00D LIAP_N9 = 0x1FF0 For #=1..6, SetPhSWJfetVSS#(0x007F) 8(404+#)007F For #=1..4, SetPhMWJfetVSS#(0x007F) 8(40A+#)007F For #=1..2, SetPhLWJfetVSS#(0x007F) 8(40E+#)007F SetPhTCJfetVSS1(0x007F) 8414007F SetPhSWJfetPwr(0x003F) 8412003F SetPhMLWJfetPwr(0x007F) 8413007F SetPhotoHeaterBias(0x007F) 8411007F Wait 1 min SetPhotoHeaterBias(0x0000) 84110000 SetPhotoBiasFreq(0x0062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetPhotoBiasMode(0x00FF) 840000FF SetPhotoBiasAmplSW(0x004D) 841A004D SetPhotoBiasAmplMW(0x004D) 841B004D SetPhotoBiasAmplLW(0x004D) 841C004D SetPhotoBiasAmplTC(0x004D) 841D004D SetPhotoDemodSW(0x007F) 841A007F SetPhotoDemodMW(0x007F) 841B007F SetPhotoDemodLW(0x007F) 841C007F SetPhotoDemodTC(0x007F) 841D007F SetDataMode(0x0010) 843C0010 Wait 10 sec |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set photo bias mode to “off” Set Vss to 0V on all channel Switch off drain voltages Switch of LIA card | SetPhotoBiasMode(0x0000) 84000000 For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 / For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 / SetPhTCJfetVSS1(0x0000) 8414000 SetPhSWJfetPwr(0x0000) 84120000 SetPhMLWJfetPwr(0x0000) 84130000 SetDRelOnOff(0x0000) A0870000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)

## FUNC-DCU-11S, DCU Spectrometer detector switch on

| ID: | FUNC-DCU-11S |
| --- | --- |
| Purpose | Switching on the all chain for reading detectors signal. This procedure must be performed before test FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10 – see flowchart p17 |
| Description of test: | Switch on spectrometer LIAs and JFETs, set sampling frequency, bias amplitude, bias frequency, demodulation phase to their operating values |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | detectors off |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | none |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Switch on spectrometer LIAs card Check HK parameters Switch on JFET source voltages to half-level Switch drain voltages to high level Switch on JFET heaters to half power during 1 min Switch off JFET heaters Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set spectro bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetDRelOnOff(0x0002) A0870002 or SetDRelOnOff(0x0006) A087006 if MCU is on PWR_STATUS = 0x01FF LIAP_P5 = 0xB554 LIAP_P9 = 0xE00D LIAP_N9 = 0x1FF0 For #=1..2, SetSpSWJfetVSS#(0x007F) 8(434+#)007F SetSpLWJfetVSS(0x007F) 8434007F SetSpSLWJfetPwr(0x0007) 84370007 SetSpectroHeaterBias(0x007F) 8433007F Wait 1 min SetSpectroHeaterBias(0x0000) 84330000 SetSpectroBiasFreq(0x0062) 84390062 SetSpectroSampFreq(0x000C) 8438000C SetSpectroBiasMode(0x00FF) 843000FF SetSpectroBiasAmplSW(0x004D) 843A004D SetSpectroBiasAmplLW(0x004D) 843B004D SetSpectroDemodSW(0x007F) 843A007F SetSpectroDemodMW(0x007F) 843B007F SetDataMode(0x0014) 843C0014 Wait 10 sec |  |

- Note: for QM1 the command to switch on LIA cards is SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A087005 if MCU is on

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set photo bias mode to “off” Set Vss to 0V on all channel Switch off drain voltages Switch of LIA card | SetSpectroBiasMode(0x0000) 84300000 For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 SetSpSLWJfetPwr(0x0000) 84370000 SetDRelOnOff(0x0000) A0870000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)

## FUNC-DCU-12P, DCU Photometer detector settings

| ID: | FUNC-DCU-12P |
| --- | --- |
| Purpose | Setting sampling frequency, bias frequency, bias amplitude and demodulation phase when LIA and JFET are already on. This procedure must be performed after FUNC-DCU-04 and FUNC-DCU-07 before carrying on with tests FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10. – see flowchart p17 |
| Description of test: | Set sampling frequency, bias amplitude, bias frequency, demodulation phase on photometer side to their operating values. Activate auto-offset procedure |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | LIA, JFET on. |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | FUNC-DCU-04 and FUNC-DCU-07 successfully passd |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set photo bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetPhotoBiasFreq(0x0062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetPhotoBiasMode(0x00FF) 840000FF SetPhotoBiasAmplSW(0x004D) 8401004D SetPhotoBiasAmplMW(0x004D) 8402004D SetPhotoBiasAmplLW(0x004D) 8403004D SetPhotoBiasAmplTC(0x004D) 8404004D SetPhotoDemodSW(0x007F) 841A007F SetPhotoDemodMW(0x007F) 841B007F SetPhotoDemodLW(0x007F) 841C007F SetPhotoDemodTC(0x007F) 841D007F SetDataMode(0x0010) 843C0010 Wait xx sec |  |

Switch off procedure: see test FUNC-DCU-11P

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)

## FUNC-DCU-12S, DCU Detector settings

| ID: | FUNC-DCU-12S |
| --- | --- |
| Purpose | Setting sampling frequency, bias frequency, bias amplitude and demodulation phase when LIA and JFET are already on. This procedure must be performed after FUNC-DCU-04 and FUNC-DCU-07 before carrying on with tests FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10. – see flowchart p17 |
| Description of test: | Set sampling frequency, bias amplitude, bias frequency, demodulation phase on spectrometer side to their operating values. Activate auto-offset procedure |
| Test Type: | integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | no |
| Initial configuration: | LIA, JFET on. |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | FUNC-DCU-04 and FUNC-DCU-07 successfully passd |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set spectro bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetSpectroBiasFreq(0x0062) 84390062 SetSpectroSampFreq(0x000C) 8438000C SetSpectroBiasMode(0x00FF) 843000FF SetSpectroBiasAmplSW(0x004D) 843A004D SetSpectroBiasAmplLW(0x004D) 843B004D SetSpectroDemodSW(0x007F) 843A007F SetSpectroDemodMW(0x007F) 843B007F SetDataMode(0x0014) 843C0014 Wait xx sec |  |

Switch off procedure: see test FUNC-DCU-11S

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)