| SUBJECT: | SPIRE Functional test specificationsSPIRE Functional test specifications |
| --- | --- |
| PREPARED BY: | Samuel RonayetteSamuel Ronayette (U of L/RAL) Asier Abreu (U of L/RAL) |
| DOCUMENT No: | SPIRE-RAL-DOC-001652SPIRE-RAL-DOC-001652 |
| ISSUE: | Issue 1.2Issue 1.4 | Date: 20/06/2005 |
| CHECKED BY: | Asier Abreu (U o L/ RAL) | Date: 15/06/2005 |
| APPROVED BY: |  | Date: |

Distribution

| RAL | Dave Smith |
| --- | --- |
|  | Ken King |
|  | Bruce Swinyard |
|  | Tanya Lim |
|  | Sunil Sidher |
|  | Eric Sawyer |
|  | Doug Griffin |
|  | Eric |
| UofCardiff | Tim Waskett |
|  | Peter Hargrave |
| Imperial College | Davide Rizzo |

Change Record

| ISSUE | DATE | Changes |
| --- | --- | --- |
| 0.1 | 08th April 2003 | 1st Draft for comment |
| 0.2 | 28th April 2003 | Updated Draft incorporating comments |
| 0.3 | 23rd May 2003 | test FUNC-MCU-05 changed into test FUNC-SMEC-02 10 DCU test specification added |
| 0.3_1 | 2nd June 2003 | Page Headers Corrected |
| 0.4 | 20th June 2003 | test FUNC-DCU-06 updated |
| 0.5 | 25th July 2003 | LEDs test added (FUNC-SMEC-04) DCU tests modified, new numbering |
| 0.6 |  | Characterisation tests described in more details FUNC-SMEC-01 includes initialisation and SMEC initialisation removed Definition of a switch on detector “test” Test condition summary added Test flow added |
| 0.7 | 2nd Sept. 2003 | New format Command names added Values of parameters added |
| 1.0 (draft 1) | 22/09/2003 | BSM-04 modified FUNC-SCAL-02 modified FUNC-SCU-07 modified FUNC-DCU-04 to FUNC-DCU-11 split in photometer and spectrometer tests FUNC-DCU-08 split in short and full test “switch off” procedure added in each test concerned Use of the command SetDRelOnOff corrected |
| 1.0 (draft 2) | 05/12/2003 | New definition of BSM tests (open loop) |
| 1.0 | 23/12/2003 | New BSM test (BSM-05c and j) Other BSM tests revised Spectro sampling frequency corrected (in DCU tests) Flow chart updated SMEC launch latch test modified Other SMEC test revised (cf use of open loop command) |
| 1.1 | ? | ? |
| 1.2 | 05/07/2004 06/08/2004 | Correct current values for SCU-07 Cooler Heater Check DCU Frequency test: new frequencies BBID and STEPS added (Asier) Several updates to the functional tests specification. Functional tests affected : SCU-FUNC-03 comments added (see comments) SCU-FUNC-07 command specification corrected DCU-FUNC-04 comments added DCU-FUNC-05 to 08(P/S) step setting corrected |
| 1.3 | 11/11/04 20/12/04 | MCU, SMEC and BSM tests updated according to the new MCU IDC (V4.0) BBIDs and STEPs added to MCU,SMEC and BSM functionsl tests SCU-04 and SCU-05 (PCAL and SCAL): current parameter modified |
| 1.4 | 16/05/05-15/06/05 | Changes to the BSM functional test sequence. Changes to the commanding sequence definition |

Table of Contents

1.	Introduction	6

1.1	Scope	6

1.2	Structure of Document	6

1.3	Documents	6

2.	Test Summaries	7

2.1	SCU tests	7

2.2	MCU tests	7

2.3	SMEC tests	8

2.4	BSM tests	9

2.5	PCAL test	9

2.6	SCAL tests	9

2.7	DCU tests	9

3.	Summary of test condition and test flow	11

3.1	Summary of test condition	11

3.2	Test flow	12

4.	Test Specifications	19

4.1	FUNC-SCU-01, SCU science packet generation check	19

4.2	FUNC-SCU-02, SCU Science data check	21

4.3	FUNC-SCU-03, SCU DC thermometry check	23

4.4	FUNC-SCU-04, SCU PCAL check	25

4.5	FUNC-SCU-05, SCU SCAL check	27

4.6	FUNC-SCU-06, SCU AC thermometry check	29

4.7	FUNC-SCU-07, SCU cooler heaters check	31

4.8	FUNC-SCU-08, SCU Test pattern test	33

4.9	FUNC-MCU-01, MCU power on	35

4.10	FUNC-MCU-02, MCU Science Packet Generation Test	38

4.11	FUNC-MCU-03, MCU Science data check	42

4.12	FUNC-MCU-04, MCU test pattern test	45

4.13	FUNC-SMEC-01, SMEC switch on and initialisation	47

4.14	FUNC-SMEC-02, SMEC launch latch check – not to perform during PFM1	51

4.15	FUNC-SMEC-03, SMEC LEDs test	53

4.16	FUNC-SMEC-04, SMEC position test	55

4.17	FUNC-SMEC-05, SMEC step and look scan test	58

4.18	FUNC-SMEC-06, SMEC saw tooth scan test	61

4.19	FUNC-SMEC-07, SMEC triangular scan test	65

4.20	FUNC-SMEC-08, SMEC open loop position test	69

4.21	FUNC-SMEC-09, SMEC open loop scan test	72

4.22	FUNC-BSM-01c, BSM power on chop, perform small step	76

4.23	FUNC-BSM-01j, BSM power on jiggle, perform small step	78

4.24	FUNC-BSM-02c, BSM chop axis in open loop	80

4.25	FUNC-BSM-02j, BSM jiggle axis in open loop	83

4.26	FUNC-BSM-03c, BSM position test with closed loop, chop axis	86

4.27	FUNC-BSM-03j, BSM position test with closed loop, jiggle axis	89

4.28	FUNC-BSM-04c, BSM scan with closed loop, chop axis	92

4.29	FUNC-BSM-04j, BSM scan with closed loop, jiggle axis	94

4.30	FUNC-BSM-05c, BSM scan with open loop, chop axis	97

4.31	FUNC-BSM-05j, BSM scan with open loop, jiggle axis	100

4.32	FUNC-BSM-06, BSM operating mode test	103

4.33	FUNC-PCAL-01, PCAL characterisation test	107

4.34	FUNC-SCAL-01, SCAL characterisation test	110

4.35	FUNC-SCAL-02, SCAL PID test	113

4.36	FUNC-DCU-01, DCU Science Packet generation check	115

4.37	FUNC-DCU-02, DCU Science data check	117

4.38	FUNC-DCU-03, DCU Test pattern test	120

4.39	FUNC-DCU-04P, DCU Photometer LIAs switch on	123

4.40	FUNC-DCU-04S, DCU Spectrometer LIAs switch on	125

4.41	FUNC-DCU-05P, DCU Photometer Offset test	127

4.42	FUNC-DCU-05S, DCU Spectrometer Offset test	131

4.43	FUNC-DCU-06P, DCU Photometer JFET heaters	134

4.44	FUNC-DCU-06S, DCU Spectrometer JFET heaters	137

4.45	FUNC-DCU-07P, DCU Photometer JFET test	140

4.46	FUNC-DCU-07S, DCU Spectrometer JFET test	144

4.47	FUNC-DCU-08P_full, DCU Full photometer phase shift test	148

4.48	FUNC-DCU-08P_short, DCU Short Photometer phase shift test	151

4.49	FUNC-DCU-08S_full, DCU Full spectrometer Phase shift test	154

4.50	FUNC-DCU-08S_short, DCU Short spectrometer Phase shift test	157

4.51	FUNC-DCU-09P, DCU Photometer bias frequency test	160

4.52	FUNC-DCU-09S, DCU Spectrometer bias frequency test	162

4.53	FUNC-DCU-10P, DCU Photometer bias amplitude test	164

4.54	FUNC-DCU-10S, DCU Spectrometer bias amplitude test	168

4.55	FUNC-DCU-11P, DCU Photometer detector switch on	172

4.56	FUNC-DCU-11S, DCU Spectrometer detector switch on	175

4.57	FUNC-DCU-12P, DCU Photometer detector settings	178

4.58	FUNC-DCU-12S, DCU Detector settings	180

Figures

Test flow	12

Tables

Summary of test condition	11

Glossary

| AD | Applicable Document |
| --- | --- |
| AIV | Assembly Integration and Verification |
| AVM | Avionics Model |
| BSM | Beam Steering Mechanism |
| CQM | Cryogenic Qualification Model |
| DCU | Detector Control Unit |
| DPU | Digital Processing Unit |
| DRCU | Detector Readout and Control Unit |
| EGSE | Electrical Ground Support Equipment |
| FLT | Flight Level Test |
| FTS | Fourier Transform Spectrometer |
| HOB | Herschel Optical Bench |
| ILT | Instrument Level Test |
| IST | Integrated System Test |
| MCU | Mechanism Control Unit |
| MGSE | Mechanical Ground Support Equipment |
| N/A | Not Applicable |
| PCAL | Photometer Calibrator source |
| PFM | Proto-Flight Model |
| PLT | Payload Level Test |
| RAL | Rutherford Appleton Laboratory |
| RD | Reference Document |
| SCAL | Spectrometer Calibrator source |
| SCU | Subsystem Control Unit |
| SMEC | Spectrometer Mechanism |
| SPIRE | Spectral and Photometric Imaging Receiver |
| STM | Structural Test Model |
| TBC | To Be Confirmed |
| TBD | To Be Defined |
| TFCS | Test Facility Control System |
| AD | Applicable Document |
| AIV | Assembly Integration and Verification |
| AVM | Avionics Model |

# Introduction

## Scope

This document contains the functional test specification for SPIRE. These are tests of different subsystems (SCU, MCU, SMEC, BSM, PCAL, SCAL, DCU) that will be performed on different model of the instrument (AVM, CQM, PFM, FS). They consist of integrity check and characterisation tests. These tests are not calibration or performance tests.

## Structure of Document

A summary of each test is given first (name, short description) and in the following section, we give the specification of each test. They are grouped together per subsystem (respectively SCU, MCU, SMEC, BSM, PCAL, SCAL, and DCU)

## Documents

Applicable Documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| AD 1 | SPIRE Data ICD | K.J. King | SPIRE-RAL-PRJ-001078, issue 1.0 | 15/01/2003 |
| AD2 | MCU/DCU Command list ICD an user manual | Didier Ferrand | LAM/ELE/SPI/011011, issue 3.0 | 15/01/2003 |
| AD3 | DRCU/DPU ICD | Pinsard, Ferrand, Mur | Sap-SPIRE-Cca-076-02, issue 1.0 | 14/02/2003 |

Reference Documents

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| RD 1 | Subsystem reaction for specification for the instrument simulator | Bruce Swinyard | SPIRE-RAL-NOT-001715, issue 1.0 | 11/06/2003 |
| RD 2 | Operating modes for the SPIRE instrument | Bruce Swinyard, Matt Griffin | SPIRE-RAL-DOC-000320, issue 3.0 | 04/01/2002 |
| RD 3 | MCU/BSM DM Mechanism, Integration and test report | Didier Ferrand | LAM/ELE/SPI/031020, issue 1.0 | 20/10/2003 |
| RD4 | Instrument Requirements Document | Bruce Swinyard | SPIRE-RAL-PRJ-000034 Issue 1.1 | 10/01/2002 |

# Test Summaries

## SCU tests

### FUNC-SCU-01, SCU Science packet generation check

To check the integrity of SCU science packet interface and that the SCU generates the nominal science frame data. Request SCU science frame and check that all the frames are received correctly by the DPU

### FUNC-SCU-02, SCU Science data check

Checking the integrity of the SCU science data. Request science frames and housekeeping data and compare content

### FUNC-SCU-03, SCU DC thermometry check

To check the integrity of thermometers. Switch on thermometers all at once and check values

### FUNC-SCU-04, SCU PCAL check

Checking the integrity of PCAL. Switch on PCAL, set a current and check voltage and current, then switch off

### FUNC-SCU-05, SCU SCAL check

Checking the integrity of SCAL. Switch on SCAL, set a current and check voltage, current and temperature changes, then switch off

### FUNC-SCU-06, SCU AC thermometry check

Checking the integrity of the AC thermometer. Switch on AC thermometers and check values

### FUNC-SCU-07, SCU cooler heater check

Checking the integrity of the cooler heaters. Switch on heaters, check voltage

### FUNC-SCU-08, SCU Test pattern test

Checking the integrity of SCU test pattern packet. Load test pattern and check content of packets

## MCU tests

### FUNC-MCU-01, MCU power on

To power on the MCU into a state ready to execute SMEC or BSM commands.

This procedure also tests the low speed interface to the MCU, used for command and housekeeping data transfer. Apply power to MCU from SCU and boot the MCU DSP ROM software, checking voltages and status

### FUNC-MCU-02, MCU Science packet generation check

To check the integrity of MCU science packet interface and that the MCU generates all the nominal science data frames. Generate science data from the SMEC and BSM at typical data rates and check that all the frames are received correctly by the DPU.

### FUNC-MCU-03, MCU Science data check

Checking the integrity of MCU science data. Compare science data with HK. Perform test for each type of packet. (SMEC, BSM, ENG)

### FUNC-MCU-04, MCU test pattern test

Checking the integrity of MCU test pattern packets. Load MCU test pattern and check contents of packets.

## SMEC tests

### FUNC-SMEC-01, SMEC switch on and initialisation

Switching on and initialising SMEC. SMEC ends up at home position and is ready for following tests. Switch on various part of the SMEC system then initialise it. Check parameters

### FUNC-SMEC-02, SMEC launch latch check

Checking that the launch latch works properly. Open launch latch, move SMEC short distance and check response

### FUNC-SMEC-03, SMEC LEDs test

Checking integrity of the LEDs. Switch on LEDs and step up voltages

### FUNC-SMEC-04, SMEC position test

Checking that SMEC moves to the position commanded. Move to set positions and measure read back position, currents, back emf

### FUNC-SMEC-05, SMEC multiple position test

Characterising SMEC movement for a step-and-look scan. Perform a step-and-look scan, measure positions, current, back emf

### FUNC-SMEC-06, SMEC saw tooth scan test

Characterising SMEC movement during a saw tooth scan. Scan over range and measure positions, current, back-emf – repeat for different scans

### FUNC-SMEC-07, SMEC triangular scan test

Characterising SMEC movement during a triangular scan. Scan over range and measure positions, current, back-emf – repeat for different scans

### FUNC-SMEC-08, SMEC open loop position test

Characterisation test. Move SMEC to a position in open loop. Scale BEMF. Move to a new position

### FUNC-SMEC-09, SMEC open loop scan test

Identical to FUNC-SMEC-07 but with control lop open

## BSM tests

### FUNC-BSM-01c/j, BSM power on motor and sensor

Checking the integrity of BSM motors and sensors. Switch BSM motors on. Check telemetry responses, consumed power. Perform a small step

### FUNC-BSM-02c/j, position test, open loop

Observing the behaviour of the BSM in open loop. Determining sensor sensibility and scaling the back emf

### FUNC-BSM-03c/j, BSM raster test, open loop

### FUNC-BSM-04c/j, BSM position test, close loop

### FUNC-BSM-05c/j, BSM scan test, close loop

### FUNC-BSM-06, BSM operating mode test

Characterising BSM movement on chop axis for various positions on jiggle axis. Set BSM to each mode (Chop, jiggle…). Check telemetry response, consumed power

## PCAL test

### FUNC-PCAL-01, PCAL characterisation test

Characterising PCAL. Go through each setting and measure telemetry response

## SCAL tests

### FUNC-SCAL-01, SCAL characterisation test

Characterising SCAL. Set current and record temperatures

### FUNC-SCAL-02, SCAL PID test

Verifying that the PID controller works properly. Set temperatures and record temperatures

## DCU tests

### FUNC-DCU-01, DCU Science Packet generation check

Checking the integrity of the DCU science packet generation. Request DCU science packet and check that the number of frames generated matches the number commanded

### FUNC-DCU-02, DCU Science data check

Checking the integrity of the DCU science packet. Request DCU science packet and check some values

### FUNC-DCU-03, DCU Test pattern test

Checking the integrity of the DCU test pattern. Load test pattern and check content of packets. Perform for each type of packet (Photometer and Spectrometer)

### FUNC-DCU-04P/S, DCU LIAs switch on

Checking the integrity of the LIAs. Switch on LIAs. Check parameter values

### FUNC-DCU-05P/S, DCU Offset test

Characterising the detector output for various values of offset. Change offsets up and down, record output and analyse off line

### FUNC-DCU-06P/S, DCU JFET heaters

Characterising influence of JFET heaters on detectors output. Switch on JFET heaters. Record detector signal, derive noise

### FUNC-DCU-07P/S, DCU JFET test

Characterising detector signal for various JFET drain voltages. Switch on JFETS with Vdd on. Activate auto-offset control. Measure noise. Vary Vss and measure noise

### FUNC-DCU-08P/S_full, DCU Phase shift test

Adjusting the demodulation phase. Adjust phase by small amounts until maximum signal is found

### FUNC-DCU-08P/S_short, DCU Phase shift test

Shorter version of the previous

### FUNC-DCU-09P/S, DCU Bias frequency test

Characterising detector signal for various bias frequencies. Set up bias frequency. Measure noise

### FUNC-DCU-10P/S, DCU Bias amplitude test

Characterising detector signal for various bias amplitudes. Set bias amplitude. Measure change in signal level, measure noise, read out offsets

### FUNC-DCU-11P/S, DCU detectors switch on

Switching on the all chain for reading detectors signal. This procedure must be performed before test FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10 – see flowchart p17

### FUNC-DCU-12P/S, DCU detectors settings

Setting sampling frequency, bias frequency, bias amplitude and demodulation phase when LIA and JFET are already on. This procedure must be performed after FUNC-DCU-04 and FUNC-DCU-07 before carrying on with tests FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10. – see flowchart p17

# Summary of test condition and test flow

In this section is table that sum up the condition in which each test is performed, i.e. instrument configuration, if the test is performed during short functional test, during warm functional test. It is followed by the test flow

## Summary of test condition

S: performed during Short functional test

F: performed during Full functional test

W: performed during Warm functional test

C: performed during Cold functional test

*: test that have a short and a full version (refer to procedure for details)

!: different script according to instrument model (refer  to procedure for details)

?: TBD or TBC

| TEST | TEST CONDITION | Estimated dur. |
| --- | --- | --- |
| FUNC-SCU-01 | S F W C |  |
| FUNC-SCU-02 | F W C |  |
| FUNC-SCU-03 | S F W C |  |
| FUNC-SCU-04 | S F W C |  |
| FUNC-SCU-05 | S F W C |  |
| FUNC-SCU-06 | S F W C |  |
| FUNC-SCU-07 | S F W C |  |
| FUNC-SCU-08 | F W C |  |
| FUNC-MCU-01 | S F W C |  |
| FUNC-MCU-02 | S F W C |  |
| FUNC-MCU-03 | F W C |  |
| FUNC-MCU-04 | F W C |  |
| FUNC-SMEC-01 | S F W C |  |
| FUNC-SMEC-02 | S F W C |  |
| FUNC-SMEC-03 | S F W C |  |
| FUNC-SMEC-04 | S F W C |  |
| FUNC-SMEC-05 | F C |  |
| FUNC-SMEC-06 | F C |  |
| FUNC-SMEC-07 | F C |  |
| FUNC-SMEC-08 | F C |  |
| FUNC-SMEC-09 | F C |  |
| FUNC-BSM-01 | S F W C |  |
| FUNC-BSM-02 | F C |  |
| FUNC-BSM-03 | F C |  |
| FUNC-BSM-04 | F C |  |
| FUNC-BSM-05 | F C |  |
| FUNC-BSM-06 | F C |  |
| FUNC-PCAL-01 | F W C |  |
| FUNC-SCAL-01 | F W? C |  |
| FUNC-SCAL-02 | F W? C |  |
| FUNC-DCU-01 | S F W C * |  |
| FUNC-DCU-02 | F W C * |  |
| FUNC-DCU-03 | F W C * |  |
| FUNC-DCU-04P/S | S F W C |  |
| FUNC-DCU-05P/S | F C ! |  |
| FUNC-DCU-06P/S | F C ! |  |
| FUNC-DCU-07P/S | F C ! |  |
| FUNC-DCU-08P/S_full | F C ! |  |
| FUNC-DCU-08P/S_short | F C ! |  |
| FUNC-DCU-09P/S | F C ! |  |
| FUNC-DCU-10P/S | F C ! |  |
| FUNC-DCU-11P/S | F C ! |  |
| FUNC-DCU-12P/S |  |  |

##### Total duration:

##### Comment/open issues

-There might be other tests that deserve a * (different short and full), TBD

## Test flow

Warm short functional test: perform test in yellow boxes only

Warm full functional test: perform test in yellow and pink boxes only

Cold short functional test: perform test in yellow and blue boxes only

Cold full functional test: perform all tests

Test into brackets: cannot be performed with the CQM

Comment/open issues:

-Maybe FUNC-SCU-06 (AC thermometry) should be repeated when the detectors are at 300 mK?

# Test Specifications

## FUNC-SCU-01, SCU science packet generation check

| ID: | FUNC-SCU-01 |
| --- | --- |
| Purpose | To check the integrity of SCU science packet interface and that the SCU generates the nominal science frame data |
| Description of test: | Request SCU science frame and check that all the frames are received correctly by the DPU |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration: | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | < 1 min |
| Constraints: | None |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
|  | Select SCOS display (DPU and OBS PARAMETER) |  |  |
|  | Request nominal SCU science frames Check frame counter Set OBS STEP 1 Request 31 SCU TM frames, sampled at 80Hz Flush SCU FIFO Set OBS STEP(0xffff) | Read SCUFRAMECNT Set_OBS_STEP(0x0001) SetFrameConf(0x0000) A0830000 SetSeqLength(0x001F) A084001F SetFrameCtrl(0x0001) A0820001 TM(8,4: CA-02) (0x4000) Set_OBS_STEP(0xffff) |  |
|  | Check generation of packets Check that frame counter has incremented by ~31 Check No Science Packet Events occurred | Read SCUFRAMECNT ? |  |

Success/Failure Criteria: Test passed if correct number of frame generated

Comments/Open issues:

## FUNC-SCU-02, SCU Science data check

| ID: | FUNC-SCU-02 |
| --- | --- |
| Purpose | Checking the integrity of the SCU science data |
| Description of test: | Request science frames and housekeeping data and compare content |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial Configuration: | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 2 min |
| Constraints: | SCU thermometry needs to be switched on – see test FUNC-SCU-03 and FUNC-SCU-06 FUNC-SCU-01 successfully passed |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
|  | Run QLA script FUNC-SCU-02 |  |  |
|  | With SCOS, request nominal SCU science frames Mark beginning of test ( BBId+Step) Request 31 SCUTM frames, sampled at 80Hz Flush SCU FIFO Mark end of data stream (step to 0xffff) Once all packet are received, mark end of test (reset BBId to 0x8000000) | BBTYPE =0x8A00 ,STEP=0x0001 SetFrameConf(0x0000) A0830000 SetSeqLength(0x001F) A084001F SetFrameCtrl(0x0001) A0820001 TM(8,4: CA-02) (0x4000) STEP=0xffff BBTYPE =0x8000 |  |
|  | Display parameters with QLA Display the last science packet received Display the HK parameters received just after BBId changed to 0x8000000 | TCHeaterVolt, PhCalCur, SCal4Cur, SCal2Cur, PhCalVolt, SCal4Volt, SCal2Volt, CPHPtemp, CPHStemp, CEHStemp, CSHTtemp, SOBtemp, SL0temp, PL0temp, SUBtemp, BAFtemp, BSMStemp, SCL2temp, SCL4temp, SCSTtemp, FTSStemp, FTSMtemp, BSMMtemp, CEVTemp |  |
|  | Analyse with QLA Compare that the values of the SCU science frame are similar to those in the SCU HK. Write in log file the result of the comparison | Check values on the ‘clock’ displays on the screen |  |

Success/Failure Criteria: Test passed if values in the science frame are identical to those in the HK nominal report.

Comments/Open issues:

## FUNC-SCU-03, SCU DC thermometry check

| ID: | FUNC-SCU-03 |
| --- | --- |
| Purpose | To check the integrity of thermometers |
| Description of test: | Switch on thermometers all at once and check values |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode - DC thermometers off |
| Final configuration | SPIRE in REDY mode - DC thermometers on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
|  | Select SCOS display (SCU parameters) |  |  |
|  | Switch On thermometers Switch on SCU DC Thermometers, all at once Wait 5 seconds (time for sensor to activate) | TempOnOff(0xFFFF) A085FFFF Wait 5 seconds |  |
|  | Check parameters values | SCUTEMPSTAT =0xffff CPHPtemp= CPHStemp= CEHStemp= CSHTtemp= SOBtemp = SL0temp = PL0temp = SUBtemp= BAFtemp = BSMStemp= SCAL2TEMP= SCAL4TEMP = SCSTtemp = FTSStemp= FTSMtemp= BSMMtemp= | Readings will vary depending on FPU temperature. |

Switch off procedure:

| Action | Command |
| --- | --- |
| Switch off thermometers | TempOnOff(0x0000) A0850000 |

Success/Failure Criteria: Test passed if all thermometer channels are reading expected values.

Comments/Open issues: Have to check with Dave what would be the indication for a sensor that is not working properly

A not connected thermometer will read a very small raw value. A non working thermometer will read a wrong value?.

## FUNC-SCU-04, SCU PCAL check

| ID: | FUNC-SCU-04 |
| --- | --- |
| Purpose | Checking the integrity of PCAL |
| Description of test: | Switch on PCAL, set a current and check voltage and current, then switch off |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | < 1 min |
| Constraints: | None |

Procedure and analysis:

| Steps | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU parameters) |  |  |
| 2 | Switch On PCAL Set a small current (0.1 mA) Wait 2 seconds | SetPhCalBias(0x003D) A0C8003D Wait 2 seconds |  |
| 3 | Check parameters values | PCALCURR = 0x0191 (0.1 mA) PCALV = 0x01EE (25.2 mV) |  |
| 4 | Switch off PCAL | SetPhCalbias(0x0000) A0C80000 |  |

Success/Failure Criteria: Test passed if value of current read = current set and voltage has value expected

Comments/Open issues:

This test needs a calibration curve. Voltage expected is dependent on instrument model and version of MIB

## FUNC-SCU-05, SCU SCAL check

| ID: | FUNC-SCU-05 |
| --- | --- |
| Purpose | Checking the integrity of SCAL |
| Description of test: | Switch on SCAL, set a current and check voltage, current and temperature changes, then switch off |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 3 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU parameters) |  |  |
| 2 | Switch on SCAL-4 Set a small current (0.1 mA) Wait 10 sec | SetSCal4Bias(0x004D) A0CC004D Wait 10 sec |  |
| 3 | Check parameters values for SCAL-4 | SCAL4CURR = 0x0232 (0.1mA) SCAL4V = 0x01F3 (50mV) SCAL4TEMP = 0x |  |
| 4 | Switch off SCAL-4 Set Scal4Bias to 0 | a) SetScal4Bias(0x0000) A0CC0000 |  |
| 5 | Switch on SCAL-2 Set a small current (0.1 mA) Wait 10 sec | SetSCal2Bias(0x004D) A0CA004D Wait 10 sec |  |
| 6 | Check parameters values for SCAL-2 | SCAL2CURR = 0x0233 (0.1mA) SCAL2V = 0x0x01F3 (50mV) SCAL2TEMP = 0x |  |
|  | Switch off SCAL-2 Set Scal2Bias to 0 | a) SetScal2Bias(0x0000) A0CA0000 |  |

Success/Failure Criteria Test passed if value of current read = current set and voltage and temperature have values expected

Comments/Open issues:

This test needs a calibration curve.

Time constant is long and may be we do not wait enough for the stable temperature to be achieved, but it is not the purpose of this test. We just make sure that the wires are OK (shall we wait even less than 1 min then?)

## FUNC-SCU-06, SCU AC thermometry check

| ID: | FUNC-SCU-06 |
| --- | --- |
| Purpose | Checking the integrity of the AC thermometer |
| Description of test: | Switch on AC thermometers and check values |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode – AC thermometers off |
| Final configuration | SPIRE in REDY mode – AC thermometers on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | < 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU PARAMETERS) |  |  |
| 2 | Switch on thermometer Switch on AC thermometer Wait 5 seconds | SubKOnOff(0x0001) A0860001 Wait 5 seconds |  |
| 3 | Check parameter value | CEVtemp (SUBKTEMP) |  |

Switch off procedure:

| Action | Command |
| --- | --- |
| Switch off thermometer | SubKOnOff(0x0000) A0860000 |

Success/Failure Criteria: Test passed if thermometer channel is reading expected value

Comments/Open issues:

Temperature is dependent on cooler condition. If SPIRE is in REDY mode, temperature expected is > 4K If in STAND-BY mode, temperature expected is 300 mK

## FUNC-SCU-07, SCU cooler heaters check

| ID: | FUNC-SCU-07 |
| --- | --- |
| Purpose | Checking the integrity of the cooler heaters |
| Description of test: | Switch on heaters, check voltage |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | FUNC-SCU-06 should be completed successfully |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (SCU PARAMETERS) |  |  |
| 2 | Heaters switch on Apply current to SPHS heater Wait 45 sec Swich it off Apply current to EVHS heater Wait 45 sec Switch it off Apply current to SP heater Wait 45 sec Switch it off | SetSPHSHeatB(0x0800) A0C40800 Wait 45 sec SetSPHSHeatB(0x0000) A0C40000 SetEVHSHeatB(0x0800) A0C50800 Wait 45 sec SetEVHSHeatB(0x0000) A0C50000 SetSPHeaterB(0x0708) A0C70708 Wait 45 sec SetSPHeaterB(0x0000) A0C70000 |  |
| 3 | Check parameter values after each voltage setting | SPHSHeatVolt = 0x31CC (324.38 mV) EVHSHeatVolt = 0x31CC (324.38 mV) SPHeaterVolt = 0x388D (8.828 V) |  |

Success/Failure Criteria: Test passed if voltages are read and have the value expected

Comments/Open issues:

Waiting 1 minute is not enough for the voltage to stabilize but it is not the purpose of this test. It is long enough for the voltage to move. We just check that we can measure something

## FUNC-SCU-08, SCU Test pattern test

| ID: | FUNC-SCU-08 |
| --- | --- |
| Purpose | Checking the integrity of SCU test pattern packet |
| Description of test: | Load test pattern and check content of packets |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode |
| Final Configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | warm and cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-SCU-08 |  |  |
| 2 | With SCOS, request nominal SCU science frames Mark beginning of test (change BBId) Request 31 SCUTEST frames, sampled at 80Hz Flush SCU FIFO Mark end of data stream (step to 0xffff) Once all packet are received, mark end of test (reset BBId to 0x8000000) Display contents of last test packet | BBTYPE =0x8A07 ,STEP=0x0001 SetFrameConf(0x0000) A0838000 SetSeqLength(0x001F) A084001F SetFrameCtrl(0x0001) A0820001 TM(8,4: CA-02) (0x4000) STEP=0xffff BBTYPE =0x8000 |  |
| 3 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-SCU-08 Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |

Success/Failure criteria: Test passed if test pattern generated is similar to the one generated during previous run of this test.

Comments/Open issues:

## FUNC-MCU-01, MCU power on

| ID: | FUNC-MCU-01 |
| --- | --- |
| Purpose | To power on the MCU into a state ready to execute SMEC or BSM commands. This procedure also tests the low speed interface to the MCU, used for command and housekeeping data transfer |
| Description of test: | Apply power to MCU from SCU and boot the MCU DSP ROM software, checking voltages and status |
| Test Type: | Configuration |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode, MCU off |
| Final configuration | SPIRE in REDY mode, MCU on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (MCU_PARAMETERS) |  |  |
| 1 | Mark beginning of building block | SET_BBID(0x89000001) |  |
| 2 | Power on MCU from SCU | If LIAs are OFF SetDRelOnOff(0x0004) A0870004 If LIAs are ON SetDRelOnOff(0x0005) A087005 | If nominal configuration of the instrument is followed then LIA will be switched OFF later than the MCU |
| 3 | Check MCU power status Check parameters values | SCUDCDCSTAT = 4 (or 5) MCU+5V = ~ + 5V MCU+15V = ~ +15V FUNC-MCU-15V = ~ -15V MCU+13V = ~ +13V FUNC-MCU-13V = ~-13V MCUMACTEMP = 0x1000 MCUSMECTEMP = 0x1000 MCUBSMTEMP = 0x1000 | All these vales should be recorded in the Functional Test Procedure document . |
| 4 | Check threshold flag |  |  |
| 5 | Reset MCU Subsystem Reset on Wait 5 seconds Reset off Wait 5 seconds Reset on Wait 5 seconds Reset off Wait 5 seconds Check Boot Status Check Interface Status | SetMCUCmdIfCtrl(0x0005) 90010005 Wait 5 seconds SetMCUCmdIfCtrl(0x0007) 90010007 Wait 5 seconds SetMCUCmdIfCtrl(0x0005) 90010003 Wait 5 seconds SetMCUCmdIfCtrl(0x0005) 90010007 Wait 5 seconds MCUBOOTSTAT = 0x0000 MCUIFSTAT = 0x0000 | MCUIFSTAT parameter may NOT equal to zero after the interface reset process, but this does not mean that the process was not successful. |
| 6 | Boot MCU Download Code to RAM Wait 5 seconds Check Boot Status Start Code in RAM | SetDownloadConfig(0xC000) 9021C000 Wait 5 seconds MCUBOOTSTAT = 0x0001 SetBootRam(0x0001) 90240001 |  |
| 7 | Check MCU status | MCUBOOTSTAT = 0 MCUERR = 0x0000 MCUSCHEDCNTLSW = TBD MCUSCHEDCNTMSW = TBD | MCUBOOTSTAT will ONLY have value 1 DURING boot. When boot is finished MCUBOOTSTAT=0 |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Switch off MCU from SCU | SetDRelOnOff(0x0000) A0870000 |

Success/Failure Criteria Procedure completed with no errors

Comment/Open Issue:

How does this change for testing the redundant MCU?

If DSP is not booted correctly, then should the OBS cope with this?

## FUNC-MCU-02, MCU Science Packet Generation Test

| ID: | FUNC-MCU-02 |
| --- | --- |
| Purpose | To check the integrity of MCU science packet interface and that the MCU generates all the nominal science data frames |
| Description of test: | Generate science data from the SMEC and BSM at typical data rates and check that all the frames are received correctly by the DPU. |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode -MCU Powered on (following FUNC-MCU-01) |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 3 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display(s) (DPU PARAMETERS, OBS PARAMETERS) |  |  |
| 2 | Request ENG Frames Check Frame Counter Mark beginning of building block Set step 1 Set ENG TM Packet Sampling to 64.1Hz Set no SMEC, no BSM, no TEST to be sent Set number of frames to be sent to N=0xffff (continuous generation) Start sending ENG TM Packet Wait 10 sec Stop Packet generation Flush MCU FIFO Set step 0xffff Check Frame counter has incremented by ~640 Check No Science Packet Events occurred | Read MCUFRAMECNT SET_BBID(0x89010001) SET_OBS_STEP(0x0001) SetTP14SampFreq(0x002C) 91C40027 91C00000, 91C2000, 91C50000 SetFrameNumber(0xFFFF) 91C3FFFF SetFrameStart(0x0001) 91C10001 Wait 10 seconds SetFrameStart(0x0000) 91C10000 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) Read MCUFRAMECNT ???? |  |
| 3 | Request SMEC Frames Check Frame Counter Mark beginning of building block Set step 1 Set SMEC TM Packet Sampling to 250 Hz Set no ENG, no BSM, no TEST to be sent Set number of frames to be sent to N=0xFFFF (continuous generation) Start sending SMEC TM Packet Wait 10 seconds Stop Packet generation Flush MCU FIFO Set step 0xffff Check Frame counter has incremented by ~2500 Check No Science Packet Events occurred | Read MCUFRAMECNT SET_BBID(0x89020001) SET_OBS_STEP(0x0001) SetTP10SampFreq(0x002C) 91C0000B 91C40000, 91C2000, 91C50000 SetFrameNumber(0xFFFF) 91C3FFFF SetFrameStart (0x0001) 91C10001 Wait seconds SetFrameStart(0x0000) 91C10000 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) Read MCUFRAMECNT ???? |  |
| 4 | Request BSM Frames Check Frame Counter Mark beginning of building block Set step to 1 Set BSM TM Packet Sampling to 64.1 Hz Set no SMEC, no ENG, no TEST to be sent Set number of frames to be sent to N=0xFFFF (continuous generation) Start sending BSM TM Packet Wait Stop Packet generation Flush MCU FIFO Set step to 0xffff Check Frame counter has incremented by ~640 Check No Science Packet Events occurred | Read MCUFRAMECNT SET_BBID(0x89030001) SET_OBS_STEP(0x0001) SetTP12SampFreq(0x002C) 91C20039 91C00000, 91C40000, 91C50000 SetFrameNumber(0xFFFF) 91C3FFFF SetFrameStart (0x0001) 91C10001 Wait 10 seconds SetFrameStart(0x0000) 91C10000 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) Read MCUFRAMECNT ???? |  |
| 5 | Request SMEC+BSM Frames Check Frame Counter Mark beginning of building block Set step to 1 Set BSM TM Packet Sampling to 50 Hz Set SMEC TM Packet Sampling to 250 Hz Set no ENG, no TEST to be sent Set number of frames to be sent to N=0xFFFF (continuous generation) Start sending TM Packet Wait 10 seconds Stop Packet generation Flush MCU FIFO Set step to 0xffff Check Frame counter has incremented by ~ 3000 Check No Science Packet Events occurred | Read MCUFRAMECNT SET_BBID(0x89040001) SET_OBS_STEP(0x0001) SetTP12SampFreq(0x0032) 91C20032 SetTP10SampFreq(0x000A) 91C0000A 91C40000, 91C50000 SetFrameNumber(0xFFFF) 91C3FFFF SetFrameStart (0x0001) 91C10001 Wait 10 seconds SetFrameStart(0x0000) 91C10000 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) Read MCUFRAMECNT ???? |  |

Success/Failure Criteria: Correct number of frames generated for each frame type

Comment/Open Issue:

## FUNC-MCU-03, MCU Science data check

| ID: | FUNC-MCU-03 |
| --- | --- |
| Purpose of test | Checking the integrity of MCU science data |
| Description of test: | Compare science data with HK. Perform test for each type of packet. (SMEC, BSM, ENG) |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode- MCU Powered on (following FUNC-MCU-01) |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm and cold |
| Total duration | 8 min |
| Constraints: | FUNC-MCU-01 and FUNC-MCU-02 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-MCU-03 |  |  |
| 2 | Request ENG Frames Mark beginning of ENG building block Set step to 1 Request ENG TM as in step 2 of test MCU-02 Flush MCU FIFO Set step 0xffff Display science parameters from last frame Display HK parameters received just after BBId changed to value3 | SET_BBID(0x89010001) SET_OBS_STEP(0x0001) see step 2 of MCU-02 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) SMECENCSIG1, 2 and 3, SMECLVDTACSIG, SMECLVDTDCSIG, SMECMOTORCURR, CHOPSENSSIG, CHOPMOTORCURR, JIGGSENSSIG, JIGGMOTORCURR |  |
| 3 | Request SMEC science Frames Mark beginning of SMEC building block Set step to 1 Request SMEC TM frames, as in step3 of test MCU-02 Flush MCU FIFO Set step to 0xffff Display science parameters from last frame Display HK parameters received just after BBId changed to value5 | SET_BBID(0x89020001) SET_OBS_STEP(0x0001) see step 3 of MCU-02 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) SMECENCPOSN, SMECENCFINEPOSN, SMECLVDTPOSN, SMECMOTORBEMF |  |
| 4 | Request BSM Frames Mark beginning of BSM building block Set step to 1 Request SMEC TM frames, as in step3 of test MCU-02 Flush MCU FIFO Set step 0xffff Display science parameters from last frame Display HK parameters received just after BBId changed to value7 | SET_BBID(0x89030001) SET_OBS_STEP(0x0001) see step 4 of test MCU-02 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) CHOPSENSSIG, CHOPMOTORCURR, JIGGSENSSIG, JIGGMOTORCURR |  |
| 5 | Analyse with QLA Compare science frame with HK Write in a file the result of the comparison | Compare the parameters in the two ‘clock displays on the screen for each type of packet (ENG, SMEC and BSM) Write the results of the comparison to a file using the ‘save’ button on the FUNC-MCU-03 GUI |  |

Success/failure criteria: test passed if values in HK are identical to those in the science frames

Comments/Open issues: Jigg, Chop, Smec motor voltages are in science frame but not in HK. BEMFs are in HK but not in science frame, except for SMEC

## FUNC-MCU-04, MCU test pattern test

| ID: | FUNC-MCU-04 |
| --- | --- |
| Purpose | Checking the integrity of MCU test pattern packets |
| Description of test: | Load MCU test pattern and check contents of packets |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode- MCU Powered on (following FUNC-MCU-01) |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-MCU-04 |  |  |
| 2 | Request TEST Frames Check Frame Counter Mark beginning of TEST frames building block Set step to 1 Set TEST TM Packet Sampling to 64.1 Hz Set no SMEC, no BSM, no ENG to be sent Set number of frames to be sent to N=100 Start sending TEST TM Packet Wait Stop Packet generation Flush MCU FIFO Set step 0xffff Check Frame counter has incremented by ~100 Display test pattern with QLA | Read MCUFRAMECNT SET_BBID(0x89050001) SET_OBS_STEP(0x0001) SetTP15SampFreq(0x002C) 91C50039 91C00000, 91C20000, 91C40000 SetFrameNumber(0x0064) 91C30064 SetFrameStart (0x0001) 91C10001 Wait N/frequency SetFrameStart(0x0000) 91C10000 TM(8,4: CA-02) (0x2000) SET_OBS_STEP(0xffff) Read MCUFRAMECNT The content of the test data is TBD |  |
| 3 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-MCU-04 Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |

Success/Failure criteria: test passed is test pattern is properly generated

Comments/Open issues:

## FUNC-SMEC-01, SMEC switch on and initialisation

| ID: | FUNC-SMEC-01 |
| --- | --- |
| Purpose | Switching on and initialising SMEC. SMEC ends up at home position and is ready for following tests |
| Description of test: | Switch on various part of the SMEC system then initialise it. Check parameters |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode, SMEC off |
| Final configuration | SPIRE in REDY mode , SMEC on and initialised, at home position |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2-3 min |
| Constraints: | FUNC-MCU-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS Alpha Numeric Display SMEC PARAMETERS |  |  |
| 1 | Mark beginning of building block | SET_BBID(0x82000001) | BBTYPE=0x8200 |
|  | Set up open loop feed forward gain and feed forward offset | If cold: 0 ?? if warm: 0x90541899 90558000 |  |
| 2 | Encoder set-up Set encoder LED power Set signal 1 offset Set signal2 offset Check SMECStatus Check SMECENCSIG1 Check SMECENCSIG2 | If COLD If WARM 0x90400004 0x90400006 0x90587A44 0x90586590 0x905A927C 0x905A7918 SMECSTAT = 0x0002 SMECENCSIG1=0x7FFF SMECENCSIG2=0xC9E8 |  |
| 3 | LVDT set-up (only needed for degraded mode operation) Power on LVDT Set LVDT scale Set LVDT Offset Check LVDTAC Check LVDTDC Check LVDT position | If COLD If WARM 0x90410001 0x90410001 0x905F9157 0x905F161C 0x905E1F40 0x905E1F40 SMECLVDTACSIG = 0x1DB7 SMECLVDTDCSIG = 0x1DBA SMECLVDTPOSN = 0x0000 |  |
| 4 | Set PID Control parameters Set Kp Set Kd Set Ki Set Integration limit Set Rate limit | 0x904A07D0 0x904B02BC 0x904D03E8 0x904E07D0 0x9051012C |  |
| 5 | Initialise SMEC Set SMEC to hold mode Apply –50mA current on SMEC Wait 15 seconds Set end traj to 0 Set start trajectory to 0 Move SMEC to 0 Read trajectory position Set traj modeEncoder initialisation Set Traj mode to 1 Apply 0mA to release Wait 15 seconds Close loop Check SMECStatus Check SencoderSignal1 Check SencoderSignal2 Check SencoderSignal3 Check LVDTAC Check LVDTDC Check LVDT position Check motor current Check back emf Check encoder increment position Check encoder fine position | SetSTrajMode(0x0000), 0x90490000 SetSFeedForwardOffset(0x4000), 0x90554000 Wait(15) 0x90450000 0x90460000 0x90490001 cf param on HK 90490004 90490001 SetSFeedForwardOffset(0x8000), 0x90558000 Wait(15) 0x90440001 SMECSTAT = 0x0188 SMECENCSIG1=0xD527 SMECENCSIG2=0x5A4F SMECENCSIG3=0x5085 SMECLVDTACSIG = 0x315F SMECLVDTDCSIG = 0x3165 SMECLVDTPOSN = 0x0709 SMECMOTORCURR = 0x03E8 SMECBEMF = 0x03E8 SMECENCPOSN = 0x0708 SMECENCFINEPOS = 0x0000 |  |

Success/Failure Criteria Test passed if SMEC ends up at home position ,loop is closed ,and parameters values reflect this

Comment/Open Issue: Need conversion curve between signals and physical position. D. Ferrand will give actual values for all the parameters after all tests at ambient and 4K are done

Switch off procedure:

| Action | Command |
| --- | --- |
| Send SMEC to mechanical limit Engage launch latch Apply 0 mA Switch off LVDT and encoder | SetSTrajMode(0x0001) 0x90490001 SetSTrajEnd(0x0000), 0x90450000 SetSFeedForwardOffset(0x4000), 0x90554000 SetSLaunchLatch(0x0001) 0x90430001 SetSFeedForwardOffset(0x8000), 0x90558000 SetSLVDTPwr(0x0000) 0x90410000 and SetSEncoderPwr(0x0000) 0x90400000 |

## FUNC-SMEC-02, SMEC launch latch check – not to perform during PFM1

| ID: | FUNC-SMEC-02 |
| --- | --- |
| Purpose | Checking that the launch latch works properly |
| Description of test: | Open launch latch, move SMEC short distance and check response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 |
| Level | ILT/PLT/IST |
| Test Conditions: | Warm and cold |
| Total duration | 2-3 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Mark beginning of building block | SET_BBID(0x82040001) | BBTYPE=0x8204 |
| 2 | Send SMEC to position for launch latch Move SMEC to position Read SMEC position | SetTrajEndPosition(0x0000) 90450000 SetSTrajMode(0x0001) 90490001 SMECENCPOSN=0x???? SMECENCFINEPOSN=0x???? SMECLVDTACSIG=0x???? SMECLVDTPOSN=0x???? SMECPOSNDELTA=0x???? |  |
| 3 | Test launch latch Engage launch latch Check latch status Try and move SMEC by a short distance (500 um) Make sure SMEC hasn’t moved Go back to 0 Disengage launch latch Check latch status | SetSLaunchLatch(0x0001) 0x90430001 SMECLATCHSTAT=0x0001 SetTrajEndPosition(0x01F4) 0x904501F4 Check values on SCOS display SetTrajEndPosition(0x01F4) 0x90450000 SetSLaunchLatch(0x0002) 0x90430002 SMECLATCHSTAT=0x0002 |  |
| 4 | Check Move SMEC by a short distance (500 um) Read SMEC position Make sure that SMEC moved by the distance requested | SetTrajEndPosition(0x01F4) 0x904501F4 SMECENCPOSN=0x???? SMECENCFINEPOSN=0x???? SMECLVDTACSIG=0x???? SMECLVDTPOSN=0x???? SMECPOSNDELTA=0x???? Check values on SCOS display |  |
| 5 | Drive SMEC back to its home position | SetTrajEndPosition(0x0000) 0x90450000 |  |

Success/Failure Criteria: Test passed if launch latch is successfully engaged/disengaged

Comments/Open issues:

Is the actuator for the latch redundant?

Would it be safer to move SMEC to position using SetFeedForwardOffset rather than SetTrajEndPosition? What happen if SMEC is not in position and the command Engage Latch issued?

## FUNC-SMEC-03, SMEC LEDs test

| ID: | FUNC-SMEC-03 |
| --- | --- |
| Purpose | Checking integrity of the LEDs |
| Description of test: | Switch on LEDs and step up voltages |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode, SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 2 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Request MCU ENG frames at 64Hz | SetTP14SampFreq(0x0027) 0x91C40027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 2 | Mark beginning of bb for QLA triggering Set Step to 1 | SET_BBID(0x82050001) SET_OBS_STEP(0x0001) | BBTYPE=0x8205 |
| 3 | Test LED Apply power to LEDs by stepping up voltages from 0 to 5V increasing the step number each time (this voltage is dependent on temperature) Same bias is applied to resistor for a given LED level but the resistance drops exponentially approaching cold temperatures. Record MCUSMECENCSIG1, MCUSMECENCSIG2 frame parameters FOR EACH STEP Make 3 QLA plots at the end of the sequence of : Plot 1: MCUSMECENCSIG1 and 2 against encoder level Plot 2 : SMECTEMP and SMECIFTEMP HK parameters against time Plot 3 : MCUSMECENCSIG1 and 2 against SMECTEMP HK Stop data generation. | a) SET_OBS_STEP(0x0001) Qla operation SetSEncoderPwr(0x0001) 0x90400001 Wait 5 sec SET_OBS_STEP(0xFFFF) SET_OBS_STEP(0x0002) Qla operation SetSEncoderPwr(0x0001) 0x90400002 Wait 5 sec SET_OBS_STEP(0xFFFF) … SET_OBS_STEP(0x0001) Qla operation SetSEncoderPwr(0x0006) 0x90400006 Wait 5 sec SET_OBS_STEP(0xFFFF) b) Qla operation c) SetFrameStart(0x0000) 0x91C10000 FLUSH_FIFO ( Fifo flags =0x2000) |  |
| 3 | Power off LED or leave at TBD level if necessary | SetSEncoderPwr(0x0000) 0x90400000 |  |

Success/Failure Criteria: Test passed if encoder signal increases as expected, i.e. with the power applied to the LEDs

Comment/open issues: Need the expected SMEC encoder signals 1 and 2 levels from Dominique Poulenquin

## FUNC-SMEC-04a, SMEC open loop position test

| ID: | FUNC-SMEC-04a |
| --- | --- |
| Purpose | Checking that SMEC moves to the position commanded, in open loop |
| Description of test: | Move to set positions and measure read back position, currents, back emf |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + MCU on + SMEC on and initialised for open loop operation at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 3 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Run QLA script in order to: Record SMEC frame parameters SMECSCANLVDTPOSN SMECSCANMOTORBEMF SMECCMDMOTORCURR Record MCU ENG frame parameters MCUENGSMECLVDTACSIG MCUENGSMECLVDTDCSIG |  |  |
| 2 | Start SMEC frame generation @ 250Hz and MCU ENG frame generation @ 64Hz | SetTelemetryPacket10Sampl(10) 0x91C0000A SetTelemetryPacket14Sampl(39) 0x91C00027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetStartFrame 0x91C10001 |  |
| 3 | Mark beginning of building block | SET_BBID(0x82060001) | BBTYPE=0x8206 |
| 4 | Move SMEC Force open loop operation Set Step to 1 Start recording the SMEC frame parameters described in 1 Set a position to reach (e.g. +1mm from mechanical stop) Move SMEC to first position Wait n sec Set trajectory end position to next position Wait n sec …Repeat steps f) to h) for a set of positions Set step to 0xffff Stop recording parameters | SetSLoopMode(0x0001) 0x90440006 SET_OSB_STEP(0x0001) Qla operation SetTrajEndPosition(0x03E8) 0x904503E8 SetSTrajMode(0x0001) 0x90490001 Wait n sec SetTrajEndPosition(0x1B58) 0x90451B58 Wait n sec … SET_OBS_STEP(0xffff) Qla operation SMECENCPOSN=0x03E8 SMECENCFINEPOSN=0x???? SMECLVDTACSIG=0x1DB7 SMECLVDTPOSN=0x03E8 SMECPOSNDELTA=0x0000 SMECMOTORCURR = 0x???? SMECMOTORBEMF = 0x???? | We don’t set Step to 0xffff each SMEC step because we need a continuous graphic at the end. |
| 5 | Set of position If QLA is available, send SMEC science frames and check sensor signal of graph display | SetTrajEndPosition 0x90451B58 SetTrajEndPosition 0x90451F40 SetTrajEndPosition 0x90452328 SetTrajEndPosition 0x90452710 SetTrajEndPosition 0x90452AF8 SetTrajEndPosition 0x90452EE0 SetTrajEndPosition 0x904532C8 SetTrajEndPosition 0x904536B0 SetTrajEndPosition 0x90453A98 SetTrajEndPosition 0x90453E80 |  |
| 6 | Drive SMEC back to the start position | SetTrajEndPosition(0x0000) 0x90450000 |  |

Success/Failure Criteria: Test passed if SMEC moves to the set position and current and back emf have expected values

Comment/Open issue: Need conversion curve between signal and physical position. I need to specify some of the values expected.  I use conversion curves in Bruce’s document

## FUNC-SMEC-04b, SMEC closed loop position test

| ID: | FUNC-SMEC-04b |
| --- | --- |
| Purpose | Checking that SMEC moves to the position commanded, in open loop |
| Description of test: | Move to set positions and measure read back position, currents, back emf |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + MCU on + SMEC on, initialised for close loop operation at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 3 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Run QLA script in order to: Record SMEC frame parameters : SMECENCPOSN SMECSCANLVDTPOSN SMECSCANMOTORBEMF SMECCMDMOTORCURR Record MCU ENG frame parameters : MCUENGSMECLVDTACSIG MCUENGSMECLVDTDCSIG |  |  |
| 2 | Start SMEC frame generation @ 250Hz and MCU ENG frame generation @ 64Hz | SetTelemetryPacket10Sampl(10) 0x91C0000A SetTelemetryPacket14Sampl(39) 0x91C00027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetStartFrame 0x91C10001 |  |
| 3 | Mark beginning of building block | SET_BBID(0x82060001) | BBTYPE=0x8206 |
| 4 | Move SMEC Force open loop operation Set Step to 1 Start recording the SMEC frame parameters described in 1 Set a position to reach (e.g. +1mm from mechanical stop) Move SMEC to first position Wait n sec | SetSLoopMode(0x0001) 0x90440006 SET_OSB_STEP(0x0001) Qla operation SetTrajEndPosition(0x03E8) 0x904503E8 SetSTrajMode(0x0001) 0x90490001 Wait n sec | We don’t set Step to 0xffff each SMEC step because we need a continuous graphic at the end. |
| 5 | Repeat for a small set of 2mm from mechanical stop 3mm from mechanical stop 4mm from mechanical stop 5mm from mechanical stop 6mm from mechanical stop 7mm from mechanical stop 8mm from mechanical stop 9mm from mechanical stop 10mm from mechanical stop 11mm from mechanical stop | a) SetTrajEndPosition (0x904507D0) b) SetTrajEndPosition (0x90450BB8) c) SetTrajEndPosition (0x90450FA0) d) SetTrajEndPosition (0x90451388) e) SetTrajEndPosition (0x90451770) f) SetTrajEndPosition (0x90451F40) g) SetTrajEndPosition (0x90452328) h) SetTrajEndPosition (0x90452710) i) SetTrajEndPosition (0x90452AF8) j) SetTrajEndPosition (0x90452EE0) |  |
| 6 | a) Set STEP to 0xffff b) Stop MCU frame generation | a ) SET_OBS_STEP(0xFFFF) b ) SetStartFrame (0x91C10000) FLUSH_FIFO ( Fifo flags = 0x2000 ) |  |
| 5 | Drive SMEC back to its home position | SetTrajEndPosition(0x????) 0x9045???? |  |

Success/Failure Criteria: Test passed if SMEC moves to the set position and current and back emf have expected values

Comment/Open issue: Need conversion curve between signal and physical position. I need to specify some of the values expected

## FUNC-SMEC-05, SMEC step and look scan test

| ID: | FUNC-SMEC-05 |
| --- | --- |
| Purpose | Characterising SMEC movement for a step-and-look scan |
| Description of test: | Perform a step-and-look scan, measure positions, current, back emf |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + SMEC on and initialised for closed loop operation at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | ~18 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Mark beginning of building block Set Step to 1 | SET_BBID(0x8207000) SET_OBS_STEP(0x0001) | BBTYPE=0x8207 |
| 2 | Start MCU ENG and SMEC packet sampling at 64Hz and 250Hz | SetTP10SampFreq(0x000B) 0x91C0000A SetTP14SampFreq(0x0027) 0x91C40027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 3 | Run QLA script FUNC-SMEC-05 in order to: Display SMEC parameters Record a time series of the same SMEC frame parameters Record a time series of Full Spectrometer detector data | a) b) SMECENCPOSN, SMECENCFINEPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, SMECMOTORBEMF, (+others?) c)SPECFARRAY01…SPECFARRAY294 |  |
| 4 | Move SMEC Set SMEC to hold mode Set scan speed to max (2000 um/sec) Set speed for fly back (2000 um/sec) Set start position of the scan (e.g. –1.0 mm from ZPD) Move to position Set step to 1 Record parameters for 10 sec Set step to 0xffff For i=1..100{ increment position to reach by 10 um move to position Set step to i+1 Record parameters for 10 sec Set step to 0xffff } | SetSTrajMode(0x0000) 0x90490000 SetScanFwdSpeed 0x90474E20 Set Scan RevSpeed 0x90564E20 SetSTrajEndPosition(0x1B58) 0x90451B58 SetSTrajMode(0x0001) 0x90490001 SET_OBS_STEP(0x0001) Wait 10 sec SET_OBS_STEP(0xffff) pos = 0x1B58 + i*0x000A SetSTrajEndPosition(0x(pos)) 0x9045(pos) SET_OBS_STEP(0x000i+1) Wait 10 sec SET_OBS_STEP(x0xffff) |  |
| 5 | Write in a file the parameters recorded with QLA |  |  |
| 6 | Drive SMEC back to its home position | SetTrajEndPosition(0x????) 0x9045???? |  |
| 7 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 0x91C00000 |  |
| 8 | Offline analysis Characterise SMEC scans | Derive Average value of each parameter for each step and plot average value vs position compare LVDT and encoder signal (plot LVDT vs encoder) consumed power |  |

Success/Failure Criteria Test passed if SMEC does the scan required

Comment/Open issue:

Need conversion curve between signals and physical position. I use conversion curves in Bruce’s document

## FUNC-SMEC-06, SMEC saw tooth scan test

| ID: | FUNC-SMEC-06 |
| --- | --- |
| Purpose | Characterising SMEC movement during a saw tooth scan |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scans |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised for closed loop operation at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 10 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Mark beginning of building block Set Step to 1 | SET_BBID(0x82080000) SET_OBS_STEP(0x0001) | BBTYPE=0x8208 |
| 1 | Run QLA script FUNC-SMEC-06 in order to: Display SMEC frame parameters Record a time series of all the SMEC frame parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECLVDTACSIG, SMECOTORCURR, SMECMOTORBEMF,SMECENCSIG1,SMECENCSIG2 |  |
| 2 | Start MCU ENG and SMEC packet sampling at 64Hz and 250Hz | SetTP10SampFreq(0x000B) 0x91C0000A SetTP14SampFreq(0x0027) 0x91C40027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 3 | Move SMEC Set SMEC to hold mode Set a scan speed (500 um/s) Set fly back speed (2000 um/sec) Set number of scan to 2 Set a start position (-1mm from ZPD) Set end position (+1 mm from ZPD) Set step to 1 Start the scan Wait for scan to finish Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetSScanSpeedForward(0x1388) 0x90471388 SetSScanRevSpeed(0x4E20) 0x90564E20 SetSScanNumber(0x0002) 0x90480002 SetStrajStartPosition(0x1B58) 0x90461B58 SetSTrajEndPosition(0x2328) 0x90452328 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 3’ | Repeat scan for different scan distance and speed. Each new scan will be assigned a new obsid. |  |  |
| 4 | Repeat for different scan amplitude Set SMEC to hold mode Set scan amplitude from –4 to +4 mm Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetStrajStartPosition(0x0FA0) 0x90460FA0 SetSTrajEndPosition(0x2EE0) 0x90452EE0 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 5 | Repeat for different scan speed Set SMEC to hold mode Reset scan amplitude +/- 1mm Set the speed to 1mm/s Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step to 0xffff Set SMEC to hold mode Set the speed to 0.2mm/s Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetStrajStartPosition(0x1B58) 0x90461B58 SetSTrajEndPosition(0x2328) 0x90452328 SetSScanSpeedForward(0x2710) 0x90472710 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) SetSTrajMode(0x0000) 0x90490000 SetSScanSpeedForward(0x07D0) 0x904707D0 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 6 | Write in a file the parameters recorded with QLA |  |  |
| 7 | Drive SMEC back to its home position | SetSTrajMode(0x0000) 0x90490000 SetSTrajEndPosition(0x0000) 0x90450000 SetSTrajMode(0x0001) 0x90490001 |  |
| 8 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 0x91C00000 |  |
| 9 | Offline analysis Characterise SMEC motion | Derive: scan length, accuracy of position reached velocity, velocity stability consumed power |  |

Success/Failure Criteria: Test passed if SMEC does the scan required

Comment/Open issue:

## FUNC-SMEC-07, SMEC triangular scan test

| ID: | FUNC-SMEC-07 |
| --- | --- |
| Purpose | Characterising SMEC movement during a triangular scan |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scans |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised for closed loop operation, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Mark beginning of building block Set Step to 1 | SET_BBID(0x82090000) SET_OBS_STEP(0x0001) | BBTYPE=0x8209 |
| 2 | Run QLA script FUNC-SMEC-06 in order to: Display SMEC parameters Record a time series of the parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, SMECENCSIG1,SMECENCSIG2, (+others???) |  |
| 3 | Start MCU ENG and SMEC packet sampling at 64Hz and 250Hz | SetTP10SampFreq(0x000B) 0x91C0000A SetTP14SampFreq(0x0027) 0x91C40027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 4 | Move SMEC Set SMEC to hold mode Set a scan speed (500 um/s) Set fly back speed (500 um/sec) Set number of scan to 2 Set a start position (-1mm from ZPD) Set end position (+1 mm from ZPD) Set step to 1 Run the scan Wait for scan to finish Set step 0xffff | SetSTrajMode(0x0000) 0x90490000 SetSScanSpeedForward(0x1388) 0x90471388 SetSScanRevSpeed(0x1388) 0x90561388 SetSScanNumber(0x0002) 0x90480002 SetStrajStartPosition(0x1B58) 0x90461B58 SetSTrajEndPosition(0x2328) 0x90452328 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 4’ | Repeat scan for different scan distance and speed. Each new scan will be assigned a new obsid. |  |  |
| 5 | Repeat for different scan amplitude Set SMEC to hold mode Set scan amplitude from –4 to +4 mm Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step 0xffff | SetSTrajMode(0x0000) 0x90490000 SetStrajStartPosition(0x0FA0) 0x90460FA0 SetSTrajEndPosition(0x2EE0) 0x90452EE0 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 6 | Repeat for different scan speed Set SMEC to hold mode Reset scan amplitude +/- 1mm Set the speed to 1mm/s Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step 0xffff Set SMEC to hold mode Set the speed to 0.2mm/s Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step 0xffff | SetSTrajMode(0x0000) 0x90490000 SetStrajStartPosition(0x1B58) 0x90461B58 SetSTrajEndPosition(0x2328) 0x90452328 SetSScanSpeedForward(0x2710) 0x90472710 SetSScanRevSpeed(0x2710) 0x90562710 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) SetSTrajMode(0x0000) 0x90490000 SetSScanSpeedForward(0x07D0) 0x904707D0 SetSScanRevSpeed(0x07D0) 0x905607D0 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 7 | Write in a file the parameters recorded with QLA |  |  |
| 8 | Drive SMEC back to its home position | SetSTrajMode(0x0000) 0x90490000 SetSTrajEndPosition(0x0000) 0x90450000 SetSTrajMode(0x0001) 0x90490001 |  |
| 9 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 0x91C00000 |  |
| 10 | Offline analysis Characterise SMEC motion | Derive: scan length, accuracy of position reached velocity, velocity stability consumed power |  |

Success/Failure Criteria: Test passed if SMEC does the scan required

Comment/Open issue:

I use conversion curves in Bruce’s document

## FUNC-SMEC-08, SMEC open loop position test

| ID: | FUNC-SMEC-08 |
| --- | --- |
| Purpose | Moving SMEC in open loop mode. Scaling back emf |
| Description of test: | Move SMEC short distance and check response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA |
| Level | ILT/PLT/IST |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Mark beginning of building block Set Step to 1 | SET_BBID(0x820A0000) SET_OBS_STEP (0x0001) | BBTYPE=0x820A |
| 1 | Start MCU ENG and SMEC packet sampling at 250Hz and 250Hz | SetTP10SampFreq(0x000B) 0x91C0000A SetTP14SampFreq(0x0027) 0x91C4000A SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 2 | Move SMEC Set SMEC to Hold mode Open loop Set a position to reach (ZPD -1mm) Start recording parameters with QLA Set step to 1 Move SMEC Wait for position to be reached Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetSLoopMode 0x90440006 SetTrajEndPosition 0x90451B58 SMECLVDTDCSIG, SMECENCPOSN, SMECMOTORBEMF SET_OBS_STEP(0x0001) SetTrajMode 0x90490001 Wait SET_OB_STEP(0xffff) |  |
| 3 | Scale back emf TBD | SetMotorResistance |  |
| 4 | Move SMEC to a new position Set SMEC to Hold mode Set a position to reach (ZPD+1mm) Start recording parameters with QLA Set step to 2 Move SMEC Wait for position to be reached Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetTrajEndPosition 0x90452328 SMECLVDTDCSIG, SMECENCPOSN, SMECMOTORBEMF SET_OBS_STEP(0x0002) SetTrajMode 0x90490001 Wait SET_OBS_STEP(0xffff) |  |
| 7 | Drive SMEC back to its home position | SetSTrajMode(0x0000) 0x90490000 SetSTrajEndPosition(0x0000) 0x90450000 SetSTrajMode(0x0001) 0x90490001 |  |

## FUNC-SMEC-09, SMEC open loop scan test

| ID: | FUNC-SMEC-09 |
| --- | --- |
| Purpose | Scanning in open loop mode (feedback on back emf only) |
| Description of test: | Scan over range and measure positions, current, back-emf – repeat for different scans. Test identical to SMEC-07 but with open loop |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + MCU on + SMEC on and initialised, at home position |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01 and FUNC-SMEC-01 and SMEC-08 successfully passed |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 0 | Select SCOS display (SMEC PARAMETERS) |  |  |
| 1 | Mark beginning of building block Set Step to 1 | SET_BBID(0x820B0000) SET_OBS_STEP (0x0001) | BBTYPE=0x820B |
| 1 | Start MCU ENG and SMEC packet sampling at 64Hz and 250Hz | SetTP10SampFreq(0x000B) 0x91C0000A SetTP14SampFreq(0x0027) 0x91C40027 SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 2 | Run QLA script FUNC-SMEC-07 in order to: Display SMEC parameters Record a time series of the parameters | SMECENCPOSN, SMECENCFINEPOSN, SMECENCPOSN, SMECLVDTACSIG, SMECMOTORBEMF, SMECOTORCURR, (+others???) |  |
| 3 | Move SMEC Set SMEC to hold mode Open loop (no sensor) Set a scan speed (500 um/s) Set fly back speed (500 um/sec) Set number of scan to 2 Set a start position (ZPD-1mm) Set end position (ZPD+1 mm) Set step to 1 Run the scan Wait for scan to finish Set step 0xffff | SetSTrajMode(0x0000) 0x90490000 SetSLoopMode(0x0006) 0x90440006 SetSScanSpeedForward(0x1388) 0x90471388 SetSScanRevSpeed(0x1388) 0x90561388 SetSScanNumber(0x0002) 0x90480002 SetStrajStartPosition(0x1B58) 0x90461B58 SetSTrajEndPosition(0x2328) 0x90452328 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 3’ | Repeat for different scan distances and speed. Each new scan will have a new obsid |  |  |
| 4 | Repeat for different scan amplitude Set SMEC to hold mode Set scan amplitude (+/-4mm) Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetStrajStartPosition(0x0FA0) 0x90460FA0 SetSTrajEndPosition(0x2EE0) 0x90452EE0 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) |  |
| 5 | Repeat for different scan speed Set SMEC to hold mode Reset scan amplitude +/- 1mm Set the speed to 1mm/s Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step to 0xffff Set SMEC to hold mode Set the speed to 0.2mm/s Set number of scan to 2 Set step to 1 Run the scan Wait for scan to finish Set step to 0xffff | SetSTrajMode(0x0000) 0x90490000 SetStrajStartPosition(0x1B58) 0x90461B58 SetSTrajEndPosition(0x2328) 0x90452328 SetSScanSpeedForward(0x2710) 0x90472710 SetSScanRevSpeed(0x2710) 0x90562710 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OBS_STEP(0xffff) SetSTrajMode(0x0000) 0x90490000 SetSScanSpeedForward(0x07D0) 0x904707D0 SetSScanRevSpeed(0x07D0) 0x905607D0 SetSScanNumber(0x0002) 0x90480002 SET_OBS_STEP(0x0001) SetSTrajMode(0x0002) 0x90490002 Wait SET_OSB_STEP(0xffff) |  |
| 6 | Write in a file the parameters recorded with QLA | Use save button on QLA GUI |  |
| 7 | Drive SMEC back to its home position | SetSTrajMode(0x0000) 0x90490000 SetSTrajEndPosition(0x0000) 0x90450000 SetSTrajMode(0x0001) 0x90490001 |  |
| 8 | Stop TM packet sampling | SetTP10SampFreq(0x0000) 0x91C00000 |  |
| 9 | Offline analysis Characterise SMEC movement: | Derive: scan length, accuracy of position reached velocity, velocity stability consumed power |  |

Success/Failure Criteria: Test passed if SMEC does the scan required

Comment/Open issue:

I use conversion curves in Bruce’s document

## FUNC-BSM-01c, BSM power on chop, perform small step

| ID: | FUNC-BSM-01c |
| --- | --- |
| Purpose | Checking the integrity of BSM motors and sensors on chop axis. |
| Description of test: | Switch BSM chop motors on. Check telemetry responses, consumed power – perform a small step in open loop |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + BSM chop sensor OFF |
| Final configuration | SPIRE in REDY mode + BSM chop sensor ON |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Mark beginning of building block | SET_BBID(0x81000001) |  |
| 3 | Switch on BSM chop axis Power on chop MR sensor Force open loop in chop axis Set Chop and Jiggle to move independently Record chop status Record chop magneto resistive signal Record chop motor current Record chop motor voltage Record Chop DAC value | SetCSensorPwr(0x0001) 0x90C00001 SetChopLoopMode 0x90C20003 SetBSMMove(0x0000), 0x90C60000 BSMSTAT CHOPSENSSIG CHOPMOTORCURR CHOPMOTORVOLT CHOPDACVAL |  |
| 4 | Evaluate consumed power | Proportional to (CHOPMOTORCURR)^2 | There is an offset between measured current and DAC value. In order to do a correct evaluation so substract the offset read when no current is applied. This offset is the value of getMotorcurrent when dac value is set to 0x8000. We can know the coil resistance because we measure both current and voltage. The motor coil voltage read by mean of the ADC is (Motor Voltage ADCvalue -32768)/32768*10Volts/19.71 (19.71 is the voltage amplification in BSM board which amplifies the motor coil voltage. The resistance is given by measured voltage/measured current (take many points to get a mean value). |
| 5 | Perform a small step in open loop Send to position 35000 (<0.2 degree) Check parameters | SetChopTargetPosition 0x90C388B8 CHOPSENSSIG, CHOPMOTORCURR, CHOPMOTORVOLT, CHOPDACVAL | The conversion between position on the sky and target position is dependent on version of the electronics, BSM model and prime or redundant channel use. The commanded target position that appears here is just a guide. The current which is sent is: The DAC value sent is: DAC Input= (ChopTarget Position-32768)*CHOPFFGAIN/3051.757813 + 32768 The DAC output voltage (Volt) is: DAC output volts = (DACinput-32768)/32768*10 Volts The current sent in coil is: DAC output volts*60mA/10V So a target position of 35000 sends a current of 4.1 mA |
| 6 | Return to 0 | SetChopTargetPosition 0x90C38000 |  |

Success/Failure Criteria:

Test passed if:

-      BSM sensors are switched on and the small angular step is performed.

DF Comments: test is passed if:

- The measured current given by CHOPMOTORCURRENT is close from CHOPDACVALUE (there is a small offset and little gain difference between desired current which is the DAC value and the actual measured current by ADC.

- The magnetoresistive signal evolves in the same direction as the motor current (but gains are different because the scale of the sensor is different and depends on magnetoresistors themselves. THIS TESTS THE POLARITY COMPATIBILITY.

Comment/Open issue:

Need conversion curve of angle vs target position.

## FUNC-BSM-01j, BSM power on jiggle, perform small step

| ID: | FUNC-BSM-01j |
| --- | --- |
| Purpose | Checking the integrity of BSM motors and sensors in jiggle axis |
| Description of test: | Switch BSM jiggle motors on. Check telemetry responses, consumed power – perform a small step in open loop |
| Test Type: | Integrity |
| Instrument Models: | CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + BSM jiggle sensor OFF |
| Final configuration | SPIRE in REDY mode + BSM sensor ON |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm and cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGGLE PARAMETERS) |  |  |
| 2 | Mark beginning of building block | SET_BBID(0x81010001) |  |
| 3 | Switch on BSM jiggle axis Power on jiggle sensors Force open loop in jiggle axis Set Chop and Jiggle to move independently Record jiggle status Record jiggle magneto resistive signal Record jiggle motor current Record jiggle motor voltage Record Jiggle DAC value | SetJSensorPwr(0x0001) 0x91400001 SetJigLoopMode 0x91420003 SetBSMMove(0x0000), 0x90C60000 BSMSTAT JIGGSENSSIG JIGGMOTORCURR JIGGMOTORVOLT JIGGDACVAL |  |
| 4 | Evaluate consumed power | Proportional to (JIGGMOTORCURR)^2 |  |
| 5 | Perform a small step in open loop Send to position 35000 (<0.05 degree) Check parameters | SetJiggTargetPosition 0x914388B8 JIGGSENSSIG, JIGGMOTORCURR, JIGGMOTORVOLT, JIGGDACVAL | Same comments as for chop axis |
| 6 | Return to 0 | SetJiggTargetPosition 0x91438000 |  |

Success/Failure Criteria:

Test passed if:

-      BSM sensors are switched on and the small angular step is performed.

DF Comments: test is passed if:

- The measured current given by JIGGMOTORCURRENT is close from JIGGDACVALUE (there is a small offset and little gain difference between desired current which is the DAC value and the actual measured current by ADC.

- The magnetoresistive signal evolves in the same direction as the motor current (but gains are different because the scale of the sensor is different and depends on magnetoresistors themselves. THIS TESTS THE POLARITY COMPATIBILITY.

Comment/Open issue:

Need conversion curve of angle vs target position.

## FUNC-BSM-02c, BSM chop axis in open loop

| ID: | BSM-02c |
| --- | --- |
| Purpose | Scaling the back emf to observe the BSM behaviour on degraded mode operation. |
| Description of test: | Perform a step of 1 deg, record sensor signal and plot it with QLA. Set the chop motor resistance in order to damp the natural axis oscillation. Set a new target value, record and plot sensor signal. |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: |  |
| Initial configuration: | SPIRE in REDY mode + BSM ON |
| Final configuration | SPIRE in REDY mode + BSM ON |
| EGSE Configuration: | SCOS-2000 QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Cold |
| Total duration | 2 min |
| Constraints: | MCU-01, BSM-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-02c |  |  |
| 3 | Mark beginning of building block | SET_BBID(0x81020001) |  |
| 4 | Setup BSM movement in open loop with backemf use a) Set Chop and Jiggle to move independently b) Set the open loop constant between target position and DAC value = 1/32768 c) Set slew rate limit d) Set mode to open loop (no sensor) e) Scaling the back emf Set Cmotor resistance to : the mean value of R(calculated on BSM03c)*10000 Apply some gain on the feedback | a) SetBSMMove(0x0000) (0x90C60000) b) SetCPositionScaleFactor(S=1/32768) (0x90D60BEB) c) SetCRateLimit(C=1000) (0x90D103E8) d) SetChopLoopMode(Mode= 3) (0x90C20003) e) SetCmotor resistance 0x90D3(0xValue) SetCmotorbackEmfGain 0x90D2???? |  |
| 5 | Start BSM frame generation @ 250Hz | SetTelemetryPacket12Sample(s=10) (0x91C2000A) SetFrameNumber(0xFFFF) (0x91C3FFFF) SetFrameStart(0x0001) (0x91C10001). |  |
|  | Move BSM in chop axis Set step to 1 Start recording BSMCHOPSENSSIG with QLA Wait 1 sec Perform a step to 20000 (~-1 degree) Wait 1 sec Set step to 0xffff Stop recording BSMCHOPSENSSIG Make a time series QLA plot of BSMCHOPSENSSIG BSM frame parameter Set step to 2 Start recording BSMCHOPSENSIG with QLA Wait 1 sec Perform a step to 40000 ( ~+0.25 degree) Wait 1 sec Set step to 0xffff Stop recording BSMCHOPSENSSIG Make a time series QLA plot of BSMCHOPSENSSIG BSM frame parameter Move BSM to 0 Stop BSM packet telemetry Flush MCU FIFO | SET_OBS_STEP(0x0001) Qla operation Wait 1 sec SetChopTargetPosition(20000) (0x90C34E20) Wait 1 sec SET_OBS_STEP(0xffff) Qla operation Make a time series QLA plot of BSMCHOPSENSSIG BSM frame parameter SET_OBS_STEP(0x0002) Qla operation Wait 1 sec SetChopTargetPosition (0x90C39C40) Wait 1 sec SET_OBS_STEP(0xffff) Qla operation Make a time series QLA plot of BSMCHOPSENSSIG BSM frame parameter SetChopTargetPosition (0x90C38000) SetFrameStart(0x0000) (0x91C10000) FLUSH_FIFO(0x2000) | Plot should show large oscillation (during ~1sec) when backemf is not used. Oscillations should be much smaller (during ~100ms) when backemf is used. |

Test passed if:

-    BSM moves as required and back emf use is effective in damping the natural axis oscillation.

Comment/open issues: Need conversion curve of angle vs target position.

## FUNC-BSM-02j, BSM jiggle axis in open loop

| ID: | BSM-02j |
| --- | --- |
| Purpose | Scaling the back emf to observe the BSM behaviour on degraded mode operation. |
| Description of test: | Perform a step of 0.25 deg, record sensor signal and plot it with QLA. Then set the jiggle motor resistance in order to damp the natural axis oscillation. Set a new target value, record and plot sensor signal. |
| Test Type: | Characterisation |
| Instrument Models: | PFM/FS |
| Redundancy: |  |
| Initial configuration: | SPIRE in REDY mode + BSM ON |
| Final configuration | SPIRE in REDY mode + BSM ON |
| EGSE Configuration: | SCOS-2000 QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Cold |
| Total duration | 2 min |
| Constraints: | MCU-01, BSM-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGGLE PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-02j |  |  |
| 3 | Mark beginning of building block | SET_BBID(0x81030001) |  |
| 4 | Setup BSM movement in open loop with backemf use a) Set Chop and Jiggle to move independently b) Set the open loop constant between target position and DAC value = 1/32768 c) Set slew rate limit d) Set mode to open loop (no sensor) e) Scaling the back emf Set Jmotor resistance to : the mean value of R(calculated on BSM03j)*10000 Apply some gain on the feedback | a) SetBSMMove(0x0000) (0x90C60000) b) SetJPositionScaleFactor(S=1/32768) (0x914D0BEB) c) SetJRateLimit(C=1000) (0x915103E8) d) SetJiggLoopMode(Mode= 3) (0x91420003) e) SetJmotor resistance 0x9153(0xValue) SetJmotorbackEmfGain 0x9152???? |  |
| 5 | Start BSM frame generation @ 250Hz | SetTelemetryPacket12Sample(s=10) (0x91C2000A) SetFrameNumber(0xFFFF) (0x91C3FFFF) SetFrameStart(0x0001) (0x91C10001). |  |
| 6 | Move BSM in jiggle axis Set step to 1 Start recording BSMJIGGSENSSIG with QLA Wait 1 sec Perform a step to 20000 (~-1 degree) Wait 1 sec Set step to 0xffff Stop recording BSMJIGGSENSSIG Make a time series QLA plot of BSMJIGGSSENSSIG BSM frame parameter Set step to 2 Start recording BSMJIGGSENSSIG with QLA Wait 1 sec Perform a step to 40000 (~+ 0.25degree) Wait 1 sec Set step to 0xffff Stop recording BSMJIGGSENSSIG Make a time series QLA plot of BSMCHOPSENSSIG BSM frame parameter Move BSM to 0 Stop BSM packet telemetry Flush MCU FIFO | SET_OBS_STEP(0x0001) Qla operation Wait 1 sec SetJIggTargetPosition(20000) (0x91434E20) Wait 1 sec SET_OBS_STEP(0xffff) Qla operation Make a time series QLA plot of BSMJIGGSENSSIG BSM frame parameter SET_OBS_STEP(0x0002) Qla operation Wait 1 sec SetJIggTargetPosition (0x91439C40) Wait 1 sec SET_OBS_STEP(0xffff) Qla operation Make a time series QLA plot of BSMCHOPSENSSIG BSM frame parameter SetJIggTargetPosition (0x91438000) SetFrameStart(0x0000) (0x91C10000) FLUSH_FIFO(0x2000) | Plot should show large oscillation (during ~1sec) when backemf is not used. Oscillations should be much smaller (during ~100ms) when backemf is used. |

Test passed if:

-    BSM moves as required and back emf use is effective in damping the natural axis oscillation.

Comment/open issues: Need conversion curve of angle vs target position.

## FUNC-BSM-03c, BSM open loop raster, chop axis

| ID: | FUNC-BSM-03c |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on chop axis in open loop. |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, evaluate consumed power, evaluate chop motor resistance value necessary to test BSM degraded mode. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + BSM ON |
| Final configuration | SPIRE in REDY mode + BSM ON |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01c, 02c successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-03c in order to: Record a time series of BSM HK parameters Record a time series of BSM frame parameters | CHOPMOTORCURR, CHOPMOTORVOLT, CHOPSENSSIG, JIGGMOTORCURR, JIGGMOTORVOLT, JIGGSENSSIG BSMCHOPMOTORCURR, BSMCHOPMOTORVOLT, BSMCHOPSENSSIG |  |
| 3 | Mark beginning of building block | SET_BBID(0x81080001) |  |
| 4 | Start TM packet sampling of BSM science frames at 64Hz | SetTP12SampFreq(0x002C) 0x91C2002C SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 5 | Move BSM on chop axis Set Chop and Jiggle to move independently Open loop (no sensor) Send BSM to one end of the chop axis travel Set step to 1 Record BSM HK and frame params specified in 2 parameters for 3 sec Set step to 0xffff Perform a step of 0x1000 Set step to 2 Record BSM HK and frame params specified in 2 parameters for 3 sec Set step to 0xffff … Perform a step of 0x1000 Set step to 16 Record BSM HK and frame params specified in 2 parameters for 3 sec Set step to 0xffff | SetBSMMove(0x0000), 0x90C60000 SetChopLoopMode 0x90C20003 SetChopTargetPosition 0x90C30000 SET_OBS_STEP(0x0001) Wait 3 sec SET_OBS_STEP(0xffff) SetChopTargetPosition 0x90C31000 SET_OBS_STEP(0x0002) Wait 3 sec SET_OBS_STEP(0xffff) … SetChopTargetPosition 0x90C3F000 SET_OBS_STEP(0x0010) Wait 3 sec SET_OBS_STEP(0xffff) |  |
| 6 | With QLA make 2 time series plots FOR EACH STEP of: BSMCHOPSENSSIG against time. (BSMCHOPMOTORVOLT-32768) / (BSMCHOPMOTORCURR-32768) against time. At the end of the sequence step up sequence, make a QLA plot of HK params: CHOPSENSSIG vs CHOPDACVAL Also write in a file a table with the maximum of each of the ratios given in 2 against commanded target position and then calculate the mean value and standard deviation. |  |  |
| 5 | Stop TM packet sampling Flush MCU FIFO | SetFrameStart(0x0000) 0x91C10000 FLUSH_FIFO (FIFO flags = 0x2000) |  |
| 6 | Offline analysis: Characterise BSM movement | Analysis: Consumed power. comm.Pos Vs current read – check is hysteresis occurs current Vs position read Find position where current = 0 Find current when MR sensor value is 32768 | You can determine actual power by calculation of motor coil resistance see above. This shall be used for sending a current offset on chopper axis to center the closed loop. |

Test passed if:

Test passed if BSM does scan required.

DF Comments: test is passed if:

- The measured current given by CHOPMOTORCURRENT is close from CHOPDACVALUE (there is a small offset and little gain difference between desired current which is the DAC value and the actual measured current by ADC.

- The magnetoresistive signal evolves in the same direction as the motor current (but gains are different because the scale of the sensor is different and depends on magnetoresistors themselves. THIS TESTS THE POLARITY COMPATIBILITY.

Comment/open issues: Need conversion curve of angle vs target position.

FUNC-BSM-03j, BSM scan with open loop, jiggle axis

| ID: | FUNC-BSM-03j |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on jiggle axis in open loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + BSM ON |
| Final configuration | SPIRE in REDY mode + BSM ON |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01j, 02j successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGGLE PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-03j in order to: Record a time series of BSM HK parameters Record a time series of BSM frame parameters | JIGGMOTORCURR, JIGGMOTORVOLT, JIGGSENSSIG, CHOPMOTORCURR, CHOPMOTORVOLT, CHOPSENSSIG BSMJIGGMOTORCURR, BSMJIGGMOTORVOLT, BSMJIGGSENSSIG |  |
| 3 | Mark beginning of building block | SET_BBID(0x81090001) |  |
| 4 | Start TM packet sampling of BSM science frames at 64Hz | SetTP12SampFreq(0x002C) 0x91C2002C SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 5 | Move BSM on jiggle axis Set Chop and Jiggle to move independently Open loop (no sensor) Send BSM to one end of the jiggle axis travel Set step to 1 Record BSM HK and frame params specified in 2 parameters for 3 sec Set step to 0xffff Perform a step of 0x1000 Set step to 2 Record BSM HK and frame params specified in 2 parameters for 3 sec Set step to 0xffff … Perform a step of 0x1000 Set step to 16 Record BSM HK and frame params specified in 2 parameters for 3 sec Set step to 0xffff | SetBSMMove(0x0000), 0x90C60000 SetJigLoopMode(0x0003) 0x91420003 SetJiggleTargetPosition(0x0000) 0x91430000 SET_OBS_STEP(0x0001) Wait 3 sec SET_OBS_STEP(0xffff) SetJiggleTargetPosition (0x1000) 0x91431000 SET_OBS_STEP(0x0002) Wait 3 sec SET_OBS_STEP(0xffff) … SetChopTargetPosition 0x90C3F000 SET_OBS_STEP(0x0010) Wait 3 sec SET_OBS_STEP(0xffff) |  |
| 6 | With QLA make 2 time series plots FOR EACH STEP of: BSMJIGGPSENSSIG against time. (BSMJIGGMOTORVOLT-32768) / (BSMJIGGMOTORCURR-32768) against time. At the end of the sequence step up sequence, make a QLA plot of HK params: JIGGSENSSIG vs JIGGDACVAL Also write in a file a table with the maximum of each of the ratios given in 2 against commanded target position and then calculate the mean value and standard deviation. |  |  |
| 7 | Stop TM packet sampling Flush MCU FIFO | SetFrameStart(0x0000) 0x91C10000 FLUSH_FIFO(0x2000) |  |
| 8 | Offline analysis: Characterise BSM movement | Analysis: Consumed power. comm.Pos Vs current read – check if hysteresis occurs current Vs position read Find position where current = 0 Find current when MR sensor value is 32768 | You can determine actual power by calculation of motor coil resistance see above. This shall be used for sending a current offset on jiggle axis to center the closed loop. |

Success/Failure Criteria:

Test passed if BSM does the scan required

DF Comments: test is passed if:

- The measured current given by JIGGMOTORCURRENT is close from JIGGDACVALUE (there is a small offset and little gain difference between desired current which is the DAC value and the actual measured current by ADC.

- The magnetoresistive signal evolves in the same direction as the motor current (but gains are different because the scale of the sensor is different and depends on magnetoresistors themselves. THIS TESTS THE POLARITY COMPATIBILITY.

Comment/open issues: Need conversion curve of angle vs target position.

## FUNC-BSM-04c, BSM position test with closed loop, chop axis

| ID: | FUNC-BSM-04c |
| --- | --- |
| Purpose | Checking that BSM moves to the position commanded and that control loop on MRS works, for chop axis |
| Description of test: | Move BSM to a set position. Check BSMChopSensSig parameter to verify the step response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + BSM ON |
| Final configuration | SPIRE in REDY mode + BSM ON |
| EGSE Configuration: | SCOS-2000 required, QLA |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01, FUNC-BSM-01c, 02c, 03c successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Run QLA script |  |  |
| 3 | Mark beginning of building block | SET_BBID(0x81040001) |  |
| 4 | Set Chop and Jiggle to move independently | SetBSMMove(0x0000), 0x90C60000 |  |
| 5 | Move BSM on chop axis Start BSM packet telemetry sampling @ 250 Hz Set step to 1 Start recording BSMCHOPSENSIG,BSMCHOPMOTORVOLT,BSMCHOPMOTORCURR with QLA Wait 1 s Perform a step to 20000 Wait 1sec Set step to 0xffff Stop recording the previous parameters with QLA | SetTelemetryPacket12Sample 0x91C2000A SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 SET_OBS_STEP(0x0001) Qla operation Wait 1 sec SetChopTargetPosition 0x90C34E20 Wait 1 sec SET_OBS_STEP(0xffff) Qla operation |  |
| 6 | Show plot of BSMCHOPSENSSIG vs time in ms for the step up including requirement. Show plot of chop position error vs time in ms for the step up including requirement. Evaluate consumed power | Plot should show a settling time of 20-30 ms Proportional to (CHOPMOTORCURR)^2 |  |
| 7 | Wait 5 seconds Set step to 2 BSMCHOPSENSIG,BSMCHOPMOTORVOLT,BSMCHOPMOTORCURR with QLA with QLA Wait 1 sec Go to position 40000 Wait 1 sec Set step to 0xffff Stop recording the previous parameters with QLA Go to position 0 Stop BSM packet telemetry | Wait 5 seconds SET_OBS_STEP(0x0002) Qla operation Wait 1 sec SetChopTargetPosition 0x90C39C40 Wait 1 sec SET_OBS_STEP(0xffff) Qla operation SetChopTargetPosition 0x90C38000 SetFrameStart(0x0000) 0x91C10000 |  |
| 8 | Show plot of BSMCHOPSENSSIG vs time in ms for the step up including requirement. Show plot of chop position error vs time in ms for the step up including requirement. Evaluate consumed power | Plot should show a settling time of 20-30 ms Proportional to (CHOPMOTORCURR)^2 |  |

Success/Failure Criteria: test passed if BSM moves to the requested position meeting the requirements over chop step response time

Comment/open issues: Need conversion curve of angle vs target position. This test is more a performance test rather than a functional test so it could be skipped in the BSM functional tests sequence.

## FUNC-BSM-04j, BSM position test with closed loop, jiggle axis

| ID: | FUNC-BSM-04j |
| --- | --- |
| Purpose | Checking that BSM moves to the position commanded and that control loop on MRS works, for jiggle axis |
| Description of test: | Move BSM to a set position. Check telemetry response |
| Test Type: | Integrity |
| Instrument Models: | PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + BSM ON |
| Final configuration | SPIRE in REDY mode + BSM ON |
| EGSE Configuration: | SCOS-2000 required QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Cold |
| Total duration | 1 min |
| Constraints: | FUNC-MCU-01, FUNC-BSM-01j, 02j,03j successful |

Procedure and analysis:

| Step | Action | Command/TM | comment |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGGLE PARAMETERS) |  |  |
| 2 | Run QLA script |  |  |
| 3 | Mark beginning of building block | SET_BBID(0x81050001) |  |
| 4 | Move BSM on jiggle axis Start BSM packet telemetry Set step to 1 Start recording BSMJIGGSENSSIG,BSMJIGGMOTORVOLT,BSMJIGGMOTORCURR with QLA Wait 1 sec Perform a step to 20000 Wait 1 sec Set step 0xffff Stop recording the previous parameters with QLA | SetTelemetryPacket12Sample 0x91C20008 SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 SET_OBS_STEP(0x0001) Qla operation Wait 1 sec SetJiggTargetPosition 0x91534E20 Wait 1 sec SET_OBS_STEP(0xffff) Qla operation |  |
| 5 | Show QLA plot of BSMJIGGSENSSIG vs time in ms for the step up including requirement. Show plot of jiggle position error vs time in ms for the step up including requirement. Evaluate consumed power | Plot should show a settling time of <100 ms Proportional to (JIGGMOTORCURR)^2 |  |
| 6 | Wait 5 seconds Set step to 2 Start recording BSMJIGGSENSSIG,BSMJIGGMOTORVOLT,BSMJIGGMOTORCURR with QLA Wait 1 sec Go to position 40000 Wait 1 sec Set step to 0xffff Stop recording the previous parameters with QLA Go to position 0 Stop BSM packet telemetry | Wait 5 seconds SET_OBS_STEP(0x0002) Qla operation Wait 1 sec SetJiggTargetPosition 0x91439C40 Wait 1 sec SET_OBS_STEP(0xffff) Qla operation SetJiggTargetPosition 0x91438000 SetFrameStart(0x0000) 0x91C10000 |  |
| 7 | Show plot of BSMJIGGSENSSIG vs time in ms for the step up including requirement. Show plot of jiggle position error vs time in ms for the step up including requirement. Evaluate consumed power | Plot should show a settling time of <100 ms Proportional to (JIGGMOTORCURR)^2 |  |

Success/Failure Criteria: test passed if BSM moves to the requested position meeting the requirements over jiggle step response time

Comment/open issues: Need conversion curve of angle vs target position. This test is more a performance test rather than a functional test so it could be skipped in the BSM functional tests sequence.

## FUNC-BSM-05c, BSM scan with closed loop, chop axis

| ID: | FUNC-BSM-05c |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on chop axis, in close loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + BSM ON |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01c, 02c, 03c , 04c successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-05c in order to: Display a time series plot of BSM frame parameter Record BSM frame parameters | BSMCHOPSENSSIG BSMCHOPMOTORCURR, BSMCHOPMOTORVOLT, BSMCHOPSENSSIG,BSMJIGGMOTORCURR, BSMJIGGMOTORVOLT, BSMJIGGSENSSIG |  |
| 3 | Mark beginning of building block | SET_BBID(0x81060001) |  |
| 4 | Move BSM on chop axis Set Chop and Jiggle to move independently Close loop Start BSM packet Telemetry @ 64Hz Set step to 1 Start recording frame parameters specified in 2 Go to BSM raster start position Wait for 2 sec Set step 0xffff Stop recording Set step to 2 Start recording frame parameters specified in 2 Perform a step (e.g. +800hex) Wait for 2 sec Set step to 0xffff Stop recording Repeat steps j), to o), incrementing the position by TBD step size until BSM raster end position | SetBSMMove(0x0000), 0x90C60000 SetChopLoopMode 0x90C20001 SetTP12SampFreq(0x002C) 0x91C2002C SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 SET_OBS_STEP(0x0001) Qla operation SetChopTargetPosition 0x90C30000+xxxx Wait 2 sec SET_OBS_STEP(0xffff) Qla operation SET_OBS_STEP(0x0002) Qla operation SetChopTargetPosition 0x90C3xxxx+0x800 Wait 2 sec SET_OBS_STEP(0xffff) Stop recording SetChopTargetPosition, xxxx SET_OBS_STEP(0x0002+i) where i=1…n Qla operation Wait 2 seconds SET_OBS_STEP(0xffff) Qla operation |  |
|  | Stop MCU telemetry Flush MCU FIFO | SetFrameStart(0x0000) 0x91C10000 FLUSH_FIFO (Fifo flags = 0x2000 MCU fifo) |  |
| 5 | Write in a file the parameters recorded with QLA |  |  |
| 6 | Offline analysis: Characterise BSM movement | Analysis: Consumed power. comm.Pos Vs current read – check is hysteresis occurs current Vs position read Find position where current = 0 |  |
| 7 | Switch off (if required) Send BSM to 0 Power off Chop sensors Open loop (DAC forced to 0) | SetChopTargetPosition(0x0000) 0x90C38000 SetCSensorPwr(0x0000) 0x90C00000 0x90C20000 |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/open issues: Need conversion curve of angle vs target position.Values for target position are “inspired” by the values in  MCU/BSM integration and Test report LAM/ELE/SPI/031020

## FUNC-BSM-05j, BSM scan with closed loop, jiggle axis

| ID: | FUNC-BSM-05j |
| --- | --- |
| Purpose | Characterising BSM movement for a single scan on jiggle axis, in close loop |
| Description of test: | Move BSM to multiple set positions. Check telemetry response, consumed power. |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in REDY mode + BSM ON |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 5 min |
| Constraints: | FUNC-MCU-01, BSM-01j, 02j, 03j ,04j successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (CHOP PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-05j in order to: Display a time series plot of BSM frame parameter Record BSM frame parameters | BSMJIGGSENSSIG BSMCHOPMOTORCURR, BSMCHOPMOTORVOLT, BSMCHOPSENSSIG,BSMJIGGMOTORCURR, BSMJIGGMOTORVOLT, BSMJIGGSENSSIG |  |
| 3 | Mark beginning of building block | SET_BBID(0x81060001) |  |
| 4 | Move BSM on jiggle axis Set Chop and Jiggle to move independently Close loop Start BSM packet Telemetry @ 64Hz Set step to 1 Start recording frame parameters specified in 2 Go to BSM raster start position Wait for 2 sec Set step 0xffff Stop recording Set step to 2 Start recording frame parameters specified in 2 Perform a step (e.g. +800hex) Wait for 2 sec Set step to 0xffff Stop recording Repeat steps j), to o), incrementing the position by TBD step size until BSM raster end position | SetBSMMove(0x0000), 0x90C60000 SetChopLoopMode 0x90C20001 SetTP12SampFreq(0x002C) 0x91C2002C SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 SET_OBS_STEP(0x0001) Qla operation SetJIggTargetPosition 0x91430000+xxxx Wait 2 sec SET_OBS_STEP(0xffff) Qla operation SET_OBS_STEP(0x0002) Qla operation SetChopTargetPosition 0x9143xxxx+0x800 Wait 2 sec SET_OBS_STEP(0xffff) Stop recording SetJiggTargetPosition, xxxx SET_OBS_STEP(0x0002+i) where i=1…n Qla operation Wait 2 seconds SET_OBS_STEP(0xffff) Qla operation |  |
| 5 | Write in a file the parameters recorded with QLA |  |  |
| 6 | Offline analysis: Characterise BSM movement | Analysis: Consumed power. comm.Pos Vs current read – check is hysteresis occurs current Vs position read Find position where current = 0 |  |
| 7 | Switch off Send BSM to 0 Power off jiggle sensors Open loop (DAC forced to 0) | SetJiggTargetPosition(0x0000) 0x91438000 SetJSensorPwr(0x0000) 0x91400000 SetJigLoopMode(0x0000) 0x91420000 |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/open issues: Need conversion curve of angle vs target position.Values for target position are “inspired” by the values in  MCU/BSM integration and Test report LAM/ELE/SPI/031020

## FUNC-BSM-06, BSM operating mode test

| ID: | FUNC-BSM-06 |
| --- | --- |
| Purpose | Characterising BSM movement on chop axis for various position on jiggle axis |
| Description of test: | Chop. Vary position on jiggle axis and chop again. Check telemetry response, consumed power |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in REDY mode + BSM ON |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 45 min |
| Constraints: | FUNC-MCU-01 and BSM-01c/j, 02c/j successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (JIGGLE PARAMETERS) |  |  |
| 2 | Run QLA script FUNC-BSM-06 in order to: Display BSM parameters Record a time series of the parameters | CHOPMOTORCURR, CHOPMOTORVOLT, CHOPSENSSIG, JIGGMOTORCURR, JIGGBMOTORVOLT, JIGGSENSSIG |  |
| 3 | Mark beginning of building block | SET_BBID(0x81090001) |  |
| 4 | Start TM packet sampling for BSM parameters at 64Hz | SetTP12SampFreq(0x002C) 0x91C2002C SetFrameNumber(0xFFFF) 0x91C3FFFF SetFrameStart(0x0001) 0x91C10001 |  |
| 5 | Set the axis to move independently Send command Close loop on both axis | SetBSMMove(0x0000) 0x90C60000 0x90C20001, 0x91420001 |  |
| 6 | Set position on Jiggle axis | SetJiggTarget Position(0x8000) 0x91438000 |  |
| 7 | Chop Set a chop frequency Set a chop amplitude (e.g. 126” on-sky) Set a number of chop throw to do Set building block Id 0x81090001 Set step to 1 Start chopping Wait for Nchops to finish Set step 0xffff Change the chop throw Set step tp 2 Start chopping Wait for Nchop to finish Set step 0xffff Change chop throw … Repeat this sequence incrementing the step number for the list of chop throws | Fchop = 2Hz Nchop = 200 SET_BBID(0x81090001) SET_OBS_STEP(0x0001) SET_OBS_STEP(0xffff) SET_OBS_STEP(0x0002) SET_OBS_STEP(0xffff) |  |
| 8 | Repeat step 7 for different chop frequencies. Each different frequency must be recorded as a different obsid. |  |  |
| 10 | Vary position on jiggle axis and repeat steps 4 to 7 30” 15” -15” -30” | SetJiggTarget Position(0x9555) 0x91439555 SetJiggTarget Position(0x8AAA) 0x91438AAA SetJiggTarget Position(0x7555) 0x91437555 SetJiggTarget Position(0x6AAA) 0x91436AAA |  |
| 11 | Write in a file the parameters recorded with QLA |  |  |
| 12 | Stop TM packet sampling | SetFrameStart(0x0000) 0x91C10000 |  |
| 13 | Switch off Send BSM to 0 Power off sensors Open loop (DAC forced to 0) | SetJiggTargetPosition(0x8000) 0x91438000 SetChopTargetPosition(0x8000) 0x90C38000 SetJSensorPwr(0x0000) 0x91400000 SetChopSensorPwr(0x0000) 0x90C00000 SetJigLoopMode(0x0000) 0x91420000 SetChopLoopMode(0x0000) 0x90C20000 |  |
| 12 | Offline analysis: Characterise BSM movement | Analysis: current Vs position on chop axis for each position on the jiggle axis |  |

Success/Failure Criteria: test passed if BSM does the scan required

Comment/open issues: Need conversion curve of angle vs target position.Values for target position are “inspired” by the values in  MCU/BSM integration and Test report LAM/ELE/SPI/031020

## FUNC-PCAL-01, PCAL characterisation test

| ID: | FUNC-PCAL-01 |
| --- | --- |
| Purpose | Characterising PCAL |
| Description of test: | Step through each setting and measure telemetry response |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in SPIRE in REDY mode |
| Final Configuration | idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 2 min |
| Constraints: | FUNC-SCU-04 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-PCAL-01 in order to: Display PCAL parameters Record parameters at each current setting | PCALCURR, PCALV |  |
| 2 | Start SCU science packet sampling at 80 Hz mark beginning of PCAL bblock | a) BBTYPE=0x8300 ,STEP =0x0001 SetFrameConf(0x0000) A0830000 SetSequenceLength(0x0000) A0840000 SetFrameCtrl(0x0001) A0820001 |  |
| 3 | Apply current to PCAL Set current to 1 mA Set STEP to 0x0002 (Read voltage) record during 5 sec Set STEP to 0xffff Set current to 2.5 mA Set STEP to 0x0003 (Read voltage) record during 5 sec Set STEP 0xffff Set current to 4 mA Set STEP to 0x0004 (Read voltage) record during 5 sec Set STEP to 0xffff Set current to 5.5 mA Set STEP to 0x0005 (Read voltage) record during 5 sec Set STEP to 0xffff Set current to 7 mA Set STEP to 0x0006 (Read voltage) record during 5 sec Set STEP to 0xffff Set current to 0 mA Set STEP 0x0007 (Read voltage) record during 5 sec Set STEP to 0xffff | SetPhCalBias(0x0240) A0C80240 Set STEP= 0x0002 PCALV = 0x135B (252mV) Wait 5 sec Set STEP =0xffff SetPhCalBias(0x0599) A0C80599 Set STEP=0x0003 PCALV = 0x3066 (630mV) Wait 5 sec Set STEP =0xffff SetPhCalBias(0x08F2) A0C808F2 Set STEP =0x0004 PCALV = 0x4D70 (1.01V) Wait 5 sec Set STEP =0xffff SetPhCalBias(0x0C4B) A0C80C4B Set STEP=0x0005 PCALV = 0x6A7A (1.39V) Wait 5 sec Set STEP =0xffff SetPhCalBias(0x0FA5) A0C80FA5 Set STEP=0x0006 PCALV = 0x8784 (1.76V) Wait 5 sec STEP =0xffff SetPhCalBias(0x0000) A0C80000 Set STEP=0x0007 PCALV = 0x0000 Wait 5 sec STEP =0xffff |  |
| 4 | Write in a file parameters recorded |  |  |
| 5 | Stop TM packet sampling | SetFrameCtrl(0x0000) A0820000 |  |
| 6 | Offline analysis Characterise PCAL | Derive for each setting: Voltage stablity Time constant (cooling and warming) Consumed power |  |

##### Success/Failure Criteria: Test passed if each bias current is set as required

Comment/Open issue:

## FUNC-SCAL-01, SCAL characterisation test

| ID: | FUNC-SCAL-01 |
| --- | --- |
| Purpose | Characterising SCAL |
| Description of test: | Set current and record temperatures |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration | SPIRE in SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 15 min |
| Constraints: | FUNC-SCU-05 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-SCAL-01 in order to: Display SCAL parameters a) Mark beginning od SCAL2 bblock | SCAL2CURR, SCAL2V, SCAL2TEMP, SCAL4CURR, SCAL4V, SCAL4TEMP a) BBTYPE =0x8400 ,STEP=0x0001 |  |
| 2 | Apply current to SCAL2 Set current to 1 mA Set STEP 0x0002 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 2 mA Set STEP 0x0003 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 3 mA Set STEP 0x0004 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 4 mA Set STEP 0x0005 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 5 mA Set STEP 0x0006 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 5.5 mA Set STEP 0x0007 (Read voltage) Record during 1 min then set STEP 0xffff Set current to 0 mA Set STEP 0x0008 (Read voltage) Record during 1 min then set STEP 0xffff | SetSCal2Bias(0x02E2) A0CA02E2 STEP =0x0002 SCAL2V = 0x138E (0.5V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x05C4) A0CA05C4 STEP=0x0003 SCAL2V = 0x271E (1V) Wait 1 min then STEP =0xffff SetSCal2Bias(0x08A6) A0CA08A6 STEP=0x0004 SCAL2V = 0x3AAE (1.5V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0B88) A0CA0B88 STEP=0x0005 SCAL2V = 0x4E3D (2V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0E6A) A0CA0E6A STEP =0x0006 SCAL2V = 0x61CD (2.5V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0FC7) A0CA0FC7 STEP =0x0007 SCAL2V = 0x6B95 (2.75V) Wait 1 min then STEP=0xffff SetSCal2Bias(0x0000) A0CA0000 STEP 0x0008 SCAL2V = 0x0000 Wait 1 min then STEP 0xffff |  |
| 3 | Mark beginning of SCAL4 bblock Repeat test for SCAL4 with the same step layout as SCAL2 Set current to 1 mA (Read voltage) Record during 1 min Set current to 2 mA (Read voltage) Record during 1 min Set current to 3 mA (Read voltage) Record during 1 min Set current to 4 mA (Read voltage) Record during 1 min Set current to 5 mA (Read voltage) Record during 1 min Set current to 5.5 mA (Read voltage) Record during 1 min Set current to 0 mA (Read voltage) Record during 1 min | BBTYPE=0x8401 ,STEP=0x0001 SetSCal4Bias(0x02E2) A0CC02E2 SCAL2V = 0x138A (0.5V) Wait 1 min SetSCal4Bias(0x05C0) A0CC05C0 SCAL2V = 0x2714 (1V) Wait 1 min SetSCal4Bias(0x089E) A0CC089E SCAL2V = 0x3A9F (1.5V) Wait 1 min SetSCal4Bias(0x0B7B) A0CC0B7B SCAL2V = 0x4E2A (2V) Wait 1 min SetSCal4Bias(0x0E59) A0CC0E59 SCAL2V = 0x61B5 (2.5V) Wait 1 min SetSCal4Bias(0x0FC8) A0CC0FC8 SCAL2V = 0x6B7A (2.75V) Wait 1 min SetSCal4Bias(0x0000) A0CC0000 SCAL2V = 0x0000 Wait 1 min |  |
| 4 | Write in a file parameters recorded |  |  |
| 5 | Offline analysis Characterise SCAL | Derive for each setting: Temperature stablity Time constant (cooling and warming) Consumed power |  |

Success/Failure Criteria: Test passed if each bias current is set as required

Comment/Open issue: 1 min< SCAL time constant. Shall we record data during 10 or 15 min instead of 1?

## FUNC-SCAL-02, SCAL PID test

| ID: | FUNC-SCAL-02 |
| --- | --- |
| Purpose | Verifying that the PID controller works properly |
| Description of test: | Set temperatures and record temperatures |
| Test Type: | Characterisation |
| Instrument Models: | CQM/PFM |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 2 hours |
| Constraints: | FUNC-SCU-05 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA Display SCAL parameters (in HK) Record a time series of the parameters | SCAL2CURR, SCAL2V, SCAL2TEMP, SCAL4CURR, SCAL4V, SCAL4TEMP |  |
| 2 | Apply current to SCAL2 Set current to 2.5 mA Start PID Read voltage, temperature Record during 1 hour Switch off SCAL2 | SetSCal2Bias(0x072E) A0CA072E SCAL2V = 0x30E6 (1.25V) SCAL2TEMP Wait 1 hour SetSCal2Bias(0x0000) A0CC0000 |  |
| 3 | Apply current to SCAL4 Set current to 2.5 mA Start PID Read voltage, temperature Record during 1 hour Switch off SCAL4 | SetSCal4Bias(0x072F) A0CC072F SCAL4V = 0x30DA (1.25V) SCAL4TEMP Wait 1 hour SetSCal4Bias(0x0000) A0CC0000 |  |
| 4 | Write in a file parameters recorded |  |  |
| 5 | Offline analysis Characterise SCAL PID | Derive standard deviation around nominal temperature |  |

Success/Failure Criteria Test passed if bias current is set as required

Comment/Open issue:

Can’t this test and the previous one be merged into a single test?

Do we have to wait until the stable temperature is achieved? I guess if the purpose of this test is indeed to see if the PID controller works well, we have to record temperature for a long time (…1 hr?)

## FUNC-DCU-01, DCU Science Packet generation check

| ID: | FUNC-DCU-01 |
| --- | --- |
| Purpose | Checking the integrity of the DCU science packet generation |
| Description of test: | Request DCU science packet and check that the number of frames generated matches the number commanded |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min (short) – 3 min (full) |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (DPU and OBS PARAMETER) |  |  |
| 2 | Request nominal DCU photometer science frames Set mode to photometer full array Check frame counter Request 200 frames at 15.3Hz Flush DCU FIFO Check that frame counter has incremented by ~200 Check No Science Packet Events occurred | SetDataMode(0x0000) 843C0000 Read DCUFRAMECNT SetFrameCount(0x00C8) 843D00C8 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) Read DCUFRAMECNT ? |  |
| 3 | Request nominal DCU spectrometer science frames Set mode to spectrometer full array Check frame counter Request 200 frames, sampled at 80Hz Flush DCU FIFO Check that frame counter has incremented by ~200 Check No Science Packet Events occurred | SetDataMode(0x0004) 843C0004 Read DCUFRAMECNT SetFrameCount(0x00C8) 843D0008 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) Read DCUFRAMECNT ? |  |
| 4 | For full functional test, repeat step 3 for: PSW PMW&T/C PLW SLW SSW | SetDataMode(0x0001) 843C0001 SetDataMode(0x0002) 843C0002 SetDataMode(0x0003) 843C0003 SetDataMode(0x0005) 843C0005 SetDataMode(0x0006) 843C0006 |  |

Success/Failure Criteria: Correct number of frames generated for each frame type

Comment/Open issue:

## FUNC-DCU-02, DCU Science data check

| ID: | FUNC-DCU-02 |
| --- | --- |
| Purpose | Checking the integrity of the DCU science packet |
| Description of test: | Request DCU science packet and check some values |
| Test Type: | Integrity Check |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 2 min (short) – 5 min( full) |
| Constraints: | FUNC-DCU-01 successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-DCU-02 |  |  |
| 2 | Request nominal Photometer science frame a) Mark beginning of data bblobk Request 200 frames, sampled at 15.3Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) | a) BBTYPE =0x8800, STEP=0x0001 SetDataMode(0x0000) 843C0000 SetFrameCount(0x00C8) 843D00C8 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff |  |
| 3 | Analyse with QLA Check reception of packets | Not sure what kind of analysis can be done, but for example: Verify regularity of the reception: time difference between two consecutive frame should be constant, or time of each frame vs frame# should be linear |  |
| 4 | Repeat for spectrometer frames a) Mark beginning of data bblobk Request 200 frames, sampled at 80Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) | a) BBTYPE =0x8801, STEP=0x0001 SetDataMode(0x0004) 843C0004 SetFrameCount(0x00C8) 843D00C8 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff |  |
| 5 | Same analysis with QLA |  |  |
| 6 | For full functional test, repeat step 2 for: PSW BBTYPE =0x8802 PMW&T/C BBTYPE =0x8803 PLW BBTYPE =0x8804 SLW BBTYPE =0x8805 SSW BBTYPE =0x8806 | SetDataMode(0x0001) 843C0001 SetDataMode(0x0002) 843C0002 SetDataMode(0x0003) 843C0003 SetDataMode(0x0005) 843C0005 SetDataMode(0x0006) 843C0006 |  |

Success/Failure Criteria: Test passed if frames are properly generated

Comment/Open issue:

Require different scripts for CQM (not all BDAs) and PFM

## FUNC-DCU-03, DCU Test pattern test

| ID: | FUNC-DCU-03 |
| --- | --- |
| Purpose | Checking the integrity of the DCU test pattern |
| Description of test: | Load test pattern and check content of packets. Perform for each type of packet (Photometer and Spectrometer) |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 2 min (short) – 5 min (full) |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Run QLA script FUNC-DCU-03 |  |  |
| 2 | Request DCU photometer test frames a) Mark beginning of data bblobk Set mode to photometer test pattern Request 100 frames at 15.3Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) Display test pattern with QLA | a) BBTYPE= 0x8807 , STEP =0x0001 SetDataMode(0x0008) 843C0008 SetFrameCount(0x00C8) 843D00C8 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff The content is TBD |  |
| 3 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-DCU-03-PHOT-FULL Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |
| 4 | Request DCU spectrometer test frames a) Mark beginning of data bblobk Set mode to spectrometer test pattern Request 200 frames at 80Hz Flush DCU FIFO Mark end of data stream (step to 0xffff) Display test pattern with QLA | a) BBTYPE= 0x8808 , STEP =0x0001 SetDataMode(0x000C) 843C000C SetFrameCount(0x00C8) 843D0008 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 TM(8,4: CA-02) (0x1000) STEP =0xffff The content is TBD |  |
| 5 | Analysis with QLA Compare test pattern with the one obtained during previous run of this test Write in log file the result of the comparison | Read previous file stored on QLA computer e.g. using Notepad Run QLA script FUNC-DCU-03-SPEC-FULL Wait for telemetry stream to stop Compare values on the screen with those from the previous file Use ‘save’ button to save an ascii file containing the values of the parameters in the last test frame |  |
| 6 | For full functional test, repeat step 2 and 3 for PSW BBTYPE =0x880B PMW&T/C BBTYPE =0x880C PLW BBTYPE =0x880D SLW BBTYPE =0x880E SSW BBTYPE =0x880F | SetDataMode(0x0009) 843C0009 SetDataMode(0x000A) 843C000A SetDataMode(0x000B) 843C000B SetDataMode(0x000D) 843C000D SetDataMode(0x000E) 843C000E |  |

Success/failure criteria: Test passed if test pattern generated is similar to the one generated during previous run of this test

Comment/Open issue:

Require different scripts for CQM (not all BDAs) and PFM

## FUNC-DCU-04P, DCU Photometer LIAs switch on

| ID: | FUNC-DCU-04P |
| --- | --- |
| Purpose | Checking the integrity of the Photometer LIAs |
| Description of test: | Switch on photometer LIAs. Check parameter values |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode – Photometer LIAs cards off |
| Final configuration: | SPIRE in REDY mode + Photometer LIAs cards on |
| EGSE Configuration: | SCOS 2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (DCU PARAMETERS) |  |  |
| 2 | Switch on photometer LIAs Switch on Photometer LIA cards (command sent to SCU) Check status Read voltages Read temperatures | SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A0870005 if MCU is on PWR_STATUS = 0x01FF LIAP_P5 = 0xB554 LIAP_P9 = 0xE00D LIAP_N9 = 0x1FF0 LIA01TEMP = ? LIA02TEMP = ? LIA03TEMP = ? LIA04TEMP = ? LIA05TEMP = ? LIA06TEMP = ? LIA07TEMP = ? LIA08TEMP = ? LIA09TEMP = ? | According to document DRCU QM1 Acceptance Data Package (DRAFT) Date 03/11/2003 Page 11 NCR 126 – No temperature probes on BIAS board. Correspondent HK channels returns unpredictable values. |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Switch off Photometer LIA cards | SetDRelOnOff(0x0000) A0870000 or SetDRelOnOff(0x0004) A0870004 if MCU is on |

Success/Failure Criteria: Test passed if parameters have expected values

Comment/Open issue: By using QLA, we could also look at the detector output and evaluate noise, which should be very low…should it be another test?

## FUNC-DCU-04S, DCU Spectrometer LIAs switch on

| ID: | FUNC-DCU-04S |
| --- | --- |
| Purpose | Checking the integrity of the spectrometer LIAs |
| Description of test: | Switch on spectrometer LIAs. Check parameter values |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode + spectrometer LIAs cards off |
| Final configuration: | SPIRE in REDY mode – spectrometer LIAs cards on |
| EGSE Configuration: | SCOS 2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Select SCOS display (DCU PARAMETERS) |  |  |
| 2 | Switch on spectrometer LIA Switch on spectrometer LIA cards (command sent to SCU) Check status Read voltages Read temperatures | SetDRelOnOff(0x0002) A0870002 or SetDRelOnOff(0x0006) A0870006 if MCU is on PWR_STATUS = 0x0E00 LIAS_P5 = 0xB554 LIAS_P9 = 0xE00D LIAS_N9 = 0x1FF0 LIA09TEMP = ? LIA10TEMP = ? LIA11TEMP = ? | See previous test |

Note: for QM1, the command to switch on LIA cards is SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A0870005 if MCU is on

##### Switch off procedure

| Action | Command |
| --- | --- |
| Switch of spectrometer LIA cards | SetDRelOnOff(0x0000) A0870000 Or SetDRelOnOff(0x0004) A0870004 if MCU is on |

Success/Failure Criteria: Test passed if parameters have expected values

Comment/Open issue: By using QLA, we could also look at the detector output and evaluate noise, which should be very low…should it be another test?

## FUNC-DCU-05P, DCU Photometer Offset test

| ID: | FUNC-DCU-05P |
| --- | --- |
| Purpose | Characterising the detector output for various values of offset, on the photometer side |
| Description of test: | Change offsets up and down, record output and analyse off line. |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode + photometer LIAs cards on |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | FUNC-DCU-04P successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8816, STEP=0x0001 SetDataMode(0x0000) 843C0018 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA script FUNC-DCU-05P Look at the first detector of each channel to verify the signal is changing Record detector signals |  |  |
| 3 | Set photometer LIAs offsets Set offset to 0V Set STEP to 0x0002 Record output during 20 sec Set STEP to 0xffff Set offset to 1V Set STEP to 0x0003 Record output during 20 sec Set STEP to 0xffff Set offset to 2V Set STEP to 0x0004 Record output during 20 sec Set STEP to 0xffff Set offset to 3V Set STEP to 0x0005 Record output during 20 sec Set STEP to 0xffff Set offset to 4V Set STEP to 0x0006 Record output during 20 sec Set STEP to 0xffff Set offset to 5V Set STEP to 0x0007 Record output during 20 sec Set STEP to 0xffff Reset offset to 0V Set STEP to 0x0008 Record output during 20 sec Set STEP to 0xffff | For (i=0…8) 0x84200000+i*10000 STEP=0x0002 Wait 20 sec STEP=0xffff For (i=0…8) 0x84200000+i*10003 STEP=0x0003 Wait 20 sec STEP=0xffff For (i=0…8) 0x84200000+i*10006 STEP=0x0004 Wait 20 sec STEP=0xffff For (i=0…8) 0x84200000+i*10009 STEP=0x0005 Wait 20 sec STEP=0xffff For (i=0…8) 0x84200000+i*1000c STEP=0x0006 Wait 20 sec STEP=0xffff For (i=0…8) 0x84200000+i*1000f STEP=0x0007 Wait 20 sec STEP=0xffff For (i=0…8) 0x84200000+i*10000 STEP=0x0008 Wait 20 sec STEP=0xffff |  |
| 4 | Write in a file signal recorded |  |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 6 | Offline analysis | Verify that detector signal follows the offsets Derive noise for each setting? Analyse noise Vs offset value? |  |

##### Success/Failure Criteria: Test passed if the analysis shows relevant results

Comment/Open issue:

Require different scripts for CQM and PFM. Is it the way to test it (same voltage on all channel at the same time)? Or should we do it channel par channel?

## FUNC-DCU-05S, DCU Spectrometer Offset test

| ID: | FUNC-DCU-05S |
| --- | --- |
| Purpose | Characterising the detector output for various values of offset, on the spectrometer side |
| Description of test: | Change offsets up and down, record output and analyse off line. |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode + spectrometer LIAs cards on |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | FUNC-DCU-04S successful |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8817, STEP=0x0001 SetDataMode(0x0004) 843C001C SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA script FUNC-DCU-05S Look at the first detector of each channel to verify the signal is changing Record detector signals |  |  |
| 3 | Set spectrometer LIAs offsets Set offset to 0V Set STEP to 0x0002 Record output during 20 sec Set STEP to 0xffff Set offset to 1V Set STEP to 0x0003 Record output during 20 sec Set STEP to 0xffff Set offset to 2V Set STEP to 0x0004 Record output during 20 sec Set STEP to 0xffff Set offset to 3V Set STEP to 0x0005 Record output during 20 sec Set STEP to 0xffff Set offset to 4V Set STEP to 0x0006 Record output during 20 sec Set STEP to 0xffff Set offset to 5V Set STEP to 0x0007 Record output during 20 sec Set STEP to 0xffff Reset offset to 0V Set STEP to 0x0008 Record output during 20 sec Set STEP to 0xffff | For (i=0…2) 0x842c0000+i*10000 STEP=0x0002 Wait 20 sec STEP=0xffff For (i=0…2) 0x842c0000+i*10003 STEP=0x0003 Wait 20 sec STEP=0xffff For (i=0…2) 0x842c0000+i*10006 STEP=0x0004 Wait 20 sec STEP=0xffff For (i=0…2) 0x842c0000+i*10009 STEP=0x0005 Wait 20 sec STEP=0xffff For (i=0…2) 0x842c0000+i*1000c STEP=0x0006 Wait 20 sec STEP=0xffff For (i=0…2) 0x842c0000+i*1000f STEP=0x0007 Wait 20 sec STEP=0xffff For (i=0…2) 0x842c0000+i*10000 STEP=0x0008 Wait 20 sec STEP=0xffff |  |
| 4 | Write in a file signal recorded |  |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 6 | Offline analysis | Verify that detector signal follows the offsets Derive noise for each setting? Analyse noise Vs offset value? |  |

##### Success/Failure Criteria: Test passed if the analysis shows relevant results

Comment/Open issue:

Require different scripts for CQM and PFM. Is it the way to test it (same voltage on all channel at the same time)? Or should we do it channel par channel?

## FUNC-DCU-06P, DCU Photometer JFET heaters

| ID: | FUNC-DCU-06P |
| --- | --- |
| Purpose | Characterising influence of JFET heaters on detectors output, for the photometer |
| Description of test: | Switch on photometer JFET heaters. Record detector signal, derive noise |
| Test Type: | Characterisation |
| Instrument model | STM/CQM |
| Redundancy: | No |
| Initial configuration | SPIRE in REDY mode – Photometer JFET heaters off |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 3 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8818, STEP=0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-06P Look at detector signals Record detector signals Display parameter | PHOTHTRV |  |
| 3 | apply power to photometer JFET heaters Set bias voltage to 0V Record output for 20 sec Set bias voltage to -1V Record output for 20 sec Set bias voltage to -2V Record output for 20 sec Set bias voltage to -3V Record output for 20 sec Set bias voltage to -4V Record output for 20 sec Set bias voltage to -5V Record output for 20 sec Set voltage down to 0 | SetPhotoHeaterBias(0x0000) 84110000, then STEP=0x0002 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0033) 84110033, then STEP =0x0003 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0066) 84110066, then STEP=0x0004 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0099) 84110099, then STEP=0x0005 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x00CC) 841100CC, then STEP=0x0006 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x00FF) 841100FF, then STEP=0x0007 Wait 20 sec then STEP=0xffff SetPhotoHeaterBias(0x0000) 84110000, then STEP=0x0008 Wait 20 sec then STEP=0xffff |  |
| 4 | Write in a file the signal recorded |  |  |
| 5 | Offline analysis | Derive noise for each setting Analyse noise Vs voltage |  |

##### Success/Failure Criteria: Test passed is bias voltages are set as required

Comment/Open issue:

Require different scripts for CQM and PFM

A Test Facility thermometer might be checked as well (TBC)

There should be thermometers on the instrument to check as well? But which ones?

## FUNC-DCU-06S, DCU Spectrometer JFET heaters

| ID: | FUNC-DCU-06S |
| --- | --- |
| Purpose | Characterising influence of JFET heaters on detectors output, for the spectrometer |
| Description of test: | Switch on Spectrometer JFET heaters. Record detector signal, derive noise |
| Test Type: | Characterisation |
| Instrument model | STM/CQM |
| Redundancy: | No |
| Initial configuration | SPIRE in REDY – Spectrometer JFET heaters off |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Cold |
| Total duration | 3 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE = 0x8819, STEP=0x0001 SetDataMode(0x0000) 843C0004 SetSpectroBiasFreq(0x062) 84390062 SetSpectroSampFreq(0x000C) 8438000C SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-06S Look at detector signals Record detector signals Display parameter | SPECHTRV |  |
| 3 | Apply power to spectrometer JFET heaters Set bias voltage to 0V Record output for 20 sec Set bias voltage to -1V Record output for 20 sec Set bias voltage to -2V Record output for 20 sec Set bias voltage to -3V Record output for 20 sec Set bias voltage to -4V Record output for 20 sec Set bias voltage to -5V Record output for 20 sec Set voltage down to 0 | SetSpectroHeaterBias(0x0000) 84330000, then STEP=0x0002 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0033) 84330033, then STEP =0x0003 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0066) 84330066, then STEP=0x0004 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0099) 84330099, then STEP=0x0005 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x00CC)843300CC then STEP=0x0006 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x00FF)843300FF, then STEP=0x0007 Wait 20 sec then STEP=0xffff SetSpectroHeaterBias(0x0000) 84330000, then STEP=0x0008 Wait 20 sec then STEP=0xffff |  |
| 4 | Write in a file the signal recorded |  |  |
| 5 | Offline analysis | Derive noise for each setting Analyse noise Vs voltage |  |

##### Success/Failure Criteria: Test passed is bias voltages are set as required

##### Comment/Open issue:

Require different scripts for CQM and PFM

A Test Facility thermometer might be checked as well (TBC)

There should be thermometers on the instrument to check as well? But which ones?

## FUNC-DCU-07P, DCU Photometer JFET test

| ID: | FUNC-DCU-07P |
| --- | --- |
| Purpose | Characterising detector signal for various photometer JFET drain voltages |
| Description of test: | Switch on Photometer JFETs with Vdd on. Activate auto-offset control. Measure noise. Vary Vss and measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode – Photometer JFETs off |
| Final configuration | SPIRE in REDY mode – Photometer JFETs on and warm |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3-4 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | a) BBTYPE=0x881A , STEP=0x0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-07P Look at detector signal Record detector signal Display parameters | PSWJFETSTAT, PMLWJFETSTAT, PSWJFET#V (#=1..6), PMWJFET#V (#=1..4), PLWJFET#V (#=1..2), TCJFETV |  |
| 3 | Switch on photometer JFETs Switch drain voltages to high level Switch on JFET heaters to half power during 1 min This setting has been changed to full poswer during one minute Switch off heaters Activate auto-offset control | SetPhSWJfetPwr(0x003F) 8412003F SetPhMLWJfetPwr(0x007F) 8413007F SetPhotoHeaterBias(0x007F) 8411007F Wait 1 min 841100FF SetPhotoHeaterBias(0x0000) 84110000 SetDataMode(0x0010) 843C0010 Wait xx sec |  |
| 4 | Vary Vss and record signal Set Vss to 0V on all channel Set STEP 0x0002 Record signal during 20 sec Set STEP 0xffff Set Vss to -1V on all channel Set STEP 0x0003 Record signal during 20 sec Set STEP 0xffff Set Vss to -2V on all channel Set STEP 0x0004 Record signal during 20 sec Set STEP 0xffff Set Vss to -3V on all channel Set STEP 0x0005 Record signal during 20 sec Set STEP 0xffff Set Vss to -4V on all channel Set STEP =0x0006 Record signal during 20 sec Set STEP =0xffff Set Vss to -5V on all channel Set STEP =0x0007 Record signal during 20 sec Set STEP =0xffff | For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 / For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 / SetPhTCJfetVSS1(0x0000) 8414000 Set STEP =0x0002 Wait 20 sec Set STEP =0xffff For #=1..6, SetPhSWJfetVSS#(0x0033) 8(404+#)0033 / For #=1..4, SetPhMWJfetVSS#(0x0033) 8(40A+#)0033 For #=1..2, SetPhLWJfetVSS#(0x0033) 8(40E+#)0033 / SetPhTCJfetVSS1(0x0033) 84140033 Set STEP =0x0003 Wait 20 sec Set STEP =0xffff For #=1..6, SetPhSWJfetVSS#(0x0066) 8(404+#)0066 / For #=1..4, SetPhMWJfetVSS#(0x0066) 8(40A+#)0066 For #=1..2, SetPhLWJfetVSS#(0x0066) 8(40E+#)0066 / SetPhTCJfetVSS1(0x0066) 84140066 Set STEP =0x0004 Wait 20 sec Set STEP =0xffff For #=1..6, SetPhSWJfetVSS#(0x0099) 8(404+#)0099 / For #=1..4, SetPhMWJfetVSS#(0x0099) 8(40A+#)0099 For #=1..2, SetPhLWJfetVSS#(0x0099) 8(40E+#)0099 / SetPhTCJfetVSS1(0x0099) 84140099 Set STEP =0x0005 Wait 20 sec Set STEP =0xffff For #=1..6, SetPhSWJfetVSS#(0x00CC) 8(404+#)00CC / For #=1..4, SetPhMWJfetVSS#(0x00CC) 8(40A+#)00CC For #=1..2, SetPhLWJfetVSS#(0x00CC) 8(40E+#)00CC / SetPhTCJfetVSS1(0x00CC) 841400CC Set STEP =0x0006 Wait 20 sec Set STEP =0xffff For #=1..6, SetPhSWJfetVSS#(0x00FF) 8(404+#)00FF / For #=1..4, SetPhMWJfetVSS#(0x00FF) 8(40A+#)00FF For #=1..2, SetPhLWJfetVSS#(0x00FF) 8(40E+#)00FF / SetPhTCJfetVSS1(0x00FF) 841400FF Set STEP =0x0007 Wait 20 sec STEP =0xffff |  |
| 5 | Set Vss back to 0 Set STEP =0x0008 Record signal during 20 sec aa) Set STEP =0xffff | For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 SetPhTCJfetVSS1(0x007F) 84140000 Set STEP =0x0008 Record signal during 20 sec Set STEP =0xffff |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each value of Vss Analyse noise vs Vss |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set Vss to 0V on all channel Switch off drain voltages Switch off LIA cards | For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 / For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 / SetPhTCJfetVSS1(0x0000) 8414000 SetPhSWJfetPwr(0x0000) 84120000 SetPhMLWJfetPwr(0x0000) 84130000 SetDRelOnOff(0x0000) A087000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

##### Success/Failure Criteria: Test passed if each voltage is set as required

Comment/Open issue:

All JFETs are tested together

Refer also to performance test ILT-PERF-DAN

Require different scripts for CQM and PFM

Is that the way to turn on JFETs (drain then source)? Does the order matter?

## FUNC-DCU-07S, DCU Spectrometer JFET test

| ID: | FUNC-DCU-07S |
| --- | --- |
| Purpose | Characterising detector signal for various spectrometer JFET drain voltages |
| Description of test: | Switch on spectrometer JFETS with Vdd on. Activate auto-offset control. Measure noise. Vary Vss and measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | SPIRE in REDY mode – spectrometer JFETs off |
| Final configuration | SPIRE in REDY mode – spectrometer JFETs on and warm |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3-4 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a) BBTYPE=0x881B, STEP=0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 2 | Run QLA FUNC-DCU-07S Look at detector signal Record detector signal Display parameters | SPECJFETSTAT, SSWJFET#V (#=1..2), SLWJFETV |  |
| 3 | Switch on spectrometer JFETs Switch drain voltages to high level Switch on JFET heaters to half power during 1 min Switch off heaters Activate auto-offset control | SetSpSLWJfetPwr(0x0007) 84370007 SetPhotoHeaterBias(0x007F) 8433007F Wait 1 min SetPhotoHeaterBias(0x0000) 84330000 SetDataMode(0x0014) 843C0014 Wait xx sec |  |
| 4 | Vary Vss and record signal (Same step setting as previous test) Set Vss to 0V on all channel Record signal during 20 sec Set Vss to -1V on all channel Record signal during 20 sec Set Vss to -2V on all channel Record signal during 20 sec Set Vss to -3V on all channel Record signal during 20 sec Set Vss to -4V on all channel Record signal during 20 sec Set Vss to -5V on all channel Record signal during 20 sec | For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x0033) 8(434+#)0033 SetSpLWJfetVSS(0x0033) 84340033 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x0033) 8(434+#)0066 SetSpLWJfetVSS(0x0033) 84340066 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x0099) 8(434+#)0099 SetSpLWJfetVSS(0x0099) 84340099 Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x00CC) 8(434+#)00CC SetSpLWJfetVSS(0x00CC) 843400CC Wait 20 sec For #=1..2, SetSpSWJfetVSS#(0x00FF) 8(434+#)00FF SetSpLWJfetVSS(0x00FF) 843400FF Wait 20 sec |  |
| 5 | Write in a file the data recorded |  |  |
| 6 | Set Vss back to 0 | For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each value of Vss Analyse noise vs Vss |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set Vss to 0V on all channel Switch off drain voltages Switch off LIA cards | For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 SetSpSLWJfetPwr(0x0000) 84370000 SetDRelOnOff(0x0000) A087000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

##### Success/Failure Criteria: Test passed if each voltage is set as required

Comment/Open issue:

All JFETs are tested together

Refer also to performance test ILT-PERF-DAN

Require different scripts for CQM and PFM

Is that the way to turn on JFETs (drain then source)? Does the order matter?

## FUNC-DCU-08P_full, DCU Full photometer phase shift test

| ID: | FUNC-DCU-08P_full |
| --- | --- |
| Purpose | Adjusting the demodulation phase |
| Description of test: | Set phase to 0 degrees + duntill maximum signal is found |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in PHOT-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in PHOT-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 5 min |
| Constraints: | Photometer detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometers detectors |  |
| 2 | a) Mark beginning of test Set_BBID(0x881C0001) With SCOS, request bolometer array science packet at 15.3Hz | a) Set BBID =0x881C0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08P_full Look at detector signal Record detector signal Display parameter | PSWPHASE, PMWPHASE, PLWPHASE, TCPHASE |  |
| 4 | a) Set demodulation phase to : ( Δ di where  init_phase test procedure input parameter ΔRAW delta phase test procedure input parameter dRAW step phase test procedure input parameter i = 0 … (2* Δd b) Set automatic offsets c) Set Step 1+i d) Wait 10 seconds e) Set STEP to 0xffff f) Repeat steps a) to e) until the list of phases is | SetPhotoDemodSW(0x0040) 841AΔ0000+ d SetPhotoDemodMW(0x0040) 841B0000+ d SetPhotoDemodLW(0x0040) 841C0000+ d SetPhotoDemodTC(0x0040) 841D0000+ d SetAutomaticOffsets Set _OBS_STEP(0x0001) Wait 10 seconds Set_OBS_STEP(0xffff) Repeat |  |
| 6 | Write in file data recorded and value of d |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

Require different scripts for CQM and PFM

## FUNC-DCU-08P_short, DCU Short Photometer phase shift test

| ID: | FUNC-DCU-08P_short |
| --- | --- |
| Purpose | Adjusting the demodulation phase? |
| Description of test: | Set phase to 0 degrees + duntill maximum signal is found. Same as previous but these one is used to refine the search of the peak phase found roughly with the previous test |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in PHOT-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in PHOT-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | Photometer detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | comment |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometers detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of test | a) BBID =0x881C0001 SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08P_short Look at detector signal Record detector signal Display parameter | PSWPHASE, PMWPHASE, PLWPHASE, TCPHASE |  |
| 4 | a) Set demodulation phase to : ( Δ di where  init_phase test procedure input parameter ΔRAW delta phase test procedure input parameter dRAW step phase test procedure input parameter i = 0 … (2* Δd b) Set automatic offsets c) Set Step 1+i d) Wait 10 seconds e) Set STEP to 0xffff f) Repeat steps a) to e) until the list of phases is | SetPhotoDemodSW(0x0040) 841AΔ0000+ d SetPhotoDemodMW(0x0040) 841B0000+ d SetPhotoDemodLW(0x0040) 841C0000+ d SetPhotoDemodTC(0x0040) 841D0000+ d SetAutomaticOffsets Set _OBS_STEP(0x0001) Wait 10 seconds Set_OBS_STEP(0xffff) Repeat |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

The value of dis found by FUNC-DCU-08P_full

Require different scripts for CQM and PFM

## FUNC-DCU-08S_full, DCU Full spectrometer Phase shift test

| ID: | FUNC-DCU-08S_full |
| --- | --- |
| Purpose | Adjusting the demodulation phase? |
| Description of test: | Set phase to 0 degrees + duntill maximum signal is found |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in SPEC-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in SPEC-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 5 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of test | a) BBID=0x881D0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08S_full Look at detector signal Record detector signal Display parameter | SSWPHASE, SLWPHASE |  |
| 4 | a) Set demodulation phase to : ( Δ di where  init_phase test procedure input parameter ΔRAW delta phase test procedure input parameter dRAW step phase test procedure input parameter i = 0 … (2* Δd b) Set automatic offsets c) Set Step 1+i d) Wait 10 seconds e) Set STEP to 0xffff f) Repeat steps a) to e) until the list of phases is | SetSpectroDemodSW(0x0040) 843A0000+ d SetSpectroDemodMW(0x0040) 843B0000+ d SetAutomaticOffsets Set _OBS_STEP(0x0001) Wait 10 seconds Set_OBS_STEP(0xffff) Repeat |  |
| 6 | Write in file data recorded and value of d |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

Require different scripts for CQM and PFM

## FUNC-DCU-08S_short, DCU Short spectrometer Phase shift test

| ID: | FUNC-DCU-08S_short |
| --- | --- |
| Purpose | Adjusting the demodulation phase? |
| Description of test: | Set phase to 0 degrees + duntill maximum signal is found. Same as previous but these one is used to refine the search of the peak phase found roughly with the previous test |
| Test Type: |  |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | As applicable |
| Initial configuration: | SPIRE in SPEC-STBY mode – detectors on – demodulation phase not adjusted |
| Final configuration | SPIRE in SPEC-STBY mode – detectors on – demodulation phase adjusted |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 80Hz a) Mark the beginning of test | a) BBID=0x881D0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-08S_short Look at detector signal Record detector signal Display parameter | SSWPHASE, SLWPHASE |  |
| 4 | a) Set demodulation phase to : ( Δ di where  init_phase test procedure input parameter ΔRAW delta phase test procedure input parameter dRAW step phase test procedure input parameter i = 0 … (2* Δd b) Set automatic offsets c) Set Step 1+i d) Wait 10 seconds e) Set STEP to 0xffff f) Repeat steps a) to e) until the list of phases is | SetSpectroDemodSW(0x0040) 843A0000+ d SetSpectroDemodMW(0x0040) 843B0000+ d SetAutomaticOffsets Set _OBS_STEP(0x0001) Wait 10 seconds Set_OBS_STEP(0xffff) Repeat |  |
| 5 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |

Success/Failure Criteria: Test passed if demodulation phase is properly set to each value and detector signal is as expected.

##### Comment/Open issue

Require different scripts for CQM and PFM

## FUNC-DCU-09P, DCU Photometer bias frequency test

| ID: | FUNC-DCU-09P |
| --- | --- |
| Purpose | Characterising detector signal for various bias frequencies on photometer side |
| Description of test: | Set up bias frequency. Measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in PHOT-STBY mode |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometers detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz a)Mark beginning of data bblock | a) SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-09P Look at detector signal Record detector signal Display parameters | PHOTOBIASFREQ |  |
| 4 | Vary bias frequency Mark beginning of BiasFreq block Set bias frequency at 55Hz Mark beginning of Phase Shift block Execute phase shift full procedure. Mark beginning of BiasFreq block Set bias frequency at 130Hz Mark beginning of Phase Shift block Execute phase shift full procedure. Mark beginning of BiasFreq block Set bias frequency at 190Hz Mark beginning of Phase Shift block Execute phase shift full procedure. | BBTYPE=0x88200001,STEP=0x0001 SetPhotoBiasFreq(0x01FF) 84190163 BBTYPE=0x881C0001,STEP=0x0001 BBTYPE=0x88200002,STEP=0x0001 SetPhotoBiasFreq(0x00C3) 84190096 BBTYPE=0x881C0002,STEP=0x0001 BBTYPE=0x88200003,STEP=0x0001 SetPhotoBiasFreq(0x007A) 84190066 BBTYPE=0x881C0003,STEP=0x0001 |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each frequency Analyse noise vs frequency |  |

Success/Failure Criteria: Test passed if bias frequency is set to each value as required

Comment/Open issue:

Require different scripts for CQM and PFM

## FUNC-DCU-09S, DCU Spectrometer bias frequency test

| ID: | FUNC-DCU-09S |
| --- | --- |
| Purpose | Characterising detector signal for various bias frequencies on spectrometer side |
| Description of test: | Set up bias frequency. Measure noise |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in SPEC-STBY |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 80Hz a) Mark beginning of data bblock | a)BBTYPE=0x8821,STEP=0x0001 SetDataMode(0x0004) 843C0004 SetSpectroBiasFreq(0x0062) 8439007A SetSpectroSampFreq(0x000C) 84380001 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA FUNC-DCU-09S Look at detector signal Record detector signal Display parameters | SPECBIASFREQ |  |
| 4 | Vary bias frequency with same step setting as the photometer one Set bias frequency at 55Hz Execute phase shift test (FUNC-DCU-08S_ full) Record detector signal during 20 sec Set bias frequency at 130Hz Execute phase shift test (FUNC-DCU-08S_ full) Record detector signal during 20 sec Set bias frequency at 190Hz Execute phase shift test (FUNC-DCU-08S_ full) Record detector signal during 20 sec | SetSpectroBiasFreq(0x01FF) 84390163 Wait 20 sec SetSpectroBiasFreq(0x00C3) 84390096 Wait 20 sec SetSpectroBiasFreq(0x007A) 84390066 Wait 20 sec |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Derive noise for each frequency Analyse noise vs frequency |  |

Success/Failure Criteria: Test passed if bias frequency is set to each value as required

Comment/Open issue:

Require different scripts for CQM and PFM

## FUNC-DCU-10P, DCU Photometer bias amplitude test

| ID: | FUNC-DCU-10P |
| --- | --- |
| Purpose | Characterising detector signal for various bias amplitude, on photometer side |
| Description of test: | Set bias amplitude. Measure change in signal level, measure noise, read out offsets |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in PHOT-STBY |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on photometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz a) Mark beginning of data bblock | SetDataMode(0x0000) 843C0000 SetPhotoBiasFreq(0x062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetStarFrame(0x0001) 843E0001 |  |
| 3 | Run QLA script FUNC-DCU-10P Look at detector signal Record detector signal Display parameters | PSWBIAS, PLWBIAS, PMWBIAS, TCBIAS |  |
| 4 | Vary bias amplitude a) Mark beginning of bias ampl bblock Set bias amplitude to 0mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 40mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 80mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 120mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 160mV Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff Mark beginning of bias ampl bblock Set bias amplitude to 200m Mark beginning of DCU offset bblock Run set offset procedure Mark beginning of data bblock Record signal during 20 sec then step STEP 0xffff | a) BBID=0x88400001,STEP=0x0001 SetPhotoBiasAmplSW(0x0000) 841A0000; SetPhotoBiasAmplMW(0x0000) 841B0000; SetPhotoBiasAmplLW(0x0000) 841C0000; SetPhotoBiasAmplTC(0x0000) 841D0000 BBID=0x881E0001,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040001,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88400002,STEP=0x0001 SetPhotoBiasAmplSW(0x0033) 841A0033; SetPhotoBiasAmplMW(0x0033) 841B0033; SetPhotoBiasAmplLW(0x0033) 841C0033; SetPhotoBiasAmplTC(0x0033) 841D0033 BBID=0x881E0002,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040002,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88400003,STEP=0x0001 SetPhotoBiasAmplSW(0x0066) 841A0066; SetPhotoBiasAmplMW(0x0066) 841B0066 SetPhotoBiasAmplLW(0x0066) 841C0066; SetPhotoBiasAmplTC(0x0066) 841D0066 BBID=0x881E0003,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040003,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88400004,STEP=0x0001 SetPhotoBiasAmplSW(0x0099) 841A0099; SetPhotoBiasAmplMW(0x0099) 841B0099 SetPhotoBiasAmplLW(0x0099) 841C0099; SetPhotoBiasAmplTC(0x0099) 841D0099 BBID=0x881E0004,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040004,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88400005,STEP=0x0001 SetPhotoBiasAmplSW(0x00CC) 841A00CC; SetPhotoBiasAmplMW(0x00CC) 841B00CC SetPhotoBiasAmplLW(0x00CC) 841C00CC; SetPhotoBiasAmplTC(0x00CC) 841D00CC BBID=0x881E0005,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040005,STEP=0x0001 Wait 20 sec then STEP=0xffff BBID=0x88400006,STEP=0x0001 SetPhotoBiasAmplSW(0x00FF) 841A00FF; SetPhotoBiasAmplMW(0x00FF) 841B00FF SetPhotoBiasAmplLW(0x00FF) 841C00FF; SetPhotoBiasAmplTC(0x00FF) 841D00FF BBID=0x881E0006,STEP=0x0001 SetDataMode(0x0010) 843C0010 Wait 10 sec BBID=0x88040006,STEP=0x0001 Wait 20 sec then STEP=0xffff |  |
| 5 | Set Bias amplitude back to (0mV) | SetPhotoBiasAmplSW(0x0000) 841A0000 SetPhotoBiasAmplMW(0x0000) 841B0000 SetPhotoBiasAmplLW(0x0000) 841C0000 SetPhotoBiasAmplTC(0x0000) 841D0000 |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Evaluate change in signal level and derive noise for each value of bias amplitude Analyse signal level and noise Vs bias amplitude |  |

Success/Failure Criteria: Test passed if bias amplitude is set to each value as required and offset procedure works each time it is called.

Comment/Open issue:

How can we check that the offset procedure has done the job properly?

Require different scripts for CQM and PFM

What is the set offset procedure?

## FUNC-DCU-10S, DCU Spectrometer bias amplitude test

| ID: | FUNC-DCU-10S |
| --- | --- |
| Purpose | Characterising detector signal for various bias amplitude on spectrometer side |
| Description of test: | Set bias amplitude. Measure change in signal level, measure noise, read out offsets |
| Test Type: | Characterisation |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | prime and redundant |
| Initial configuration: | SPIRE in SPEC-STBY |
| Final configuration | Idem |
| EGSE Configuration: | SCOS-2000 and QLA required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 3 min |
| Constraints: | Detectors on (bias frequency, bias amplitude, phase are set to operating values) – see FUNC-DCU-11 |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | If not already switch on detectors with operating values | Execute test FUNC-DCU-11 to switch on spectrometer detectors |  |
| 2 | With SCOS, request bolometer array science packet at 15.3Hz Same layout as previous test but with Spectrometer Bias Amplitud building block type 0x8841 And spec data bblock 0x8801 if Specfull , or 0x8805 if SSW or 0x8806 if SLW | SetDataMode(0x0004) 843C0004 SetStartFrame(0x0001) 843E0001 |  |
| 3 | Run QLA script FUNC-DCU-10S Look at detector signal Record detector signal Display parameters | SSWBIAS, SLWBIAS |  |
| 4 | Vary bias amplitude Set bias amplitude to 0mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 40mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 80mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 120mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 160mV Run set offset procedure Record detector signal during 20 sec Set bias amplitude to 200mV Run set offset procedure Record detector signal during 20 sec | SetSpectroBiasAmplSW(0x0000) 843A0000 SetSpectroBiasAmplLW(0x0000) 843B0000 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x0033) 843A0033 SetSpectroBiasAmplLW(0x0033) 843B0033 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x0066) 843A0066 SetSpectroBiasAmplLW(0x0066) 843B0066 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x0099) 843A0099 SetSpectroBiasAmplLW(0x0099) 843B0099 SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x00CC) 843A00CC SetSpectroBiasAmplLW(0x00CC) 843B00CC SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec SetSpectroBiasAmplSW(0x00FF) 843A00FF SetSpectroBiasAmplLW(0x00FF) 843B00FF SetDataMode(0x0014) 843C0014 Wait 10 sec Wait 20 sec |  |
| 5 | Set Bias amplitude back to (0mV) | SetSpectroBiasAmplSW(0x0000) 843A0000 SetSpectroBiasAmplLW(0x0000) 843B0000 |  |
| 6 | Write in a file the data recorded |  |  |
| 7 | Stop TM packet sampling | SetStartFrame(0x0000) 843E0000 |  |
| 8 | Offline analysis | Evaluate change in signal level and derive noise for each value of bias amplitude Analyse signal level and noise Vs bias amplitude |  |

Success/Failure Criteria: Test passed if bias amplitude is set to each value as required and offset procedure works each time it is called.

Comment/Open issue:

How can we check that the offset procedure has done the job properly?

Require different scripts for CQM and PFM

What is the set offset procedure?

## FUNC-DCU-11P, DCU Photometer detector switch on

| ID: | FUNC-DCU-11P |
| --- | --- |
| Purpose | Switching on the all chain for reading detectors signal. This procedure must be performed before test FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10 – see flowchart p17 |
| Description of test: | Switch on photometer LIAs and JFETs, set sampling frequency, bias amplitude, bias frequency, demodulation phase to their operating values |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | detectors off |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Switch on photometer LIAs card Check HK parameters Switch on JFET source voltages to level found in test FUNC-DCU-07P Switch drain voltages to high level Switch on JFET heaters to max power during 1 min Switch off JFET heaters Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set photo bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A087005 if MCU is on PWR_STATUS = 0x01FF LIAP_P5 = 0xB554 LIAP_P9 = 0xE00D LIAP_N9 = 0x1FF0 For #=1..6, SetPhSWJfetVSS#(0x00xx) 8(404+#)007F For #=1..4, SetPhMWJfetVSS#(0x00xx)8(40A+#)00xxFor #=1..2, SetPhLWJfetVSS#(0x00xx) 8(40E+#)00xx SetPhTCJfetVSS1(0x00xx) 841400xx SetPhSWJfetPwr(0x003F) 8412003F SetPhMLWJfetPwr(0x007F) 8413007F SetPhotoHeaterBias(0x007F) 841100FF Wait 1 min SetPhotoHeaterBias(0x0000) 84110000 SetPhotoBiasFreq(0x0062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetPhotoBiasMode(0x00FF) 840000FF SetPhotoBiasAmplSW(0x004D) 841A004D SetPhotoBiasAmplMW(0x004D) 841B004D SetPhotoBiasAmplLW(0x004D) 841C004D SetPhotoBiasAmplTC(0x004D) 841D004D SetPhotoDemodSW(0x007F) 841A007F SetPhotoDemodMW(0x007F) 841B007F SetPhotoDemodLW(0x007F) 841C007F SetPhotoDemodTC(0x007F) 841D007F SetDataMode(0x0010) 843C0010 Wait 10 sec |  |

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set photo bias mode to “off” Set Vss to 0V on all channel Switch off drain voltages Switch of LIA card | SetPhotoBiasMode(0x0000) 84000000 For #=1..6, SetPhSWJfetVSS#(0x0000) 8(404+#)0000 / For #=1..4, SetPhMWJfetVSS#(0x0000) 8(40A+#)0000 For #=1..2, SetPhLWJfetVSS#(0x0000) 8(40E+#)0000 / SetPhTCJfetVSS1(0x0000) 8414000 SetPhSWJfetPwr(0x0000) 84120000 SetPhMLWJfetPwr(0x0000) 84130000 SetDRelOnOff(0x0000) A0870000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)

## FUNC-DCU-11S, DCU Spectrometer detector switch on

| ID: | FUNC-DCU-11S |
| --- | --- |
| Purpose | Switching on the all chain for reading detectors signal. This procedure must be performed before test FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10 – see flowchart p17 |
| Description of test: | Switch on spectrometer LIAs and JFETs, set sampling frequency, bias amplitude, bias frequency, demodulation phase to their operating values |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | detectors off |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | None |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Switch on spectrometer LIAs card Check HK parameters Switch on JFET source voltages to level found in FUNC-DCU-07S Switch drain voltages to high level Switch on JFET heaters to max power during 1 min Switch off JFET heaters Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set spectro bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetDRelOnOff(0x0002) A0870002 or SetDRelOnOff(0x0006) A087006 if MCU is on PWR_STATUS = 0x01FF LIAP_P5 = 0xB554 LIAP_P9 = 0xE00D LIAP_N9 = 0x1FF0 For #=1..2, SetSpSWJfetVSS#(0x00xx) 8(434+#)00xx SetSpLWJfetVSS(0x00xx) 843400xx SetSpSLWJfetPwr(0x0007) 84370007 SetSpectroHeaterBias(0x007F) 843300FF Wait 1 min SetSpectroHeaterBias(0x0000) 84330000 SetSpectroBiasFreq(0x0062) 84390062 SetSpectroSampFreq(0x000C) 8438000C SetSpectroBiasMode(0x00FF) 843000FF SetSpectroBiasAmplSW(0x004D) 843A004D SetSpectroBiasAmplLW(0x004D) 843B004D SetSpectroDemodSW(0x007F) 843A007F SetSpectroDemodMW(0x007F) 843B007F SetDataMode(0x0014) 843C0014 Wait 10 sec |  |

- Note: for QM1 the command to switch on LIA cards is SetDRelOnOff(0x0001) A0870001 or SetDRelOnOff(0x0005) A087005 if MCU is on

##### Switch off procedure

| Action | Command |
| --- | --- |
| Set photo bias mode to “off” Set Vss to 0V on all channel Switch off drain voltages Switch of LIA card | SetSpectroBiasMode(0x0000) 84300000 For #=1..2, SetSpSWJfetVSS#(0x0000) 8(434+#)0000 SetSpLWJfetVSS(0x0000) 84340000 SetSpSLWJfetPwr(0x0000) 84370000 SetDRelOnOff(0x0000) A0870000 or SetDRelOnOff(0x0004) A087004 if MCU is on |

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)

## FUNC-DCU-12P, DCU Photometer detector settings

| ID: | FUNC-DCU-12P |
| --- | --- |
| Purpose | Setting sampling frequency, bias frequency, bias amplitude and demodulation phase when LIA and JFET are already on. This procedure must be performed after FUNC-DCU-04 and FUNC-DCU-07 before carrying on with tests FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10. – see flowchart p17 |
| Description of test: | Set sampling frequency, bias amplitude, bias frequency, demodulation phase on photometer side to their operating values. Activate auto-offset procedure |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | LIA, JFET on. |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | FUNC-DCU-04 and FUNC-DCU-07 successfully passd |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set photo bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetPhotoBiasFreq(0x0062) 84190062 SetPhotoSampFreq(0x000C) 8418000C SetPhotoBiasMode(0x00FF) 840000FF SetPhotoBiasAmplSW(0x004D) 8401004D SetPhotoBiasAmplMW(0x004D) 8402004D SetPhotoBiasAmplLW(0x004D) 8403004D SetPhotoBiasAmplTC(0x004D) 8404004D SetPhotoDemodSW(0x007F) 841A007F SetPhotoDemodMW(0x007F) 841B007F SetPhotoDemodLW(0x007F) 841C007F SetPhotoDemodTC(0x007F) 841D007F SetDataMode(0x0010) 843C0010 Wait xx sec |  |

Switch off procedure: see test FUNC-DCU-11P

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)

## FUNC-DCU-12S, DCU Detector settings

| ID: | FUNC-DCU-12S |
| --- | --- |
| Purpose | Setting sampling frequency, bias frequency, bias amplitude and demodulation phase when LIA and JFET are already on. This procedure must be performed after FUNC-DCU-04 and FUNC-DCU-07 before carrying on with tests FUNC-DCU-08, FUNC-DCU-09 and FUNC-DCU-10. – see flowchart p17 |
| Description of test: | Set sampling frequency, bias amplitude, bias frequency, demodulation phase on spectrometer side to their operating values. Activate auto-offset procedure |
| Test Type: | Integrity |
| Instrument Models: | AVM/CQM/PFM/FS |
| Redundancy: | No |
| Initial configuration: | LIA, JFET on. |
| Final configuration | detectors on |
| EGSE Configuration: | SCOS-2000 required |
| Level | ILT/PLT/IST/FLT |
| Test Conditions: | Warm or Cold |
| Total duration | 1 min |
| Constraints: | FUNC-DCU-04 and FUNC-DCU-07 successfully passd |

Procedure and analysis:

| Step | Action | Command/TM | Comment/notes |
| --- | --- | --- | --- |
| 1 | Set bias frequency to 200 Hz Set the sampling frequency to 15.3Hz Set spectro bias mode to “run” Set bias amplitude to 60 mV Set demodulation phase to 180 degrees (or 180+ d, found by test FUNC-DCU-08 ?) Activate auto-offset control | SetSpectroBiasFreq(0x0062) 84390062 SetSpectroSampFreq(0x000C) 8438000C SetSpectroBiasMode(0x00FF) 843000FF SetSpectroBiasAmplSW(0x004D) 843A004D SetSpectroBiasAmplLW(0x004D) 843B004D SetSpectroDemodSW(0x007F) 843A007F SetSpectroDemodMW(0x007F) 843B007F SetDataMode(0x0014) 843C0014 Wait xx sec |  |

Switch off procedure: see test FUNC-DCU-11S

Success/Failure Criteria: Test passed if each parameter is set to the required value

Comment/Open issues:

I use the values in Bruce’s document (Subsystem reaction for specification for the instrument simulator)