Work Package:	BSM                                                                    SPIRE-ATC-REP-001780

| 1. Subsystem Progress Since Project Inception |
| --- |
| Baseline design accepted at Delta-PDR, including digital electronics control by MCU. CPP parts requested. Single axis prototyping complete. DDR held. 2 axis prototype tests complete. IBDR complete. Development Model testing |
| 2. Subsystem Progress This Month |
| STM Acceptance review Board held 17.Oct.02: minor adjustments in hand, delivery due next ~1 wk. DM-2 assembly continues : alignment of chop mirror performed. LAM electronics board checked out (stand-alone). |
| 3. Problem Areas | Remedial Action |
| Full checks on repeatability/linearity in dewar to be achieved Launch Latch relocated on design and problem noted with DM-2 motor wiring. Tayco flexi-tape delayed 8wks by export paperwork Zeiss motor quotation above budget price Cash flow | high efficiency anti-reflection coating window received and tested warm. Will test further in conjunction with DM-2. Change to proposed ATC mounting in hand - will need to tie up with LAM. DM-2 motor wiring not a problem as can route away from LL Investigated JPL route but determined would not help. Agreed with PI & PPARC - extra funding to be provided. Discussions in hand with funding body : additional funds authorised |
| 4. Engineering Activities |
| Completed re-calibration of Cernox thermometer. Wiring layout and specification Discussed with RAL (D.Griffin) : use of STQ thought to be OK. Solder buckets will be OK for SPIRE.. |
| 5. Design Changes |
| Problem routing wiring for DM-2 jiggle frame sensor is being worked on, minor design changes will result |
| 6. PA/QA Activities |
| Improvements to the drawing configuration data base are being sought. Useful discussion with RAL PA on requirements here. A number of RID responses have been worked on, and the backlog of open ECRs will be addressed also. Two 'major' NCRs raised: (1) mirror contamination on STM (traced to dew formation on warm up where Cernox thermistor indicated warm enough to open, but mirror surface remained cold). Several procedural changes introduced as a result. (2) tapped holes missing on STM for PCAL (originated as original DM-1 parts did not need to mount PCAL, and drawings were not updated) : has been rectified. |
| 7. Subsystem Management Issues |
| Funding for motors. Late delivery of flexi-tape. |
| 8. Actions Requiring Immediate Attention |
| SW11MAR02.02 IP Place contract for Zeiss motors … ongoing. |
| 9. Status of Previous Actions |
| IP/D.Griffin : ATC PO placed with Tekdata for central procurement of harness. |
| 10. Activities Yet to be Achieved |
| STM shipment has slipped a few weeks but is in hand. Overall slippage of DM-2 assembly and test : impacts liaison with LAM and possible knock on to CQM delivery for Jan03. |
| 11. Milestones | Status |
| Milestone | Baseline | Scheduled | Actual | Change |
| 20K test dewar available | N/A | Nov.02 |  | +6 months |
| Integrate DM 2 at LAM | 30 Mar 01 | Nov.02 |  |  |
| DM release for manufacture | Oct.01 | Nov.01 | Dec.01 |  |
| DM-2 release for manufacture | N/A | Apr.02 | Apr.02 |  |
| 2 axis prototyping at ATC complete | N/A | Mar.02 | Apr.02 |  |
| SPIRE IBDR | N/A | Mar.02 | Mar.02 |  |
| DM-2 tests complete | 30.Apr 02 | Nov.02 |  | + 4 weeks |
| BSM STM del'y to RAL | 01 Feb 00 | 5.Nov 02 |  | +3 weeks |
| BSM CQM delivery to LAM | Jul.02 | 20.Dec.02 |  |  |
| BSM CQM del'y to RAL | 13 Mar 02 | Feb.03 |  |  |
| BSM QM tests complete (flight design valid) | 25.Apr.03 | 20.Jun.03 |  |  |
| BSM PFM delivery to LAM | 03.Jun 03 | 04.Sep.03 |  |  |