The background detector used for the ISOLWS cryostat has been fitted to the MTD to assess the FIR background in the SPIRE cryostat.  The Spectrometer bias tail has been used to connect the existing JF4 amplifier through to the outside of the cryostat – see attached figures.

The table below is a record of the checks carried out when the detector was installed and connected to the cryoharness C1 P10 on 9 September 2003.

| Connection Name | Pin on Detector Connector | Pin on MDM on detector lead | Pin on D-Type (after cryoharness installed) | Checked end to end on cryostat (S1 J32) |
| --- | --- | --- | --- | --- |
| Reset | 1 | 17 | 17 | 140  |
| Gnd | 2 | 26 | 26 | 7  |
| VB | 3 | 15 | 15 | 146  |
| VSS | 4 | 35 | 35 | 144  |
| Thermometer | 5 | 1 | 1 | 140  |
| SigOut | 6 | 13 | 13 | 141  |
| Heater | 7 | 22 | 22 | 141  |
| VDD | 8 | 32 | 32 | 144  |
| Thermometer | 9 | 20 | 20 | 148  |
| Compensation | 10 | 10 | 10 | 144  |
| Screen | N/C | 12 | 12 | 7  |

Pin to pin checks

| Description | Pins | Measured Resistance One way round | The other way round |
| --- | --- | --- | --- |
| Heater to VDD | 22 – 32 | 6.28 k | Same |
| Sigout-VDD | 13 - 32 | 50.4 k | Same |
| Sigout - VSS | 13 - 35 | 50.3 k | Same |
| Reset-Gnd | 26-17 | 850  | O/C |

Rear view of free socket fitted to back of photo-detector block.  MDM referred to is the spectrometer bias tail on C1 P10 (see Harness definition documents  SPIRE-RAL-PRJ-000608 v1.1 section 4.4.1)