# SPIRE

| SUBJECT: | Cold vibration test procedure CQM |
| --- | --- |
| PREPARED BY: | E Sawyer |
| DOCUMENT No: | SPIRE-RAL-PRC-001956 |
| ISSUE: | 2 | Date: | 26/3/04 |
| CHECKED BY: | …………………………………. | Date: | …………………………. |
| APPROVED BY: | …………………………………. | Date: | …………………………. |

Distribution

# Change Record

| ISSUE | DATE |  |
| --- | --- | --- |
| 1 | 10/3/04 | New issue |
| 2 | 26/3/04 | Modifications noted during TRR. Times added Torques added Accelerometer locations added Temperature sensor locations added |

# Table of Contents

1.	SCOPE	7

2.	Accelerometer allocation	7

3.	Temperature sensor locations	7

4.	Step by step test sequence record	8

# Glossary

| SPIRE | Spectral and Photometric Imaging REceiver |
| --- | --- |

# References

# Applicable Documents

| No | Document | Ref |
| --- | --- | --- |
| AD1 | SPIRE cold vibration test plan | SPIRE-RAL-PRC-001955 |

Reference Documents

# SCOPE

This document describes the detailed procedure to be followed during the cold vibration testing on the FPU to be carried out at CSL, Liege.

It should be used in conjunction with AD1

# Accelerometer allocation

| Channel No | Location | Code | Axis |
| --- | --- | --- | --- |
| F66 | Photometer detector (side) | PHS.XY | X and Y |
| F72 | Photometer detector (side) | PHS.Z | Z |
| F58 | Photometer detector (end) | PHE.XY | X and Y |
| F56 | Cooler | COOL | X |
| F73 | FPU top of optics bench over cone | FPUX | X |
| F71 | FPU top | FPUY | Y |
| F63 | FPU top | FPUZ | Z |
| F51 | Spectrometer detector | SPECX | X |
| F52 | Spectrometer detector | SPECZ | Z |
| F53 | Spectrometer detector | SPECY | Y |
| F61 | Optics bench near SMEC | OBY | Y |
| F54 | SMEC moving carriage | SMECLX | X |
| 11030 | Cold tip | CTIPY | Y |
| 11606 | SMEC top | SMECUX | X |

# Temperature sensor locations

| Sensor number | Location | CODE |  |
| --- | --- | --- | --- |
| 1 | RF filter | RF |  |
| 2 | Spectrometer 2K | SPEC |  |
| 3 | Photometer 2K | PHOT |  |
| 4 | Optical sub bench | OPSB |  |
| 5 | Input baffle | BAF |  |
| 6 | BSM/SOB interface | BSM |  |

# Step by step test sequence record

| No | Activity | Time | Date | Comments | Resp. |
| --- | --- | --- | --- | --- | --- |
| 1 | Test preparation |  |  |  |  |
| 1.1 | Unpacking from container: As per AD1 Store container until the end of the test | 3 hrs | 29/3/04 |  |  |
| 1.2 | Visual inspection of the FPU Cover with cover provided while awaiting integration Accelerometers and temperature sensors will be fitted before delivery. | 2 hrs |  | Record any anomalies |  |
| 2 | Vibration Test in TBD axis |  |  |  |  |
| 2.1 | Test preparation Assuming test adaptor is fitted to the shaker Fit FPU and JFETs to test adaptor in accordance with AD1 Check torque of mounting feet 8.1 Nm plus running torque for M6 8.25 plus running torque or M8 Kaylock nut. 2.1 Nm for the JFET bolts. 2.5 nm for the L3 straps to JFETS Remove MGSE Connect accelerometers carry out continuity check. Connect temperature sensors Note These sensors do not record accurate temperatures at room temperature. Accurate only below 40K. Individual calibration curves will be supplied by SPIRE Carry out very low level test in ambient 0.2g max. Note The Kevlar suspended detectors will not tolerate significant levels when warm. This test is to verify the instrumentation only. Shrouds installation, connection of pumps and fluid lines | 3 hrs ½ hrs ½ hrs ½ hrs |  |  |  |
| 2.2 | TRR | 4 hrs | 30/304 |  |  |
| 2.2 | Evacuation Monitor pressure during evacuation and check for consistence with evacuation rate requirement in AD1 Note pressure values or attach record |  | 01/4/04 |  |  |
| 2.3 | Cool down Monitor temperature sensor during cool down and check for consistence with cool down requirements in AD1 Note Temperature or attach temperature record of temperature sensors |  | 03/04/04 |  |  |
| 2.4 | Test readiness review |  |  |  |  |
| 2.5 | TBD axis test |  |  |  |  |
| 2.6 | Resonance search Check results against those expected. Confirm levels for next test |  |  |  |  |
| 2.7 | Intermediate sine test Optional test, SPIRE to decide if intermediate level is required. |  |  |  |  |
| 2.8 | Full level sine test As per requirements in AD1 Check results |  |  |  |  |
| 2.9 | Resonance search Cary out resonance search as before Compare with previous results |  |  |  |  |
| 2.10 | Low level Random test Optional test, SPIRE to decide if low level is required. If required test to AD1 at –12dB Check results |  |  |  |  |
| 2.11 | Intermediate level Random test Optional test, SPIRE to decide if intermediate level is required. If required test to AD1 at –6dB Check results |  |  |  |  |
| 2.12 | Qualification level random test Carry out test to AD1 Check results |  |  |  |  |
| 2.13 | Resonance search Cary out resonance search as before Compare with previous results |  |  |  |  |
| 2.14 | Post test analysis Check completeness of data Carry out initial assessment |  |  |  |  |

| 3. | Test in TBD axis |  |  |  |
| --- | --- | --- | --- | --- |
| 3.2 | Evacuation if required Monitor pressure during evacuation and check for consistence with evacuation rate requirement in AD1 Note pressure values or attach record |  |  |  |
| 3.3 | Cool down if required Monitor temperature sensor during cool down and check for consistence with cool down requirements in AD1 Note Temperature or attach temperature record of temperature sensors |  |  |  |
| 3.4 | Test readiness review |  |  |  |
| 3.5 | TBD axis test |  |  |  |
| 3.6 | Resonance search Check results against those expected. Confirm levels for next test |  |  |  |
| 3.7 | Intermediate sine test Optional test, SPIRE to decide if intermediate level is required. |  |  |  |
| 3.8 | Full level sine test As per requirements in AD1 Check results |  |  |  |
| 3.9 | Resonance search Cary out resonance search as before Compare with previous results |  |  |  |
| 3.10 | Low level Random test Optional test, SPIRE to decide if low level is required. If required test to AD1 at –12dB Check results |  |  |  |
| 3.11 | Intermediate level Random test Optional test, SPIRE to decide if intermediate level is required. If required test to AD1 at –6dB Check results |  |  |  |
| 3.12 | Qualification level random test Carry out test to AD1 Check results |  |  |  |
| 3.13 | Resonance search Cary out resonance search as before Compare with previous results |  |  |  |
| 3.14 | Post test analysis Check completeness of data Carry out initial assessment |  |  |  |

| 4 | TBD axis test |  |  |  |
| --- | --- | --- | --- | --- |
| 4.2 | Evacuation if required Monitor pressure during evacuation and check for consistence with evacuation rate requirement in AD1 Note pressure values or attach record |  |  |  |
| 4.3 | Cool down if required Monitor temperature sensor during cool down and check for consistence with cool down requirements in AD1 Note Temperature or attach temperature record of temperature sensors |  |  |  |
| 4.4 | Test readiness review |  |  |  |
| 4.5 | TBD axis test |  |  |  |
| 4.6 | Resonance search Check results against those expected. Confirm levels for next test |  |  |  |
| 4.7 | Intermediate sine test Optional test, SPIRE to decide if intermediate level is required. |  |  |  |
| 4.8 | Full level sine test As per requirements in AD1 Check results |  |  |  |
| 4.9 | Resonance search Cary out resonance search as before Compare with previous results |  |  |  |
| 4.10 | Low level Random test Optional test, SPIRE to decide if low level is required. If required test to AD1 at –12dB Check results |  |  |  |
| 4.11 | Intermediate level Random test Optional test, SPIRE to decide if intermediate level is required. If required test to AD1 at –6dB Check results |  |  |  |
| 4.12 | Qualification level random test Carry out test to AD1 Check results |  |  |  |
| 4.13 | Resonance search Cary out resonance search as before Compare with previous results |  |  |  |
| 4.14 | Post test analysis Check completeness of data Carry out initial assessment |  |  |  |
| 5 | Post Test Activities • check completeness of test data • Warm up Monitor temperature sensor during warm up and check for consistence with cool down requirements in AD1 Note Temperature or attach temperature record of temperature sensors Venting Monitor pressure during venting and check for consistence with evacuation rate requirement in AD1 Note pressure values or attach record Open cryostat. dismounting of shrouds • Visual inspection of test specimen • Removal of instrument (TBC) and cleanliness inspection with UV light • Installation of the instrument inside its own container in accordance with AD1 |  |  |  |

| 6 | Post Test Review |  |  |  |
| --- | --- | --- | --- | --- |