Issue status:

Draft 27 July 2004 – Initial draft to flag problem – I missed the fact the there was some treatment of the LOU hole straylight

1.0 – 31 August 2004 – Revised version with the same conclusions but now with discussion of the limitations of modelling straylight diffusion.

The analysis presented in HP-2-ASED-TN-0023 iss 3 deals, as far as I can tell, with straylight from the LOU windows for three cases:

- The light that scatters around the gap between the cryostat inner shield and the instrument shield

- The light that gets through HIFI and scatters back from the surfaces forward of the instrument apertures.

- The emission from the LOU windows themselves that enters the instrument shield via the holes for the LOU path and arrives at the instrument aperture

The third item here is discussed in section 5.2 of the report but the results are not explicitly given for this path (see table 1) hence my confusion in the previous issue of this note.  On the basis of this analysis it seems that no baffle is provided on the inside of the instrument shield to prevent straylight from the LOU holes reaching PACS and SPIRE – this is despite our explicit request for such a shield and its inclusion in the original cryostat study – see figure 1.

We have severe reservations about accepting the proposed design based simply on analysis:

- By admission the analysis only deals with specific identified paths as the software cannot deal with large numbers of ray paths.

- Only scattering and reflection are dealt with – the effects of diffraction are not (and cannot I suspect) be dealt with.

- In the end the process light scattering and diffraction at these wavelengths and temperatures is virtually one of diffusion and very difficult to model.

- The emission from the cavity between the inner shield and the first cryostat shield (at about 34 K) is not dealt with at all in any case.

All these concerns lead us to the question: “where has the LOU inner shield baffle it gone?”  The CDR “design” no longer appears to show this baffle.  This is completely unacceptable and SPIRE rejects this design forthwith.  We indicate on the sketch provided an indication of where the baffle must be located in order to prevent unwanted radiation entering the instrument shield environment.

All our experience leads us to a “belt-and-braces” approach to preventing straylight problems.  We have done our very best to seal our instrument box against straylight, but we cannot be certain that we have no leaks: we depend on each stage of the system doing its part in baffling against unwanted radiation and that includes baffling the LOU path properly.

| Emitting object | temperature / emissivity | scattering(Sc)/ diffraction(D) on | PACS detector | SPIRE detector |
| --- | --- | --- | --- | --- |
| Sunshade | 204 K / 0.05 | Sc: M1/M2, +spec. | 2.326 | 0.611 |
| gap between Sunshade and M1 | 204 K / 0.10 | only Sc: M1 / M2 | 0.193 | 0.092 |
| gap between Sunshade and M1 | 204 K / 0.10 | only D: M2(rim) | 0.5 | 1.3 |
| Hexapod | 70 K / 0.02 |  | 3.122 | 3.977 |
| M1+M2 (without reference path) | 70 K / 0.015 |  | 2.488 | 1.873 |
| scattercone | 70 K / 0.015 |  | 0.615 | 0 |
| M1-Baffle flat | 75 K / 0.05 |  | 1.382 | 1.098 |
| M1-Baffle cone / cylinder | 75 K / 0.05 |  | 3.537 | 0.499 |
| gap (12 mm) between M1-Baffle cone and cylinder | 75 K / 0.90 |  | 1.542 | 0.330 |
| Cryocover mirrors | 75 K / 0.05 |  | 0.676 | 0.025 |
| other reflecting parts of Cryocover | 75 K / 0.05 |  | 0.069 | 0.020 |
| Cryocover black rim | 75 K / 0.80 |  | 1.749 | 0.246 |
| reflecting objects near Cryocover | 75 K / 0.05 |  | 0.462 | 0.070 |
| black gaps around Cryocover / M1-Baffle | 75 K / 0.90 |  | 2.796 | 0.414 |
| CVV top | 75 K / 0.05 |  | 1.237 | 0.077 |
| gap betw. CVV / Thermal Shield 2 Baffle | 75 K / 0.90 |  | 0.500 | 0.119 |
| Thermal Shield 2 Baffle black | 43 K / 0.80 | only specular | 1.478 | 1.391 |
| Thermal Shield 2 Baffle black | 43 K / 0.80 | only scattering in instrument. | 0.071 | 0.805 |
| Instrument Shield Baffle | 12 K / 0.05 |  | 0.002 | 0.002 |
| Gap below Instrument Shield Baffle | 12 K / 0.90 |  | 0.076 | 0.033 |
| LOU windows via HiFi | 150 K / 0.90 |  | 0.05 | 0.04 |
| LOU windows via heat shield gaps | 150 K / 0.90 |  | 0.231 | 0.020 |
| Holes in OB for cooling straps (worst case consideration only) | 34 K / 0.90 |  | 0.2 | 0.02 |
| sum |  |  | 25.3 | 13.1 |

Table 1: From the straylight report – there is no explicit result for the path from the LOU holes in the inner shield to the instrument.  Do we assume the results are negligible?  Also the straylight from the 34-K environment via the LOU holes is not treated although the cooling strap holes are.

# Text boxes

Radiation from the cryostat shield 34-K environment will also enter the LOU holes

Figure 2: Sketch of design for LOU baffling from Astrium – the instrument shield baffle has gone missing – the red lines indicate where the baffle must be placed to minimise the straylight into the instrument shield.  Also missing is any treatment from the 34-K environment indicated by the wavy line.

This path MUST be baffled to prevent straylight into instrument shield environment

Figure 1:  Extract from Cryostat Study final report from 2000 showing that industry accepted the need for an LOU baffle and that there was a notional design for it.