Procedure for connecting PFM1 FPU to Warm Electronics in Cryolab

| Prepared by: |  |  |
| --- | --- | --- |
|  | D. Griffin (RAL) | Date |
| Checked: |  |  |
|  | D.L. Smith (RAL) | Date |
| Approval: |  |  |
|  | E. Sawyer (RAL) | Date |

| Host system | Windows 2000 SP2 |
| --- | --- |
| Word Processor | Microsoft Word 2000 SR1 |
| File | Cryoharness Integration Procedure Draft 0.4.Doc |

# Document Change Record

| Date | Index | Affected Pages | Changes |
| --- | --- | --- | --- |
|  | 0.1 | All | First Draft |

# Scope of Document

This document describes the correct sequence to connect and disconnect the SPIRE FPU to the warm electronics.

NOTE: This procedure applies only to the PFM1 build which only has spectrometer detectors mounted.

# Applicable Document

|  | Title | Author | Reference | Date |
| --- | --- | --- | --- | --- |
| RD 1 | Making SPIRE ESD Safe | Doug Griffin | SPIRE-RAL-NOT-002028 Issue 2 | 28-Oct-2004 |

# Constraints

- Personnel carrying out this operation are to wear ESD wrist straps at all times when carrying out this procedure.

- Only the following harnesses are present

  - C1 – Spect. Bias and SLW (Warm and Cold)

  - C2  - SSW (Warm and Cold)

  - C10 – Prime S/S A (Warm and Cold)

  - C11 – Prime S/S B (Warm and Cold)

  - C12 – Red S/S A (Warm and Cold)

  - C4 – PMW and PTC readout (Warm only)

# Procedure

## Transport of FPU

- During transport of the FPU, the JFET Bias shorting plugs and must be in place

- The FPU harnesses are to be mated to the gates of the JFET modules

- A grounding strap shall connect the chassis of the FPU to the HOB plate

## Verify correct installation of the Cryoharnesses

Pre-requisites

- The internal-cryoharnesses are routed on the back of the HOB plate in the cryostat and any obvious contact points between the cryoharness and the chassis of the cryostat are removed.

- The internal and external Cryoharnesses are mated at the vacuum feed-through

- The FPU Faraday Shield Links are not mated

- The HOB Plate is mounted on the external MGSE with the instrument mounted in the vertical position

- The 10K and 77K shields are not installed (as this would imply that the instrument has been placed inside the cryostat

- The cryoharness is not mated with the warm units

Verify that the shield of the internal cryoharness and the chassis of the cryostat are isolated.

Checked:

Who:

When:

Comments:

## Mating Cryoharnesses to DRCU

- Check that the chasses of the DCU and FCU are bonded to local lab ground

- Route C1, C2, C4, C10 and C11 from cryostat to DRCU

- Mate J22 link between C2 and C4

- In the sequence outlined in Table 1, mate the cryoharness to the DRCU. In each case, place a conductive foil over the contacts of the harness connector immediately prior to mating the harness to discharge any static charge on the harness

Table 1 – Sequence for mating Cryoharnesses to DRCU

| Harness | Connector |
| --- | --- |
| C1 | DCU J31 – Bias Prime |
|  | DCU J32 - Bias Red. |
|  | DCU J27 – SLW LIA |
|  | DCU J28 – SLW LIA |
| C2 | DCU J23 – SSW LIA |
|  | DCU J24 – SSW LIA |
|  | DCU J25 – SSW LIA |
|  | DCU J26 – SSW LIA |
| C4 | DCU J22 – PTC LIA + PWM LIA |
| C10 | FCU J11 – Prime Cooler Tail |
|  | FCU J25 – Thermometry Tail B (Plus S-Cal) |
|  | FCU J23 – Thermometry Tail A |
| C11 | FCU J13 – P-Cal Prime |
|  | MCU J17 – SMEC Control Prime |
|  | MCU J19 – BSM Prime |
|  | FCU J21 – Thermometry C |
|  | MCU J29 – SMEC Launch Tail |

Who:

When:

Comments:

## Mating Cryoharnesses to FPU

Pre-requisites

- FPU is mounted on MGSE in the horizontal position

Notes

- Ensure that when the harnesses are mated, that the MGSE can be removed and the instrument inserted into the cryostat without having the cryoharness tangled

Procedure

- Verify that the chassis of the FPU is electrically connected to the chassis of the cryostat

- Verify that the cryostat is electrically connected to the local lab ground

- Mate all FPU Faraday Shield Link pins in the cryoharness

- Remove Safeing plug from JFS J09

- Mate C1 JFS P09 to JFS J09

- Mate C1 JFS P05 to JFS J05

- Mate C1 JFS P06 to JFS J06

- Mate C1 JFS P07 to JFS J07

- Mate C2 JFS P01 to JFS J01

- Mate C2 JFS P02 to JFS J02

- Mate C2 JFS P03 to JFS J03

- Mate C2 JFS P04 to JFS J04

- Remove safeing plug from JFS J10

- Mate C1 JFS P10 to JFS J10

- Mate C10 FPU P19 to FPU J19 (Cooler prime)

- Mate C10 FPU P21 to FPU J21 (S-Cal Prime)

- Mate C10 FPU P23 to FPU J24 (sic.) (Thermometry harness plus PTC. There are no thermistors connected to this harness only the thermometry….)

- Mate C11 FPU P25 to FPU J25 (BSM Prime)

- Mate C11 FPU P27 to FPU J27 (SMEC Launch Prime)

- Mate C11 FPU P29 to FPU J29 (SMEC Control Prime)

- Mate C12 FPU P24 to FPU J23 (sic.) (Thermometry harness read out by Lakeshore units)

- Remove the grounding strap between the FPU and the cryostat chassis

- Demate all the FPU Faraday Shield Links

- Verify that the harness is not shorting to chassis

- Mate all the FPU Faraday Shield Links

- Replace the grounding strap between the FPU and the cryostat chassis

Who:

When:

Comments:

## Warm functional Test

- Rotate MGSE so that the instrument is in the vertical position

- Route harness on the back of the HOB

- Remove the grounding strap between the FPU and the cryostat chassis

- Demate all the FPU Faraday Shield Links

- Verify that the harness is not shorting to chassis

- Carry out Warm Functional Test (ref SPIRE-RAL-PRC-002223 step 4)

- Mate all the FPU Faraday Shield Links

- Replace the grounding strap between the FPU and the cryostat chassis

## Roll instrument into cryostat

Move FPU into cryostat (SPIRE-RAL-PRC-002223 step 5)

## Install thermal straps and then check isolation

Pre-requisites

- Thermal straps have been connected to the interfaces (ref SPIRE-RAL-PRC-002223 step 5)

Procedure

- Remove the grounding strap between the FPU and the cryostat chassis

- Demate all the FPU Faraday Shield Links

- Verify that the harness is not shorting to chassis

- Mate all the FPU Faraday Shield Links

- Replace the grounding strap between the FPU and the cryostat chassis

Who:

When:

Comments:

## Complete integration activities

Perform optical alignment of telescope simulator (SPIRE-RAL-PRC-002223 step 6)

Install Cold Blackbody (SPIRE-RAL-PRC-002223 step 7)

Perform short functional test (SPIRE-RAL-PRC-002223 step 8)

Perform cryostat functional test (SPIRE-RAL-PRC-002223 step 9)

## Install 10K Shield

Ref: SPIRE-RAL-PRC-002223 step 10

- Remove the grounding strap between the FPU and the cryostat chassis

- Demate all the FPU Faraday Shield Links

- Carry out a final check of the FPU isolation

- Install 10K shield

- Route harnesses on 10K

- Check FPU isolation

Who:

When:

Comments:

## Install 77K Shield

Ref: SPIRE-RAL-PRC-002223 step 10

- Install 77K shield

- Route harnesses on 77K

- Check FPU isolation

Who:

When:

Comments:

## Close cryostat

Ref: SPIRE-RAL-PRC-002223 step 10

- Check FPU isolation on FPU Faraday Shield Links

- Mate all FPU Faraday Shield Links

Who:

When:

Comments: