# Introduction

The purpose of this test is to check the interaction between the CDMS and the SPIRE instrument when performing a SPIRE Peak-up operation (Note this is not the same as the satellite peak-up mode).

The On-board software provided with the SPIRE AVM does not contain the code necessary for performing the SPIRE Peak-up mode and so the operation of this will be simulated by a Command List in the DPU which will act as if the peak-up operation had occurred.

The normal sequence of events for a Peak-up mode operation is:

- Send peak-up mode command to SPIRE

- SPIRE executes a set of operations to determine the offsets to bring the source back onto the central pixel (takes approximately 3 mins)

- SPIRE issues an event packet, TM(5,1) containing the offsets in y and z (format defined in AD01)

- The CDMS acts on this event packet and sends the appropriate TC(s) to the AOCS.

The test sequence will be:

- Load the test Command List into the DPU

- Send a command to SPIRE to execute the command list (equivalent to the command to perform the peak-up mode). The command List will wait and then issue the required event packet.

- The CDMS acts on this event packet and sends the appropriate TC(s) to the AOCS.

The test Command List is specified in Appendix A

## References

### Applicable Documents

| AD01 | IID Part A (SCI-PT-IIDA-04624), Issue 3.1 |
| --- | --- |
| AD02 | SPIRE Data ICD (SPIRE-RAL-PRJ-001078), Issue 1.2k |

### Reference Documents

| RD01 | IID Part B (SCI-PT-IIDB-02124), Issue 3.11 |
| --- | --- |

# Test Specification

## Prerequisites

Before execution of this procedure it will be necessary to specify the following information (identified in AD01, Section 5.12.5.1:

- The 3 components of Yfocal plane in the Satellite frame (this is to be taken from RD01)

- The 3 components of Zfocal plane in the Satellite frame (this is to be taken from RD01)

- The rotation angle conversion factor, kY = 4.8481368110954e-8

- The rotation angle conversion factor, kZ = 4.8481368110954e-8

## Test Parameters

For this test the SPIRE Peak-up Mode simulation Command List is passed the following parameters (see Appendix A)

- the time to wait before issuing the event packet (in microseconds)

- the offset value in y (Yangle) to be placed in the event packet

- the offset value in z (Zangle) to be placed in the event packet.

Offset values are specified as integers, i, where rotation angle (in radians) = k * I

For the value of k see section 2.1

This procedure will be repeated with several sets of parameters to the test Command List to verify full compliance to the interface:

| Parameter Set | Delay | Yangle | Zangle | Comment |
| --- | --- | --- | --- | --- |
| 1 | 180 | 500 | 500 | Move +5 arcsec in Y and Z |
| 2 | 10 | -500 | -500 | Move -5 arcsec in Y and Y |
| 3 | 10 | 200 | -200 | Move +2 arcsec in Y -2 arcsec in Z |
| 4 | 10 | -200 | 200 | Move -2 arcsec in Y +2 arcsec in Z |
| 5 | 10 | 10100 | 100 | Move +10.1 arcsec in Y 1 arcsec in Z Should be flagged as an error |
| 6 | 10 | 100 | 10100 | Move 1 arcsec in Y -10.1 arcsec in Z Should be flagged as an error |

## Configuration

The instrument should be in the REDY mode before the procedure starts.

It remains in the REDY mode after test execution

## Procedure

| Step | Description | Action(s) | Comments |
| --- | --- | --- | --- |
| 1 | Load Peak-up Mode Simulation Command List |  |  |
| 2 | Execute Peak-up Mode (delay = 180s Offsety = 200 Offsetz = 300) | RUN_VM (92,0,3,Delay, Yangle, Zangle) | The command used for this is not the same as that Peak-up mode command |
| 3 | Check event packet issued with correct contents |  | On SPIRE EGSE or SCOS |
|  | Check CDMS issues correct commands to AOCS |  | Post test check by AOCS experts |

Appendix A – Peak-up Mode Command List

TABLE	92

NAME    PeakupTest

VERSION 1.0

CVSID   $ID$

; This table contains the Peakup Mode test Command List

;

; Author: Ken King

;

; Version 1.0:   2nd March 2005

INC	..\\Includes\\SPIRE.inc	; include SPIRE constants

; ******************************************************

; PeakupTest

;

; This function generates an event packet in the form expected from the

; DPU during a Peakup Mode operation

;

; Input State (assumed):

;	Int3 timer is set to interrupt at the minimum command intervals (TimF)

;	The Mutex is unlocked

;

; Inputs: the following registers should be set before calling the Cmd List

;	Reg 00 = Delay before issuing the event packet (microsecs)

; 	Reg 01 = The value to be put in the Y rotational angle field of the event packet

;	Reg 02 = The value to be put in the Z rotational angle field of the event packet

;

; Uses: the following registers

;	Regs 10 to 13

;

; Outputs:

;	Event packet

;

; Output State:

;	Int3 timer is set to interrupt at the minimum command intervals (TimF)

;	The Mutex is unlocked

;

; ------------------------------------------------------

_put

; wait for delay time

rtim	00

NOP

; create Event packet at register 10

rset	10, 0x0504		; Event ID

rset	11, 0x0002		; Instrument ID

rreq	12, 01			; Y angle

rreq	13, 02			; Z angle

; send Event packet

NOP

evnt	4, 10

end