1.	INTRODUCTION	2

1.1	Change Record	2

1.2	Applicable Documents	2

1.3	Reference Documents	2

1.4	General instructions for executing test procedures	2

1.5	Assumptions	2

1.6	Open Issues	3

2.	COLD FUNCTIONAL TEST PROCEDURES	4

2.1	Procedure SPIRE-IST-DPU-ON	4

2.2	Procedure SPIRE-IST-DRCU-ON	5

2.3	Procedure SPIRE-IST-FUNC-SCU-02	6

2.4	Procedure SPIRE-IST-FUNC-SCU-03	7

2.5	Procedure SPIRE-IST-FUNC-SCU-06	8

2.6	Procedure SPIRE-IST-FUNC-SCU-07	9

2.7	Procedure SPIRE-IST-FUNC-SCU-04	10

2.8	Procedure SPIRE-IST-FUNC-PCAL-01	11

2.9	Procedure SPIRE-IST-FUNC-SCU-05	12

2.10	Procedure SPIRE-IST-FUNC-SCAL-01	13

2.11	Procedure SPIRE-IST-FUNC-MCU-01	14

2.12	Procedure: SPIRE-IST-FUNC-MCU-03	15

2.13	Procedure SPIRE-IST-FUNC-BSM-01	16

2.14	Procedure SPIRE-IST-FUNC-BSM-03	17

2.15	Procedure SPIRE-IST-FUNC-BSM-05a	18

2.16	Procedure SPIRE-IST-FUNC-BSM-05b	19

2.17	Procedure SPIRE-IST-BSM-OFF	20

2.18	Procedure SPIRE-IST-FUNC-SMEC-02a	21

2.19	Procedure SPIRE-IST-FUNC-SMEC-01	22

2.20	Procedure SPIRE-IST-FUNC-SMEC-03	23

2.21	Procedure SPIRE-IST-FUNC-SMEC-04a	24

2.22	Procedure SPIRE-IST-FUNC-SMEC-09	25

2.23	Procedure SPIRE-IST-FUNC-SMEC-04b	26

2.24	Procedure SPIRE-IST-FUNC-SMEC-07	27

2.25	Procedure SPIRE-IST-FUNC-SMEC-06	28

2.26	Procedure SPIRE-IST-SMEC-OFF	29

2.27	Procedure SPIRE-IST-FUNC-SMEC-02b	30

2.28	Procedure SPIRE-IST-FUNC-DCU-01	31

2.29	Procedure SPIRE-IST-FUNC-DCU-11-P	32

2.30	Procedure SPIRE-IST-FUNC-DCU-13-P	33

2.31	Procedure SPIRE-IST-PDET-OFF	34

2.32	Procedure SPIRE-IST-FUNC-DCU-11-S	35

2.33	Procedure SPIRE-IST-FUNC-DCU-13-S	36

2.34	Procedure SPIRE-IST-SDET-OFF	37

2.35	Procedure SPIRE-IST-MCU-OFF	38

2.36	Procedure SPIRE-IST-SCU-OFF	39

2.37	Procedure SPIRE-IST-DRCU-OFF	40

2.38	Procedure SPIRE-IST-DPU-OFF	41

# INTRODUCTION

This document contains the SPIRE Cold Functional Test Procedures to be executed during the Herschel IST. These procedures, which are valid for both He I and He II temperatures, are listed here in the order in which they are expected to be performed. These procedures give step-by-step instructions on how to perform each test. Some procedures can only be run after integration – where appropriate this is clearly indicated in the preconditions section of each procedure.

## Change Record

Issue 0.1, 27/04/2005 – Draft version for EQM

Issue 2.0, 12/06/2006 – First version for IST

## Applicable Documents

AD01 SPIRE Functional Test Specification, Issue 1.4, SPIRE-RAL-DOC-001652, 22/07/2005

AD02 SPIRE ILT Cold Functional Test Procedure, Issue 1.1, SPIRE-RAL-PRC-002585, 22/03/2006

## Reference Documents

RD01 SPIRE Instrument User Manual, Issue 1.0, SPIRE-RAL-PRJ-002395, 08/04/2005

## General instructions for executing test procedures

- Before executing any of the procedures please always check with the I-EGSE staff

- Any text in boldface in the procedural steps generally indicates an action which has to be performed manually by the Instrument EGSE (I-EGSE) staff.

- The procedure tables include blank boxes where the actual values of parameters can be noted. Based on the comparison with the expected values the success or failure of a step should be recorded in the final column of the table.

- The last row in a procedure table should be used to record the overall Pass/Fail result of each test.

## Assumptions

- The CCS are only required to check changes in instrument configuration related HK parameters.

- For these functional tests the instrument will not always be in a pre-defined mode as listed in the IUM (RD01).

- For the SPIRE spectrometer mechanism (SMECm) tests it is assumed that the Herschel cryostat will be tilted (TBD).

- These procedures should be suitable for operation of both the Prime and Redundant side of the instrument (TBD).

- The converted TM parameter values are extracted from the MIB in use for PFM ILT. These values are subject to change for both prime and redundant operations.

- The SPIRE thermometry channel values quoted in the procedures are dependent on the Herschel cryostat achieving nominal Level 0 and Level 1 temperatures.

## Open Issues

- Procedures for dealing with instrument contingencies will be addressed in the next version of this document.

- Names of the Herschel Satellite procedures for powering on/off the SPIRE DPU and DRCU are to be filled in the next version. In this version they are marked as procedure XXXXXX.

## Duration

The estimated duration for executing the entire CFT sequence of procedures, including switch off of the SPIRE instrument afterwards is estimated to be about 4 hours.

# COLD FUNCTIONAL TEST PROCEDURES

## Procedure SPIRE-IST-DPU-ON

Version: 1.0

Date: 12th June 2006

Purpose: To switch on the SPIRE DPU and start generating housekeeping

Duration: 2 minutes plus time to execute CCS procedure XXXXXX

Preconditions:

- Procedure to supply 28V Power Supply from the satellite to the SPIRE DPU is available

- SPIRE MIB is imported in the CCS database.

- CCS is up and running (SCOS, TOPE and the CDMU)

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

- The I-EGSE is up and running

Initial Configuration: SPIRE Warm Electronics (DPU and DRCU) are switched off

| Step | Description | Parameter | Expected Values Before/After | Actual Values Before/After | Pass/ Fail |
| --- | --- | --- | --- | --- | --- |
| 1 | Using CCS procedure XXXXX Power on the SPIRE DPU 28V Power Supply | — | — | — |  |
| 2 | Wait for instruction from I-EGSE staff to continue with the procedure | — | — | — |  |
| 3 | Execute TCL script SPIRE-IST-DPU-ON.tcl | — | — | — |  |
| 4 | Check that THSK parameter is refreshing every second | — | — | — |  |
| 5 | Check that TM2N parameter is incrementing every second | — | — | — |  |
| Test Result (Pass/Fail): |

Procedure Steps:

Final Configuration: SPIRE DPU is on but the DRCU is still off

## Procedure SPIRE-IST-DRCU-ON

Version: 1.0

Date: 12th June 2006

Purpose: To switch on the SPIRE DRCU and start generating housekeeping

Duration: 4 minutes

Preconditions:

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched off

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

Procedure Steps:

| Step | Description | Parameter | Expected Values Before/After | Actual Values Before/After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-DRCU-ON-STEP1.tcl | — | — | — |  |
| 2 | Check that THSK parameter is not refreshing anymore | — | — | — |  |
| 3 | Check that TM2N parameter is not incrementing anymore | — | — | — |  |
| 4 | When instructed by the I-EGSE staff Power on the SPIRE DRCU using the CCS procedure XXXXXX | — | — | — |  |
| 5 | Execute TCL script SPIRE-IST-DRCU-ON-STEP2.tcl | — | — | — |  |
| 6 | Check that THSK parameter is again refreshing every 4 seconds | — | — | — |  |
| 7 | Check that TM2N parameter is again incrementing every 4 seconds | — | — | — |  |
| Test Result (Pass/Fail): |

Final Configuration:

- SPIRE DPU and DRCU are both on

- HK generation is on

## Procedure SPIRE-IST-FUNC-SCU-02

Version: 1.0

Date: 12th June 2006

Purpose: SCU science contents check by the I-EGSE

Duration: 5 minutes

Preconditions:

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

Procedure Steps:

| Step | Description | Parameter | Expected Values Before/ After | Actual Values Before /After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCU-02.tcl | SCUFRAMECNT TM5N | 31/62 1/3 |  |  |
| 2 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SCU-03

Version: 1.0

Date: 21st June 2006

Purpose: SCU DC thermometry check

Duration: 6 minutes

Preconditions: SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCU-03.tcl | — | — | — |  |
| 2 | Wait for the parameter BBFULLTYPE to get set to SCU_DC_Therm | BBFULLTYPE | -/SCU_DC_Therm/- |  |  |
| 3 | A few seconds later record the value of parameter SCUTEMPSTAT | SCUTEMPSTAT | 0/0xFFFF/0xFFFF |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: SCU DC thermometry is switched on.

## Procedure SPIRE-IST-FUNC-SCU-06

Version: 1.0

Date: 12th June 2006

Purpose: SCU AC thermometry check

Duration: 2 minutes

Preconditions: SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCU-06.tcl | — | — | — |  |
| 2 | Wait for the parameter BBFULLTYPE to get set to SCU_AC_Therm | BBFULLTYPE | -/SCU_AC_Therm/- |  |  |
| 3 | A few seconds later record the value of parameter SUBKSTAT | SUBKSTAT | 0/1/1 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: SCU AC thermometry is switched on.

## Procedure SPIRE-IST-FUNC-SCU-07

Version: 1.0

Date: 11th June 2005

Purpose: SCU cooler heaters check

Duration: 3 minutes

Preconditions: SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCU-07.tcl | — | — | — |  |
| 2 | Wait for the parameter BBFULLTYPE to get set to Cooler_Htr_Chk | BBFULLTYPE | Cooler_Htr_Chk |  |  |
| 3 | A few seconds later record the value of parameter EVHSV – the Evaporator Heat Switch Voltage. This voltage stays on for ~20 seconds. | EVHSV - mV | 0/~323/0 |  |  |
| 4 | A few seconds after the EVHSV parameter has been set back to 0, record the value of parameter SPHSV – the Sorption Pump Heat Switch Voltage. This voltage stays on for ~20 seconds. | SPHSV - mV | 0/~323/0 |  |  |
| 5 | A few seconds after the SPHSV parameter has been set back to 0, record the value of parameter SPHTRV – the Sorption Pump Heater Voltage. This voltage stays on for ~20 seconds. | SPHTRV - V | 0/~8.8/0 |  |  |
| 6 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SCU-04

Version: 1.0

Date: 12th June 2006

Purpose: SCU PCAL check

Duration: 2 minutes

Preconditions: SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter Name - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCU-04.tcl The expected values during the test should be monitored when parameter BBFULLTYPE in the FUNCTIONAL TEST PARAMETERS display is set to PCAL_Check This usually happens about 30 seconds from the start of test execution. | PCALCURR - mA PCALV – V BBFULLTYPE | 0.0/0.1/0.0 0.0/0.026/0.0 PCAL_Check |  |  |
| 2 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-PCAL-01

Version: 1.0

Date: 12th June 2006

Purpose: SCU PCAL characterisation

Duration: 8 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter Name - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-PCAL-01.tcl | — | — | — |  |
| 2 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SCU-05

Version: 1.0

Date: 21st June 2006

Purpose: SCU Spectrometer SCAL4 and SCAL2 check

Duration: 4 minutes

Preconditions: SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCU-05.tcl | — | — | — |  |
| 2 | Wait for the parameter BBFULLTYPE to get set to SCAL4_Check | BBFULLTYPE | SCAL4_Check |  |  |
| 3 | A few seconds later record the value of parameters SCAL4CURR and SCAL4V These parameters are set back to 0 after ~30 seconds | SCAL4CURR – mA SCAL4V – V | 0.0/0.10/0.0 0.0/0.05/0.0 |  |  |
| 4 | Wait for the parameter BBFULLTYPE to get set to SCAL2_Check | BBFULLTYPE | SCAL2_Check |  |  |
| 5 | A few seconds later record the values of parameters SCAL2CURR and SCAL2V These parameters are set back to 0 after ~30 seconds | SCAL2CURR – mA SCAL2V – V | 0.0/0.10/0.0 0.0/0.05/0.0 |  |  |
| 6 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SCAL-01

Version: 1.0

Date: 12th June 2006

Purpose: SCU SCAL4 and SCAL2 characterisation

Duration: 16 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter Name - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-SCAL-01.tcl | — | — | — |  |
| 2 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-MCU-01

Version: 1.0

Date: 12th June 2006

Purpose: To boot up the MCU

Duration: 3 minutes

Preconditions:

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is DRCU_ON | MODE | DRCU_ON |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-MCU-01.tcl | — | — | — |  |
| 3 | Check that the mode parameter is REDY | MODE | DRCU_ON/-/REDY |  |  |
| 4 | Check that the MCU is booted up successfully | MCUBITSTAT | 0/1/1 |  |  |
| 5 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration:

- MCU is switched on and booted up

- SPIRE is in REDY mode

## Procedure: SPIRE-IST-FUNC-MCU-03

Version: 1.0

Date: 12th June 2006

Purpose: MCU science data contents check

Duration: 5 minutes

Preconditions:

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY | REDY |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-MCU-03.tcl | — | — | — | — |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-BSM-01

Version: 1.0

Date: 12th June 2006

Purpose: BSM switch on check

Duration: 3 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-BSM-01.tcl | — | — | — |  |
| 3 | Check that the Chop and Jiggle sensors have switched on | CHOPSENSPWR JIGGSENSPWR | 0/1/1 0/1/1 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: BSM is switched on.

## Procedure SPIRE-IST-FUNC-BSM-03

Version: 1.0

Date: 12th June 2006

Purpose: BSM open loop dynamics check

Duration: 6 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-BSM-03.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-BSM-05a

Version: 1.0

Date: 12th June 2006

Purpose: BSM open loop chop test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-BSM-05a.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-BSM-05b

Version: 1.0

Date: 12th June 2006

Purpose: BSM closed loop chop test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute SPIRE-IST-BSM-INIT.tcl | — | — |  |  |
| 3 | Execute TCL script SPIRE-IST-FUNC-BSM-03.tcl | — | — |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-BSM-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off the BSM

Duration: 2 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute SPIRE-IST-BSM-OFF.tcl | — | — | — |  |
| 3 | Check that the power to the BSM sensors is switched off | CHOPSENSPWR JIGGSENSPWR | 1/-/0 1/-/0 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: BSM is switched off.

## Procedure SPIRE-IST-FUNC-SMEC-02a

Version: 1.0

Date: 12th June 2006

Purpose: Open the SMECm launch latch

Duration: 5 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-02a.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: SMECm is unlatched.

## Procedure SPIRE-IST-FUNC-SMEC-01

Version: 1.0

Date: 12th June 2006

Purpose: SMECm switch on check

Duration: 5 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-01.tcl | — | — | — |  |
| 3 | Check that power to the SMEC LED and LVDT sensor is on | SMECENCPWR SMECLVDTPWR | 0/-/4(TBD) 0/1/1 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: SMECm is switched on.

## Procedure SPIRE-IST-FUNC-SMEC-03

Version: 1.0

Date: 12th June 2006

Purpose: SMEC LED Optical Encoder LED Check

Duration: 5 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-03.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SMEC-04a

Version: 1.0

Date: 12th June 2006

Purpose: SMECm open loop position test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-04a.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SMEC-09

Version: 1.0

Date: 12th June 2006

Purpose: SMECm open loop scan test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Check the SMECm is in open loop | SMECLOOPMODE | 6/6/6 |  |  |
| 3 | Execute TCL script SPIRE-IST-FUNC-SMEC-09 .tcl | — | — | — |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SMEC-04b

Version: 1.0

Date: 12th June 2006

Purpose: SMECm closed loop position test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-INIT.tcl | SMECLOOPMODE | 6/1/1 |  |  |
| 3 | Execute TCL script SPIRE-IST-FUNC-SMEC-04B .tcl | — | — | — |  |
| 4 | Check that SMECm is still in closed loop | SMECLOOPMODE | 1 |  |  |
| 5 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: The SMECm is in closed loop.

## Procedure SPIRE-IST-FUNC-SMEC-07

Version: 1.0

Date: 12th June 2006

Purpose: SMECm closed loop scan test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-INIT.tcl | SMECLOOPMODE | -/1/1 |  |  |
| 3 | Execute TCL script SPIRE-IST-FUNC-SMEC-07 .tcl | — | — | — |  |
| 4 | Check that SMECm is still in closed loop | SMECLOOPMODE | 1 |  |  |
| 5 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-SMEC-06

Version: 1.0

Date: 12th June 2006

Purpose: SMECm saw-tooth test

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-INIT.tcl | SMECLOOPMODE | -/1/1 |  |  |
| 3 | Execute TCL script SPIRE-IST-FUNC-SMEC-06 .tcl | — | — | — |  |
| 4 | Check that SMECm is still in closed loop | SMECLOOPMODE | 1 |  |  |
| 5 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-SMEC-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off the SMEC

Duration: 2 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute SPIRE-IST-SMEC-OFF.tcl | — | — | — |  |
| 3 | Check that the power to the SMEC sensors is switched off | SMECENCPWR SMECLVDTPWR | 4(TBD)/-/0 1/-/0 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — |  |  |  |
| Test Result (Pass/Fail): |

Final Configuration: SMECm is switched off.

## Procedure SPIRE-IST-FUNC-SMEC-02b

Version: 1.0

Date: 12th June 2006

Purpose: Close the SMECm launch latch

Duration: 2 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-SMEC-02b.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: SMECm is latched.

## Procedure SPIRE-IST-FUNC-DCU-01

Version: 1.0

Date: 12th June 2006

Purpose: DCU science packet generation check for all Photometer and Spectrometer packet types (PF, PSW, PMW, PLW, SF, SSW and SLW)

Duration: 5 minutes

Preconditions:

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter | Expected Values Before/ After | Actual Values Before /After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-FUNC-DCU-01.tcl | DCUFRAMECNT | 0/700 |  |  |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-FUNC-DCU-11-P

Version: 1.0

Date: 12th June 2006

Purpose: Photometer detectors switch on

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that Photometer LIAs are switched on | PLIABITSTAT | 1 |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-DCU-13-P.tcl | — | — | — |  |
| 3 | Check that the Photometer detectors are switched on | PSWJFETSTAT PMLWJFETSTAT PSWJFET1V PSWJFET2V PSWJFET3V PSWJFET4V PSWJFET5V PSWJFET6V PMWJFET1V PMWJFET2V PMWJFET3V PMWJFET4V PLWJFET1V PLWJFET2V TCJFETV PHOTHTRV | 0/0x3F/0x3F 0/-/0x7F -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V -1.49V 0.0V |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Photometer detectors are switched on.

## Procedure SPIRE-IST-FUNC-DCU-13-P

Version: 1.0

Date: 12th June 2006

Purpose: Perform a Photometer Load Curve

Duration: 20 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- Procedure SPIRE-IST-FUNC-DCU-11-P has been executed

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- The Photometer detectors and LIAs are on

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that Photometer LIAs are switched on | PLIABITSTAT | 1 |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-DCU-13-P.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-PDET-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off photometer detectors

Duration: 2 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- SPIRE-IST-FUNC-DCU-11-P has been executed

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- Photometer detectors and LIAs are switched on

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-PDET-OFF.tcl | — | — | — |  |
| 2 | Check that the Photometer detectors are switched off | PSWJFETSTAT PMLWJFETSTAT PSWJFET1V PSWJFET2V PSWJFET3V PSWJFET4V PSWJFET5V PSWJFET6V PMWJFET1V PMWJFET2V PMWJFET3V PMWJFET4V PLWJFET1V PLWJFET2V TCJFETV PHOTHTRV | 0x3F/-/0 0x7F/-/0 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0.0V 0. |  |  |
| 3 | Check that the Photometer LIAs are switched off | PLIABITSTAT | 1/-/0 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Photometer detectors are switched off.

## Procedure SPIRE-IST-FUNC-DCU-11-S

Version: 1.0

Date: 12th June 2006

Purpose: Spectrometer detectors switch on

Duration: 10 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that Spectrometer LIAs are switched on | SLIABITSTAT | 1 |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-DCU-11-S.tcl | — | — | — |  |
| 3 | Check that the Spectrometer detectors are switched on | SPECJFETSTAT SSWJFET1V SSWJFET2V SLWJFET1V SPECHTRV | 0/7/7 -1.49V -1.49V -1.49V 0.0V |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Spectrometer detectors switched on.

## Procedure SPIRE-IST-FUNC-DCU-13-S

Version: 1.0

Date: 15th June 2006

Purpose: Perform a Spectrometer Load Curve

Duration: 20 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- Procedure SPIRE-IST-FUNC-DCU-11-S has been executed

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- Spectrometer detectors and LIAs are switched on

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that Spectrometer LIAs are switched on | SLIABITSTAT | 1 |  |  |
| 2 | Execute TCL script SPIRE-IST-FUNC-DCU-13-S.tcl | — | — | — |  |
| 3 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Unchanged

## Procedure SPIRE-IST-SDET-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off Spectrometer detectors

Duration: 2 minutes

Preconditions:

- SPIRE FM is electrically integrated with the Herschel Satellite

- Procedure SPIRE-IST-FUNC-DCU-11-S has been executed

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- Spectrometer detectors and LIAs are switched on

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-SDET-OFF.tcl | — | — | — |  |
| 2 | Check that the Spectrometer detectors are switched off | SPECJFETSTAT SSWJFET1V SSWJFET2V SLWJFET1V SPECHTRV | 7/-/0 0.0V 0.0V 0.0V 0.0V |  |  |
| 3 | Check that the Spectrometer LIAs are switched off | SLIABITSTAT | 1/-/0 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: Spectrometer detectors are switched off.

## Procedure SPIRE-IST-MCU-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off the MCU – if necessary

Duration: 2 minutes

Preconditions:

- SPIRE is in REDY mode

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- SPIRE FUNCTIONAL PARAMETERS display is selected on the CCS

| Step | Description | Parameter – Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Check that the mode parameter is REDY | MODE | REDY | REDY |  |
| 2 | Execute SPIRE-IST-MCU-OFF.tcl | — | — | — |  |
| 3 | Check that the MCU is switched off | MCUBITSTAT | 1/-/0 |  |  |
| 4 | Wait for the I-EGSE staff to confirm the success or failure of this test | — |  |  |  |
| Test Result (Pass/Fail): |

Final Configuration: MCU switched off.

## Procedure SPIRE-IST-SCU-OFF

Version: 1.1

Date: 12th June 2006

Purpose: Switch off SCU DC and AC thermometry – if necessary

Duration: 2 minutes

Preconditions:

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-SCU-OFF.tcl | — | — | — |  |
| 2 | A few seconds later record the value of parameter SCUTEMPSTAT | SCUTEMPSTAT | FFFF/-/0 |  |  |
| 3 | A few seconds later record the value of parameter SUBKSTAT | SUBKSTAT | 1/-/0 |  |  |
| 4 | Check that SPIRE is in DRCU_ON mode | MODE | REDY/-/DRCU_ON |  |  |
| 5 | Wait for the I-EGSE staff to confirm the success or failure of this test | — | — | — | — |
| Test Result (Pass/Fail): |

Final Configuration: SPIRE in DRCU_ON mode.

## Procedure SPIRE-IST-DRCU-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off the DRCU – if necessary

Duration: 2 minutes

Preconditions:

- Procedure SPIRE-IST-SCU-OFF has been successfully executed

- SPIRE is electrically integrated with the Herschel EQM.

Initial Configuration:

- SPIRE DPU is on and generating HK

- DRCU is switched ON

- FUNCTIONAL TEST PARAMETERS display is selected on the CCS

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | Execute TCL script SPIRE-IST-DRCU-OFF.tcl | — | — | — |  |
| 2 | Check that THSK parameter is not refreshing anymore | — | — | — |  |
| 3 | Check that TM2N parameter is not incrementing anymore | — | — | — |  |
| 4 | When instructed by the I-EGSE staff Power off the SPIRE DRCU using CCS procedure XXXXXX | — | — | — |  |
| Test Result (Pass/Fail): |

Final Configuration:

- DRCU is switched off

- SPIRE DPU is on but not generating HK

## Procedure SPIRE-IST-DPU-OFF

Version: 1.0

Date: 12th June 2006

Purpose: Switch off the DPU – if necessary

Duration: 2 minutes

Preconditions: SPIRE-IST-DRCU-OFF has been successfully executed.

Initial Configuration:

- SPIRE DPU is on but not generating any HK

- DRCU is switched OFF

| Step | Description | Parameter - Unit | Expected Values Before/ During/ After | Actual Values Before/ During/ After | Success/ Failure |
| --- | --- | --- | --- | --- | --- |
| 1 | When instructed by the I-EGSE staff Power off the SPIRE DRCU using the CCS procedure XXXXXX | — | — | — |  |
| Test Result (Pass/Fail): |

Final Configuration: SPIRE DPU is switched off and the SPIRE instrument is OFF.