| Subject: PFM1 Cold Verification – Post Test Review |
| --- |
| AGENDA: 1. Equipment Description 2. Design vs. Build Reconciliation 3. Mandatory Inspection Points (MIP) References & Closure 4. Review of Actions from Test Readiness Review (TRR) & any Previous Meetings 5. Non-Conformances / open Items & Actions 6. Test Procedure & Changes since TRR 7. Request for Waiver (RFW) 8. Hardware Inspection 9. Preliminary Summary of Test Results - Overview of test campaign - Test facility status - Functional Tests - Instrument Operations - Performance Tests 10. Any Other relevant Business |
| Conclusion Main objectives of the PFM1 cold tests were carried out. Full analysis of test data to be performed. |
| Meeting Participants |
| RAL: Dave Smith, Bruce Swinyard, Tanya Lim, Doug Griffin, Sunil Sidher, Samuel Ronayette, Eric Clark, Asier Aramburu IC Davide Rizzo Cardiff: Pete Hargrave, Tim Waskett, Adam Woodcraft, Matt Griffin (part time) |
| Written By: Dave Smith | Chaired By: Dave Smith |
| DISTRIBUTION Participants + Eric Sawyer, Ken King, Dave Clements(IC), Carsten Schamberg(ESTEC), Didier Ferrand (LAM) |

| Equipment Description |
| --- |
| Spacecraft / Project | Herschel / SPIRE |
| Instrument / Model | PFM 1 |
| Sub System / Serial No. | N/A |
| As Built Status | SPIRE-RAL-DOC-002326 Issue 2.1 |
| List of all items /Parts |  |

| Test Documentation |
| --- |
| Type of Test | Cold Functional and Performance test |
| AIV Facility Test No. | N/A |
| Location & Date(s) of Testing | SPIRE Calibration Facility 18-Feb-2005 to 15-Apr-2005 |
| Applicable Test Specification (Document No. & Issue) | Test Plan SPIRE-RAL-DOC-2055 issue 1 |
| Applicable Test Procedure (Document No. & Issue) | SPIRE-RAL-PRC-02223 issue 1.0 (formally released at start of test) |
| AIV Facility Test Plan (if applicable?) | NA |
| TRR MoM | SPIRE-RAL-MoM-002306 phase 3 |

| NCRs, Waivers, Open Work |
| --- |
| NCR’s | See continuation sheets and appendix 1 |
| Waivers | None |
| Open Work | All open work from TRR closed. |
| Inspection Points | KIP final inspection of FPU, 17/1/05 MIP before closing tank, final TRR 27/1/05 Inspection of FPU to be performed (KIP) |

CONTINUATION SHEET

Review of Actions: AI-002306-1 - Pete Hargrave to review the test procedure for PCAL and SCAL, warm functional test – not done. NCRS See attachment 1 – NCR spreadsheet NCRs raised before Test 92.2 N/A for this test 100 open, remedial action taken 101 open, remedial action taken 102 open – no effect on test 103 open - no effect on test 104 open – flight spare unit swapped in – original unit returned 105 open - simultaneous generation BSM & SMEC packets – A update to EEPROM on MCU has been prepared. 106 open - electrical short between FPU and cryostat due to incorrect harness wiring – used as is, did not have significant effect on testing – analysis of noise tests to be done NCRs raised during test 107 open – dynamic range of SCU thermometers not matched to flight thermometers – fix to SCAL temperature channels to allow tests to continue 108 open – Automatic shutdown of LIAs – to be discussed 109 closed – N/A for this test.

Preliminary Summary of Test Results Overview of test campaign Presentation by Dave Smith – Attachment 2 Staffing Cryostat operations – OK Instrument operations – Worked well with Tim and Davide supporting Sunil and Asier Test directors – Bruce/Tanya – need 3rd person to direct tests – could have Locke, Marc or Pete Cold tests at ASTRIUM in July AVM tests in Alenia Test Facility Status No major problems occurred but some issues need to be resolved before next test campaign. Some failing thermometers – caused by poor wiring quality – to be investigated. Thermometer calibrations at instrument I/F – new thermometers needed? Cold blackbody – cannot get to 4K – may need a further NIR filter – NCR-113 Air leak – possible location identified – propose to seal suspect joint with stycast – NCR114 Pump controller shutting down - NCR-085 Vacuum gauge head failed during pump-down – NCR 111 Laser gas – faulty regulator – NCR 112 - should be checked before each test campaign Could not command TFTS from SCOS – NCR115 – to be investigated before next test campaign. Calibration curves in SCOS did not match actual thermometers – need to match to serial no. For PFM2 – QM2 does not have redundant side – hence could read redundant thermometers via lakeshore unit – to be investigated

Functional Tests & Operations Presentation by Sunil Action #1 : Safe switch on procedure needs document number - Bruce Action #2: Test reports for functional tests to be given individual document numbers - Sunil Problems occurred with simultaneous generation of SMEC and BSM data – an update to the MCU EEPROMs has been generated by Didier Sequence follows master test procedure – comments added – some changes to be made for PFM2 Power bench needs resetting when MCU is switched on – NCR has been raised – should not be an issue for PFM2 as PSU will be used. Missing SMEC parameters from HK report – temporary fix implemented BSM – oscillations in chop and jiggle axes – possibly caused by missing capacitors on QM1 board – NCR-116 JFETs turned on without having to use heaters PCAL and SCAL tests – plots show resistances as expected PCAL & SCAL temp vs. current plots – test ok although could be simplified with fewer power steps – e.g. low and high SMEC scans – show expected behaviour AMV2 DPU did not work with DRCU – separate investigation being conducted Problems encountered – LIA shutdown Action #3: Provide a log of power cycles of JFETs – DS Anomalous HK parameters for LIA-P and LIA-S at specific bias/sample frequencies – NCR-117 to be raised. Inability to put SMEC into closed loop at first attempt following SMEC switch on – under investigation Separate commands are needed to switch on MCU, spectrometer and photometer LIAs – NCR-118 It is easy to send an unknown command to the DRCU – NCR-119 SPR to be raised on TM ingestion in HCSS – can grow to 42,000 – should be investigated before PFM2 QLA frequently ‘hangs’ – possibly due to amount of memory – should be resolved for PFM2 Should have more people trained in QLA Lessons to be learned Avoid unnecessary changes to instrument configuration Warm electronics need to be protected – NOTE configuration will be different for next tests Need point of contact for contingency – i.e. difficult to contact Test scripts need to be rationalised – old test scripts need to be moved off Some effort is needed to write up test sequences as scripts to avoid manual commanding Fault recognition document – need secure folder containing key documents Action # 4: Compile hard copy folder of critical documents – incl. list of documents , summary sheet – SS/DS Need a printer Label plugs Successes Experience gained from CQM tests enabled problems to be anticipated and solved Sufficient human resources to conduct tests without stress on test team Most tests successfully performed Future tests could be conducted Plans for future Test specifications to be updated Creation of HCSS scripts for BSM and SMEC tests Tests with AVM FPU simulator needs updating Update data ICD New version of the OBS which supports autonomy functions

SMEC & BSM Operations Presentation by Asier Need written procedures from Didier to tune PID settings Need to analyse SMEC data to understand relationship between power, temperature, settings

Performance Testing Tanya Data Test Log, Data Download Status Excel spreadsheet – to be made available on web site Log complete up to 473 lines – should be complete by end of next week Includes functional tests and operator procedures Data have been downloaded for early tests – remaining will be extracted ASAP Times in UTC – some discrepancy between actual test times and times recorded in word logs Additional information for web site pdf’d daily logs, TC history Web address -> http://scott1.bnsc.rl.ac.uk:8080/hcss/test_area/index.htm

Test Plan Most tests in plan completed with the following exceptions: DNC-S noise test with tel sim background DLA-S Beam attenuation using laser CPF-S PCAL loaded frequency characterisation CPN PCAL Flash while nulling OBL/L-S Beam scan using hot BB – type 1 (one was done but not intended) OPI-S Pupil scan using HBB – (done with laser) SRN – room temp nulling – done but failed due to high background SFL – Fringe contrast step and look SML – Mirror carriage characterisation step and look OBS – OOB radiation using SMEC Issues Need for separate settings for each array Need to search for optimum operating frequency Settings for VMs to be better documented – folder in Q drive – should be in HCSS Unable to do jiggle tests Variable background for BSM tests – most likely due to loading from window More work on pupil scanning – may need test without using actuators - TS not really designed for this test Changes for next tests Change noise tests to peakup at each bias Need better SCAL warm up procedure Internal filtering – possibly need more IDL toolkit – pass tools through Tanya to put on website

Beam Scans Preliminary plots presented by Samuel For further investigation – look at how beam width varies with wavelength – at least on nominal pixel Peak up maps – measured pixel positions to be compared with optical model

Detector performance Bruce presented preliminary results Measured spectral resolution is as predicted

Plan for data analysis Action # 5: Tanya to formalise plan for data processing – who does what and when.

Date of Next Test Campaign Nominally end of June – Eric to update Schedule

Summary of Actions Action #1 : Safe switch on procedure needs document number - Bruce Action #2: Test reports for functional tests to be given individual document numbers – Sunil Action #3: Provide a log of power cycles of JFETs – DS Action # 4: Compile hard copy folder of critical documents – incl. list of documents , summary sheet – SS/DS Action # 5: Tanya to formalise plan for data processing – who does what and when.