# Introduction

This document specifies the necessary steps to execute a COLD SPIRE functional test during ILT test campaigns. It currently applies to the PFM model of SPIRE but it can be applied to further models as a minimum set of tests to verify the instrument correct functioning prior to cold performance testing.

## Scope

To judge the success or failure of a warm functional test by checking that:

- The telecommand sequence generated for a particular functional test is correctly received and executed on board by the SPIRE DPU.

- No error/event reports or command failures are generated during the execution of these commands.

- Telemetry is generated by the instrument as a result of telemetry requests to its different subunits.

- Particular telemetry parameters for each functional test change in an expected manner.

- A particular success criterion (specified in this document) is met.

## Reference Documents

| Ref | Document | Name | Version/Issue Nb. |
| --- | --- | --- | --- |
| RD01 | SPIRE-RAL-DOC-001652 | SPIRE Functional Tests Specification | Issue 1.4 |
| RD02 | SPIRE-RAL-DOC-001630 | SPIRE EGSE-ILT Start-Up Procedures | Issue 0.7 |
| RD03 | SPIRE-RAL-PRC-002222 | DRCU Switch ON Procedure | Issue 1.0 |
| RD04 | SPIRE-RAL-PRJ-001078 | SPIRE Data ICD | Issue 2.0 |
| RD05 | Sap-SPIRE-CCa-076-02 | DRCU/DPU Interface Control Document | Issue 1.2 |
| RD06 | LAM.PJT.SPI.NOT.011011 | MCU/DPU Command List ICD | Issue 5.0 |
| RD07 | SPIRE-IFS-PRJ-001391 | SPIRE OBS User Manual | Issue 2.1 |

## Change Record

| Document | Change date | Changes |
| --- | --- | --- |
| Issue 1.0 | 27/02/06 | Created from Warm Functional Test Procedure |
| Issue1.1 | 22/03/06 | Some updates following the warm functional report results |

# Table of Contents

1.	Introduction	1

1.1	Scope	1

1.2	Reference Documents	1

1.3	Change Record	1

2.	FUNCTIONAL TEST CONFIGURATION	3

2.1	Software Configuration	3

2.2	EGSE Configuration Checks	3

2.3	SPIRE Instrument Configuration	4

3.	Test procedure	6

3.1	GENERAL PASS/FAIL CRITERIA	6

3.2	GENERAL TEST PROCEDURE LAYOUT	6

3.3	DETAILED TEST PROCEDURE	7

3.3.1	FUNC-SCU-01	7

3.3.2	FUNC-SCU-03	9

3.3.3	FUNC-SCU-06	10

3.3.4	FUNC-SCU-07	12

3.3.5	FUNC-PCAL-01	13

3.3.6	FUNC-SCAL-01	14

3.3.7	FUNC-MCU-01	16

3.3.8	FUNC-MCU-02	17

3.3.9	FUNC-BSM-01	18

3.3.10	FUNC-BSM-03	19

3.3.11	FUNC-BSM-05A	20

3.3.12	FUNC-BSM-05B	21

3.3.13	FUNC-BSM-06	22

3.3.14	FUNC-SMEC-01	23

3.3.15	FUNC-SMEC-04A	25

3.3.16	FUNC-SMEC-09	26

3.3.17	FUNC-SMEC-03	27

3.3.18	FUNC-SMEC-04B	29

3.3.19	FUNC-SMEC-07	30

3.3.20	FUNC-SMEC-06	31

3.3.21	FUNC-DCU-01	33

3.3.22	FUNC-DCU-04P	35

3.3.23	FUNC-DCU-11P	36

3.3.24	FUNC-DCU-13P	37

3.3.25	FUNC-DCU-04S	38

3.3.26	FUNC-DCU-11S	39

3.3.27	FUNC-DCU-13S	40

4.	END TEST SEQUENCE	41

4.1	NORMAL END TEST SEQUENCE	41

4.2	END TEST SEQUENCE WHEN THE FUNCTIONAL TEST HAS FAILED	41

5.	Annexe1 (Instrument nominal configuration/modes)	44

# FUNCTIONAL TEST CONFIGURATION

## Software Configuration

Write down in the table the current EGSE software configuration for the tests:

| EGSE component | Version/Build number | Comment |
| --- | --- | --- |
| SCOS2000 |  |  |
| HCSS |  |  |
| QLA |  |  |
| QLA scripts |  |  |
| Test Control scripts |  |  |
| CUS Scripts |  |  |

## EGSE Configuration Checks

To check for the success of failure of a functional test, the real time telemetry of the instrument has to be monitored. The following applications must be running to do so. Before the test sequence starts, make the following checks:

| Workstation | EGSE component | Status | Check |
| --- | --- | --- | --- |
| Lichfield | EGSE router | Started | ( |
| Lichfield | EGSE Gateway | Started |  |
| Lichfield | Telemetry Ingestion | Started |  |
| Lichfield | Packet Display | Started |  |
| Lincoln | SCOS2000 | Started |  |
| Lincoln | EXIF + TOPE | Started |  |
| Lincoln | Manual Stack | Started |  |
| Gordon | CDMS Simulator | Started |  |
| Lincoln | Test Control Server | Started |  |

Note: Despite restarting the CDMS simulator there is still ~ 30 seconds disagreement between the the THSK value that SCOS shows and the time reported by QLA (QLA is 31 seconds behind) on its packet receiver window (this time time actually agrees with the time that PacketDisplay shows on the time column) ?? two java applications showing 31 seconds delay (= current difference between UTC and TAI).

## SPIRE Instrument Configuration

The functional test flow is such that at a given time during the procedure is possible that the instrument configuration does not conform to any nominal configuration as specified in Annexe 1.  Nevertheless the configuration prior and after a test is specified for each functional test, and were applicable and the instrument configuration is in accordance to Anexe1, this will be denoted by showing the main configuration in blue bold font and any additional information in black.

The initial instrument configuration for these tests should be SPIRE DRCU_ON + DC and AC Thermometry ON.

Perform the following actions to ensure that the instrument is in the correct configuration for the tests.

| Step# | Action | Comments | Check |
| --- | --- | --- | --- |
| 1 | In SCOS open DPU_AND_OBS_PARAMETERS display Check if : TM2N is incrementing by one @1Hz. TM1N is incrementing by one @0.5Hz. - If they are, go to step 5. - If only TM2N is incrementing it means that the critical HK report request has been stopped, go to step5. - If they are not, go to step 2. |  |  |
| 2 | Check if the DPU is powered ON: - If ON, the DPU power supply LCD will show ~ 28V and 0.40A, go to step 4. - If not ON, refer to RD2, then go to step 3. |  |  |
| 3 | In SCOS open Boot_ROM_Memory_Check display and check no errors are reported: - If no errors are reported, execute DPU_ON from HCSS Test Procedures. Then repeat step 1. - If the (5,2) contains errors: Check the error code in RD07. Then switch OFF the DPU and repeat step 2 |  |  |
| 4 | Execute define_new_HK_report.tcl HCSS Test procedure. Repeat step 1. |  |  |
| 5 | In SCOS open SCU_PARAMETERS display Check SCUP5V/P9V/M9V In SCOS open BIAS_PARAMETERS display - Check BIASTEMP, BIASP9V,BIASM9V,BIASP5V Go to step 6. |  |  |
| 6 | In SCOS open DPU_AND_OBS_PARAMETERS display Check MODE HK parameter, it should be DRCU_ON (RAW=0x100) |  |  |

Table 1. Initial configuration check

# Test procedure

The following two sections describe general pass/fail criteria (Section 3.1), the general test layout (Section 3.1) and the detailed procedure for each functional test (Section 3.2).

## GENERAL PASS/FAIL CRITERIA

The general criteria for declaring a single test failed is the repeated failure of 2 consecutive runs of this test. In that case the functional test procedure should be aborted and the overall functional testing declared FAILED.

In the case of a ‘first run’ failure followed by a successful execution a third run of the same test should be performed and in the unlikely event of this third run being a failure the test procedure should be also aborted and the overall functional testing declared FAILED, as this would imply a not reliable operability of the instrument.

As a general remark ANY failure should be closely analyzed.

Note:  If the functional test is declared FAILED refer to section 4.1 for instrument switch OFF.

## GENERAL TEST PROCEDURE LAYOUT

The table below shows the general CFT sequence as it should be performed. In each step of this procedure the operator should refer to the detailed procedure in Section 3.2 .Test Control TCL scripts are available to invoke the correspondent CUS script stored in the HCSS database for each functional test. These CUS scripts will generate the appropriate command sequence for the particular functional test.

| Step | Subsystem tested | Test Id | Test Purpose |
| --- | --- | --- | --- |
| 1 | SCU | FUNC-SCU-01 | SCU Nominal Science Generation Check |
| 2 |  | FUNC-SCU-03 | FPU DC Thermometry Check |
| 3 |  | FUNC-SCU-06 | FPU AC Thermometry Check |
| 4 |  | FUNC-SCU-07 | Sorption Cooler Check |
| 5 |  | FUNC-PCAL-01 | Photometer Calibrator Characterisation |
| 6 |  | FUNC-SCAL-01 | Spectrometer Calibrators Characterisation |
| 7 | MCU | FUNC-MCU-01 | MCU Boot Check |
| 8 |  | FUNC-MCU-02 | MCU Nominal Science Generation Check |
| 9 | BSMm | FUNC-BSM-01 | BSM Switch ON Check |
| 10 |  | FUNC-BSM-03 | BSM Open Loop dynamics Check |
| 11 |  | FUNC-BSM-05A | BSM Open Loop chop test |
| 12 |  | BSM-INIT | BSM Initialisation Procedure |
| 13 |  | FUNC-BSM-05B | BSM Close Loop chop test |
| 14 |  | FUNC-BSM-06 | BSM Operational Mode Check |
| 15 | SMECm | FUNC-SMEC-01 | SMECm Switch ON Check |
| 16 |  | FUNC-SMEC-04A | SMEC Open Loop Position Test |
| 17 |  | FUNC-SMEC-09 | SMEC Open Loop Scan Test |
| 18 |  | SMEC-INIT* | SMEC Initialisation Procedure |
| 19 |  | FUNC-SMEC-03 | SMEC LED Output Characterisation |
| 20 |  | FUNC-SMEC-04B | SMEC Close Loop Position Test |
| 21 |  | FUNC-SMEC-07 | SMEC Close Loop Scan Test |
| 22 |  | FUNC-SMEC-06 | SMEC Close Loop Saw Tooth Scan Test |
| 23 | DCU | FUNC-DCU-01 | DCU Nominal Science Generation Check |
| 24 | Photometer LIAs | FUNC-DCU-04P | Photometer LIAs Check |
| 25 |  | FUNC-DCU-05P | Photometer Manual Offset Setting Check |
| 26 |  | FUNC-DCU-11P | Photometer Detectors Switch ON Check |
| 27 | Photometer BDAs | FUNC-DCU-013P | Photometer Detectors Check |
| 28 | Spectrometer LIAs | FUNC-DCU-04S | Spectrometer LIAs Check |
| 29 |  | FUNC-DCU-05S | Spectrometer Manual Offset Setting Check |
| 30 |  | FUNC-DCU-11S | Spectrometer Detectors Switch ON Check |
| 31 | Spectrometer BDAs | FUNC-DCU-013S | Spectrometer Detectors Check |

Table 2. General CFT sequence

*Note 1: This procedure is not a functional test, is a close loop initialisation procedure required to test the close loop operability of the BSM.

*Note 2: This procedure is not a functional test, is a close loop initialisation procedure required to test the close loop operability of the SMEC.

## DETAILED TEST PROCEDURE

The following is a detailed (test by test) procedure including the steps required to perform each functional test individually.

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open SCU_PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-SCU-01

| Test Id: | FUNC-SCU-01 |
| --- | --- |
| Test Purpose: | SCU Nominal Science Generation Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : Two SCU Nominal Science telemetry packets are received on QLA with the following characteristics: APID Type Subtype SID FrameID Frame length 0x508 21 1 0xA20 0x20 0x1E The frame time difference between consecutive SCU frames within these packets corresponds to the sampling rate. Nominal SCU sampling rate is 80Hz ( Δt = 12.5 ms The SPIRE HK parameter SCUFRAMECNT increments by 31. No events are generated during the frame generation. QLA to give go ahead. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Write down the initial value of SCUFRAMECNT parameter located in SCU_PARAMETERS display. |
| 2 | Run QLA script FUNC-SCU-01.py on QLA console. |
| 3 | Run FUNC-SCU-01 test procedure from the HCSS Test Procedure window on TOPE |
| 4 | Write down the final value of SCUFRAMECNT. |
| 5 | Contingency: If test fails repeat steps 1 to 4. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-01 | SCUFRAMECNT | n/ n+ 31 |  |  |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SCU-03

| Test Id: | FUNC-SCU-03 |
| --- | --- |
| Test Purpose: | FPU DC Thermometry Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if thermometry channels show temperature values indicating a correct functioning of the sensor, not open/short-circuited. If ANY reading is anomalous check RAW sensor reading. Open Circuit Criterion: RAW reading in the range [0, -100] Short Circuit Criterion: RAW reading of -32768 |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Write down the current value of SCUTEMPSTAT and the values of the 16 FPU temperatures located in SCU_PARAMETERS display. |
| 2 | Contingency: If test fails: Execute SCU_OFF procedure. Execute FUNC-SCU-03 procedure. Repeat step 1 of the Test Procedure. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-03 | SCUTEMPSTAT | 0xFFFF/0xFFFF |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SCU-06

| Test Id: | FUNC-SCU-06 |
| --- | --- |
| Test Purpose: | FPU AC Thermometry Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | At ~ 2K the SUBKTEMP reading should calibration should start being in range. Open Circuit Criterion: RAW reading in the range 0 -100 Short Circuit Criterion: RAW reading of -32768 |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-SCU-06 test procedure from the HCSS Test Procedure window on TOPE. |
| 2 | When the test is finished Write down the current value of SUBKSTAT located in SCU_PARAMETERS display. And write down the RAW reading of the sensor. |
| 3 | Contingency: If test fails : Send manual command: SEND_DRCU_COMMAND Parameter1 = 0xA0860000 Parameter2 = 0 Then repeat steps 1 and 2 of the Test Procedure. Note: If the test fails and the SUBKTEMP channel is switched OFF manually, the expected value before/after execution of FUNC-SCU-06 for SUBKSTAT is 0/1 |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-06 | SUBKSTAT SUBKTEMP | 1/1 ? |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SCU-07

| Test Id: | FUNC-SCU-07 |
| --- | --- |
| Test Purpose: | Sorption Cooler Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if during the execution of the test the following SCU HK parameters give correspondent readings of: SCU HK parameter RAW Converted SPHSV ~12715 ~323mV EVHSV ~12715 ~323mV SPHTRV ~14390 ~ 8 V |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-SCU-07 test procedure from the HCSS Test Procedure window on TOPE. |
| 2 | While the test is running Write down the values of current values of SPHSV, EVHSV, SPHTRV located in SCU_PARAMETERS display. (RAW and CONVERTED) |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-07 | SPHSV EVHSV SPHTRV | 0/ ~ 323 mV 0/ ~ 323 mV 0/ ~ 8 V |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-PCAL-01

| Test Id: | FUNC-PCAL-01 |
| --- | --- |
| Test Purpose: | Photometer Calibrator Characterisation |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : PCALCURR HK parameter shows the commanded current. PCALV parameter shows a linear increase proportional to the bias applied. (the proportionality constant in this case should be the PCAL resistor value) |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | Run QLA script FUNC-PCAL-01.py on QLA console. |
| 2 | Run FUNC-PCAL-01 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-05 | PCALCURR PCALV | 0/0.5,2,3.5,5,6.5,7mA |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SCAL-01

| Test Id: | FUNC-PCAL-01 |
| --- | --- |
| Test Purpose: | Spectrometer Calibrator Characterisation |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : SCAL4CURR HK parameter shows the commanded current sequence (1,2,3,4,5,5.5mA) SCAL2CURR HK parameter shows the commanded current sequence(1,2,3,4,5,5.5mA) SCA2LV parameter shows a linear increase proportional to the bias applied. (the proportionality constant in this case should be the SCAL2V resistor value) SCAL4V parameter shows a linear increase proportional to the bias applied. (the proportionality constant in this case should be the SCAL4V resistor value) SCAL2TEMP and SCAL4TEMP values follow the increased bias settings |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | Run QLA script FUNC-SCAL-01.py on QLA console. |
| 2 | Run FUNC-SCAL-01 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-06 | SCAL2CURR SCAL4CURR SCAL2V SCAL4V SCAL2TEMP SCAL4TEMP |  |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open MCU_PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-MCU-01

| Test Id: | FUNC-MCU-01 |
| --- | --- |
| Test Purpose: | MCU Boot Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if: MCU boots. MCU voltages show expected values. MAC Board Temperature Reading shows ambient temperature. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-MCU-01 test procedure from the HCSS Test Procedure window on TOPE |
| 2 | When procedure is finished Write down the values of the MCU voltages. |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-01 | MCUP5V MCUP15V MCUP14V MCUM14V MCUM15V MCUMACTEMP MCUBSMSTEMP MCUSMECTEMP | N/A/ ~ 5V N/A / ~15V N/A / ~ 14V N/A / ~ -14V N/A / ~ -15V N/A / ~ 300K N/A / ~ 300K N/A / ~ 300K |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-MCU-02

| Test Id: | FUNC-MCU-02 |
| --- | --- |
| Test Purpose: | MCU Nominal Science Generation Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if : MCU produces each type of the frames requested and with the following characteristics. Frame APID Type Subtype SID FrameID Frame length Eng. 0x508 21 3 0x814 0x14 0x15 BSM 0x508 21 1 0x612 0x12 0xD SMEC 0x508 21 1 0x410 0x10 0xC BSM +SMEC No events are generated during the different frames generation. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Write down the current value of MCUFRAMECNT located in MCU_PARAMETERS display |
| 2 | Run QLA script FUNC-MCU-02.py on QLA console. |
| 3 | Run FUNC-MCU-02 test procedure from the HCSS Test Procedure window on TOPE |
| 4 | When test is finished Write down the current value of MCUFRAMECNT. |
| 5 | Contingency: If test fails repeat steps 1 to 4. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-02 | MCUFRAMECNT | 0/ ~ 6600 |  |  |  |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open CHOP PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-BSM-01

| Test Id: | FUNC-BSM-01 |
| --- | --- |
| Test Purpose: | BSM Switch ON Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if: CHOPSENSPWR HK parameter goes from 0 to 1 CHOPSENSIG HK parameter changes JIGGSENSPWR HK parameter goes from 0 to 1 JIGGSENSSIG HK parameter changes |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-BSM-01 test procedure from the HCSS Test Procedure window on TOPE |
| 2 | When the test is finished record all the Key parameters noted bellow |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-01 | CHOPSENSPWR CHOPLOOPMODE CHOPFFGAIN CHOPSENSSIG JIGGSENSPWR JIGGLOOPMODE JIGGFFGAIN JIGGSENSSIG | 0/1 3/3 0xBEB/0xBEB ?/0x8E5A 0/1 3/3 0xBEB/0xBEB ?/ 0x9164 |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-BSM-03

| Test Id: | FUNC-BSM-03 |
| --- | --- |
| Test Purpose: | BSM Open Loop dynamics Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if the chop sensor signal evolves in the same way as the positions set.(i.e if choppos2 > choppos1 ( chopsenssig2 > chopsenssig1) for each jiggle position. |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | On QLA open up 2 time series display with the following HK parameters in each display: Display 1: HK : CHOPPOSN BSM Nominal Science: BSMCHOPMOTORCURR BSMCHOPSENSSIG BSMCHOPMOTORVOLT Display2: HK : JIGGPOSN BSM Nominal Science: BSMJIGGMOTORCURR BSMJIGGSENSSIG BSMJIGGMOTORVOLT |
| 2 | Run FUNC-BSM-03 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat step 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-03 |  |  |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-BSM-05A

| Test Id: | FUNC-BSM-05A |
| --- | --- |
| Test Purpose: | BSM Open Loop chop test (Degraded operational mode check) |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Note: The purpose of this test is to check the effectiveness of the dumping of the natural oscillations of the BSM in chop axis via motor bemf, through the use of the commendable motor resistance value. As it is now is just a check of the default motor resistance value. NO adjusting of this value should be attempted during this test. The success criteria are therefore not applicable. |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | On QLA open up a time series display of HK parameters: BSMCHOPSENSSIG BSMCHOPMOTORCURR BSMCHOPMOTORVOLT |
| 2 | Run FUNC-BSM-05A test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: None contemplated. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-05A | BSMCHOPSENSSIG BSMCHOPMOTORCURR BSMCHOPMOTORVOLT | ?/? ?/? ?/? |  | N/A | N/A |
| Start time @: End time @: OBSID: Comments: |

### FUNC-BSM-05B

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Execute BSM_INIT from HCSS Test Procedures |  |

| Test Id: | FUNC-BSM-05B |
| --- | --- |
| Test Purpose: | BSM Close Loop chop test |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Note: Currently this test does not differ at ALL from the next one. In any case the success/fail criteria are NOT applicable for this test. |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | On QLA open up a time series display of HK parameters: BSMCHOPSENSSIG BSMCHOPMOTORCURR BSMCHOPMOTORVOLT |
| 2 | Run FUNC-BSM-05B test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: None contemplated. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-06 | BSMCHOPSENSSIG BSMCHOPMOTORCURR BSMCHOPMOTORVOLT | ?/? ?/? ?/? |  | N/A | N/A |

### FUNC-BSM-06

| Test Id: | FUNC-BSM-06 |
| --- | --- |
| Test Purpose: | BSM Operational Mode Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Note: The purpose of this test is to check the effectiveness of the BSM close loop initialisation procedure and the default PID parameters. If the dynamical behaviour of the BSM during chopping with these PID parameters is close or within requirements this indicates that the PID parameters used can be applied to cold testing with certain adjustment. If NOT these indicates that the PID parameters need further tuning BUT NOT TO BE DONE DURING THESE TEST. In any case the success/fail criteria are NOT applicable for this test. |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | On QLA open up a time series display of HK parameters: BSMCHOPSENSSIG BSMCHOPMOTORCURR BSMCHOPMOTORVOLT |
| 2 | Run FUNC-BSM-06 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: None contemplated. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-06 | BSMCHOPSENSSIG BSMCHOPMOTORCURR BSMCHOPMOTORVOLT | ?/? ?/? ?/? |  | N/A | N/A |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute BSM_OFF from HCSS Test Procedures |  |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open SMEC PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-SMEC-01

| Test Id: | FUNC-SMEC-01 |
| --- | --- |
| Test Purpose: | SMECm Switch ON Check.. Encoder and LVDT alive check. |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if : SMECENCPWR HK parameter changes from 0 to 6. SMEC encoder signals 1 and 2 show variation when encoder is switched ON. SMEC LVDT is switched ON. SMEC LVDT DC and AC signals show variation when LVDT is switched ON. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a display of the following HK parameters: SMECENCPWR SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG |  |
| 2 | Run FUNC-SMEC-01 test procedure from the HCSS Test Procedure window on TOPE |  |
|  | Contingency: If test fails repeat steps 1. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-01 | SMECENCPWR SMECLVDTPWR SMECENCSIG1 SMECENCSIG2 |  |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SMEC-04A

| Test Id: | FUNC-SMEC-04A |
| --- | --- |
| Test Purpose: | SMEC Open Loop Position Test. SMEC displacement check. |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if SMECLVDTDCSIG parameter shows a variation according to the different positions set. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following MCU Engineering block parameters: MCUENGSMECENCSIG1 MCUENGSMECENCSIG2 MCUENGSMECLVDTDCSIG MCUENGSMECLVDTACSIG MCUENGSMECMOTORCURR |
| 2 | Run FUNC-SMEC-04a test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-04a | All above mentioned in step 2 | N/A | N/A | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SMEC-09

| Test Id: | FUNC-SMEC-09 |
| --- | --- |
| Test Purpose: | SMEC Open Loop Scan Test. |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if SMECLVDTDCSIG parameter shows a variation according to the different position of the SMEC along the scan. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following SMEC nominal science parameters: SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG SMECMOTORCURR |
| 2 | Run FUNC-SMEC-09 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-09 | All above mentioned in step 2 | N/A | N/A | N/A |  |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Execute SMEC_INIT from HCSS Test Procedures |  |

### FUNC-SMEC-03

| Test Id: | FUNC-SMEC-03 |
| --- | --- |
| Test Purpose: | SMEC LED Output Characterisation |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if: SMEC encoder signals 1 and 2 show a variation on their amplitudes from one LED illumination level to another. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following Nominal HK parameters: SMECENCSIGAMP1 SMECENCSIGAMP2 |
| 2 | Run FUNC-SMEC-03 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-03 | SMECENCPWR SMECENCSIG1AMP SMECENCSIG2AMP |  |  |  |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SMEC-04B

| Test Id: | FUNC-SMEC-04B |
| --- | --- |
| Test Purpose: | SMEC Close Loop Position Test. SMEC close loop operation check. |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if : SMECENCPOS HK parameter shows identical values as those of the SMECTRAJPOSN HK parameter for the different commanded SMEC positions.( 1mm to 25mm in steps of 1mm) |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up 2 time series displays: Display 1 of the following MCU Engineering block parameters: MCUENGSMECENCSIG1 MCUENGSMECENCSIG2 MCUENGSMECLVDTDCSIG MCUENGSMECLVDTACSIG MCUENGSMECMOTORCURR Display 1 of the following SMEC HK parameters: SMECENCPOSN SMECTRAJPOSN |
| 2 | Run FUNC-SMEC-04B test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-04B | All above mentioned in step 1 | N/A | N/A | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SMEC-07

| Test Id: | FUNC-SMEC-07 |
| --- | --- |
| Test Purpose: | SMEC Close Loop Scan Test. |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if: SMECENCPOS HK parameter shows identical values as those of the SMECTRAJPOSN HK parameter during the scan. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following SMEC nominal science parameters: SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG SMECMOTORCURR |
| 2 | Run FUNC-SMEC-07 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-07 | All above mentioned in step 1 | N/A | N/A | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-SMEC-06

| Test Id: | FUNC-SMEC-06 |
| --- | --- |
| Test Purpose: | SMEC Close Loop Saw Tooth Scan Test. |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test Passed if: If SMEC can perform a saw tooth scan. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following SMEC nominal science parameters: SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG SMECMOTORCURR |
| 2 | Run FUNC-SMEC-06 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-06 | All above mentioned in step 1 | N/A | N/A | N/A |  |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute SMEC_OFF from HCSS Test Procedures |  |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open DCU PARAMETERS SCOS Alpha Numeric Display |  |

### FUNC-DCU-01

| Test Id: | FUNC-DCU-01 |
| --- | --- |
| Test Purpose: | DCU Nominal Science Generation Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Success Criteria: | Test passed if DCU produces each type of DCU nominal science frame: APID Type S.type SID Frame ID Frame type Nb. Of frames Nb. of pkts. 0x504 21 1 0x200 0 PF 100 100 0x506 21 1 0x201 1 SF 100 17 0x504 21 2 0x102 2 PSW 100 34 0x504 21 2 0x103 3 PMW 100 25 0x504 21 2 0x104 4 PLW 100 12 0x506 21 2 0x105 5 SSW 100 12 0x506 21 2 0x106 6 SLW 100 7 |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Write down the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |
| 2 | Run FUNC-DCU-01 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Write down the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |
| 4 | Contingency: If test fails repeat steps 1 to 3. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-01 | DCUFRAMECNT | 0/700 |  |  |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-DCU-04P

| Test Id: | FUNC-DCU-04P |
| --- | --- |
| Test Purpose: | Photometer LIA Integrity Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Photometer LIAs ON |
| Success Criteria: | Test passed if : SCUDCDCSTAT parameter goes from 4 to 5. Photometer LIA card voltages are showing correct readings of +5V,+9V,-9V. Photometer LIA temperatures show an increase indicating that they are ON. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the HK parameters: PLIAP5V PLIAP9V PLIAM9V LIAP1/2/3/4/5/6/7/8/9TEMP |
| 2 | Run FUNC-DCU-04P test procedure from the HCSS Test Procedure window on TOPE |
| 5 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-04P | SCUDCDCSTAT PLIAP5V PLIAP9V PLIAM9V LIA1/2/3/4/5/6/7/9TEMP | 4/5 0/~ +5V 0/~+9V 0/~-9V N/A/ [290-300]K |  |  |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-DCU-11P

| Test Id: | FUNC-DCU-11P |
| --- | --- |
| Test Purpose: | Photometer BDAs Switch ON Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON + Photometer BIAS ON +Photometer JFETs ON |
| Success Criteria: | Test passed if Photometer JFET source and drain voltages are correct: PSWJFETVSS1/2/3/4/5/6 (RAW: 0x4C CONVERTED ~ -1.5V). PMLWJFETVSS1/2/3/4 (RAW: 0x4C CONVERTED ~ -1.5V). PSWJFETSTAT = 0x3F PMLWJFETSTAT = 0x3F |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-DCU-11P test procedure from the HCSS Test Procedure window on TOPE |
| 2 | After the test Write down the values RAW and converted values of: PSWJFETSTAT,PMLWJFETSTAT, PSWJFET1/2/3/4/5/6V PMWJFET1/2/3/4V PLWJFET1/2V located in DCU PARAMETERS AND |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-11P | PSWJFETSTAT PMLWJFETSTAT PSWJFET1/2/3/4/5/6V PMWJFET1/2/3/4V PLWJFET1/2V | 0/0x3f 0/0x3f 0V/-1.5V 0V/-1.5V 0V/-1.5V |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-DCU-13P

| Test Id: | FUNC-DCU-13P |
| --- | --- |
| Test Purpose: | Photometer Detectors Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON + Photometer BIAS ON +Photometer JFETs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON + Photometer BIAS ON +Photometer JFETs ON |
| Success Criteria: | Test passed if : The photometer detectors show a small linear variation on the output voltage when different bias is applied through the load curve. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of a couple of pixels on each of the photometer BDAs |
| 2 | Run FUNC-DCU-13P test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat step 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-13P |  |  |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | From TOPE HCSS Test Procedures run PDET-OFF |  |

### FUNC-DCU-04S

| Test Id: | FUNC-DCU-04S |
| --- | --- |
| Test Purpose: | Spectrometer LIAs Integrity Check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Success Criteria: | Test passed if : SCUDCDCSTAT parameter goes from 4 to 6. Spectrometer LIA card voltages are showing correct readings of +5V,+9V,-9V. Spectrometer LIA temperatures show an increase indicating that they are ON. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the HK parameters: SLIAP5V SLIAP9V SLIAM9V LIAS1/2/3TEMP |
| 2 | Run FUNC-DCU-04S test procedure from the HCSS Test Procedure window on TOPE |
| 5 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-04S | SCUDCDCSTAT SLIAP5V SLIAP9V SLIAM9V LIA1/2/3TEMP | 4/6 0/~ +5V 0/~+9V 0/~-9V N/A/ [290-300]K |  |  |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-DCU-11S

| Test Id: | FUNC-DCU-11S |
| --- | --- |
| Test Purpose: | Spectrometer BDAs switch ON check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON + Spectrometer BIAS ON + Spectrometer JFETs ON |
| Success Criteria: | Test passed if SCUDCDCSTAT goes from 4 to 6, Spectrometer LIAs voltages are correct and SJFET voltages are also correct. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-DCU-11S test procedure from the HCSS Test Procedure window on TOPE |
| 2 | After the test Write down the values RAW and converted values of: LIASTAT SLIAP5V, SLIAP9V, SLIAN9V, SSWJFETSTAT,SLWJFETSTAT, SSWJFET1V,SLWJFET2V located in DCU PARAMETERS AND |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-11S | SSWJFETSTAT SLWJFETSTAT SSWJFET1/2V SLWJFET1/2V | 0/7 0/7 0V/-1.5V 0V/-1.5V |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

### FUNC-DCU-13S

| Test Id: | FUNC-DCU-13S |
| --- | --- |
| Test Purpose: | Spectrometer detectors check |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON + Spectrometer BIAS ON + Spectrometer JFETs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON + Spectrometer BIAS ON + Spectrometer JFETs ON |
| Success Criteria: | Test passed if : The spectrometer detectors show a small linear variation on the output voltage when different bias is applied through the load curve. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of a couple of pixels on each of the spectrometer BDAs |  |
| 2 | Run FUNC-DCU-13S test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat steps 1 and 2 |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-13S |  |  |  | N/A |  |
| Start time @: End time @: OBSID: Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | From TOPE HCSS Test Procedures run SDET-OFF |  |

# END TEST SEQUENCE

## NORMAL END TEST SEQUENCE

The following table shows the necessary steps to be followed to end the warm functional test sequence.

| Step# | Action | HK parameters | Expected Value | Comments | Check |
| --- | --- | --- | --- | --- | --- |
| 1 | Check BSM is OFF | CHOPSENSPWR JIGGSENSPWR | 0 0 |  |  |
| 2 | Check SMEC is OFF | SMECENCPWR | 0 |  |  |
| 3 | Run MCU-OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT | 0 |  |  |
| 4 | Check Instrument Configuration | MODE | 0x100/ DRCU_ON |  |  |
| 5* | Run SCU_OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT SUBKSTAT | 0 0 |  |  |
| 6* | Run DRCU_OFF from the HCSS Test Procedure window on TOPE | MODE | 0/ DPU_ON |  |  |

* These two steps should ONLY be executed if the functional test takes place BEFORE the instrument is placed in the test cryostat.  If the instrument is already in the cryostat and the cryo-harness are connected this steps should NOT be executed.

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS 1 to 4 are executed is DRCU_ON

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS 1 to 6 are executed is DPU_ON

## END TEST SEQUENCE WHEN THE FUNCTIONAL TEST HAS FAILED

The following table shows the necessary steps to be followed to end the warm functional test sequence when this has been declared failed.

| Step# | Action | HK parameters | Expected Value | Comments | Check |
| --- | --- | --- | --- | --- | --- |
| 1 | Check BSM is OFF | CHOPSENSPWR JIGGSENSPWR | 0 0 |  |  |
| 2 | Check SMEC is OFF | SMECENCPWR | 0 |  |  |
| 3 | Run MCU-OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT | 0 |  |  |
| 4 | Check Instrument Configuration | MODE | 0x100/ DRCU_ON |  |  |
| 5 | Run SCU_OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT SUBKSTAT | 0 0 |  |  |
| 6 | Run DRCU_OFF from the HCSS Test Procedure window on TOPE | MODE | 0/ DPU_ON |  |  |

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS DPU_ON

# Annexe1 (Instrument nominal configuration/modes)

|  | OFF | INIT | DPU_ON | DRCU_ON | REDY | PHOT STBY | SPEC STBY | CREC | SAFE |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| DPU |  | ON | ON | ON | ON | ON | ON | ON | ON |
| Essential Hsk packets |  |  | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz |
| Normal Hsk packets |  |  | 1.0Hz | 1.0Hz | 0.25Hz | 0.25Hz | 0.25Hz | 1.0Hz | 0.25Hz |
| TC Acceptance |  | TBC | Yes | Yes | Yes | Yes | Yes | Yes | Yes |
| Event packets |  | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes |
| Science packets |  |  |  |  |  |  |  |  |  |
| VM |  |  |  |  |  |  |  | Cooler_Recycle |  |
| VM1 |  |  |  |  |  | Det_Temp_Ctrl |  |  |  |
| VM2 |  |  |  |  |  |  |  |  |  |
| VM3 |  |  |  |  |  |  |  |  |  |
| MODE |  |  | 0x0000 | 0x0100 | 0x0200 | 0x0300 | 0x0400 | 0x0600 | 0x0900 |
| DRCU |  |  | ON | ON | ON | ON | ON | ON |  |
| SCU |  |  |  |  |  |  |  |  |  |
| Temp Channels powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| SubK Channel powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| PCAL source powered |  |  |  |  |  |  |  |  |  |
| SCAL sources powered |  |  |  |  |  |  | TBD |  |  |
| TC Heater powered |  |  |  |  |  | Yes |  |  |  |
| Cooler SP Heater powered |  |  |  |  |  |  |  | Yes |  |
| Cooler EV HS powered |  |  |  |  |  |  |  | Yes |  |
| Cooler SP HS powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| DCU |  |  |  |  |  |  |  |  |  |
| Photometer BIAS |  |  |  |  |  | Yes |  |  |  |
| Photometer JFETS |  |  |  |  |  | Yes |  |  |  |
| Photometer LIAs |  |  |  |  |  | Yes |  |  |  |
| TC BIAS |  |  |  |  |  | Yes |  |  |  |
| TC JFETS |  |  |  |  |  | Yes |  |  |  |
| TC LIAs |  |  |  |  |  | Yes |  |  |  |
| Spectrometer BIAS |  |  |  |  |  |  | Yes |  |  |
| Spectrometer JFETS |  |  |  |  |  |  | Yes |  |  |
| Spectrometer LIAs |  |  |  |  |  |  | Yes |  |  |
| MCU |  |  |  |  |  |  |  |  |  |
| DSP |  |  |  |  | On | On | On | On |  |
| BSM |  |  |  |  |  | Hold | Hold |  |  |
| SMEC |  |  |  |  |  |  | Hold |  |  |