# Introduction

This document reports on the WARM FUNCTIONAL TESTS carried out (started) on the 15th March 2006 for  SPIRE to verify the instrument’s correct functioning.

## Scope

To judge the success or failure of a warm functional test by checking that:

- The telecommand sequence generated for a particular functional test is correctly received and executed on board by the SPIRE DPU.

- No error/event reports or command failures are generated during the execution of these commands.

- Telemetry is generated by the instrument as a result of telemetry requests to its different subunits.

- Particular telemetry parameters for each functional test change in an expected manner.

- A particular success criterion/criteria (specified in this document) is/are met.

## Reference Documents

| Ref | Document | Name | Version/Issue Nb. |
| --- | --- | --- | --- |
| RD01 | SPIRE-RAL-DOC-001652 | SPIRE Functional Tests Specification | Issue 1.4 |
| RD02 | SPIRE-RAL-DOC-001630 | SPIRE EGSE-ILT Start-Up Procedures | Issue 0.7 |
| RD03 | SPIRE-RAL-PRC-002222 | DRCU Switch ON Procedure | Issue 1.0 |
| RD04 | SPIRE-RAL-PRJ-001078 | SPIRE Data ICD | Issue 2.0 |
| RD05 | Sap-SPIRE-CCa-076-02 | DRCU/DPU Interface Control Document | Issue 1.2 |
| RD06 | LAM.PJT.SPI.NOT.011011 | MCU/DPU Command List ICD | Issue 5.0 |
| RD07 | SPIRE-IFS-PRJ-001391 | SPIRE OBS User Manual | Issue 2.1 |

## Change Record

| Document | Change date | Changes |
| --- | --- | --- |
| Issue 1.0 | 15th March 2006 | Created |

# Table of Contents

1.	Introduction	1

1.1	Scope	1

1.2	Reference Documents	1

1.3	Change Record	1

2.	FUNCTIONAL TEST CONFIGURATION	3

2.1	Software Configuration	3

2.2	EGSE Configuration Checks	3

2.3	SPIRE Instrument Configuration	4

3.	Test procedure	5

3.1	General Pass/Fail Criteria	5

3.2	Test Procedure Layout	5

3.3	Detailed Test Procedure	6

3.3.1	FUNC-SCU-01	6

3.3.2	FUNC-SCU-03	8

3.3.3	FUNC-SCU-06	9

3.3.4	FUNC-SCU-02	10

3.3.5	FUNC-SCU-08	11

3.3.6	FUNC-SCU-04	12

3.3.7	FUNC-SCU-05	14

3.3.8	FUNC-SCU-07	15

3.3.9	FUNC-MCU-01	16

3.3.10	FUNC-MCU-02	17

3.3.11	FUNC-MCU-03	18

3.3.12	FUNC-MCU-04	19

3.3.13	FUNC-BSM-01	21

3.3.14	FUNC-BSM-02c	22

3.3.15	FUNC-BSM-02j	23

3.3.16	FUNC-BSM-03	24

3.3.17	FUNC-BSM-06	25

3.3.18	FUNC-SMEC-02A/B	27

3.3.19	FUNC-SMEC-01	28

3.3.20	FUNC-SMEC-03	29

3.3.21	FUNC-SMEC-04a	30

3.3.22	FUNC-DCU-01	31

3.3.23	FUNC-DCU-02	32

3.3.24	FUNC-DCU-03	33

3.3.25	FUNC-DCU-04P	34

3.3.26	FUNC-DCU-05P	35

3.3.27	FUNC-DCU-11P	36

3.3.28	FUNC-DCU-13P	37

3.3.29	FUNC-DCU-04S	38

3.3.30	FUNC-DCU-05S	39

3.3.31	FUNC-DCU-11S	40

3.3.32	FUNC-DCU-13S	41

4.	END TEST SEQUENCE	42

4.1	Nominal End Test Sequence	42

4.2	Non-Nominal End Test Sequence	42

5.	Annexe1 (Instrument nominal configuration/modes)	44

# FUNCTIONAL TEST CONFIGURATION

## Software Configuration

Annotate in the table the current EGSE software configuration for the tests:

| EGSE component | Version/Build number | Comment |
| --- | --- | --- |
| SCOS2000 | SCOS2.3e Patch 5 | Operations done on the 14th March 2006: SCOS archives for PFM3_TEST1 under /spired/hfiles and /spired/TMD (12 GB free on /spired – OK to proceed with WFTs) Imported MIB PFM3_Issue2.1.G1 – success (Previous MIB PFM3_Issue2.1.E1_14032006.tar.gz, as used for the DRCU QM2 AT, is in ~/SCOS2.3eP5/data) Restarted SCOS Started Test Control Server on Lichfield – success |
| HCSS | Build #812 |  |
| QLA | 3.0 Alpha |  |
| QLA scripts | ? |  |
| Test Control scripts | See Annexe2 |  |
| CUS Scripts | See Annexe 2 |  |

## EGSE Configuration Checks

To check for the success of failure of a functional test, the real time telemetry of the instrument has to be monitored. The following applications must be running to do so. Before the test sequence starts, make the following checks:

| Workstation | EGSE component | Status | Check | Comments |
| --- | --- | --- | --- | --- |
| Lichfield | EGSE router | Started | ( | Started by typing router R |
| Lichfield | EGSE Gateway | Started | ( | Started by typing router S –scosapids 1280,1282,2036-2038 |
| Lichfield | Telemetry Ingestion | Started | ( |  |
| Lichfield | Packet Display | Started | ( |  |
| Lincoln | SCOS2000 | Started | ( |  |
| Lincoln | EXIF + TOPE | Started | ( |  |
| Lincoln | Manual Stack | Started | ( |  |
| Gordon | CDMS Simulator | Started | ( | CDMS telemetry appearing in PacketDisplay |
| Lichfiehd | Test Control Server | Started | ( |  |

## SPIRE Instrument Configuration

The functional test flow is such that at a given time during the procedure is possible that the instrument configuration does not conform to any nominal configuration as specified in Annexe 1.  Nevertheless the configuration prior and after a test is specified for each functional test, and were applicable and the instrument configuration is in accordance to Anexe1, this will be denoted by showing the main configuration in blue bold font and any additional subsystems configuration in black.

The initial instrument configuration for WFT does conform to a ‘known’ configuration and should be SPIRE DRCU_ON.

Perform the following actions to ensure that the instrument is in the correct configuration for the tests.

| Step# | Action | Comments | Check |
| --- | --- | --- | --- |
| 1 | In SCOS open DPU_AND_OBS_PARAMETERS display Check if : TM2N is incrementing by one @1Hz. TM1N is incrementing by one @0.5Hz. - If they are, go to step 5. - If they are not, go to step 2. | 15/03/06: No telemetry currently appears in the displays as a new archive was created yesterday. Will go to step 2. | ( |
| 2 | Check if the DPU is powered ON: - If ON, the DPU power supply LCD will show ~ 28V and 0.40A, go to step 4. - If not ON, power ON DPU see comments on the right, then go to step 3. | 15/03/06: DPU is switched OFF. Will go to step 3. | ( |
| 3 | In SCOS open Boot_ROM_Memory_Check display and check no errors are reported: - If no errors are reported, execute DPU_ON from HCSS Test Procedures. Then repeat step 1. - If the (5,1) contains errors: Check the error code in RD07. Then switch OFF the DPU and repeat step 2 | 15/03/06: DPU produces an event packet (5,1) as expected after switch ON with the following contents: ID: 0D00 Seq: C000 Len: 0019 0000: 0D00 C000 0019 0005 0100 8000 0000 0000 8008 0003 ABAB CDCD 0051 AAAA 0000 AE6C No errors reported on the penultimate field. Note: SCOS display shows an erroneous reading of 0xABAB errors, probably due to the change of the length of this packet in the latest OBS version. Executed DPU_START procedure, the name of the script should be corrected in the procedure. Received the following events as expected: ID: 0D00 Seq: C000 Len: 0019 0000: 0D00 C000 0019 0005 0100 8000 0000 0000 8008 0003 ABAB CDCD 0051 AAAA 0000 AE6C ID: 0D00 Seq: C001 Len: 0019 0000: 0D00 C001 0019 0005 0100 8000 0000 0000 8008 0003 1D00 C000 0011 CAFE 0001 8D01 ID: 0D00 Seq: C000 Len: 001D 0000: 0D00 C000 001D 0005 0100 8000 0000 073A 0521 510F 0000 0D05 0000 0000 0000 9800 0020: 0000 2B58 ID: 0D00 Seq: C001 Len: 001D 0000: 0D00 C001 001D 0005 0100 8000 0000 0A0B 0522 5110 0000 0D05 0000 0000 0001 A800 0020: 0000 D389 ID: 0D00 Seq: C002 Len: 001D 0000: 0D00 C002 001D 0005 0100 8000 0000 0CDB 0520 510E 0000 0D05 0000 0000 0002 8C12 0020: 0000 2F97 ID: 0D00 Seq: C004 Len: 001F 0000: 0D00 C004 001F 0005 0100 8000 0000 18DA 0512 5117 0000 0D05 0000 0000 0003 0000 0020: 0000 0000 68DF Will proceed to step 1 as specified in the procedure. | ( |
| 4 | Execute define_new_HK_report.tcl HCSS Test procedure. |  |  |
| 5 | In SCOS open SCU_PARAMETERS display - If SCUP5V/P9V/M9V are jittering and BIAS_PARAMETERS display - If BIAS1TEMP,BIAS2TEMP,… show ambient temperature, the DRCU is ON.Go to step 6. - If DRCU is not ON, refer to RD03 on how to start up the DRCU. | 15/03/06: DRCU is not ON. Will clear Nominal and Critical HK reports and proceed to DRCU switch ON procedure. Run clear_HK_report.tcl procedure and received an event (5,1) with the following contents: ID: 0D00 Seq: C164 Len: 001D 0000: 0D00 C164 001D 0005 0100 5AAA 9C8A 0424 0511 5111 0000 0D05 0000 0000 0004 0000 0020: 0003 664C Used manual stack UpdateHKTable_21G.lincoln to update Nominal HK table contents. Both commands successful. Will power UP the DRCU now: 1. Powered UP main DRCU power and secondary ( DRCU power display shows 28.0V and 0.38 A) 2. Run DRCU_ON HCSS Test procedure 3. Opened SCU and DCU displays ALL SCU VOLTAGES LOOKING GOOD. SCUP5V = 5.22V SCUP9V =9.0652V SCUM9V = -9.07V ALL BIAS VOLTAGES LOOKING GOOD. BIASP5V = 5.12V BIASP9V = 8.95V BIASM9V= -9.08V | ( |
| 6 | In SCOS open DPU_AND_OBS_PARAMETERS display and check that the MODE housekeeping parameter is DRCU_ON. | 15/03/06; MODE (RAW)= 0x100 MODE (ENG) = DRCU_ON | ( |

Table 1. Initial configuration check

# Test procedure

The following two sections describe general pass/fail criteria (Section 3.1), the general test procedure layout (Section 3.1) and the detailed procedure for each functional test (Section 3.2).

## GENERAL PASS/FAIL CRITERIA

The general criterion for declaring a single test failed is the repeated failure of 2 consecutive runs of this test. In that case the functional test procedure should be aborted and the overall functional testing declared FAILED.

In the case of a ‘first run’ failure followed by a successful execution a third run of the same test should be performed and in the unlikely event of this third run being a failure the test procedure should be also aborted and the overall functional testing declared FAILED, as this would imply a not reliable operability of the instrument.

As a general remark ANY failure should be closely analyzed.

Note:  If the functional test is declared FAILED refer to section 4.1 for instrument switch OFF.

## GENERAL TEST PROCEDURE LAYOUT

The table below shows the general WFT sequence as it should be performed. In each step of this procedure the operator should refer to the detailed procedure in Section 3.2 .Test Control TCL scripts are available to invoke the correspondent CUS script stored in the HCSS database for each functional test. These CUS scripts will generate the appropriate command sequence for the particular functional test.

| Step | Subsystem tested | Test Id | Test Purpose |
| --- | --- | --- | --- |
| 1 | SCU | FUNC-SCU-01 | SCU Nominal Science Generation Check |
| 2 |  | FUNC-SCU-03 | FPU DC Thermometry Check |
| 3 |  | FUNC-SCU-06 | FPU AC Thermometry Check |
| 4 |  | FUNC-SCU-02 | SCU Nominal Science Contents Check |
| 5 |  | FUNC-SCU-08 | SCU Test Pattern Check |
| 6 |  | FUNC-SCU-07 | Sorption Cooler Check |
| 7 |  | FUNC-SCU-04 | Photometer Calibrator Check |
| 8 |  | FUNC-SCU-05 | Spectrometer Calibrators Check |
| 9 | MCU | FUNC-MCU-01 | MCU Boot Check |
| 10 |  | FUNC-MCU-02 | MCU Nominal Science Generation Check |
| 11 |  | FUNC-MCU-03 | MCU Nominal Science Contents Check |
| 12 |  | FUNC-MCU-04 | MCU Test Pattern Check |
| 13 | BSMm | FUNC-BSM-01 | BSM Switch ON Check |
| 14 |  | FUNC-BSM-02c | BSM Chop Sensor Polarity Check |
| 15 |  | FUNC-BSM-02j | BSM Jiggle Sensor Polarity Check |
| 16 |  | FUNC-BSM-03 | BSM Open Loop dynamics Check |
| 17 |  | BSM_INIT* | Close loop on chop/jiggle axis. |
| 18 |  | FUNC-BSM-06 | BSM Operational Mode Check |
| 18 | SMECm | FUNC-SMEC-02A/B* | SMEC Launch Latch Open/Close Check |
| 19 |  | FUNC-SMEC-01 | SMECm Switch ON Check |
| 20 |  | FUNC-SMEC-03 | SMEC LED Optical Encoder LED Check |
| 21 |  | FUNC-SMEC-04 | SMEC Positioning Test (Open Loop) |
| 22 | DCU | FUNC-DCU-01 | DCU Nominal Science Generation Check |
| 23 |  | FUNC-DCU-02 | DCU Nominal Science Contents Check |
| 24 |  | FUNC-DCU-03 | DCU Test Pattern Check |
| 25 | Photometer LIAs | FUNC-DCU-04P | Photometer LIAs Check |
| 26 |  | FUNC-DCU-05P | Photometer Manual Offset Setting Check |
| 27 |  | FUNC-DCU-11P | Photometer Detectors Switch ON Check |
| 28 | Photometer BDAs | FUNC-DCU-013P | Photometer Detectors Check |
| 29 | Spectrometer LIAs | FUNC-DCU-04S | Spectrometer LIAs Check |
| 30 |  | FUNC-DCU-05S | Spectrometer Manual Offset Setting Check |
| 31 |  | FUNC-DCU-11S | Spectrometer Detectors Switch ON Check |
| 32 | Spectrometer BDAs | FUNC-DCU-013S | Spectrometer Detectors Check |

Table 2. General WFT sequence

*Note: If the procedure is executed prior to an FPU cooldown FUNC-SMEC-02A should be executed. If the procedure is executed after an FPU warm up FUNC-SMEC-02B should be executed.

*Note: This procedure is not a functional test, is a close loop initialisation procedure required to test the close loop operability of the BSM.

## DETAILED TEST PROCEDURE

The following is a detailed (test by test) procedure including the steps required to perform each functional test individually.

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open SCU_PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-SCU-01

| Test Id: | FUNC-SCU-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON |
| Final Configuration: | DRCU_ON |
| Success Criteria: | Test passed if : Two SCU Nominal Science Report telemetry packets are received on QLA with the following characteristics: APID Type Subtype SID FrameID Frame length 0x508 21 1 0xA20 0x20 0x1E The frame time difference between consecutive SCU frames within these packets corresponds to the sampling rate. Nominal SCU sampling rate is 80Hz ( Δt = 12.5 ms The SPIRE HK parameter SCUFRAMECNT increments by 31. No events are generated during the frame generation. QLA to give go ahead. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the initial value of SCUFRAMECNT parameter located in SCU_PARAMETERS display and the initial value of TM1N located in DPU_AND-OBS_PARAMETERS display. | SCUFRAMCNT = 0 |
| 2 | Run QLA script FUNC-SCU-01.py on QLA console. |  |
| 3 | Run FUNC-SCU-01 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | Annotate the final value of SCUFRAMECNT and TM1N. | SCUFRAMCNT = 31 |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-01 | SCUFRAMECNT TM1N | n/ n+ 31 | 0/ 31 Not stable as critical HK is running | 31 | Success |
| Start time @:16:39 End time @: 16:40 OBSID: 0x3000E003 Comments: QLA produces QLA-SCU-01_3000E003.txt file which shows that the contents of the SCU nominal science telemetry packets are as expected according to success criterion. |

### FUNC-SCU-03

| Test Id: | FUNC-SCU-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON |
| Final Configuration: | DRCU_ON + DC thermometry ON |
| Success Criteria: | Test passed if all FPU DC thermometry sensors show temperature values indicating a correct functioning of the sensor, not open/short-circuited. If ANY reading is anomalous check RAW sensor reading. Open Circuit Criterion: RAW reading in the range [0, -100] Short Circuit Criterion: RAW reading of -32768 Note: For some parameters the calibration curve above 75K has only 2 points, thus the linearly interpolated temperature reading given by SCOS is usually not correct at T > 75K. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-SCU-03 test procedure from the HCSS Test Procedure window on TOPE |  |
| 2 | When the test is finished annotate the current value of SCUTEMPSTAT and the RAW/converted values of the 16 FPU temperatures located in SCU_PARAMETERS display. |  |
| 3 | Contingency: If test fails execute SCU_OFF procedure from HCSS Test Procedure window on TOPE and then repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-03 | SCUTEMPSTAT 0 | 0/0xFFFF 0/0xffff | 0/0xffff | N/A | Pass |
| Start time @:16:43 End time @:16:44 OBSID:0x3000E004 Commets: PUMPHTRTEMP = 55.22 K RAW=-32768 PUMPHSTEMP = 37.3 K RAW=-32768 EVAPHSTMP = 38.52 K RAW=-32768 SHUNTTEMP = 19.94 K RAW=-32768 EMCFIELDTMP = 284.03 K RAW = -29544 SL0TEMP = 19.89 K RAW =-32768 PL0TEMP= 21.3 K RAW = -32768 OPTTEMP =148.48 K RAW = -32768 BAFTEMP = 187.81 RAW =-32768 BSMIFTEMP =97.61 K RAW =-32768 SCAL2TEMP = 85.99 K RAW=-32768 SCAL4TEMP =88.89 K RAW =-32768 SCALTEMP=86.61K RAW=-32768 SMECIFTEMP=93.13K RAW=-32768 SMECTEMP=8.39K RAW=-32768 BSMTEMP=12.14 K RAW=-32768 |

### FUNC-SCU-06

| Test Id: | FUNC-SCU-06 |
| --- | --- |
| Initial Configuration: | DRCU_ON + DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if SUBKSTAT parameter went from 0 to 1. Open Circuit Criterion: RAW reading in the range 0 -100 Short Circuit Criterion: RAW reading of -32768 |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-SCU-06 test procedure from the HCSS Test Procedure window on TOPE. |  |
| 2 | When the test is finished annotate the current value of SUBKSTAT located in SCU_PARAMETERS display. Also write down the RAW value of the SUBKTEMP parameter. |  |
| 3 | Contingency: If test fails : Send manual command: SEND_DRCU_COMMAND Parameter1 = 0xA0860000 Parameter2 = 0 Then repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-06 | SUBKSTAT 0 | 0/1 1 | 0/1 | N/A | Pass |
| Start time @:16:53 End time @:16:54 OBSID:0x3000E005 Commets:Success |

### FUNC-SCU-02

| Test Id: | FUNC-SCU-02 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : Parameters in the SCU Nominal science packets and the same parameters in the Nominal HK packet have similar RAW/converted values to within 10%. The SPIRE HK parameter SCUFRAMECNT located in SCU_PARAMETERS display increments by 31. No events are generated during the frame generation. QLA to give the go ahead. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of SCUFRAMECNT located in SCU_PARAMETERS display. |  |
| 2 | Run QLA script FUNC-SCU-02.py on QLA console. |  |
| 3 | Run FUNC-SCU-02 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | When the test is finished annotate the current value of SCUFRAMCNT. |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-02 | SCUFRAMECNT TM5N | n+31/n+62 | 31/62 | 31 | Pass |
| Start time @:16:56 End time @:16:57 OBSID:0x3000E006 Commets: All parameters the same between NHK and SCUNOMINAL except SCAL4CURR and TCHTRV SCAL4CURR NHK=13 SCU=0 – Time series plot of NHK on QLA shows the ‘normal’ jumping between 0 and 65535. This value on SCOS is showing a lot of noise (raw value jumping between -31 and 31 and this is probably the cause of the discrepancy. TCHTRV NHK=0 SCU=1 – This is noise so is OK |

### FUNC-SCU-08

| Test Id: | FUNC-SCU-08 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : Two SCU Diagnostic Science Report telemetry packets are received with the following characteristics: APID Type Subtype SID FrameID Frame length 0x508 21 3 0x1120 0x21 0x1E The SCU Test Pattern agrees with the reference test pattern. QLA to give go ahead. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current values of SCUFRAMECNT located in SCU_PARAMETERS display. |  |
| 2 | Run QLA script FUNC-SCU-08.py on QLA console. |  |
| 3 | Run FUNC-SCU-08 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | When the test is finished annotate the current value of SCUFRAMCNT. |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-08 | SCUFRAMECNT and SCU test pattern frame parameters | n+62/n+93 | 62/93 | 31 | Pass |
| Start time @:17:12 End time @:17:13 OBSID:0x3000E007 Commets: QLA has written file FUNC-SCU-08_3000E007 |

### FUNC-SCU-04

| Test Id: | FUNC-SCU-04 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if PCALCURR/PCALV SCU HK parameters show the following values: PCALCURR HK parameter which shows the measured PCAL current is ~ 0.1 mA. PCALV HK parameter which shows the measured PCAL voltage is ~ 0.02V |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of PCALV and PCALCURR located in SCU_PARAMETERS display. |  |
| 2 | Run FUNC-SCU-04 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | While the test is running Annotate the values of PCALV and PCALCURR. |  |
| 4 | Contingency: If test fails repeat steps 1 to 3. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-04 | PCALCURR PCALV | 0/0.1mA 0/0.02V | PCALV=0, 0.021 V PCALCURR=0, 0.0996 mA | N/A | Passed |
| Start time @:17:15 End time @:17:16 OBSID:0x3000E008 Commets: |

### FUNC-SCU-05

| Test Id: | FUNC-SCU-05 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : SCAL2CURR ,SCAL4CURR HK parameters which show the measured current read ~ 0.1 mA SCAL2V,SCAL4V parameters which show the measured voltage read ~ 0.05V. |

Test Procedure

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of SCAL2V ,SCAL2CURR,SCAL4V,SCAL4CURR located in SCU_PARAMETERS display. |  |
| 2 | Run FUNC-SCU-05 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | While the test is running write the values of SCAL2V ,SCAL2CURR,SCAL4V,SCAL4CURR. |  |
| 4 | Contingency: If test fails repeat steps 1 to 3. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-05 | SCAL2CURR SCAL2V SCAL4CURR SCAL4V | 0/0.1mA 0/0.05V 0/0.1mA 0/0.05V | 0, 0.1 mA 0, 0.05 V 0, 0.1 mA 0, 0.05 V | N/A | Pass |
| Start time @:17:18 End time @:17:21 OBSID:0x3000E009 Commets: |

### FUNC-SCU-07

| Test Id: | FUNC-SCU-07 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if during the execution of the test the following SCU HK parameters give correspondent readings of: SCU HK parameter RAW Converted SPHSV ~12715 ~323mV EVHSV ~12715 ~323mV SPHTRV ~14390 ~ 8 V |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-SCU-07 test procedure from the HCSS Test Procedure window on TOPE. |  |
| 2 | While the test is running annotate the values of current values of SPHSV, EVHSV, SPHTRV located in SCU_PARAMETERS display. (RAW and CONVERTED) |  |
| 3 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-07 | SPHSV EVHSV SPHTRV | 0/ ~ 323 mV 0/ ~ 323 mV 0/ ~ 8 V | 0.11, 323.94 mV 0.0189,323.13 mV 0.0037, 8.83 V | N/A | Pass |
| Start time @:17:23 End time @:17:26 OBSID:0x3000E00A Commets: |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open MCU_PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-MCU-01

| Test Id: | FUNC-MCU-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if: MCU boots. MCU voltages show expected values. MAC, SMEC and BSM board temperatures shows ambient temperature. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-MCU-01 test procedure from the HCSS Test Procedure window on TOPE |  |
| 2 | When procedure is finished Annotate the values of the MCU voltages. |  |
| 3 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-01 | MCUP5V MCUP15V MCUP14V MCUM14V MCUM15V MCUMACTEMP MCUSMECTEMP MCUBSMTEMP | N/A/ ~ 5V N/A / ~15V N/A / ~ 14V N/A / ~ -14V N/A / ~ -15V N/A / ~ 300K N/A / ~ 300K N/A / ~ 300K | 20V,5.0V 75.36,15.52 75.36,14.53 75.36,-14.61 75.36,-15.60 1651.51,289.84 1651.51,293.58 1651.51,293.13 | N/A | Pass |
| Start time @:18:04 End time @:18:06 OBSID:0x3000E00B Commets:SCUDCDCSTAT went from 0 to 4 as expected |

### FUNC-MCU-02

| Test Id: | FUNC-MCU-02 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if : MCU produces each type of the frames requested and with the following characteristics. Frame APID Type Subtype SID FrameID Frame length Eng. 0x508 21 3 0x814 0x14 0x15 BSM 0x508 21 1 0x612 0x12 0xD SMEC 0x508 21 1 0x410 0x10 0xC BSM +SMEC No events are generated during the different frames generation. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of MCUFRAMECNT located in MCU_PARAMETERS display |  |
| 2 | Run QLA script FUNC-MCU-02.py on QLA console. |  |
| 3 | Run FUNC-MCU-02 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | When test is finished annotate the current value of MCUFRAMECNT. |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-02 | MCUFRAMECNT TM1N | 0/ ~ 3300 | 0/6637 |  | Pass |
| Start time @:18:09 End time @:18:11 OBSID:0x3000E00C Commets: Used all the default parameters when running the TOPE script Framecount is correct procedure needs to be changed Packet parameters checked – all OK Events checked – no events |

### FUNC-MCU-03

| Test Id: | FUNC-MCU-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if : MCU produces 100 frames of each type of frames requested with the following characteristics: Frame APID Type Subtype SID FrameID Frame length Eng. 0x508 21 3 0x814 0x14 0x15 BSM 0x508 21 1 0x612 0x12 0xD SMEC 0x508 21 1 0x410 0x10 0xC BSM +SMEC No events are generated during the different frames generation. QLA analysis results are correct.(FIXME) QLA to give go ahead. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of MCUFRAMECNT located MCU_PARAMETERS display. |  |
| 2 | Run QLA script FUNC-MCU-03.py on QLA console. |  |
| 3 | Run FUNC-MCU-03 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | When test is finished annotate the current value of MCUFRAMECNT |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-03 | MCUFRAMECNT | n/ n+300 n~3300 | 6637/6934 |  | To Check |
| Start time @:18:16 End time @:18:17 OBSID:0x3000E00D Comments: Has not produced 300, instead has produced 297 QLA has produced files: QLA-MCU-03_3000E00D_8901.txt, QLA-MCU-03_3000E00D_8902.txt, QLA-MCU-03_3000E00D_8903.txt Errors found during playback of SCOS – computePkt: Multiple packets for packet 190001559 pktTime 2006-03-15T18:13:05 also at 18:11:51 and other times This is a SCOS error but hampers being able to determine the cause of the 297 frames. For the BSM it produced 99 frames, SMEC produced only 98 frames, Eng 100 frames This is likely to be due to an incorrect setting of the DSP scheduler unit time inside the CUS building block, this does not prevent a test pass. Asier to investigate. |

### FUNC-MCU-04

| Test Id: | FUNC-MCU-04 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if: MCU produces 100 frames of Test Pattern with the following characteristics: Frame APID Type Subtype SID FrameID Frame length Test 0x508 21 3 0x915 0x15 0x15 MCU Test pattern produced is the same as the previous time this test was run. QLA to give go ahead. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of MCUFRAMECNT located in MCU_PARAMETERS display. |  |
| 2 | Run QLA script FUNC-MCU-04.py on QLA console. |  |
| 3 | On Lincoln run FUNC-MCU-04 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | When test is finished annotate the current value of MCUFRAMECNT |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-04 | MCUFRAMECNT | m/ m+100 m~3600 | 6934, 7033 |  | Pass |
| Start time @:18:31 End time @:18:32 OBSID:0x3000E00E Comments: Produced 99 frames instead of 100 QLA checked test pattern, this is OK File produced QLA-MCU-04_3000E00E_8905 |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open CHOP PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-BSM-01

| Test Id: | FUNC-BSM-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if: CHOPSENSPWR HK parameter goes from 0 to 1 CHOPDACVAL HK parameter goes from 0 to 0x8000 CHOPSENSIG HK parameter shows variation from off to on JIGGSENSPWR HK parameter goes from 0 to 1 JIGGDACVAL parameter goes from 0 to 0x8000 JIGGSENSSIG HK parameter shows variation from off to on |

Test Procedure

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the following HK parameters: CHOPSENSPWR CHOPDACVAL CHOPSENSIG JIGGSENSPWR JIGGDACVAL JIGGSENSSIG |  |
| 2 | Run FUNC-BSM-01 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | When the test is finished record all the Key parameters noted bellow |  |
|  | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-01 | CHOPSENSPWR CHOPLOOPMODE CHOPDACVAL CHOPFFGAIN CHOPSENSSIG JIGGSENSPWR JIGGLOOPMODE JIGGDACVAL JIGGFFGAIN JIGGSENSSIG | 0/1 0/3 0/0x8000 0/0xBEB 0/? 0/1 0/3 0/0x8000 0/0xBEB 0/? | 0/1 3/3 8000/8000 BEB/BEB 7FE9/8E5A 0/1 3/3 8000/8000 BEB/BEB 7FE9/0x9164 | N/A | Pass |
| Start time @:18:44 End time @: OBSID:0x3000E00F Commets: The success criteria for the DACVAL should be corrected. |

### FUNC-BSM-02c

| Test Id: | FUNC-BSM-02c |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if the chop sensor signal evolves in the same way as the positions set.(i.e if (pos1 > pos2 ( sig1 > sig2) |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA open up a time series display of HK parameter CHOPDACVAL and CHOPSENSSIG |  |
| 2 | Run FUNC-BSM-02c test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat steps 1. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-02c | CHOPDACVAL CHOPSENSSIG |  | 0x8000 0x8E60 | N/A | Pass |
| Start time @:18:49 End time @:18:50 OBSID:0x3000E10 Comments:The BSM moved! |

### FUNC-BSM-02j

| Test Id: | FUNC-BSM-02j |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if the jiggle sensor signal evolves in the same way as the positions set.(i.e if (pos1 > pos2 ( sig1 > sig2) |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA open up a time series display of HK parameter JIGGDACVAL and JIGGSENSSIG |  |
| 2 | Run FUNC-BSM-02j test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat steps 1. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-02j | JIGGDACVAL JIGGSENSSIG |  | 0x8000 0x9150 | N/A | Pass |
| Start time @:18:58 End time @:18:59 OBSID:0x3000E012 Comments: It moved 0x3000E011 was mistakenly a repetition of 02c, it was run between 18:56 and 18:57 |

### FUNC-BSM-03

| Test Id: | FUNC-BSM-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if the chop sensor signal evolves in the same way as the positions set.(i.e if (pos1 > pos2 ( sig1 > sig2) for each jiggle position. Note: During warm tests the voltages on both chop and jiggle motors are likely to be saturated (CHOP/JIGGMOTORVOLT RAW values of ~ 0xFFFF) due to the high resistance of the motor coil at ambient temperature. |

Test Procedure

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA open up a time series display of HK parameters: CHOPPOSN CHOPDACVAL CHOPMOTORCURR CHOPSENSSIG CHOPMOTORBEMF JIGGPOSN JIGGDACVAL JIGGMOTORCURR JIGGSENSSIG JIGGMOTORBEMF |  |
| 2 | Run FUNC-BSM-03 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat step 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-03 |  |  |  | N/A | Pass |
| Start time @:19:06 End time @:19:11 OBSID:0x3000E013 Comments: Procedure needs changing from BEMF to CHOPMOTORVOLT Running the procedure with 0x2000 per step instead of 0x1000 and 3seconds per step instead of 5 seconds per step Jiggle position as given by senssig shows going to position then coming back as the chop position is stepped. This needs to be investigated. |

### FUNC-BSM-06

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Execute BSM_INIT from HCSS Test Procedures | Started at 19:20 Obsid :0x3000E014 |

| Test Id: | FUNC-BSM-06 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Note: The purpose of this test is to check the effectiveness of the BSM close loop initialisation procedure and the default PID parameters. If the dynamical behaviour of the BSM during chopping with these PID parameters is close or within requirements this indicates that the PID parameters used can be applied to cold testing with certain adjustment. If NOT these indicates that the PID parameters need further tuning BUT NOT TO BE DONE DURING THESE TEST. In any case the success/fail criteria are NOT applicable for this test. |

Test Procedure

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA open up a time series display of HK parameters: CHOPDACVAL CHOPMOTORCURR CHOPSENSSIG |  |
| 2 | Run FUNC-BSM-06 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: None contemplated. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-06 | CHOPSENSSIG CHOPMOTORCURR CHOPDACVAL | ?/? ?/? ?/? |  | N/A | N/A |
| Start time @:19:30 End time @: OBSID:0x3000E015 Comments: Ran script with default values On source chop = 0xA000 On source jiggle = 0x8000 Off source chop = 0x8000 Off source chop = 0x8000 Number of chop cycles = 50 Period of chop cycle = 2 Hz = 500000 micro seconds 61.46 milliseconds per sample using ~130 Hz and div=7 DCU frame type = 0 (= Full Photometer) Number of DCU samples = 4 Leaves 4.16 milliseconds = 4160 micro seconds Number of BSM samples = 15 Crashed OBS Power cycle the DPU Running DPU Start Procedure at 19:42 – OK Producing BSM Frames – Don’t know why Stopped data generation by manual command 91C10000 Sent flush FIFO Sent reset FIFO All looks OK, voltages correct on MCU, SCU Copied the command list from Q:\OBS\VMSIM\Table071-Chop\TC and updated table 71 with that file. Re-run test with same parameters: 0x3000E017 hasn’t crashed OBS this time but it is not chopping Lots of Events: Type 5 Subtype 2 Eventid 7000 sid 5201 sent 3 times Produced BSM frames but didn’t produce DCU frames, BSM didn’t move. Setting up for Overnight Clearing critical HK report Instrument Status: DPU:ON DRCU:ON AC/DC Thermometry: ON MCU:ON BSM:ON SMEC: OFF PHOT/SPEC BDAs :OFF Minimum delay is 6 milliseconds after talking with Ken |

Re-tried FUNC-BSM-06 on 16th March after increasing sample rate to 18.60119 still bias frequency 130 Hz

On source chop = 0xA000

On source jiggle = 0x8000

Off source chop = 0x8000

Off source jiggle = 0x8000

Number of chop cycles = 50

Period of chop cycle = 2 Hz = 500000 micro seconds

53.76 milliseconds per sample using ~130 Hz and div=6

DCU frame type = 0 (= Full  Photometer)

Number of DCU samples = 4

Leaves 34.959 milliseconds =  34959 micro seconds

Number of BSM samples = 15

OBSID=0x3000E01D start at 10:58 finish at 11:00

BSM moved but is not chopping gone to 0xA000 but did not chop off it. Test ended with BSM at stuck at 0xA000 on senssig but the chopposn went back to 0x8000 after going up to 0xA000 for the duration of the test.

11:16 Commanded BSM to move to 0x8000 it didn’t move from 0xA000, no events

11:19 Running BSM-OFF OBSID = 0x3000E01F switched OFF OK moved back to nominal position

11:20 Running FUNC-BSM-01 to switch it on again OBSID = 0x3000E020 sensor at about A9A9

11:22 Run BSM-INIT to see if it goes to 0x8000, 0BSID=0x3000E021, senssig still at A000

11:24 Commanding BSM to move to 0xA000, hasn’t moved as expected

11:24 Commanding BSM to move to 0x8000 didn’t move

This requires an MCU Boot

11:25 Run BSM-OFF

11:27 Run MCU-OFF

11:29 Run FUNC-MCU-01, OBSID 0x3000E024. BSM back to 0x8000

11:30 Run FUNC-BSM-01, OBSID 0x3000E025. CHOPSENSSIG = 0x8E5A

11:35 Run BSM-INIT, OBSID 0x3000E026

11:46 Run FUNC-BSM-06, previous configuration, except number of BSM samples per chop position, now 65535. OBSID 0x3000E027. The BSM chopped but got stuck again at the end.

12:19 Run BSM_OFF, OBSID 0x3000E028.

12:20 Run MCU_OFF, OBSID 0x3000E029

12:22 Run FUNC-MCU-01, OBSID 0x3000E02A

12:23 Run FUNC-BSM-01, OBSID 0x3000E02B

12:25 Run BSM-INIT, OBSID 0x3000E02C

12:28 Run FUNC-BSM-03, OBSID 0x3000E02D

Frame rate: 64

Jiggle start, end, step: 0x8000, 0x9000, 0x1000

Chop start, end, step: 0x8000, 0x9000, 0x1000

Delay time: 5

12:32 Run FUNC-BSM-06, OBSID 0x3000E02E. Chops but gets stuck at the end.

Frame rate: 64

On source chop = 0x9000

On source jiggle = 0x8000

Off source chop = 0x8000

Off source jiggle = 0x8000

Number of chop cycles = 50

Period of chop cycle = 2 Hz = 500000 micro seconds

53.76 milliseconds per sample using ~130 Hz and div=7

DCU frame type = 0 (= Full  Photometer)

Number of DCU samples = 4

Leaves 34.959 milliseconds =  34959 micro seconds

Number of BSM samples = 65535

12:35 Run FUNC-BSM-06 again, OBSID 0x3000E20F. Same as before.

13:40 Switch Off BSM, OBSID:0x3000E030

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute BSM_OFF from HCSS Test Procedures |  |

COMMENT INSERTED: (17/03/06) Asier Abreu

The reason for the chopper to get stuck at the on source position is due to the commanding within the chop VM, the VM sets the BSM move mode to 1 which is synchronous movement but then does not reset to normal mode 0.

That is why we did not see the chopper coming back to 0x8000.

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open SMEC PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-SMEC-02A/B

| Test Id: | FUNC-SMEC-02A/B |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + SMEC ON If WFT prior to cooldown : Smec Latched If WFT after warm up : Smec Unlatched |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + SMEC ON If WFT prior to cooldown : Smec Unlatched If WFT after warm up : Smec Latched |
| Success Criteria: | Test passed if : Prior to un-latching the resistance across pins 7 and 8 of the launch latch is ~ 368 Ohms. After un-latching the resistance is 483 Ohms. Note: These resistance values were recorded for the CQM SMEC model, for the flight SMEC, these values are expected to vary. |

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Measure the resistance across pins 7 and 8 of the launch latch. |  |
| 2 | If WFT prior to cooldown: Run FUNC-SMEC-02A test procedure from the HCSS Test Procedure window on TOPE If WFT after warm up: Run FUNC-SMEC-02B test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Measure the resistance across pins 7 and 8 of the launch latch. |  |
| 4 | Contingency: If test fails repeat steps 1. |  |

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-02A/B |  |  |  | N/A | Passed |
| Start time @: End time @: OBSID: Comments: 16/03/06: Overnight Nominal HK request shows the ‘usual’ NO DCU RESPONSES Event ID (0x520) which get cleared within the second. Measured resistance between pins 7 and 8: R initial (latched) = 367.5 Ohms Executed OBSID = 0x3000E018 – didn’t unlatch so repeating the command Re-run procedure at 09:58 OBSID = 0x3000E019 Measured resistance 483.3 Ohms latch has been undone SMCLATCHSTAT did not change |

### FUNC-SMEC-01

| Test Id: | FUNC-SMEC-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if : SMECENCPWR HK parameter changes from 0 to 6. SMEC encoder signals 1 and 2 amplitude show variation when encoder is switched ON. SMEC LVDT is switched ON. SMEC LVDT DC and AC signals show variation when LVDT is switched ON. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a display of the following HK parameters: SMECENCPWR SMECENCSIG1AMP SMECENCSIG2AMP SMECLVDTDCSIG SMECLVDTACSIG |  |
| 2 | Run FUNC-SMEC-01 test procedure from the HCSS Test Procedure window on TOPE |  |
|  | Contingency: If test fails repeat steps 1. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-01 | SMECENCPWR SMECLVDTPWR SMECENCSIG1 SMECENCSIG1AMP SMECENCSIG1OFF SMECENCSIG2 SMECENCSIG2AMP SMECENCSIG2OFF |  | 0/6 0/1 3060/4833 0/0 CE20/6590 4E5E/53F0 0/0 CE20/7918 | N/A | Pass |
| Start time @:10:18 End time @:10:19 OBSID:0x3000E01A Comments: Added SMECENCOFF and AMP to procedure QLA procedure needs to add SMECENCSIG1 and SMECENCSIG2 |

### FUNC-SMEC-03

| Test Id: | FUNC-SMEC-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if: SMEC encoder signals 1 and 2 show a variation on their amplitudes from one LED illumination level to another. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the following Nominal HK parameters: SMECENCSIGAMP1 SMECENCSIGAMP2 |  |
| 2 | Run FUNC-SMEC-03 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-03 | SMECENCPWR SMECENCSIG1 SMECENCSIG2 |  | 6/4 430D/3262 5E08/4CA4 |  | Pass |
| Start time @:10:32 End time @:10:34 OBSID:0x3000E01B Comments: ENCSIGAMP1 will not change so procedure needs to be revised to be ENCSIG1 and ENCSIG2 instead |

### FUNC-SMEC-04a

| Test Id: | FUNC-SMEC-04a |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if SMECLVDTDCSIG parameter shows a variation according to the different positions set. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the following Nominal HK parameters: SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG SMECMOTORCURR |  |
| 2 | Run FUNC-SMEC-04a test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat steps 1. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-04a | SMECLVDTDCSIG |  | 0/ |  | Pass |
| Start time @:10:41 End time @:10:47 OBSID:0x3000E01C Comments: This should be run before SMEC-03 sequence needs to be changed The SMEC-03 test left the ENCPWR at 4 which is far too low for warm conditions this was manually put back to 6 before the start of this test. |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute SMEC_OFF from HCSS Test Procedures | 0x3000E01D, SMEC is Off |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open DCU PARAMETERS SCOS Alpha Numeric Display |  |

### FUNC-DCU-01

| Test Id: | FUNC-DCU-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Success Criteria: | Test passed if: DCU produces each type of DCU nominal science frame with the following characteristics. APID Type S.type SID Frame ID Frame type Nb. Of frames Nb. of pkts. 0x504 21 1 0x200 0 PF 100 100 0x506 21 1 0x201 1 SF 100 17 0x504 21 2 0x102 2 PSW 100 34 0x504 21 2 0x103 3 PMW 100 25 0x504 21 2 0x104 4 PLW 100 12 0x506 21 2 0x105 5 SSW 100 12 0x506 21 2 0x106 6 SLW 100 7 The frame time difference between consecutive DCU frames of each type corresponds to the sampling rate. Photometer Sampling rate is 15.3Hz ( Δt ~ 65.5 ms Spectrometer Sampling rate is 80Hz ( Δt = 12.5 ms The SPIRE HK parameter DCUFRAMECNT increments by 700. No events are generated during the frames generation. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |  |
| 2 | Run FUNC-DCU-01 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |  |
| 4 | Contingency: If test fails repeat steps 1 to 3. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-01 | DCUFRAMECNT | 0/700 | 3406/4806 |  | Pass |
| Start time @:13:43 End time @:13:47 OBSID:0x3000E031 Commets: Default parameters are wrong! Ran script with Phot Bias Frequency:130 Sampling Frequency:18 Spec Bias Frequency: 160 Sampling Frequency:80 Number of Frames to request:100 Realised we have a QLA Script – should be added to the procedure – re-ran test Start time @:13:49 End time @: OBSID:0x3000E032 QLA produced file QLA-DCU-01_3000E032.txt. All packet types showed 100 frames have been produced |

### FUNC-DCU-02

| Test Id: | FUNC-DCU-02 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Success Criteria: | Test passed if DCUFRAMECNT goes from 700 to 1400 and the frametime difference between consecutive frames computed by QLA script is in agreement with the expected differences based on commanded sampling rate: Photometer Sampling rate is 15.3Hz ( Δt ~ 65.5 ms Spectrometer Sampling rate is 80Hz ( Δt = 12.5 ms |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |  |
| 2 | Run QLA script FUNC-DCU-02.py on QLA console. |  |
| 3 | Run FUNC-DCU-02 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-02 | DCUFRAMECNT | 700/1400 | 4806 | 5506 | Pass |
| Start time @:13:56 End time @: 13:58 OBSID:0x3000E033 Comments: Parameters changed to values used for previous test QLA showed times as expected from parameters entered |

### FUNC-DCU-03

| Test Id: | FUNC-DCU-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Success Criteria: | Test passed if : DCU produces 100 frames of Full Photometer Test Pattern and 100 frame of Full Spectrometer Test Pattern test. QLA analysis shows that phot/spec test patterns are the same as the reference phot/spec test patterns. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |  |
| 2 | Run QLA script FUNC-DCU-03.py on QLA console. |  |
| 3 | Run FUNC-DCU-03 test procedure from the HCSS Test Procedure window on TOPE |  |
| 4 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |  |
| 5 | Contingency: If test fails repeat steps 1 to 4. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-03 | DCUFRAMECNT | 1400/2100 | 5506 | 5706 | Pass |
| Start time @: 14:04 End time @: 14:05 OBSID: 0x3000E034 Comments: Changed photometer bias frequency to 130 Hz and sample rate to 18 Hz Files produced by QLA: QLA-DCU-03_3000E034_8807.txt QLA-DCU-03_3000E034_8808.txt Both files show a test pass Note: expected after value is wrong |

### FUNC-DCU-04P

| Test Id: | FUNC-DCU-04P |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Photometer LIAs ON |
| Success Criteria: | Test passed if : SCUDCDCSTAT parameter goes from 4 to 5. Photometer LIA card voltages are showing correct readings of +5V,+9V,-9V. Photometer LIA temperatures show an increase indicating that they are ON. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the HK parameters: PLIAP5V PLIAP9V PLIAM9V LIAP1/2/3/4/5/6/7/8/9TEMP |  |
| 2 | Run FUNC-DCU-04P test procedure from the HCSS Test Procedure window on TOPE |  |
| 5 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-04P | SCUDCDCSTAT PLIAP5V PLIAP9V PLIAM9V LIA1/2/3/4/5/6/7/9TEMP | 4/5 0/~ +5V 0/~+9V 0/~-9V N/A/ [290-300]K | 4/5 0/+5 0/+5 0/-9 ~290/warming up |  | Pass |
| Start time @: 14:31 End time @: 14:31 OBSID: 0x3000E035 Commets: LIAs switched on OK |

### FUNC-DCU-05P

| Test Id: | FUNC-DCU-05P |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON |
| Success Criteria: | Test passed if : The first channel of the 32 on each photometer LIA card (16 bit word in the photometer offset packet) shows the commanded offset variation 0x1,0x3,0x6,0x9,0xc,0xf |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the Photometer Offset Frame parameters: PHOTOFF001 PHOTOFF033 PHOTOFF055 PHOTOFF087 PHOTOFF119 PHOTOFF151 PHOTOFF183 PHOTOFF215 PHOTOFF247 PHOTOFF279 |  |
| 2 | Run FUNC-DCU-05P test procedure from the HCSS Test Procedure window on TOPE |  |
| 5 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-05P |  |  |  |  | Pass |
| Start time @: 14:36 End time @: 14:39 OBSID: 0x3000E036 Comments: |

### FUNC-DCU-11P

| Test Id: | FUNC-DCU-11P |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON + Photometer BIAS ON +Photometer JFETs ON |
| Success Criteria: | Test passed if Photometer JFET source and drain voltages are correct: PSWJFETVSS1/2/3/4/5/6 (RAW: 0x4C CONVERTED ~ -1.5V). PMLWJFETVSS1/2/3/4 (RAW: 0x4C CONVERTED ~ -1.5V). PSWJFETSTAT = 0x3F PMLWJFETSTAT = 0x1F |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-DCU-11P test procedure from the HCSS Test Procedure window on TOPE |  |
| 2 | After the test Annotate the values RAW and converted values of: PSWJFETSTAT,PMLWJFETSTAT, PSWJFET1/2/3/4/5/6V PMWJFET1/2/3/4V PLWJFET1/2V located in DCU PARAMETERS AND |  |
| 3 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-11P | PSWJFETSTAT PMLWJFETSTAT PSWJFET1/2/3/4/5/6V PMWJFET1/2/3/4V PLWJFET1/2V | 0/0x3f 0/0x1f 0V/-1.5V 0V/-1.5V 0V/-1.5V | 0/0x03f 0/0x03f 0/-1.49 0/-1.49 0/-1.49 | N/A | Pass |
| Start time @: 14:43 End time @: 14:44 OBSID: 0x3000E037 Comments: Kept the default parameter values. PSW A10 and H9 appear to be dead. PMW C8, C5 and T2 seem to be dead as well. PSW G8 and C12 have the same output as the resistors. |

### FUNC-DCU-13P

| Test Id: | FUNC-DCU-13P |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON + Photometer BIAS ON +Photometer JFETs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ Photometer LIAs ON + Photometer BIAS ON +Photometer JFETs ON |
| Success Criteria: | Test passed if : The photometer detectors show a small linear variation on the output voltage when different bias is applied through the load curve. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of a couple of pixels on each of the photometer BDAs |  |
| 2 | Run FUNC-DCU-13P test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | Contingency: If test fails repeat step 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-13P |  |  |  | N/A | Pass |
| Start time @: 14:50 End time @: 15:00 OBSID: 0x3000E038 Comments: Array: PF Bias frequency: 130 Hz Sampling frequency: 18 Hz PSW phase: 187.8 deg PMW phase: 189.2 deg PLW phase: 187.8 deg Time at each level: 10 s PSW G8 and C12 are alive but are showing larger changes of output level for each bias setting than the other detectors and unlike other detectors saturate about halfway through the load curve. |

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | From TOPE HCSS Test Procedures run PDET-OFF | OBSID 0x3000E039 |

### FUNC-DCU-04S

| Test Id: | FUNC-DCU-04S |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Success Criteria: | Test passed if : SCUDCDCSTAT parameter goes from 4 to 6. Spectrometer LIA card voltages are showing correct readings of +5V,+9V,-9V. Spectrometer LIA temperatures show an increase indicating that they are ON. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the HK parameters: SLIAP5V SLIAP9V SLIAM9V LIAS1/2/3TEMP |  |
| 2 | Run FUNC-DCU-04S test procedure from the HCSS Test Procedure window on TOPE |  |
| 5 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-04S | SCUDCDCSTAT SLIAP5V SLIAP9V SLIAM9V LIA1/2/3TEMP | 4/6 0/~ +5V 0/~+9V 0/~-9V N/A/ [290-300]K | 4/6 0.17/5.23 0.003/11.53 0.003/-11.52 ~292/warming up |  | Pass |
| Start time @: 15:35 End time @: 15:35 OBSID: 0x3000E03A Comments: Note that the values for SLIAP9V and SLIAM9V are +11.5 and -11.5 V respectively. |

### FUNC-DCU-05S

| Test Id: | FUNC-DCU-05S |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Success Criteria: | Test passed if : The first channel of the 12 on each spectrometer LIA card (16 bit word in the photometer offset packet) shows the commanded offset variation 0x1,0x3,0x6,0x9,0xc,0xf |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | On QLA bring up a time series display of the Photometer Offset Frame parameters: SPECOFF01 SPECOFF13 SPECOFF25 SPECOFF37 SPECOFF49 SPECOFF61 |  |
| 2 | Run FUNC-DCU-05S test procedure from the HCSS Test Procedure window on TOPE |  |
| 5 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-05S |  |  |  |  | Pass |
| Start time @: 15:39 End time @: 15:42 OBSID: 0x3000E03B Comments: QLA gave an “Unknown packet” message: APID 500, type 5, subtype 4, SID 5420. |

### FUNC-DCU-11S

| Test Id: | FUNC-DCU-11S |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON + Spectrometer BIAS ON + Spectrometer JFETs ON |
| Success Criteria: | Test passed if SCUDCDCSTAT goes from 4 to 6, Spectrometer LIAs voltages are correct and SJFET voltages are also correct. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-DCU-11S test procedure from the HCSS Test Procedure window on TOPE |  |
| 2 | After the test Annotate the values RAW and converted values of: LIASTAT SLIAP5V, SLIAP9V, SLIAN9V, SSWJFETSTAT,SLWJFETSTAT, SSWJFET1V,SLWJFET2V located in DCU PARAMETERS AND |  |
| 3 | Contingency: If test fails repeat steps 1 and 2. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-11S | SCUDCDCSTAT LIASTAT SLIAP5V SLIAP9V SLIAN9V SSWJFETSTAT SLWJFETSTAT SSWJFET1/2V SLWJFET1/2V | 4/6 0/0 0V/ ~ 5V 0V/~11V 0V/-11V 0/0x3f 0/0x3f 0V/-1.5V 0V/-1.5V | 6/6 0/0 5.23/5.23 11.53/11.53 -11.52/-11.53 0/0x7 0/0x7 0/-1.5 0/-1.5 | N/A | Pass |
| Start time @: 15:48 End time @: 15:49 OBSID: 0x3000E03C Comments: Kept default input parameters. Note: for LIASTAT the values of LIAS1/2/3STAT have been written. For SSWJFETSTAT and SLWJFETSTAT the value of SPECJFETSTAT has been written. Some of the numbers in the “Expected Value” column are probably wrong DP2 on SSW appears to be dead. |

### FUNC-DCU-13S

| Test Id: | FUNC-DCU-13S |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON + Spectrometer BIAS ON + Spectrometer JFETs ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON + Spectrometer BIAS ON + Spectrometer JFETs ON |
| Success Criteria: | Test passed if : The spectrometer detectors show a small linear variation on the output voltage when different bias is applied through the load curve. |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Run FUNC-DCU-13S test procedure from the HCSS Test Procedure window on TOPE |  |
| 2 | Contingency: If test fails repeat step 1. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-13S |  |  |  | N/A | Pass |
| Start time @: 15:55 End time @: 16:04 OBSID: 0x3000E03D Comments: Array: SF Bias frequency: 160 Hz Sampling frequency: 80 Hz SSW phase shift: 207.5 deg SLW phase shift: 207.5 deg Time at each level: 10 s Some of the SSW pixels did not show much change. |

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | From TOPE HCSS Test Procedures run SDET-OFF | 16:08, OBSID 0x3000E03E |

# END TEST SEQUENCE

## NORMAL END TEST SEQUENCE

The following table shows the necessary steps to be followed to end the warm functional test sequence.

| Step# | Action | HK parameters | Expected Value | Comments | Check |
| --- | --- | --- | --- | --- | --- |
| 1 | Check BSM is OFF | CHOPSENSPWR JIGGSENSPWR | 0 0 |  | OK |
| 2 | Check SMEC is OFF | SMECENCPWR | 0 |  | OK |
| 3 | Run MCU-OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT | 0 | 16:10 OBSID 0x3000E03F | OK |
| 4 | Check Instrument Configuration | MODE | 0x100/ DRCU_ON |  | OK |
| 5* | Run SCU_OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT SUBKSTAT | 0 0 | Left on |  |
| 6* | Run DRCU_OFF from the HCSS Test Procedure window on TOPE | MODE | 0/ DPU_ON | Left on |  |

* These two steps should ONLY be executed if the functional test takes place BEFORE the instrument is placed in the test cryostat.  If the instrument is already in the cryostat and the cryo-harness are connected this steps should NOT be executed.

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS 1 to 4 are executed is DRCU_ON

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS 1 to 6 are executed is DPU_ON

## END TEST SEQUENCE WHEN THE FUNCTIONAL TEST HAS FAILED

The following table shows the necessary steps to be followed to end the warm functional test sequence when this has been declared failed.

| Step# | Action | HK parameters | Expected Value | Comments | Check |
| --- | --- | --- | --- | --- | --- |
| 1 | Check BSM is OFF | CHOPSENSPWR JIGGSENSPWR | 0 0 |  |  |
| 2 | Check SMEC is OFF | SMECENCPWR | 0 |  |  |
| 3 | Run MCU-OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT | 0 |  |  |
| 4 | Check Instrument Configuration | MODE | 0x100/ DRCU_ON |  |  |
| 5 | Run SCU_OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT SUBKSTAT | 0 0 |  |  |
| 6 | Run DRCU_OFF from the HCSS Test Procedure window on TOPE | MODE | 0/ DPU_ON |  |  |

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS DPU_ON

# Annexe1 (Instrument nominal configuration/modes)

|  | OFF | INIT | DPU_ON | DRCU_ON | REDY | PHOT STBY | SPEC STBY | CREC | SAFE |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| DPU |  | ON | ON | ON | ON | ON | ON | ON | ON |
| Essential Hsk packets |  |  | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz |
| Normal Hsk packets |  |  | 1.0Hz | 1.0Hz | 0.25Hz | 0.25Hz | 0.25Hz | 1.0Hz | 0.25Hz |
| TC Acceptance |  | TBC | Yes | Yes | Yes | Yes | Yes | Yes | Yes |
| Event packets |  | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes |
| Science packets |  |  |  |  |  |  |  |  |  |
| VM |  |  |  |  |  |  |  | Cooler_Recycle |  |
| VM1 |  |  |  |  |  | Det_Temp_Ctrl |  |  |  |
| VM2 |  |  |  |  |  |  |  |  |  |
| VM3 |  |  |  |  |  |  |  |  |  |
| MODE |  |  | 0x0000 | 0x0100 | 0x0200 | 0x0300 | 0x0400 | 0x0600 | 0x0900 |
| DRCU |  |  | ON | ON | ON | ON | ON | ON |  |
| SCU |  |  |  |  |  |  |  |  |  |
| Temp Channels powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| SubK Channel powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| PCAL source powered |  |  |  |  |  |  |  |  |  |
| SCAL sources powered |  |  |  |  |  |  | TBD |  |  |
| TC Heater powered |  |  |  |  |  | Yes |  |  |  |
| Cooler SP Heater powered |  |  |  |  |  |  |  | Yes |  |
| Cooler EV HS powered |  |  |  |  |  |  |  | Yes |  |
| Cooler SP HS powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| DCU |  |  |  |  |  |  |  |  |  |
| Photometer BIAS |  |  |  |  |  | Yes |  |  |  |
| Photometer JFETS |  |  |  |  |  | Yes |  |  |  |
| Photometer LIAs |  |  |  |  |  | Yes |  |  |  |
| TC BIAS |  |  |  |  |  | Yes |  |  |  |
| TC JFETS |  |  |  |  |  | Yes |  |  |  |
| TC LIAs |  |  |  |  |  | Yes |  |  |  |
| Spectrometer BIAS |  |  |  |  |  |  | Yes |  |  |
| Spectrometer JFETS |  |  |  |  |  |  | Yes |  |  |
| Spectrometer LIAs |  |  |  |  |  |  | Yes |  |  |
| MCU |  |  |  |  |  |  |  |  |  |
| DSP |  |  |  |  | On | On | On | On |  |
| BSM |  |  |  |  |  | Hold | Hold |  |  |
| SMEC |  |  |  |  |  |  | Hold |  |  |

# WARM FUCNTIONAL TEST CUS SOFTWARE CONFIGURATION

#

#$Id$

#

# INITIAL SOFTWARE CONFIGURATION FOR PFM3 BBLOCKS

#

string                        string                        string

BBNAME                        BBFULLTYPE                    VERSION

ClearObs_StandAlone           0x8021                        1.1

Block_DefineHK                0x8006                        1.3

SMEC_Off                      0x820a                        1.6

SMEC_LED_Levels               0x8205                        1.9

PTC_Test                      0x8d05                        1.2

S_Phase_Shift                 0x881d                        1.2

SCU_DC_Therm                  0x8a01                        1.2

SetChopTarget                 0x810c                        1.13

StartMCUdata                  0x8906                        1.5

Spec_LIAs_On                  0x880a                        1.3

Setup_PMW_Data                0x8829                        1.2

SetJiggleTarget               0x810d                        1.7

PSWTest_Patt                  0x880b                        1.2

P_Phase_Shift                 0x881c                        1.2

Block_HBB                     0x8b12                        1.2

PSWBeamPeakUp_SMECScan        0x8e05                        1.3

PJFET_Vss_Test                0x881a                        1.8

PLWBeamPeakUp_SMECScan        0x8e07                        1.3

PSW_BDA_Off                   0x8505                        1.3

Scan_FTS                      0x8203                        1.10

SSW_BDA_On                    0x8601                        1.8

PLW_BDA_Off                   0x8507                        1.3

DCU_Data                      0x8014                        1.2

Chop                          0x8d07                        1.2

SF_Phase                      0x8854                        1.4

PCAL_FTable0                  0x8d00                        1.2

StartObs                      0x8002                        1.4

SCU_ACDC_Therm                0x8a0a                        1.3

MCU_Off                       0x8908                        1.2

StartObs_StandAlone           0x8022                        1.1

ClearObs                      0x8001                        1.2

DPU_Switch_ON                 0x8c00                        1.3

PF_BDA_Off                    0x8504                        1.5

SF_BDA_On                     0x8600                        1.18

PF_Data_QLA                   0x8800                        1.3

PLWTest_Patt                  0x880d                        1.2

PLWJFET_Vss_Test              0x8824                        1.6

Flash_Table                   0x8d01                        1.2

MCU_SMEC_generation           0x8902                        1.3

PMW_BDA_On_No_Jfets           0x8f09                        1.2

PMWBeamPeakUp_SMECScan        0x8e06                        1.3

EndObs_StandAlone             0x8023                        1.1

PLW_BDA_On                    0x8503                        1.10

SF_BeamPeakUp                 0x8f03                        1.7

PSWJFET_Vss_Test              0x8822                        1.6

PCAL_Flash                    0x8302                        1.3

MCU_BSM_generation            0x8903                        1.3

JiggleMap                     0x8d03                        1.2

PFBias_Ampl                   0x8840                        1.4

Step_Look_Scan                0x8b14                        1.4

Phot_LIAs_On                  0x8809                        1.3

Initialise_FTS                0x8200                        1.12

PSW_Data_QLA                  0x8802                        1.4

SJFET_Vss_Test                0x881b                        1.8

PF_BeamPeakUp                 0x8e00                        1.6

Setup_PLW_Data                0x882a                        1.2

PFBeamPeakUp_SMECScan         0x8e04                        1.3

PCAL_Bias                     0x8305                        1.2

SCAL4_Check                   0x8a03                        1.2

SSWTest_Patt                  0x880e                        1.2

Switch_On_FTS                 0x8201                        1.11

Setup_PSW_Data                0x8828                        1.2

Block_CBB                     0x8b11                        1.2

Block_PCAL                    0x830a                        1.2

PMWTest_Patt                  0x880c                        1.2

BSM_Raster                    0x8120                        1.5

POffset_Manual                0x8816                        1.3

PSWBias_Ampl                  0x8841                        1.3

SCU_ACDC_Therm_OFF            0x8a10                        1.3

BSM_SwitchOn                  0x8100                        1.6

SF_BDA_Off                    0x8603                        1.3

SSWBeamPeakUp_SMECScan        0x8f05                        1.5

SLW_Phase                     0x8856                        1.4

SLWBeamPeakUp_SMECScan        0x8f06                        1.5

PLW_Data_QLA                  0x8804                        1.3

SFTest_Patt                   0x8808                        1.2

PF_BDA_On                     0x8500                        1.20

PF_BDA_On_No_Jfets            0x8f07                        1.3

PLW_BeamPeakUp                0x8e03                        1.5

SCU_PCAL_Freq                 0x8069                        1.2

MCU_ENG_generation            0x8901                        1.3

SCAL_PID                      0x8d04                        1.2

SLWTest_Patt                  0x880f                        1.2

Phot_LIAs_Check               0x8810                        1.2

InitialiseBSMchop             0x810a                        1.7

SLW_BDA_On                    0x8602                        1.11

Flip_Mirror                   0x8b10                        1.1

MCU_SMECandBSM_generation     0x8904                        1.3

PMW_BDA_On                    0x8502                        1.10

PBias_Freq                    0x8820                        1.2

PLWBias_Ampl                  0x8843                        1.3

SCUTest_Patt                  0x8a07                        1.3

PCAL_Check                    0x8a02                        1.2

PSW_BDA_On_No_Jfets           0x8f08                        1.2

Setup_PF_Data                 0x8826                        1.7

SCAL2                         0x8400                        1.3

PJFET_Htr_Test                0x8818                        1.2

SCAL4                         0x8401                        1.3

StopMCUdata                   0x8907                        1.5

PSW_BeamPeakUp                0x8e01                        1.5

SSW_Data_QLA                  0x8805                        1.3

SSW_BDA_Off                   0x8604                        1.3

SLW_BDA_Off                   0x8605                        1.3

PMW_Data_QLA                  0x8803                        1.3

SOffset_Manual                0x8817                        1.3

Step_Look_FTS                 0x820c                        1.16

CCS_StartObs                  0x8026                        1.1

SLWJFET_Vss_Test              0x882f                        1.7

CCS_ClearObs                  0x8025                        1.1

SSWBias_Ampl                  0x8845                        1.3

SLW_BeamPeakUp                0x8f02                        1.6

Phot_LIAs_Off                 0x8811                        1.2

SSWJFET_Vss_Test              0x882e                        1.8

SSW_Phase                     0x8855                        1.4

Setup_SLW_Data                0x882c                        1.2

SF_Data_QLA                   0x8801                        1.3

SCU_AC_Therm                  0x8a05                        1.2

PMWBias_Ampl                  0x8842                        1.3

POffset_Auto                  0x881e                        1.3

PLW_Phase                     0x8853                        1.4

PMWJFET_Vss_Test              0x8823                        1.6

Setup_SSW_Data                0x882b                        1.2

SMECm_Scan                    0x8210                        1.3

SLW_Data_QLA                  0x8806                        1.3

SBias_Freq                    0x8821                        1.2

SFBeamPeakUp_SMECScan         0x8f04                        1.5

TFTS_Scan                     0x8f00                        1.2

InitialiseBSMjiggle           0x810e                        1.5

SFBias_Ampl                   0x8844                        1.4

PMW_BeamPeakUp                0x8e02                        1.6

SwitchBSMoff                  0x810b                        1.4

Setup_SF_Data                 0x8827                        1.10

Start_DCU_Data                0x8835                        1.2

SJFET_Htr_Test                0x8819                        1.2

EndObs                        0x8003                        1.2

SLWBias_Ampl                  0x8846                        1.3

SSW_BeamPeakUp                0x8f01                        1.5

PF_Phase                      0x8850                        1.4

PMW_Phase                     0x8852                        1.4

Stop_DCU_Data                 0x8836                        1.2

DCUHK                         0x8d08                        1.1

Pupil_Scan                    0x8b15                        1.4

Block_ClearHK                 0x8005                        1.2

PLWJFetSwitchOff              0x8509                        1.2

BSM_Chop                      0x8110                        1.4

PMW_BDA_Off                   0x8506                        1.3

PLW_BDA_On_No_Jfets           0x8f0a                        1.2

SCAL2_Check                   0x8a04                        1.2

SCU_Nom_Data                  0x8a00                        1.3

Cooler_Htr_Check              0x8a06                        1.2

Block_Phase                   0x8814                        1.2

Start_SMEC_MCUENG_data        0x890a                        1.2

PSW_BDA_On                    0x8501                        1.10

Chop_Function                 0x8d02                        1.4

PSW_Phase                     0x8851                        1.4

PCAL_OFF                      0x8301                        1.2

MCU_TEST_generation           0x8905                        1.3

MCU_Boot                      0x8900                        1.6

CCS_EndObs                    0x8027                        1.1

SOffset_Auto                  0x881f                        1.3

PCAL                          0x8300                        1.3

PLWJFetSwitchOn               0x8508                        1.3

PFTest_Patt                   0x8807                        1.2

#

#$Id$

#

# INITIAL SOFTWARE CONFIGURATION FOR CUS MODES USED FOR PFM3 FUNCTIONAL TESTING

string                        string

FUNCIONAL_TEST_CUS_MODE       VERSION

Mode_DcuFunc06_Spec           1.2

SCU_OFF                       1.2

BSM_OFF                       1.1

Mode_DCU_HK                   1.1

DRCU_ON                       1.3

BSM_INIT                      1.1

Mode_DcuFunc11_Phot_No_Jfets  1.1

Mode_SCALFunc01               1.2

Mode_DcuFunc13_Phot           1.10

Mode_DcuFunc06_Phot           1.2

Mode_DcuFunc03_Full           1.2

Mode_DcuFunc07_Spec           1.3

Mode_SMECFunc01c              1.3

Mode_SMECFunc02a              1.1

Mode_SMECFunc02b              1.1

Mode_DcuFunc01                1.5

Mode_DcuFunc02                1.5

PDET_OFF                      1.4

Mode_SMECFunc04a              1.2

Mode_SMECFunc04b              1.5

Mode_DcuFunc07_Phot           1.3

Mode_SMECFunc01               1.5

Mode_SMECFunc03               1.4

Mode_BSMFunc01                1.1

Mode_BSMFunc03                1.3

Mode_DcuFunc11_Spec           1.2

Mode_BSMFunc06                1.1

Mode_DcuFunc08_Spec           1.2

Mode_DcuFunc04_Spec           1.2

Mode_BSMFunc02c               1.3

Mode_BSMFunc02j               1.2

Mode_BSMFunc04c               1.1

Mode_BSMFunc04j               1.1

Mode_BSMFunc05c               1.1

SMEC_OFF                      1.1

Mode_BSMFunc05j               1.1

Mode_McuFunc01                1.1

Mode_McuFunc02                1.2

Mode_McuFunc03                1.2

Mode_McuFunc04                1.2

SDET_OFF                      1.5

Mode_DcuFunc11_Phot           1.2

Mode_DcuFunc08_Phot           1.2

Mode_DcuFunc04_Phot           1.2

Mode_DcuFunc03_Short          1.6

SCU_ON                        1.1

BSM_ON                        1.3

MCU_OFF                       1.2

Mode_DcuFunc12_Spec           1.4

Mode_DcuFunc05_Spec           1.2

PDET_ON                       1.3

DRCU_OFF                      1.1

DPU_START                     1.4

MCU_BOOT                      1.1

SMEC_INIT                     1.2

Mode_DcuFunc12_Phot           1.3

Mode_PCALFunc01               1.2

Mode_DcuFunc05_Phot           1.2

Mode_ScuFunc01                1.2

Mode_ScuFunc02                1.2

Mode_ScuFunc03                1.2

SMEC_ON                       1.3

Mode_ScuFunc04                1.2

Mode_ScuFunc05                1.2

Mode_ScuFunc06                1.2

Mode_ScuFunc07                1.2

Mode_ScuFunc08                1.2

SDET_ON                       1.2

Mode_DcuFunc13_Spec           1.13

#

#$Id$

#

# INITIAL SOFTWARE CONFIGURATION FOR CUS PROCEDURES USED FOR PFM3

string                        string

PROCEDURE_NAME                VERSION

Proc_GetDCUData               1.1

Proc_SSWJFET_On               1.3

Proc_BSMRaster                1.1

Proc_BSM_Init                 1.1

Proc_Start_DCU_Data           1.1

Proc_MCU_Off                  1.2

Proc_PJFET_Heater             1.3

Proc_Stop_DCU_Data            1.1

Proc_SetObsMode               1.2

Proc_DPU_Start                1.2

Proc_SetSpecBiasAmpl          1.2

Proc_PLWJFET_On               1.3

Proc_SetupPhotData            1.3

Proc_ClearHK                  1.1

Proc_SMEC_Scan                1.1

Proc_SCU_Off                  1.1

Proc_DefineNewHK              1.1

Proc_SJFET_Heater             1.3

Proc_BSM_Off                  1.2

Proc_SLWJFET_On               1.3

Proc_SetPhotBiasAmpl          1.3

Proc_PMWJFET_On               1.3

Proc_ZERO_DRCU_PARAMS         1.4

Proc_SMEC_Latch               1.2

Proc_SetPhotPhase             1.2

Proc_LoadCommandList          1.1

Proc_SetupP                   1.6

Proc_SetupS                   1.4

Proc_PDET_Off                 1.4

raw_phase                     =

Proc_PDET_On                  1.4

Proc_SCU_On                   1.1

Proc_BSM_On                   1.2

Proc_MCU_Boot                 1.1

Proc_PSWJFET_On               1.3

Proc_SMEC_Off                 1.1

Proc_SetSpecPhase             1.1

Proc_SMEC_On                  1.3

Proc_SDET_On                  1.4

Proc_SMEC_Init                1.3

Proc_SDET_Off                 1.3