# Introduction

This document reports on the short functional tests (80K) carried out on the 2nd May 2006 on the SPIRE FM instrument during the PFM3 test campaign

## Scope

To judge the success or failure of a warm functional test by checking that:

- The telecommand sequence generated for a particular functional test is correctly received and executed on board by the SPIRE DPU.

- No error/event reports or command failures are generated during the execution of these commands.

- Telemetry is generated by the instrument as a result of telemetry requests to its different subunits.

- Particular telemetry parameters for each functional test change in an expected manner.

- A particular success criterion/criteria (specified in this document) is/are met.

## Reference Documents

| Ref | Document | Name | Version/Issue Nb. |
| --- | --- | --- | --- |
| RD01 | SPIRE-RAL-DOC-001652 | SPIRE Functional Tests Specification | Issue 1.4 |
| RD02 | SPIRE-RAL-DOC-001630 | SPIRE EGSE-ILT Start-Up Procedures | Issue 0.7 |
| RD03 | SPIRE-RAL-PRC-002222 | DRCU Switch ON Procedure | Issue 1.0 |
| RD04 | SPIRE-RAL-PRJ-001078 | SPIRE Data ICD | Issue 2.0 |
| RD05 | Sap-SPIRE-CCa-076-02 | DRCU/DPU Interface Control Document | Issue 1.2 |
| RD06 | LAM.PJT.SPI.NOT.011011 | MCU/DPU Command List ICD | Issue 5.0 |
| RD07 | SPIRE-IFS-PRJ-001391 | SPIRE OBS User Manual | Issue 2.1 |
| RD08 |  | SPIRE ILT Functional Testing Overview |  |

## Change Record

| Document | Change date | Changes |
| --- | --- | --- |
| Issue 1.0 | 02/05/06 | Document created |

# Table of Contents

1.	Introduction	1

1.1	Scope	1

1.2	Reference Documents	1

1.3	Change Record	1

2.	FUNCTIONAL TEST CONFIGURATION	3

2.1	Software Configuration	3

2.2	EGSE Configuration Checks	3

2.3	SPIRE Instrument Configuration	4

3.	Test procedure	5

3.1	GENERAL PASS/FAIL CRITERIA	5

3.2	GENERAL TEST PROCEDURE LAYOUT	5

3.3	DETAILED TEST PROCEDURE	6

3.3.1	FUNC-SCU-01	6

3.3.2	FUNC-SCU-03	8

3.3.3	FUNC-SCU-06	9

3.3.4	FUNC-SCU-07	10

3.3.5	FUNC-SCU-04	11

3.3.6	FUNC-SCU-05	12

3.3.7	FUNC-MCU-01	13

3.3.8	FUNC-MCU-02	14

3.3.9	FUNC-BSM-01	15

3.3.10	FUNC-BSM-03	17

3.3.11	FUNC-SMEC-01	19

3.3.12	FUNC-SMEC-03	21

3.3.13	FUNC-SMEC-04a	22

3.3.14	FUNC-DCU-01	23

3.3.15	FUNC-DCU-04P	25

3.3.16	FUNC-DCU-04S	26

4.	END TEST SEQUENCE	27

4.1	NORMAL END TEST SEQUENCE	27

4.2	END TEST SEQUENCE WHEN THE FUNCTIONAL TEST HAS FAILED	27

5.	Annexe1 (Instrument nominal configuration/modes)	29

# FUNCTIONAL TEST CONFIGURATION

## Software Configuration

Annotate in the table the current EGSE software configuration for the tests:

| EGSE component | Version/Build number | Comment |
| --- | --- | --- |
| SCOS2000 | 2.3eP5 | Archive PFM3_TEST2 |
| HCSS | v0.3.3 Build#812 |  |
| QLA | V3.3 |  |
| QLA scripts |  |  |
| Test Control scripts | See Annexe1 |  |
| CUS Scripts | See Annexe2 |  |

## EGSE Configuration Checks

To check for the success of failure of a functional test, the real time telemetry of the instrument has to be monitored. The following applications must be running to do so. Before the test sequence starts, make the following checks:

| Workstation | EGSE component | Status | Check |
| --- | --- | --- | --- |
| Lichfield | EGSE router | Started | ( |
| Lichfield | EGSE Gateway | Started | ( |
| Lichfield | Telemetry Ingestion | Started | ( |
| Lichfield | Packet Display | Started | ( |
| Lincoln | SCOS2000 | Started | ( |
| Lincoln | EXIF + TOPE | Started | ( |
| Lincoln | Manual Stack | Started | ( |
| Gordon | CDMS Simulator | Started | ( |
| Lincoln | Test Control Server | Started | ( |

## SPIRE Instrument Configuration

The functional test flow is such that at a given time during the procedure is possible that the instrument configuration does not conform to any nominal configuration as specified in Annexe 1.  Nevertheless the configuration prior and after a test is specified for each functional test, and were applicable and the instrument configuration is in accordance to Anexe1, this will be denoted by showing the main configuration in blue bold font and any additional subsystems configuration in black.

The initial instrument configuration for Short WFT does conform to a ‘known’ configuration and should be SPIRE DRCU_ON.

Perform the following actions to ensure that the instrument is in the correct configuration for the tests.

| Step# | Action | Comments (Write any comments under this tab) | Check |
| --- | --- | --- | --- |
| 1 | In SCOS open DPU_AND_OBS_PARAMETERS display Check if : TM2N is incrementing by one @1Hz. TM1N is incrementing by one @0.5Hz. - If they are, go to step 5. - If they are not, go to step 2. | TM2N is incrementing by 1 @ 1Hz TM1N is fixed (Critical HK has been previously cleared to avoid unnecessary filling of the SCOS archive) Comment: procedure should include this possible situation and which action to take then. Will go to step 5 | ( |
| 2 | Check if the DPU is powered ON: - If ON, the DPU power supply LCD will show ~ 28V and 0.40A, go to step 4. - If not ON, power ON DPU see comments on the right, then go to step 3. |  |  |
| 3 | In SCOS open Boot_ROM_Memory_Check display and check no errors are reported: - If no errors are reported, execute DPU_ON from HCSS Test Procedures. Then repeat step 1. - If the (5,2) contains errors: Check the error code in RD07. Then switch OFF the DPU and repeat step 2 |  |  |
| 4 | Execute define_new_HK_report.tcl HCSS Test procedure. Repeat step 1. |  |  |
| 5 | In SCOS open SCU_PARAMETERS display - If SCUP5V/P9V/M9V are jittering and BIAS_PARAMETERS display - If BIAS1TEMP,BIAS2TEMP,… show ambient temperature, the DRCU is ON.Go to step 6. - If DRCU is not ON, refer to RD03 on how to start up the DRCU. | SCUP5V/ SCUP9V/SCUM9V are jittering as specified but : There is only one BIASTEMP parameter under BIAS_PARAMETERS display. Procedure should include other parameters aswell which are relevant like BIASP5V, BIASP9V, BIASM9V Will go to step 6 | ( |
| 6 | In SCOS open DPU_AND_OBS_PARAMETERS display and check that the MODE housekeeping parameter is DRCU_ON. | MODE parameter is DRCU_ON | ( |

Table 1. Initial configuration check

# Test procedure

The following two sections describe general pass/fail criteria (Section 3.1), the general test procedure layout (Section 3.1) and the detailed procedure for each functional test (Section 3.2).

## GENERAL PASS/FAIL CRITERIA

The general criterion for declaring a single test failed is the repeated failure of 2 consecutive runs of this test. In that case the functional test procedure should be aborted and the overall functional testing declared FAILED.

In the case of a ‘first run’ failure followed by a successful execution a third run of the same test should be performed and in the unlikely event of this third run being a failure the test procedure should be also aborted and the overall functional testing declared FAILED, as this would imply a not reliable operability of the instrument.

As a general remark ANY failure should be closely analyzed.

Note:  If the functional test is declared FAILED refer to section 4.1 for instrument switch OFF.

## GENERAL TEST PROCEDURE LAYOUT

The table below shows the general Short WFT sequence as it should be performed. In each step of this procedure the operator should refer to the detailed procedure in Section 3.2 .Test Control TCL scripts are available to invoke the correspondent CUS script stored in the HCSS database for each functional test. These CUS scripts will generate the appropriate command sequence for the particular functional test.

| Step | Subsystem tested | Test Id | Test Purpose |
| --- | --- | --- | --- |
| 1 | SCU | FUNC-SCU-01 | SCU Nominal Science Generation Check |
| 2 |  | FUNC-SCU-03 | FPU DC Thermometry Check |
| 3 |  | FUNC-SCU-06 | FPU AC Thermometry Check |
| 4 |  | FUNC-SCU-07 | Sorption Cooler Check |
| 5 |  | FUNC-SCU-04 | Photometer Calibrator Check |
| 6 |  | FUNC-SCU-05 | Spectrometer Calibrators Check |
| 9 | MCU | FUNC-MCU-01 | MCU Boot Check |
| 10 |  | FUNC-MCU-02 | MCU Nominal Science Generation Check |
| 11 | BSMm | FUNC-BSM-01 | BSM Switch ON Check |
| 12 |  | FUNC-BSM-03 | BSM Open Loop dynamics Check |
| 13 | SMECm | FUNC-SMEC-02A/B* | SMEC Launch Latch Open/Close Check |
| 14 |  | FUNC-SMEC-01 | SMECm Switch ON Check |
| 15 |  | FUNC-SMEC-03 | SMEC LED Integrity Check |
| 16 |  | FUNC-SMEC-04 | SMEC Positioning Test (Open Loop) |
| 17 | DCU | FUNC-DCU-01 | DCU Nominal Science Generation Check |
| 18 | Photometer LIAs | FUNC-DCU-04P | Photometer LIAs Check |
| 19 | Spectrometer LIAs | FUNC-DCU-04S | Spectrometer LIAs Check |

Table 2. General Short WFT sequence

## DETAILED TEST PROCEDURE

The following is a detailed (test by test) procedure including the steps required to perform each functional test individually.

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open SCU_PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-SCU-01

| Test Id: | FUNC-SCU-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON |
| Final Configuration: | DRCU_ON |
| Success Criteria: | Test passed if : Two SCU Nominal Science Report telemetry packets are received on QLA with the following characteristics: APID Type Subtype SID FrameID Frame length 0x508 21 1 0xA20 0x20 0x1E The frame time difference between consecutive SCU frames within these packets corresponds to the sampling rate. Nominal SCU sampling rate is 80Hz ( Δt = 12.5 ms The SPIRE HK parameter SCUFRAMECNT increments by 31. No events are generated during the frame generation. QLA to give go ahead. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Annotate the initial value of SCUFRAMECNT parameter located in SCU_PARAMETERS display. |
| 2 | Run QLA script FUNC-SCU-01.py on QLA console. |
| 3 | Run FUNC-SCU-01 test procedure from the HCSS Test Procedure window on TOPE |
| 4 | Annotate the final value of SCUFRAMECNT. |
| 5 | Contingency: If test fails repeat steps 1 to 4. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-01 | SCUFRAMECNT | n/ n+ 31 | 0/ 31 | 31 | Passed |
| Start time @: 13:20 End time @: 13:21 OBSID: 0x3000E12E Comments: QLA script created file QLA-SCU-01_3000E12E.txt. This file contains the following: ************************************************************************* SCU: APID = 0x508, SID = 0xa20 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. SCUFRAMECNT 0 31 31 31 Packet type = 0x15 TM5N 3284 3286 2 2 subtype = 0x1 FrameTime 12.4960 12.4992 Frame ID = 0x20 Frame Len = 0x1E STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 12.49793 ms sigma = 0.00161 ms |

### FUNC-SCU-03

| Test Id: | FUNC-SCU-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON |
| Final Configuration: | DRCU_ON + DC thermometry ON |
| Success Criteria: | Test passed if all FPU DC thermometry sensors show temperature values indicating a correct functioning of the sensor, not open/short-circuited. If ANY reading is anomalous check RAW sensor reading. Open Circuit Criterion: RAW reading in the range [0, -100] Short Circuit Criterion: RAW reading of -32768 Note: At warm (what does warm mean in the context of this procedure??) temperature most of the FPU thermometers should be short circuited. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-SCU-03 test procedure from the HCSS Test Procedure window on TOPE |
| 2 | When the test is finished annotate the current value of SCUTEMPSTAT and the RAW/converted values of the 16 FPU temperatures located in SCU_PARAMETERS display. |
| 3 | Contingency: If test fails execute SCU_OFF procedure from HCSS Test Procedure window on TOPE and then repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-03 | SCUTEMPSTAT | 0/0xFFFF | 0xFFFF/0xFFFF | N/A | Passed |
| Start time @: 13:25 End time @: 13:40 OBSID: N/A Comments: Procedure should clarify that this is a check that if done during cooldown does not required the execution of a script, just the check is required. fpu temp temp(K) temp(raw) Pumphtrtemp 55.22 -32768 (still out of range) Pumphstemp 37.52 -32768 (still out of range) Evaphstemp 38.52 -32768 (still out of range) Shunttemp 19.94 -32768 (still out of range) Emcfiltemp 78.38 -13678 Sl0temp 19.89 -32768 (still out of range) Pl0temp 21.30 -32768 (still out of range) Optemp 78.11 -22142 Baftemp 79.84 -18760 Bsmiftemp 78.37 -28363 Scal2temp 77.45 -30807 Scal4temp 77.75 -30259 Scaltemp 79.47 -30600 Smeciftemp 78.71 -29661 Smectemp 8.39 -32768 (still out of range) Bsmtemp 12.14 -32768 (still out of range) |

### FUNC-SCU-06

| Test Id: | FUNC-SCU-06 |
| --- | --- |
| Initial Configuration: | DRCU_ON + DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if SUBKSTAT parameter went from 0 to 1. Open Circuit Criterion: RAW reading in the range 0 -100 Short Circuit Criterion: RAW reading of -32768 |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-SCU-06 test procedure from the HCSS Test Procedure window on TOPE. |
| 2 | When the test is finished annotate the current value of SUBKSTAT located in SCU_PARAMETERS display. Also write down the RAW value of the SUBKTEMP parameter. |
| 3 | Contingency: If test fails : Send manual command: SEND_DRCU_COMMAND Parameter1 = 0xA0860000 Parameter2 = 0 Then repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-06 | SUBKSTAT | 0/1 | 1/1 | N/A | Passed |
| Start time @: 13:41 End time @: 13:42 OBSID: N/A Comments: Procedure should clarify that this is a check that if done during cooldown does not required the execution of a script, just the check is required. fpu temp temp(K) temp(raw) Subktemp 169 32730 (still out of range) |

### FUNC-SCU-07

| Test Id: | FUNC-SCU-07 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if during the execution of the test the following SCU HK parameters give correspondent readings of: SCU HK parameter RAW Converted SPHSV ~12715 ~323mV EVHSV ~12715 ~323mV SPHTRV ~14390 ~ 8 V |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-SCU-07 test procedure from the HCSS Test Procedure window on TOPE. |
| 2 | While the test is running annotate the values of current values of SPHSV, EVHSV, SPHTRV located in SCU_PARAMETERS display. (RAW and CONVERTED) |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-07 | SPHSV EVHSV SPHTRV | 0/ ~ 323 mV 0/ ~ 323 mV 0/ ~ 8 V | 0/323.22mV 0/322.49mV 0/8.83V | N/A | PASS |
| Start time @: 13:49 End time @: 13:51 OBSID: 0x3000E12F Comments: |

### FUNC-SCU-04

| Test Id: | FUNC-SCU-04 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if PCALCURR/PCALV SCU HK parameters show the following values: PCALCURR HK parameter which shows the measured PCAL current is ~ 0.1 mA. PCALV HK parameter which shows the measured PCAL voltage is ~ 0.02V |

Test Procedure:

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of PCALV and PCALCURR located in SCU_PARAMETERS display. |  |
| 2 | Run FUNC-SCU-04 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | While the test is running Annotate the values of PCALV and PCALCURR. |  |
| 4 | Contingency: If test fails repeat steps 1 to 3. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/ During test | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-04 | PCALCURR PCALV | 0/0.1mA 0/0.02V | 0/0.0981mA 0/0.02V | N/A | PASS |
| Start time @: 13:58 End time @:13:59 OBSID: 0x3000E130 Comments: |

### FUNC-SCU-05

| Test Id: | FUNC-SCU-05 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON |
| Success Criteria: | Test passed if : SCAL2CURR ,SCAL4CURR HK parameters which show the measured current read ~ 0.1 mA SCAL2V,SCAL4V parameters which show the measured voltage read ~ 0.05V. |

Test Procedure

| Step# | Action | Comments |
| --- | --- | --- |
| 1 | Annotate the current value of SCAL2V ,SCAL2CURR,SCAL4V,SCAL4CURR located in SCU_PARAMETERS display. |  |
| 2 | Run FUNC-SCU-05 test procedure from the HCSS Test Procedure window on TOPE |  |
| 3 | While the test is running write the values of SCAL2V ,SCAL2CURR,SCAL4V,SCAL4CURR. |  |
| 4 | Contingency: If test fails repeat steps 1 to 3. |  |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/During test | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SCU-05 | SCAL2CURR SCAL2V SCAL4CURR SCAL4V | 0/0.1mA 0/0.05V 0/0.1mA 0/0.05V | 0/0.099mA 0/0.0498V 0/0.1mA 0/0.05V | N/A | PASS |
| Start time @: 13:59 End time @: 14:00 OBSID: 0x3000E131 Comments: SCAL4 followed by SCAL2 |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open MCU_PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-MCU-01

| Test Id: | FUNC-MCU-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if: MCU boots. MCU voltages show expected values. MAC Board Temperature Reading shows ambient temperature. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Run FUNC-MCU-01 test procedure from the HCSS Test Procedure window on TOPE |
| 2 | When procedure is finished Annotate the values of the MCU voltages. |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-01 | MCUP5V MCUP15V MCUP14V MCUM14V MCUM15V MCUMACTEMP MCUBSMTEMP MCUSMECTEMP | N/A/ ~ 5V N/A / ~15V N/A / ~ 14V N/A / ~ -14V N/A / ~ -15V N/A / ~ 300K | / 5.0V / 15.52V /14.53V / -14.61V /-15.6V / 289.14K /293.78K /293.23K | N/A | Pass |
| Start time @: 14:35 End time @:14:36 OBSID: 0x3000E132 Comments: |

### FUNC-MCU-02

| Test Id: | FUNC-MCU-02 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Success Criteria: | Test passed if : MCU produces each type of the frames requested and with the following characteristics. Frame APID Type Subtype SID FrameID Frame length Eng. 0x508 21 3 0x814 0x14 0x15 BSM 0x508 21 1 0x612 0x12 0xD SMEC 0x508 21 1 0x410 0x10 0xC BSM +SMEC No events are generated during the different frames generation. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Annotate the current value of MCUFRAMECNT located in MCU_PARAMETERS display |
| 2 | Run QLA script FUNC-MCU-02.py on QLA console. |
| 3 | Run FUNC-MCU-02 test procedure from the HCSS Test Procedure window on TOPE |
| 4 | When test is finished annotate the current value of MCUFRAMECNT. |
| 5 | Contingency: If test fails repeat steps 1 to 4. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-MCU-02 | MCUFRAMECNT | 0/ ~ 6600 | 10008/ 16534 | 6526 | Pass |
| Start time @: 14:40 End time @: 14:42 OBSID: 0x3000E133 Comments: QLA produced file QLA-MCU-02_3000E133.txt. It contains: ************************************************************************* MCUENG: APID = 0x508, SID = 0x814 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. MCUFRAMECNT 10008 10649 641 609 Packet type = 0x15 TM5N 3286 3314 28 27 subtype = 0x3 FrameTime 15.5776 15.5807 Frame ID = 0x14 Frame Len = 0x15 STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 15.58035 ms sigma = 0.00121 ms ************************************************************************* BSM: APID = 0x508, SID = 0x612 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. MCUFRAMECNT 10649 11290 641 609 Packet type = 0x15 TM5N 3314 3331 17 17 subtype = 0x1 FrameTime 15.5807 15.5809 Frame ID = 0x12 Frame Len = 0xD STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 15.58036 ms sigma = 0.00111 ms ************************************************************************* SMEC: APID = 0x508, SID = 0x410 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. MCUFRAMECNT 11290 13665 2375 2375 Packet type = 0x15 TM5N 3331 3389 58 58 subtype = 0x1 FrameTime 4.2112 4.2112 Frame ID = 0x10 Frame Len = 0xC STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 4.21091 ms sigma = 0.00093 ms ************************************************************************* SMEC+BSM: APID = 0x508, SID = 0x612 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. MCUFRAMECNT 13665 16534 2869 2850 Packet type = 0x15 TM5N 3389 3460 71 71 subtype = 0x1 FrameTime SMEC 4.2112 4.2112 Frame ID = 0x10, Len = 0xC FrameTime BSM 20.2146 20.2146 Frame ID = 0x12, Len = 0xD STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 4.21091 ms (SMEC), 20.21236 ms (BSM) sigma = 0.00093 ms (SMEC), 0.00154 ms (BSM) DISCREPANCY: For all MCU frames the sampling does not appear to be consistent with the duration of MCU QM2 clock cycle, viz. 420µs. |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open CHOP PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-BSM-01

| Test Id: | FUNC-BSM-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if: CHOPSENSPWR HK parameter goes from 0 to 1 CHOPDACVAL HK parameter goes from 0 to 0x8000 CHOPSENSIG HK parameter shows variation from off to on JIGGSENSPWR HK parameter goes from 0 to 1 JIGGDACVAL parameter goes from 0 to 0x8000 JIGGSENSSIG HK parameter shows variation from off to on |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following HK parameters: CHOPSENSPWR CHOPDACVAL CHOPSENSSIG JIGGSENSPWR JIGGDACVAL JIGGSENSSIG |
| 2 | Run FUNC-BSM-01 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | When the test is finished record all the Key parameters noted bellow |
|  | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-01 | CHOPSENSPWR CHOPLOOPMODE CHOPDACVAL CHOPFFGAIN CHOPSENSSIG JIGGSENSPWR JIGGLOOPMODE JIGGDACVAL JIGGFFGAIN JIGGSENSSIG | 0/1 3/3 0x8000/0x8000 0/0xBEB 0/? 0/1 3/3 0x8000/0x8000 0/0xBEB 0/? | 0/1 3/3 0x8000/0x8000 0xBEB/0xBEB 0x7FED/~ 0x92DE 0/1 3/3 0x8000/0x8000 0xBEB/0xBEB 0x7FE4/0xA0F5 | N/A | Pass |
| Start time @:15:00 End time @:15:01 OBSID:0x3000E134 Comments: BSM switched ON normally |

### FUNC-BSM-03

| Test Id: | FUNC-BSM-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON + BSM ON (open loop) |
| Success Criteria: | Test passed if the chop sensor signal evolves in the same way as the positions set.(i.e if (pos1 > pos2 ( sig1 > sig2) for each jiggle position. Note: During warm tests the voltages on both chop and jiggle motors are likely to be saturated (CHOP/JIGGMOTORVOLT RAW values of ~ 0xFFFF) due to the high resistance of the motor coil at ambient temperature. |

Test Procedure

| Step# | Action |
| --- | --- |
| 1 | On QLA open up a time series display of HK parameters: CHOPPOSN CHOPDACVAL CHOPMOTORCURR CHOPSENSSIG CHOPMOTORVOLT JIGGPOSN JIGGDACVAL JIGGMOTORCURR JIGGSENSSIG JIGGMOTORVOLT |
| 2 | Run FUNC-BSM-03 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat step 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-BSM-03 |  |  |  | N/A | Pass |
| Start time @: 15:18 End time @: 15:30 OBSID: 0x3000E135 Comments: Run the test with the following parameters: BSM sampling 125Hz Chop start position = 0x1000 Chop end position = 0xf000 Chop step = 0x1000 Jigg start position = 0x7000 Jigg end position = 0x9000 Jigg step =0x2000 Time at each position = 3 sec Must have input the wrong parameter the first time by hand so repeated with Obsid 0x3000E126 |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute BSM_OFF from HCSS Test Procedures |  |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open SMEC PARAMETERS display on SCOS Alpha Numeric Displays. |  |

### FUNC-SMEC-01

| Test Id: | FUNC-SMEC-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if : SMECENCPWR HK parameter changes from 0 to 6. SMEC encoder signals 1 and 2 amplitude show variation when encoder is switched ON. SMEC LVDT is switched ON. SMEC LVDT DC and AC signals show variation when LVDT is switched ON. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a display of the following HK parameters: SMECENCPWR SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG |
| 2 | Run FUNC-SMEC-01 test procedure from the HCSS Test Procedure window on TOPE |
|  | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-01 | SMECENCPWR SMECLVDTPWR SMECENCSIG1 SMECENCSIG2 |  | 0/ 6 0/ 1 | N/A | Pass |
| Start time @: 15:35 End time @: 15:37 OBSID: 0x3000E138 Comments: Script still asks for the SMEC orientation, this input parameter should be removed from the script. |

### FUNC-SMEC-03

| Test Id: | FUNC-SMEC-03 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON +MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if: SMEC encoder signals 1 and 2 show a variation from one LED illumination level to another. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following Nominal HK parameters: SMECENCPWR SMECENCSIG1 SMECENCSIG1 |
| 2 | Run FUNC-SMEC-03 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-03 | SMECENCPWR SMECENCSIG1 SMECENCSIG2 |  | Values changed from 4 to 5 to 6 as expected | N/A | Pass |
| Start time @: 15:42 End time @: 15:43 OBSID: 0x3000E13A Comments: |

### FUNC-SMEC-04a

| Test Id: | FUNC-SMEC-04a |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON+ SMEC ON (open loop) |
| Success Criteria: | Test passed if SMECLVDTDCSIG parameter shows a variation according to the different positions set. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the following Nominal HK parameters: SMECENCSIG1 SMECENCSIG2 SMECLVDTDCSIG SMECLVDTACSIG SMECMOTORCURR This should be MCU ENG frame parameters not HK parameters |
| 2 | Run FUNC-SMEC-04a test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-SMEC-04a | SMECLVDTDCSIG |  |  |  | Pass |
| Start time @: 15:49 End time @: 15:54 OBSID: 0x3000E13B Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute SMEC_OFF from HCSS Test Procedures |  |

| Step# | Action | Comments |
| --- | --- | --- |
| 0 | Open DCU PARAMETERS SCOS Alpha Numeric Display |  |

### FUNC-DCU-01

| Test Id: | FUNC-DCU-01 |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Success Criteria: | Test passed if: DCU produces each type of DCU nominal science frame with the following characteristics. APID Type S.type SID Frame ID Frame type Nb. Of frames Nb. of pkts. 0x504 21 1 0x200 0 PF 100 100 0x506 21 1 0x201 1 SF 100 17 0x504 21 2 0x102 2 PSW 100 34 0x504 21 2 0x103 3 PMW 100 25 0x504 21 2 0x104 4 PLW 100 12 0x506 21 2 0x105 5 SSW 100 12 0x506 21 2 0x106 6 SLW 100 7 The frame time difference between consecutive DCU frames of each type corresponds to the sampling rate. Photometer Sampling rate is 15.3Hz ( Δt ~ 65.5 ms Spectrometer Sampling rate is 80Hz ( Δt = 12.5 ms The SPIRE HK parameter DCUFRAMECNT increments by 700. No events are generated during the frames generation. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |
|  | Step missing to run QLA script! |
| 2 | Run FUNC-DCU-01 test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Annotate the current value of DCUFRAMECNT located d in DCU PARAMETERS AND |
| 4 | Contingency: If test fails repeat steps 1 to 3. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-01 | DCUFRAMECNT | 0/700 | 5215/ 5915 | 700 | Pass |
| Start time @: 16:02 End time @: 16:04 OBSID: 0x3000E13D Comments: QLA wrote the following file QLA-DCU-01_3000E13D.txt: ************************************************************************* PHOTF: APID = 0x504, SID = 0x200 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5215 5315 100 100 Packet type = 0x15 TM3N 2285 2385 100 100 subtype = 0x1 FrameTime 53.7600 53.7600 Frame ID = 0x0 Frame Len = 0x126 STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 53.75883 ms sigma = 0.00160 ms ************************************************************************* PHOTSW: APID = 0x504, SID = 0x102 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5315 5415 100 100 Packet type = 0x15 TM3N 2385 2419 34 34 subtype = 0x2 FrameTime 53.7600 53.7568 Frame ID = 0x2 Frame Len = 0x96 STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 53.75880 ms sigma = 0.00144 ms ************************************************************************* PHOTMW: APID = 0x504, SID = 0x103 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5415 5515 100 100 Packet type = 0x15 TM3N 2419 2444 25 25 subtype = 0x2 FrameTime 53.7600 53.7568 Frame ID = 0x3 Frame Len = 0x66 STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 53.75880 ms sigma = 0.00161 ms ************************************************************************* PHOTLW: APID = 0x504, SID = 0x104 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5515 5615 100 100 Packet type = 0x15 TM3N 2444 2456 12 12 subtype = 0x2 FrameTime 53.7568 53.7600 Frame ID = 0x4 Frame Len = 0x36 STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 53.75880 ms sigma = 0.00162 ms ************************************************************************* SPECF: APID = 0x506, SID = 0x201 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5615 5715 100 100 Packet type = 0x15 TM4N 203 220 17 17 subtype = 0x1 FrameTime 12.4896 12.4928 Frame ID = 0x1 Frame Len = 0x4E STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 12.49251 ms sigma = 0.00091 ms ************************************************************************* SPECSW: APID = 0x506, SID = 0x105 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5715 5815 100 100 Packet type = 0x15 TM4N 220 232 12 12 subtype = 0x2 FrameTime 12.4928 12.4896 Frame ID = 0x5 Frame Len = 0x36 STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 12.49251 ms sigma = 0.00094 ms ************************************************************************* SPECLW: APID = 0x506, SID = 0x106 _________________________________________________________________________ Parameter Initial Final Increment Expect Incre. Packet Chars. DCUFRAMECNT 5815 5915 100 100 Packet type = 0x15 TM4N 232 239 7 7 subtype = 0x2 FrameTime 12.4896 12.4928 Frame ID = 0x6 Frame Len = 0x1E STATISTICS ON TIME BETWEEN RECEPTION OF 2 CONSECUTIVE FRAMES: mean = 12.49251 ms sigma = 0.00094 ms |

### FUNC-DCU-04P

| Test Id: | FUNC-DCU-04P |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Photometer LIAs ON |
| Success Criteria: | Test passed if : SCUDCDCSTAT parameter goes from 4 to 5. Photometer LIA card voltages are showing correct readings of +5V,+9V,-9V. Photometer LIA temperatures show an increase indicating that they are ON. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the HK parameters: PLIAP5V PLIAP9V PLIAM9V LIAP1/2/3/4/5/6/7/8/9TEMP |
| 2 | Run FUNC-DCU-04P test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-04P | SCUDCDCSTAT PLIAP5V PLIAP9V PLIAM9V LIA1/2/3/4/5/6/7/9TEMP | 4/5 0/~ +5V 0/~+9V 0/~-9V N/A/ [290-300]K | 4/5 0.18V/5.19V 0.0027V/ 11.52V 0.0027V/ -11.52V ~ 290K/ 290K and warming up | N/A | Pass |
| Start time @: 16:17 End time @: 16:18 OBSID:0x3000E13E Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute PLIA_OFF from HCSS Test Procedures |  |

### FUNC-DCU-04S

| Test Id: | FUNC-DCU-04S |
| --- | --- |
| Initial Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON |
| Final Configuration: | DRCU_ON + AC/DC thermometry ON+MCU ON + Spectrometer LIAs ON |
| Success Criteria: | Test passed if : SCUDCDCSTAT parameter goes from 4 to 6. Spectrometer LIA card voltages are showing correct readings of +5V,+9V,-9V. Spectrometer LIA temperatures show an increase indicating that they are ON. |

Test Procedure:

| Step# | Action |
| --- | --- |
| 1 | On QLA bring up a time series display of the HK parameters: SLIAP5V SLIAP9V SLIAM9V LIAS1/2/3TEMP |
| 2 | Run FUNC-DCU-04S test procedure from the HCSS Test Procedure window on TOPE |
| 3 | Contingency: If test fails repeat steps 1 and 2. |

Test Log:

| Test Id | Key Parameter(s) | Expected Value Before/After | Actual Value Before/After | Nb. of frames received | Test Result |
| --- | --- | --- | --- | --- | --- |
| FUNC-DCU-04S | SCUDCDCSTAT SLIAP5V SLIAP9V SLIAM9V LIA1/2/3TEMP | 4/6 0/~ +5V 0/~+9V 0/~-9V N/A/ [290-300]K | 4/6 0.075V/ 5.22V 0.0036V/ 11.53V 0.0027V/ -11.52V ~ 290K/ 290K and warming up | N/A | Pass |
| Start time @:16:25 End time @:16:26 OBSID: 0x3000E13F Comments: |

| Step# | Action | Comments |
| --- | --- | --- |
| 4 | Execute SLIA_OFF from HCSS Test Procedures |  |

# END TEST SEQUENCE

## NORMAL END TEST SEQUENCE

The following table shows the necessary steps to be followed to end the warm functional test sequence.

| Step# | Action | HK parameters | Expected Value | Comments | Check |
| --- | --- | --- | --- | --- | --- |
| 1 | Check BSM is OFF | CHOPSENSPWR JIGGSENSPWR | 0 0 |  | ( ( |
| 2 | Check SMEC is OFF | SMECENCPWR | 0 |  | ( |
| 3 | Run MCU-OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT | 0 |  | ( |
| 4 | Check Instrument Configuration | MODE | 0x100/ DRCU_ON |  | ( |
| * These following two steps should ONLY be executed if the functional test takes place BEFORE the instrument is placed in the test cryostat. If the instrument is already in the cryostat and the cryo-harness are connected this steps should NOT be executed. |
| 5* | Run SCU_OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT SUBKSTAT | 0 0 |  |  |
| 6* | Run DRCU_OFF from the HCSS Test Procedure window on TOPE | MODE | 0/ DPU_ON |  |  |

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS 1 to 4 are executed is DRCU_ON

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS 1 to 6 are executed is DPU_ON

## END TEST SEQUENCE WHEN THE FUNCTIONAL TEST HAS FAILED

The following table shows the necessary steps to be followed to end the warm functional test sequence when this has been declared failed.

| Step# | Action | HK parameters | Expected Value | Comments | Check |
| --- | --- | --- | --- | --- | --- |
| 1 | Check BSM is OFF | CHOPSENSPWR JIGGSENSPWR | 0 0 |  |  |
| 2 | Check SMEC is OFF | SMECENCPWR | 0 |  |  |
| 3 | Run MCU-OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT | 0 |  |  |
| 4 | Check Instrument Configuration | MODE | 0x100/ DRCU_ON |  |  |
| 5 | Run SCU_OFF from the HCSS Test Procedure window on TOPE | SCUDCDCSTAT SUBKSTAT | 0 0 |  |  |
| 6 | Run DRCU_OFF from the HCSS Test Procedure window on TOPE | MODE | 0/ DPU_ON |  |  |

FINAL INSTRUMENT CONFIGURATION IN THE CASE STEPS DPU_ON

# Annexe1 (Instrument nominal configuration/modes)

|  | OFF | INIT | DPU_ON | DRCU_ON | REDY | PHOT STBY | SPEC STBY | CREC | SAFE |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| DPU |  | ON | ON | ON | ON | ON | ON | ON | ON |
| Essential Hsk packets |  |  | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz | 0.5Hz |
| Normal Hsk packets |  |  | 1.0Hz | 1.0Hz | 0.25Hz | 0.25Hz | 0.25Hz | 1.0Hz | 0.25Hz |
| TC Acceptance |  | TBC | Yes | Yes | Yes | Yes | Yes | Yes | Yes |
| Event packets |  | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes |
| Science packets |  |  |  |  |  |  |  |  |  |
| VM |  |  |  |  |  |  |  | Cooler_Recycle |  |
| VM1 |  |  |  |  |  | Det_Temp_Ctrl |  |  |  |
| VM2 |  |  |  |  |  |  |  |  |  |
| VM3 |  |  |  |  |  |  |  |  |  |
| MODE |  |  | 0x0000 | 0x0100 | 0x0200 | 0x0300 | 0x0400 | 0x0600 | 0x0900 |
| DRCU |  |  | ON | ON | ON | ON | ON | ON |  |
| SCU |  |  |  |  |  |  |  |  |  |
| Temp Channels powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| SubK Channel powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| PCAL source powered |  |  |  |  |  |  |  |  |  |
| SCAL sources powered |  |  |  |  |  |  | TBD |  |  |
| TC Heater powered |  |  |  |  |  | Yes |  |  |  |
| Cooler SP Heater powered |  |  |  |  |  |  |  | Yes |  |
| Cooler EV HS powered |  |  |  |  |  |  |  | Yes |  |
| Cooler SP HS powered |  |  |  |  | Yes | Yes | Yes | Yes |  |
| DCU |  |  |  |  |  |  |  |  |  |
| Photometer BIAS |  |  |  |  |  | Yes |  |  |  |
| Photometer JFETS |  |  |  |  |  | Yes |  |  |  |
| Photometer LIAs |  |  |  |  |  | Yes |  |  |  |
| TC BIAS |  |  |  |  |  | Yes |  |  |  |
| TC JFETS |  |  |  |  |  | Yes |  |  |  |
| TC LIAs |  |  |  |  |  | Yes |  |  |  |
| Spectrometer BIAS |  |  |  |  |  |  | Yes |  |  |
| Spectrometer JFETS |  |  |  |  |  |  | Yes |  |  |
| Spectrometer LIAs |  |  |  |  |  |  | Yes |  |  |
| MCU |  |  |  |  |  |  |  |  |  |
| DSP |  |  |  |  | On | On | On | On |  |
| BSM |  |  |  |  |  | Hold | Hold |  |  |
| SMEC |  |  |  |  |  |  | Hold |  |  |