# Scope

This document outlines the tests that will be carried out during and after integration of the SPIRE instrument into the Herschel cryostat.  This document is intended as an informational note for the purposes of the Science Verification Review (SVR) to allow the review board to comment on the completeness or otherwise of the planned tests.

# Organisation of the tests

For the system level tests we follow the test campaign organisation proposal by Alcatel vis.

EIT-WU 	– Electrical Integration Test – Warm Units

EIT-WFT	– Electrical Integration Test – Warm Functional Test

EIT-SWFT	– Electrical Integration Test – Short Warm Functional Test

UFT-WU 	– Unit Functional Test – Warm Units

IST-CFT 	– Integrated System Test – Cold Functional Test

IST-SCFT 	– Integrated System Test – Short Cold Functional Test

IST-SPT 	– Integrated System Test – Special Performance Test

IST-TV  	– Integrated System Test – Cold Functional Test

IST-EMC 	– Integrated System Test – Cold Functional Test

SVT		– System Verification Test

When and where each of these phases are carried is not considered in this document, only the tests that SPIRE is required/wishes to be done.

I have broken the SPIRE tests into four sections:

- Functional Tests
The standard battery of tests designed to ensure everything is connected and functions as seen during instrument level testing (ILT)

- Performance Verification Tests
Again a standard set of tests designed to ensure that the basic instrument performance is the same as, or can be extrapolated from, the results from ILT

- Instrument Flight Operations
Tests designed to exercise a subset of the instrument in flight operations to provide standard configurations for thermal verification and System Verification Tests (SVT)

- Special Engineering Tests
These are tests designed to evaluate the parts of SPIRE’s performance within the Herschel environment that cannot be tested during ILT – they include things like response to microphonics; EMC and thermal balance.

There is some overlap between the sections but I hope that organising the document in this way will make it more readable.  The order that the tests appear in these sections does not imply any sequence of execution.

# SPIRE System Level Tests

## Functional

| Name of test | Description | Purpose | When carried out |
| --- | --- | --- | --- |
| Warm Functional Test | Full checkout of SPIRE functionality whilst the FPU is warm and either under vacuum or at atmospheric pressure. | Basic electrical integration. instrument health check and characterisation at warm temperature | EIT-WU UFT-WU EIT-WFT |
| Cold Functional Test | Full checkout of SPIRE functionality once the FPU is cold and either under vacuum or at atmospheric pressure. | Instrument health check and functional characterisation at cold temperatures | EIT-CFT IST-SPT IST-TV |
| Short Functional Test – warm | Rapid check of SPIRE functionality to be carried out at key inspection points during integration with the FPU warm either under vacuum or at atmospheric pressure. | Functional check the does not require instrument personnel to support. | EIT-SWFT IST-SPT IST-TV |
| Short Functional Test – cold | Rapid check of SPIRE functionality to be carried out at key inspection points during test with the FPU cold and at operating temperature. Does not require instrument support. | Functional check the does not require instrument personnel to support. | EIT-SCFT IST-SPT IST-TV |

## Performance Verification

| Name of test | Description | Purpose | When carried out |
| --- | --- | --- | --- |
| Photometer Thermal Control Verification | With the instrument in PHOT STBY the photometer thermal control system will be activated using control parameters established during ILT and previous engineering tests (q.v.) | To verify thermal stabilisation performance. | IST-SPT IST-TV |
| Photometer bias optimisation | Measurement of the photometer noise performance as a function of bias amplitude and frequency | Verification of ILT results and system level noise performance | IST-SPT |
| Spectrometer bias optimisation | Measurement of the spectrometer noise performance as a function of bias amplitude and frequency | Verification of ILT results and system level noise performance | IST-SPT |
| Spectrometer SCAL check | Set each spectrometer calibration source to nominal operating temperature and take spectrum against ambient background set to close to nominal telescope power | Final verification of SCAL nulling and thermal performance | IST-SPT |

## Flight Operations

| Name of test | Description | Purpose | When carried out |
| --- | --- | --- | --- |
| SPIRE to Ready mode (REDY) | Only temperature sensors are switched on and housekeeping is collected. | Required for transition between modes and is the default mode that SPIRE will in when not prime instrument | IST-SPT IST-TV SVT |
| Cooler Recycle (CREC) | Once we have established the correct operating parameters for the cooler we can use an automatic procedure. | This will be the standard cooler recycle procedure for the majority of the IST. | IST-SPT IST-TV SVT |
| Photometer Standby (PHOT STBY) | Sets the instrument to be ready to go into one of the photometer observation modes. | This instrument setting will be used many times during IST both in transitions between test configurations and as a default condition. | IST-SPT IST-TV SVT |
| Photometer Scan Mode (POF5) | This sets the SPIRE instrument to its nominal scan mode AOT configuration. | We plan using this mode for thermal testing and operations checkout. | IST-SPT IST-TV SVT |
| Photometer 7-Point jiggle/chop with nodding (POF2) | This sets the SPIRE instrument to its nominal chop and jiggle mode AOT configuration. | We plan using this mode for thermal testing and operations checkout. | IST-SPT IST-TV SVT |
| Spectrometer Standby (SPEC STBY) | Sets the instrument to be ready to go into one of the spectrometer observation modes. | This instrument setting will be used many times during IST both in transitions between test configurations | IST-SPT IST-TV SVT |
| Spectrometer high resolution spectrum (SOF1) | This sets the SPIRE instrument to its spectrometer mode AOT configuration where the SMEC is scanned over its full operating range. | We plan using this mode for thermal testing and operations checkout | IST-SPT IST-TV SVT |
| Cooler Hold time | A full 48 hour cycle of flight like operations is conducted following a cooler recycle under close to flight thermal conditions | Check the cooler hold time performance in flight like environment | IST-TV |

## Special Engineering

| Name of test | Description | Purpose | When carried out |
| --- | --- | --- | --- |
| Ambient background verification | Loadcurves and calibrator flashes with the cryocover set to a number of different conditions. Carried out for both photometer and spectrometer | We need to test the straylight environment within the cryostat under different loading conditions. This test also checks the BDA temperatures. | IST-SPT IST-TV |
| Spectrometer mechanism micro vibration test | SPIRE set to spectrometer mode with the SMEC continuously scanning. During the test the S/C reaction wheels are activated and taken through a range of rotation speeds while the SMEC performance is monitored for interference | Final verification of the response of the SMEC to spacecraft induced microphonics | IST-SPT |
| Photometer EMI Susceptibility Tests | SPIRE set to most sensitive photometer mode. | Completion of formal radiative susceptibility tests as specified by IID-A | IST-EMC |
| Spectrometer EMI Susceptibility Tests | SPIRE set to most sensitive spectrometer mode | Completion of formal radiative susceptibility tests as specified by IID-A | IST-EMC |
| Spectrometer EMC Tests | SPIRE set to most noisy spectrometer mode | Completion of formal radiative emission tests as specified by IID-A | IST-EMC |
| PTC Headroom Characterisation | The power on the PTC heater is increased in steps whilst observing the response of the BDA and PTC thermistors. | Determine the required PTC power setting required to enable stable operation of the photometer detectors under ground nominal thermal conditions. | IST-SPT |
| 300mK Decontamination Test | The cooler pump is heated with both heat switches off in order to raise the temperature of the 300 mK stage as high as possible | To test efficiency of heating the 300 mK system using the pump in case it is needed to removed Helium condensed on the detectors. | IST-SPT |