This document gives the flowchart and parameter specification for the Automatic Cooler Recycle  Command List.

Version 4.6 for use in ILT test PFM4

| Param | Description | Setting | Current | Voltage | Hex |
| --- | --- | --- | --- | --- | --- |
| A | Heater Heat Switch ON (during Recycling) | 0.8 mW | 1.4 mA | 0.56V | 0x0DEB |
| B | Heaters OFF | 0 mW | 0.0 mA | 0V | 0x0000 |
| C | Pump Heat Switch – Actuation Temperature | 12 K | - | - | 0xBF9B |
| D | Heater Pump Dissipation 1 | 400 mW | 31.54 mA | 12.7V | 0x0A25 |
| E | Pump Temperature Condensation 1 | 45 K | - | - | 0x8E76 |
| F | Heater Pump Dissipation 2 | 100mW | 15.77mA | TBC | 0x0513 |
| G | Pump Temperature Condensation 2 | 46K | - | - | TBC |
| H | Heater Pump Dissipation 3 | 10mW | 4.987mA | 2V | 0x019C |
| I | Heater Pump Dissipation 4 | 70mW | 13.197mA | 5.3V | 0x043F |
| J | Pump Temperature Condensation Threshold | 45.1K | - | - | 0x8E49 |
| K | Evap Temperature Condensation | 2 K | - | - | 0x7EBE |
| L | Evaporator Heat Switch Actuation Temperature | 15K | - | - | 0xB764 |
| M | Pump Temperature Threshold | 2 K | - | - | 0xEFAE |
| N | Heater Heat Switch ON (during Recycling) | ~ 0.4 mW | 1.022 mA | 0.41V | 0x0A2A |
| O | Loop Sampling (secs) | 10 sec | - | - | 10 |
| P | Heat Switch Timeout2 (min) | ½ hr | - | - | 30 |
| Q | Pump Heating Timeout 12 (min) | 1hr | - | - | 60 |
| R | Pump Heating Timeout 22 (min) | 1hr | - | - | 60 |
| S | Evaporator Timeout2 (min) | 1 hr | - | - | 60 |
| T | Pump Cooling Timeout2 (min) | 1hr | - | - | 60 |
| U | Global Timeout12 (min) | 2hr | - | - | 120 |

Note 1: A global timeout variable should be implemented – it should stop the script U minutes after the first command has been sent should the script still be running.

Note 2: All timeouts are now specified in time rather than number of sample loops

# Text boxes

Timeout = R

N

Step 4

Y

Pump Temp > G

Heater Pump = F

Step 7

Issue Event F023

Issue Event F022

Issue Event F021

Issue Event F020

Step 7

Step 6

Step 5

Step 3

Step 2

Step 1

Timeout = S

Timeout = S

END

END

IDLE State

Timeout = T

Timeout = P

Timeout = Q

Timeout = P

Heater Pump = B

Heater Evap HS = B

Heater Pump HS = B

N

Heater Pump HS = N

N

Y

Pump Temp < M

Heater Pump = B

N

Y

Evap HS Temp < L

Heater Pump HS = A

Heater Evap HS = B

N

Y

Evap Temp < K

N

N

Y

Evap Temp < K

N

Y

N

Y

Y

Pump Temp > J

Heater Pump = I

Pump Temp < E

Heater Pump = H

Y

Pump Temp > E

Heater Pump = D

Pump HS Temp < C

Heater Pump HS = B

Heater Evap HS = A

Heater Evap HS = A

END

Heater Pump HS = B

Heater Evap HS = A

Heater Pump HS = B

Heater Pump HS = B

Global Timeout = U

Issue Event F024