## Summary
SPIRE-RAL-TN-002783
SUMMARY
issue 13
| BoDAC | PFM3 |  |  | PFM4 | [7 empty] | IST
| Dead | Dead | Noisy | Shorted/Open | Dead | WE Noisy? | Not Operational | Noisy |  | Offset | Swap | Microphonic | Dead | Noisy | Non Operational | Slow
PLW | 0 | 0 | 2 (A6,C9) | 1 (A2) | 0 |  | 2 (A2 DK1) | 4 (C9,A6,E2,B4) |  | A6, T2? |  |  | 0 | 2 (A6, C9) | 0 | 0
PMW | 0 | 0 | 10 (A11,A8,F2,E4,E5,F5,G6,F6,G7,B1) | 5 (C8,T2,F7,G9,B6) | 0 |  | 6 (C8,T2,G11,G9,DK1,B6) | 6 (G6,B11,E4,E5,F5,D9) |  | C9 | F7 |  | 1 (T2) | 4 (B11,D6,E8,D1) | 0 | 1 (A13)
PSW | 2 (C12,J7) | 2 (C12,G8) | 7 (D14,A14,B12,C11,B5,G13,DK2) | 3 (A10,J6,H9) | 2 (C12,G8) |  | 4 (B3,A10,J6,F8) | 6 (G11,D14,F11,J7,A8,B5) |  | A12,D12 |  |  | 2 (C12, G8) | 2 (F9,G11) | 1 (D15) | 2 (A13,A11)
SLW | 1 (DK2) | 1 (DK2) | 5 (C1,D2,E2,A3,T2) | 2 (C2,B3) | 1 (DK2) |  | 2 (C2,B3) | 0 |  |  |  |  | 0 | 0 | 0 | 0
SSW | 1 (D5) | 1 (D5) | 3 (A3,A2,E3) | 18 | 1 (D5) |  | 1 (D7) | 0 |  |  |  | 1 (F4) | 1 (D5) | 1 (F4) | 0 | 0
PTC | N/M | 0 | 0 | 1 (PTC3) | 0 |  | 0 | 0
|  | 4 | 27 | 30 | 4 |  | 15 | 16 |  | 5 | 1
[7 empty] | 13 definite cryoharness |  | IST Result | Preflight Status |  | F4 was noisy in PFM3
[7 empty] | PLW-DK1 | PLW-C9 repeats in PFM3 & PFM4 | Confirmed | Ok a bit noisy
[7 empty] | PMW-T2 | PLW-A6 repeats in PFM3 & PFM4 | Confirmed | Declare dead
[7 empty] | PMW-C8 | PLW-E2 1/f noisy | OK IST | Ok
[7 empty] | PMW-B6 | PLW-B4 anomolous measurement? | OK IST | Ok
PMW-DK1
[7 empty] | PMW-G9 | PMW-G6 repeats in PFM3 & PFM4 | OK IST
[7 empty] | PMW-G11 | PMW-B11 1/f noisy | Confirmed | I/f noisy
[7 empty] | PSW-A10 | PMW-E4 slightly noisy with PTC on | OK IST
[7 empty] | PSW-J6 | PMW-E5 slightly noisy with PTC on | OK IST
[7 empty] | PSW-F8 | PMW-F5 slightly noisy with PTC on | OK IST
[7 empty] | SLW-C2 | PMW-D9 slightly noisy with PTC on | OK IST
SLW-B3
[7 empty] | SSW-D7 | PSW-G11 known BDA intermittant | Confirmed | Declare dead
| Change record | [6 empty] | PSW-D14 intermittant, 1/f noisy | OK IST | Ok
[8 empty] | PSW-F11 OK in PFM3 | OK IST | Ok
| Date | issue | description |  |  |  |  | PSW-J7 known BDA intermittant | OK IST | Ok
[8 empty] | PSW-A8 1/f noisy | OK IST | Ok
| 2006-12-01 | 8 | First formal issue |  |  |  |  | PSW-B5 1/f noisy, noisy in PFM3 | OK IST | Ok
| 2006-12-01 | 9 | Updated by Hien.
0.3K noise data added to SSW
| 2006-12-18 | 10 | Updated by Hien. |  |  |  |  | SLIGHT > 30 nV/rtHz
|  |  | 1/f knee data added. |  |  |  |  | MODERATE > 50 nV/rtHz
|  |  | Summary sheet cleaned up |  |  |  |  | HIGH > 100 nV/rtHz
| 2007-01-01 | 11 | Updated by Bruce
Added failure mechanism for cryoharness following pixel investigation (see SPIRE-RAL-REP-002778)
Added Tim Waskett's spreadsheet summarising speed of response as appendix
| 2009-01-08 | 12 | Updated by Bruce to include results from IST
| 2009-01-09 | 12.1 | Typos changed in summary and PMW-D1 added to noisy list as sometimes seems bad
| 2009-07-03 | 13 | First flight issue
| Photometer Final preflight tally |  | 6 Unusable 9 Dodgy
| Unusable |  | 6 (PLW-A6,PMW-T2,PSW-C12, PSW-D15, PSW-G8,PSW-G11)
| Noisy but usable |  | 6 (PLW-C9,PMW-B11,PMW-D1,PMW-D6,PMW-E8, PSW-F9)
| Slow but usable |  | 3 (PMW-A13, PSW-A13, PSW-A11)
| Spectrometer Final preflight tally |  | 2 Unusable 0 Dodgy
| Unusable |  | 2 (SSW-D5,SSW-F4)
| Noisy but usable |  | 0
| Slow but usable |  | 0
| Photometer Flight tally |  | 6 Unusable 10 Noisy or Slow
| Unusable |  | 6 (PLW-A6,PMW-T2,PSW-C12, PSW-D15, PSW-G8,PSW-G11)
| Suspect |  | 1 (PMW-D1)
| Noisy but usable |  | 5 (PLW-C9,PMW-B11,PMW-D6,PMW-E8, PSW-F9)
| Slow but usable |  | 4 (PMW-A13, PSW-A13, PSW-A11, PSW A10)

## PLW
PLW | BoDAC |  |  |  | PFM3 |  |  |  | PFM4 | [10 empty] | PFM4 | IST
| Pixel Name | BDA Operability | BoDAC channel Operability | Noisy Bodac? | Channel # | PFM3 Noise | PFM3 WE Noise | JJB Comment | JFET Gain Test | Warm Resistance test | JFET Gain Test | Warm Resistance test | JFET Idd | VSS at 4K | JFET + LIA Noise (nV/rtHz) | Noise (nV/rtHz) (1-3 Hz) | Noise (nV/rtHz) (0.06 - 0.3 Hz) (20061127) | 1/f Knee (PTC Control) (20061127) | 0.3 K load curve | Noise taken with correct VSS | Noise IST | IST Comment | Noise Flight | Operation flight
| R1 | Yes | Yes | No | 145 | 1.46e-08 | 1.87e-20 |  | 0.95711 | 4.1294 | 0.95711 | 4.1294 | 17/43 | -1.68V | 2.0287e-08 | 1.465e-08 | 1.875e-08 | 0.313 |  | 14.3522 | 16.5533 |  | 11.7297 |  | 1.173e-08 | R1
| A8 | Yes | Yes | No | 146 | 2e-08 | 1.02e-08 |  | 0.96005 | 2.7689 | 0.96005 | 2.7689 | 2.54 |  | 9.1713e-09 | 1.7565e-08 | 2.3807e-08 | 0.059 |  | 16.2529 | 19.8551 |  | 16.1298 |  | 1.613e-08 | A8
| A7 | Yes | Yes | No | 147 | 2e-08 | 1.22e-08 |  | 0.958 | 23.586 | 0.958 | 23.586 | -1.35 |  | 1.0337e-08 | 1.7867e-08 | 2.3152e-08 | 0.069 |  | 16.0778 | 19.5299 |  | 14.4493 |  | 1.4449e-08 | A7
| A6 | Yes | Yes | No | 148 | 8.86e-08 | 1.22e-08 | V. Noisy, offset circuit | 0.95844 | 1.8471 | 0.95844 | 1.8471 | 0.97 mA |  | 9.5772e-09 | 1.1513e-07 | 1.5158e-07 | 0.47 | weird offset | 136.2086 | 146.9956 | Confirmed very noisy assume dead | 114.987 | V Noisy | 1.1499e-07 | A6
| A9 | Yes | Yes | No | 149 | 2.3e-08 | 1.66e-08 |  | 0.95998 | 1.9109 | 0.95998 | 1.9109 |  |  | 1.0007e-08 | 1.8241e-08 | 2.5787e-08 | 0.049 |  | 16.3679 | 20.6958 |  | 15.6435 |  | 1.5644e-08 | A9
| C9 | Yes | Yes | No | 150 | 3.37e-08 | 2.5e-08 | Noisy, JFETs | 0.9585 | 1.9959 | 0.9585 | 1.9959 |  |  | 9.318e-09 | 3.3103e-08 | 4.6803e-08 | 0.048 |  | 36.0921 | 34.6323 | Comfirmed noisy | 19.7149 | Noisy | 1.9715e-08 | C9
| B8 | Yes | Yes | No | 151 | 2.06e-08 | 1.22e-08 |  | 0.9585 | 1.9296 | 0.9585 | 1.9296 |  |  | 9.4878e-09 | 1.7339e-08 | 2.514e-08 | 0.095 |  | 16.0247 | 18.9861 |  | 16.0168 |  | 1.6017e-08 | B8
| B7 | Yes | Yes | No | 152 | 2.4e-08 | 1.82e-08 |  | 0.95806 | 1.9271 | 0.95806 | 1.9271 |  |  | 1.0758e-08 | 1.8304e-08 | 2.5428e-08 | 0.072 |  | 16.8217 | 18.6787 |  | 14.8862 |  | 1.4886e-08 | B7
| C7 | Yes | Yes | No | 153 | 2.01e-08 | 1.19e-08 |  | 0.96019 | 1.6392 | 0.96019 | 1.6392 |  |  | 9.1296e-09 | 1.7437e-08 | 2.3293e-08 | 0.052 |  | 16.1201 | 19.5578 |  | 15.4594 |  | 1.5459e-08 | C7
| B5 | Yes | Yes | No | 154 | 2.18e-08 | 1.44e-08 |  | 0.95771 | 2.1771 | 0.95771 | 2.1771 |  |  | 1.1689e-08 | 1.9378e-08 | 2.8268e-08 | 0.059 |  | 18.0404 | 20.3428 |  | 17.3028 |  | 1.7303e-08 | B5
| B6 | Yes | Yes | No | 155 | 1.96e-08 | 1.2e-08 |  | 0.958 | 2.5123 | 0.958 | 2.5123 |  |  | 8.9973e-09 | 1.6649e-08 | 2.1273e-08 | 0.052 |  | 15.0341 | 18.3606 |  | 14.3807 |  | 1.4381e-08 | B6
| A5 | Yes | Yes | No | 156 | 2.19e-08 | 1.52e-08 |  | 0.95816 | 2.2034 | 0.95816 | 2.2034 |  |  | 1.2024e-08 | 1.8801e-08 | 2.5265e-08 | 0.03 |  | 17.3195 | 19.3644 |  | 15.5305 |  | 1.5531e-08 | A5
| T1 | Yes | Yes | No | 157 | 1.59e-08 | 1.03e-08 |  | 0.95957 | 1.3161 | 0.95957 | 1.3161 |  |  | 9.6411e-09 | 2.2714e-08 | 3.1867e-08 | 0.174 |  | 15.8077 | 16.7244 |  | 22.7246 |  | 2.2725e-08 | T1
| B4 | Yes | Yes | No | 158 | 2.02e-08 | 1.16e-08 |  | 0.006087 | 49.168 | 0.006087 | 49.168 |  |  | 9.5045e-09 | 2.7934e-06 | 2.1501e-08 | 0.073 | Check again? | 16.2542 | 19.8324 |  | 15.3873 |  | 1.5387e-08 | B4
| C4 | Yes | Yes | Yes | 159 | 1.96e-08 | 1.09e-08 |  | 0.95943 | 1.2822 | 0.95943 | 1.2822 |  |  | 1.0635e-08 | 1.7927e-08 | 2.2338e-08 | 0.059 |  | 16.6667 | 18.169 |  | 14.8843 |  | 1.4884e-08 | C4
| B3 | Yes | Yes | No | 160 | 2e-08 | 1.13e-08 |  | 0.9553 | 1.2418 | 0.9553 | 1.2418 |  |  | 9.0329e-09 | 1.7029e-08 | 2.1447e-08 | 0.081 |  | 15.3177 | 17.7437 |  | 14.9452 |  | 1.4945e-08 | B3
| C2 | Yes | Yes | No | 161 | 2.03e-08 | 1.18e-08 |  | 0.95733 | 1.4912 | 0.95733 | 1.4912 |  |  | 8.9933e-09 | 1.7481e-08 | 2.3485e-08 | 0.066 |  | 15.9452 | 18.3857 |  | 15.5977 |  | 1.5598e-08 | C2
| B2 | Yes | Yes | No | 162 | 1.88e-08 | 1.03e-08 |  | 0.95984 | 1.3206 | 0.95984 | 1.3206 |  |  | 1.2779e-08 | 1.7472e-08 | 2.4114e-08 | 0.053 |  | 16.4647 | 18.875 |  | 14.5284 |  | 1.4528e-08 | B2
| B1 | Yes | Yes | No | 163 | 2.29e-08 | 2.01e-08 |  | 0.95776 | 1.7795 | 0.95776 | 1.7795 |  |  | 9.6357e-09 | 1.7856e-08 | 2.3649e-08 | 0.055 |  | 16.5559 | 19.3054 |  | 14.7237 |  | 1.4724e-08 | B1
| A3 | Yes | Yes | No | 164 | 2.01e-08 | 1.17e-08 |  | 0.95781 | 2.097 | 0.95781 | 2.097 |  |  | 1.0371e-08 | 1.8833e-08 | 2.3753e-08 | 0.054 |  | 16.6135 | 19.3835 |  | 15.1909 |  | 1.5191e-08 | A3
| A4 | Yes | Yes | No | 165 | 1.89e-08 | 1.06e-08 |  | 0.95911 | 1.3181 | 0.95911 | 1.3181 |  |  | 9.8401e-09 | 1.7279e-08 | 2.3604e-08 | 0.054 |  | 16.2388 | 18.9284 |  | 15.2757 |  | 1.5276e-08 | A4
| A1 | Yes | Yes | No | 166 | 1.94e-08 | 1.26e-08 |  | 0.95742 | 1.3462 | 0.95742 | 1.3462 |  |  | 1.0075e-08 | 1.7616e-08 | 2.2818e-08 | 0.066 |  | 16.2549 | 18.9725 |  | 15.5672 |  | 1.5567e-08 | A1
| DK1 | Yes | Yes | No | 167 | 1.8e-08 | 1.12e-08 |  | 0.95728 | 0.5069 | 0.95728 | 0.5069 |  |  | 2.468e-05 | 8.8047e-09 | 1.1986e-08 | 0.43 | low gain | 11.8616 | 17.2603 |  | 14.8438 |  | 1.4844e-08 | DK1
| A2 | Yes | Yes | No | 168 | 7.47e-09 | 7.4e-09 | Cryoharness | 0.0065688 | 47.283 | 0.0065688 | 47.283 |  |  | 6.3276e-09 | 0.00076621 | 1.2507e-05 | 0.233 | low gain | 4767.0845 | 19.0737 | Alive again | 15.0998 |  | 1.51e-08 | A2
| E1 | Yes | Yes | No | 169 | 2.01e-08 | 1.08e-08 |  | 0.95185 | 1.711 | 0.95185 | 1.711 | 17/41 | -1.5853V | 2.5612e-08 | 1.8217e-08 | 2.3484e-08 | 0.068 |  | 16.3574 | 18.297 |  | 15.3446 |  | 1.5345e-08 | E1
| E2 | Yes | Yes | No | 170 | 1.89e-08 | 1.03e-08 |  | 0.94854 | 1.592 | 0.94854 | 1.592 | 2.5 |  | 9.3465e-09 | 1.8602e-08 | 8.6937e-08 | 0.17 |  | 18.2028 | 18.8439 |  | 15.1047 |  | 1.5105e-08 | E2
| E3 | Yes | Yes | No | 171 | 1.92e-08 | 1.21e-08 |  | 0.95438 | 2.0516 | 0.95438 | 2.0516 | -1.01 |  | 9.5413e-09 | 1.6451e-08 | 2.3351e-08 | 0.059 |  | 16.9175 | 17.6851 |  | 13.8372 |  | 1.3837e-08 | E3
| E4 | Yes | Yes | No | 172 | 1.94e-08 | 9.83e-09 |  | 0.95266 | 1.1587 | 0.95266 | 1.1587 | 1.14 mA |  | 8.8324e-09 | 1.7364e-08 | 2.3322e-08 | 0.074 |  | 15.4122 | 18.6352 |  | 14.4185 |  | 1.4418e-08 | E4
| D1 | Yes | Yes | No | 173 | 2.06e-08 | 1.3e-08 |  | 0.95542 | 1.3972 | 0.95542 | 1.3972 |  |  | 1.1052e-08 | 1.9015e-08 | 2.6242e-08 | 0.063 |  | 17.7673 | 20.5747 |  | 16.3806 |  | 1.6381e-08 | D1
| D2 | Yes | Yes | No | 174 | 1.93e-08 | 1.11e-08 |  | 0.95515 | 1.6775 | 0.95515 | 1.6775 |  |  | 8.9881e-09 | 1.6541e-08 | 2.3815e-08 | 0.065 |  | 15.3625 | 18.1078 |  | 15.5913 |  | 1.5591e-08 | D2
| D3 | Yes | Yes | No | 175 | 1.9e-08 | 1.02e-08 |  | 0.95576 | 1.2493 | 0.95576 | 1.2493 |  |  | 9.8169e-09 | 1.7757e-08 | 2.4202e-08 | 0.055 |  | 16.0506 | 18.1144 |  | 15.031 |  | 1.5031e-08 | D3
| D4 | Yes | Yes | Yes | 176 | 1.91e-08 | 1.33e-08 |  | 0.95404 | 1.212 | 0.95404 | 1.212 |  |  | 9.3519e-09 | 1.6736e-08 | 2.2025e-08 | 0.063 |  | 15.5832 | 17.779 |  | 14.2127 |  | 1.4213e-08 | D4
| C1 | Yes | Yes | No | 177 | 1.85e-08 | 1.04e-08 |  | 0.95206 | 1.2599 | 0.95206 | 1.2599 |  |  | 9.7355e-09 | 1.6378e-08 | 2.1383e-08 | 0.084 |  | 14.9671 | 18.618 |  | 14.5312 |  | 1.4531e-08 | C1
| C3 | Yes | Yes | No | 178 | 2.05e-08 | 1.25e-08 |  | 0.95428 | 1.6127 | 0.95428 | 1.6127 |  |  | 1.0111e-08 | 1.9731e-08 | 2.6139e-08 | 0.053 |  | 17.6011 | 19.2416 |  | 15.0078 |  | 1.5008e-08 | C3
| C5 | Yes | Yes | No | 179 | 1.85e-08 | 1.05e-08 |  | 0.95457 | 1.3103 | 0.95457 | 1.3103 |  |  | 9.3246e-09 | 1.6257e-08 | 2.2076e-08 | 0.068 |  | 15.416 | 16.7837 |  | 13.4644 |  | 1.3464e-08 | C5
| T2 | Yes | Yes | No | 180 | 1.57e-08 | 1.07e-08 |  | 0.95246 | 1.2141 | 0.95246 | 1.2141 |  |  | 9.413e-09 | 0 | 3.1589e-08 | 0.178 | Bad data - pixel looks o.k. | 0 | 17.3871 | Alive again | 18.9429 |  | 1.8943e-08 | T2
| E5 | Yes | Yes | No | 181 | 1.86e-08 | 1.19e-08 |  | 0.95137 | 1.2513 | 0.95137 | 1.2513 |  |  | 1.2831e-08 | 1.6768e-08 | 2.2386e-08 | 0.087 |  | 15.7031 | 19.3927 |  | 15.5717 |  | 1.5572e-08 | E5
| C6 | Yes | Yes | No | 182 | 1.73e-08 | 1.02e-08 |  | 0.9544 | 1.4865 | 0.9544 | 1.4865 |  |  | 9.6178e-09 | 1.6191e-08 | 2.2855e-08 | 0.057 |  | 15.0461 | 16.7036 |  | 13.6844 |  | 1.3684e-08 | C6
| C8 | Yes | Yes | No | 183 | 1.85e-08 | 1.09e-08 |  | 0.95455 | 1.2467 | 0.95455 | 1.2467 |  |  | 9.3827e-09 | 1.6965e-08 | 2.1313e-08 | 0.058 |  | 15.5853 | 18.5199 |  | 14.5621 |  | 1.4562e-08 | C8
| D5 | Yes | Yes | No | 184 | 1.9e-08 | 1.01e-08 |  | 0.95452 | 1.4251 | 0.95452 | 1.4251 |  |  | 8.7465e-09 | 1.6839e-08 | 2.2512e-08 | 0.074 |  | 15.5668 | 17.8096 |  | 14.6757 |  | 1.4676e-08 | D5
| D6 | Yes | Yes | No | 185 | 1.88e-08 | 1.03e-08 |  | 0.95581 | 1.2649 | 0.95581 | 1.2649 |  |  | 9.1127e-09 | 1.714e-08 | 2.3272e-08 | 0.061 |  | 15.4324 | 17.5551 |  | 14.1099 |  | 1.411e-08 | D6
| D7 | Yes | Yes | No | 186 | 1.94e-08 | 1.07e-08 |  | 0.95483 | 1.3481 | 0.95483 | 1.3481 |  |  | 1.0076e-08 | 1.684e-08 | 2.3655e-08 | 0.053 |  | 15.4533 | 18.1095 |  | 15.1006 |  | 1.5101e-08 | D7
| D8 | Yes | Yes | No | 187 | 1.94e-08 | 1.26e-08 |  | 0.9526 | 1.4137 | 0.9526 | 1.4137 |  |  | 1.1125e-08 | 1.851e-08 | 2.5791e-08 | 0.051 |  | 17.6587 | 18.8735 |  | 14.3804 |  | 1.438e-08 | D8
| E7 | Yes | Yes | No | 188 | 1.94e-08 | 1.07e-08 |  | 0.95208 | 1.4136 | 0.95208 | 1.4136 |  |  | 9.2318e-09 | 1.7826e-08 | 2.343e-08 | 0.062 |  | 16.1985 | 19.6567 |  | 15.4676 |  | 1.5468e-08 | E7
| E6 | Yes | Yes | No | 189 | 2.08e-08 | 1.31e-08 |  | 0.95521 | 1.2901 | 0.95521 | 1.2901 |  |  | 1.0554e-08 | 1.9705e-08 | 2.6817e-08 | 0.053 |  | 18.1027 | 18.2031 |  | 15.2425 |  | 1.5242e-08 | E6
| E8 | Yes | Yes | No | 190 | 1.86e-08 | 1.02e-08 |  | 0.95401 | 1.3825 | 0.95401 | 1.3825 |  |  | 8.9081e-09 | 1.6756e-08 | 2.3764e-08 | 0.057 |  | 15.9222 | 20.0183 |  | 14.3719 |  | 1.4372e-08 | E8
| DK2 | Yes | Yes | No | 191 | 1.69e-08 | 9.6e-09 |  | 0.95281 | 1.231 | 0.95281 | 1.231 |  |  | 9.4146e-09 | 1.5497e-08 | 2.2304e-08 | 0.081 |  | 15.0326 | 17.3849 |  | 13.6383 |  | 1.3638e-08 | DK2
| E9 | Yes | Yes | No | 192 | 1.84e-08 | 9.71e-09 |  | 0.95357 | 1.432 | 0.95357 | 1.432 |  |  | 9.0707e-09 | 1.6711e-08 | 2.2723e-08 | 0.053 |  | 15.9239 | 19.2802 |  | 15.282 |  | 1.5282e-08 | E9

## PMW
PMW | BoDAC |  |  |  | PFM3 |  |  |  | PFM4
| Pixel Name | BDA Operability | BoDAC channel Operability | Noisy Bodac? | Channel # | PFM3 Noise | PFM3 WE Noise | JJB Comment | JFET Gain Test | Warm Resistance test | JFET ID | VSS at 4K | JFET+LIA Noise (nV/rtHz) | Noise (nV/rtHz) (1-3 Hz) | Noise (nV/rtHz) (0.06 - 0.3 Hz) (20061127) | 1/f Knee (PTC Control) (20061127) | 0.3 K load curve | Noise taken with correct VSS | Failure mode following cryoharness probe | Noise IST | IST Comment | Noise Flight | Operation flight
| A13 | Yes | Yes | No | 193 | 1.62e-08 | 1.17e-08 |  | 0.95931 | 3.117 | 15/35 | -1.5853V | 1.0176e-08 | 1.4042e-08 | 1.9334e-08 | 0.062 |  | 13.675 |  | 15.3366 | Very Slow | 11.8941 | Slow | 1.1894e-08 | A13
| T1 | Yes | Yes | No | 194 | 1.7e-08 | 2.22e-08 |  | 0.95865 | 1.4997 | 2.5 |  | 8.7726e-09 | 2.2491e-08 | 2.9576e-08 | 0.182 |  | 15.3858 |  | 15.8284 |  | 17.3099 |  | 1.731e-08 | T1
| B12 | Yes | Yes | Yes | 195 | 1.72e-08 | 1.04e-08 |  | 0.95972 | 1.3798 | -1.41 |  | 9.1857e-09 | 1.6152e-08 | 2.1103e-08 | 0.07 |  | 14.9085 |  | 17.5796 |  | 13.6722 |  | 1.3672e-08 | B12
| C13 | Yes | Yes | No | 196 | 1.66e-08 | 1.02e-08 |  | 0.95915 | 1.4395 | 1.13 mA |  | 8.673e-09 | 1.5723e-08 | 2.3136e-08 | 0.063 |  | 15.5574 |  | 17.5799 |  | 13.5819 |  | 1.3582e-08 | C13
| A12 | Yes | Yes | No | 197 | 1.95e-08 | 1.37e-08 |  | 0.9598 | 1.4426 |  |  | 9.0114e-09 | 1.5416e-08 | 2.1116e-08 | 0.073 |  | 14.8513 |  | 17.6199 |  | 14.1018 |  | 1.4102e-08 | A12
| D12 | Yes | Yes | No | 198 | 1.89e-08 | 1.28e-08 |  | 0.95903 | 2.216 |  |  | 9.0379e-09 | 1.575e-08 | 2.0831e-08 | 0.064 |  | 14.7794 |  | 17.6612 |  | 14.3664 |  | 1.4366e-08 | D12
| C12 | Yes | Yes | No | 199 | 1.78e-08 | 1.06e-08 |  | 0.95906 | 1.3345 |  |  | 9.4434e-09 | 1.5764e-08 | 2.2495e-08 | 0.082 |  | 14.6047 |  | 18.2685 |  | 15.1548 |  | 1.5155e-08 | C12
| B11 | Yes | Yes | Yes | 200 | 2.06e-08 | 1.12e-08 |  | 0.96064 | 1.8807 |  |  | 9.4463e-09 | 1.8629e-08 | 3.153e-08 | 0.129 |  | 18.6643 |  | 26.6418 | 1/f noisy | 15.5894 | O.k.? | 1.5589e-08 | B11
| A11 | Yes | Yes | No | 201 | 2.5e-08 | 2.18e-08 | M. Noisy, JFETs | 0.96011 | 1.5337 |  |  | 1.1418e-08 | 1.689e-08 | 2.2952e-08 | 0.069 |  | 15.686 |  | 18.2672 |  | 14.3361 |  | 1.4336e-08 | A11
| E13 | Yes | Yes | No | 202 | 1.7e-08 | 1.02e-08 |  | 0.95866 | 1.3952 |  |  | 9.6392e-09 | 1.6038e-08 | 2.1937e-08 | 0.057 |  | 15.0821 |  | 17.5116 |  | 14.1312 |  | 1.4131e-08 | E13
| D11 | Yes | Yes | No | 203 | 1.64e-08 | 9.9e-09 |  | 0.93595 | 1.5588 |  |  | 8.3178e-09 | 1.5273e-08 | 2.0662e-08 | 0.062 |  | 14.7174 |  | 16.6783 |  | 14.2185 |  | 1.4218e-08 | D11
| C11 | Yes | Yes | No | 204 | 1.74e-08 | 1.06e-08 |  | 0.95901 | 1.3278 |  |  | 1.0037e-08 | 1.6023e-08 | 2.1154e-08 | 0.067 |  | 14.7093 |  | 17.4866 |  | 14.2729 |  | 1.4273e-08 | C11
| B10 | Yes | Yes | No | 205 | 1.82e-08 | 1.08e-08 |  | 0.95981 | 1.4514 |  |  | 8.4962e-09 | 1.6138e-08 | 2.2851e-08 | 0.069 |  | 15.0694 |  | 17.8189 |  | 13.9512 |  | 1.3951e-08 | B10
| A10 | Yes | Yes | No | 206 | 1.78e-08 | 1.16e-08 |  | 0.96017 | 1.9872 |  |  | 9.0124e-09 | 1.6666e-08 | 2.1985e-08 | 0.06 |  | 14.5619 |  | 18.2225 |  | 13.9996 |  | 1.4e-08 | A10
| D10 | Yes | Yes | No | 207 | 1.99e-08 | 1.38e-08 |  | 0.95836 | 1.4221 |  |  | 9.5148e-09 | 1.6092e-08 | 2.1473e-08 | 0.058 |  | 14.7853 |  | 16.6671 |  | 14.0905 |  | 1.409e-08 | D10
| B9 | Yes | Yes | No | 208 | 1.7e-08 | 1.05e-08 |  | 0.95847 | 2.1382 |  |  | 9.0263e-09 | 1.5541e-08 | 2.0341e-08 | 0.08 |  | 14.3531 |  | 17.0989 |  | 13.5869 |  | 1.3587e-08 | B9
| C10 | Yes | Yes | No | 209 | 1.76e-08 | 1.01e-08 |  | 0.92326 | 1.7671 |  |  | 8.7341e-09 | 1.6775e-08 | 2.1336e-08 | 0.069 |  | 15.2013 |  | 16.5605 |  | 13.9313 |  | 1.3931e-08 | C10
| C9 | Yes | Yes | Yes | 210 | 2.03e-08 | 1.62e-08 |  | 0.96047 | 1.4468 |  |  | 1.1108e-08 | 1.569e-08 | 2.0366e-08 | 0.06 | weird offset | 14.8577 |  | 16.3943 | Still offset issue in loadcurve | 13.5107 | Looks o.k. | 1.3511e-08 | C9
| A9 | Yes | Yes | No | 211 | 1.85e-08 | 1.13e-08 |  | 0.95874 | 1.6222 |  |  | 8.8793e-09 | 1.6259e-08 | 2.1943e-08 | 0.066 |  | 14.8713 |  | 18.5163 |  | 14.0717 |  | 1.4072e-08 | A9
| B8 | Yes | Yes | No | 212 | 1.75e-08 | 1.03e-08 |  | 0.96029 | 1.5887 |  |  | 9.7144e-09 | 1.5516e-08 | 2.1407e-08 | 0.095 |  | 15.0852 |  | 17.6049 |  | 14.2313 |  | 1.4231e-08 | B8
| A8 | Yes | Yes | No | 213 | 2.57e-08 | 1.43e-08 | M. Noisy, unknown | 0.96103 | 1.546 |  |  | 1.3617e-08 | 1.9511e-08 | 2.5594e-08 | 0.05 |  | 18.3747 |  | 18.2412 |  | 14.483 |  | 1.4483e-08 | A8
| D8 | Yes | Yes | No | 214 | 1.83e-08 | 1.16e-08 |  | 0.95898 | 1.5025 |  |  | 9.2393e-09 | 1.6346e-08 | 2.2215e-08 | 0.052 |  | 16.2258 |  | 18.5519 |  | 14.8512 |  | 1.4851e-08 | D8
| C8 | Yes | Yes | No | 215 | 0 | 0 | N/C? | 0.96012 | 412.36 |  |  | 0 | 0 | 0 | ---- | zero | 0 | Broken S- | 16.6714 | Alive Again | 14.0408 |  | 1.4041e-08 | C8
| B7 | Yes | Yes | No | 216 | 1.78e-08 | 1.28e-08 |  | 0.95539 | 1.5766 |  |  | 8.9986e-09 | 1.4988e-08 | 2.2674e-08 | 0.067 |  | 14.2111 |  | 17.0859 |  | 13.5001 |  | 1.35e-08 | B7
| R1 | Yes | Yes | No | 217 | 1.38e-08 | 1.87e-20 |  | 0.96338 | 4.0366 | 12/22 | -1.88V | 2.1791e-08 | 1.4684e-08 | 1.7894e-08 | 0.37 |  | 15.3379 |  | 15.5379 |  | 11.8293 |  | 1.1829e-08 | R1
| G1 | Yes | Yes | No | 218 | 1.85e-08 | 1.29e-08 |  | 0.0082218 | 1.4749 | 2.5 |  | 8.9031e-09 | 1.9375e-06 | 2.2054e-08 | 0.064 | now fine | 15.1167 |  | 18.4671 |  | 13.944 |  | 1.3944e-08 | G1
| T2 | Yes | Yes | No | 219 | 0 | 1.14e-08 | Cryoharness | -2.125e-09 | 393.34 | -1.56 |  | 0 | 0 | 0 | ---- | zero | 0 | Broken S- | 0 | Really Dead | 2.9033e-10 |  | 2.9033e-19 | T2
| E1 | Yes | Yes | Yes | 220 | 1.77e-08 | 1.17e-08 |  | -0.00022597 | 47.466 | 1.13 mA |  | 8.9566e-09 | -6.8467e-05 | 2.2057e-08 | 0.097 | now fine | 14.3296 |  | 17.7103 |  | 13.5842 |  | 1.3584e-08 | E1
| D1 | Yes | Yes | No | 221 | 1.81e-08 | 1.08e-08 |  | 0.95275 | 2.4144 |  |  | 1.0206e-08 | 1.6521e-08 | 2.2188e-08 | 0.075 |  | 15.2188 |  | 21.6247 | Noisy - sometimes very noisy | 1.9717e-09 | A puzzle? | 1.9717e-18 | D1
| F1 | Yes | Yes | Yes | 222 | 2.15e-08 | 1.73e-08 |  | 0.95765 | 1.3142 |  |  | 1.2229e-08 | 1.8762e-08 | 2.6597e-08 | 0.077 |  | 17.5331 |  | 18.715 |  | 14.3549 |  | 1.4355e-08 | F1
| E2 | Yes | Yes | No | 223 | 1.83e-08 | 1.09e-08 |  | 0.96077 | 1.3559 |  |  | 9.3877e-09 | 1.5806e-08 | 2.2425e-08 | 0.065 |  | 14.5832 |  | 17.8475 |  | 13.9994 |  | 1.3999e-08 | E2
| G2 | Yes | Yes | Yes | 224 | 2.02e-08 | 1.55e-08 |  | 0.96242 | 1.5182 |  |  | 8.7192e-09 | 1.6241e-08 | 2.2465e-08 | 0.058 |  | 15.1482 |  | 18.1211 |  | 13.9149 |  | 1.3915e-08 | G2
| F2 | Yes | Yes | No | 225 | 2.81e-08 | 1.45e-08 | M. Noisy, unknown | 0.9622 | 1.2236 |  |  | 1.0512e-08 | 1.6649e-08 | 2.3152e-08 | 0.078 |  | 15.1516 |  | 18.9046 |  | 14.4666 |  | 1.4467e-08 | F2
| G3 | Yes | Yes | No | 226 | 1.8e-08 | 1.07e-08 |  | 0.96289 | 1.2181 |  |  | 9.077e-09 | 1.567e-08 | 2.1565e-08 | 0.086 |  | 14.2524 |  | 18.4014 |  | 14.0431 |  | 1.4043e-08 | G3
| E3 | Yes | Yes | No | 227 | 1.93e-08 | 1.35e-08 |  | 0.96293 | 1.4049 |  |  | 9.2283e-09 | 1.5578e-08 | 2.1259e-08 | 0.075 |  | 14.2657 |  | 17.8716 |  | 13.589 |  | 1.3589e-08 | E3
| D3 | Yes | Yes | No | 228 | 1.85e-08 | 1.41e-08 |  | 0.96322 | 1.8729 |  |  | 8.6474e-09 | 1.6565e-08 | 2.3462e-08 | 0.08 |  | 15.0265 |  | 18.2436 |  | 15.4965 |  | 1.5496e-08 | D3
| F3 | Yes | Yes | No | 229 | 1.77e-08 | 1.05e-08 |  | 0.96283 | 1.3924 |  |  | 9.2626e-09 | 1.6344e-08 | 2.1233e-08 | 0.07 |  | 14.9825 |  | 18.7089 |  | 14.0469 |  | 1.4047e-08 | F3
| G4 | Yes | Yes | No | 230 | 2e-08 | 1.42e-08 |  | 0.0064476 | 48.209 |  |  | 8.7626e-09 | 2.5321e-06 | 2.4716e-08 | 0.069 | now fine | 15.5973 |  | 19.2449 |  | 16.1025 |  | 1.6103e-08 | G4
| E4 | Yes | Yes | No | 231 | 4.22e-08 | 4.01e-08 | V. Noisy, JFETs | 0.96227 | 1.272 |  |  | 1.1032e-08 | 2.3663e-08 | 3.214e-08 | 0.055 |  | 22.1458 |  | 18.8585 |  | 14.5694 |  | 1.4569e-08 | E4
| F4 | Yes | Yes | No | 232 | 1.75e-08 | 1.08e-08 |  | 0.9593 | 1.6027 |  |  | 8.9929e-09 | 1.5519e-08 | 2.1191e-08 | 0.069 |  | 14.0244 |  | 17.4071 |  | 13.3608 |  | 1.3361e-08 | F4
| E5 | Yes | Yes | Yes | 233 | 5.18e-08 | 5.33e-08 | V. Noisy, JFETs | 0.95782 | 1.2875 |  |  | 9.7678e-09 | 2.2862e-08 | 3.0954e-08 | 0.067 |  | 21.8062 |  | 16.3427 |  | 12.9595 |  | 1.2959e-08 | E5
| D5 | Yes | Yes | No | 234 | 2.05e-08 | 1.58e-08 |  | 0.93548 | 1.1695 |  |  | 9.6423e-09 | 1.7979e-08 | 2.3199e-08 | 0.068 |  | 16.0107 |  | 17.2038 |  | 14.069 |  | 1.4069e-08 | D5
| F5 | Yes | Yes | No | 235 | 3.76e-08 | 3.57e-08 | M. Noisy, JFETs | 0.96368 | 1.3294 |  |  | 1.0693e-08 | 2.4981e-08 | 3.2335e-08 | 0.042 |  | 23.2767 |  | 17.2439 |  | 13.7982 |  | 1.3798e-08 | F5
| G5 | Yes | Yes | No | 236 | 1.76e-08 | 1.05e-08 |  | 0.96345 | 1.4115 |  |  | 1.0522e-08 | 1.6191e-08 | 2.2253e-08 | 0.058 |  | 14.6592 |  | 18.8312 |  | 13.9722 |  | 1.3972e-08 | G5
| E6 | Yes | Yes | No | 237 | 1.78e-08 | 1.05e-08 |  | 0.96354 | 1.3112 |  |  | 8.349e-09 | 1.5572e-08 | 2.2019e-08 | 0.069 |  | 14.3287 |  | 16.0969 |  | 13.1929 |  | 1.3193e-08 | E6
| G6 | Yes | Yes | No | 238 | 4.92e-08 | 4.81e-08 | V. Noisy, JFETs | 0.96379 | 1.4495 |  |  | 1.1066e-08 | 3.1379e-08 | 4.534e-08 | 0.049 |  | 30.0045 |  | 17.2882 | Noise o.k. now | 13.69 |  | 1.369e-08 | G6
| F6 | Yes | Yes | No | 239 | 2.98e-08 | 2.76e-08 | M. Noisy, JFETs | 0.96339 | 1.468 |  |  | 9.2199e-09 | 1.8313e-08 | 2.562e-08 | 0.047 |  | 17.8143 |  | 16.6313 |  | 13.1031 |  | 1.3103e-08 | F6
| G7 | Yes | Yes | No | 240 | 2.65e-08 | 2.3e-08 | M. Noisy, JFETs | 0.96317 | 1.3007 |  |  | 9.193e-09 | 1.6871e-08 | 2.2584e-08 | 0.059 |  | 15.4308 |  | 17.7392 |  | 13.7419 |  | 1.3742e-08 | G7
| F10 | Yes | Yes | No | 241 | 1.81e-08 | 1.1e-08 |  | 0.96288 | 5.5969 | 12/23 | -1.3726V | 9.0111e-09 | 1.7468e-08 | 2.2731e-08 | 0.065 |  | 17.2945 |  | 18.4097 |  | 13.7506 |  | 1.3751e-08 | F10
| E11 | Yes | Partial | Yes | 242 | 1.75e-08 | 9.96e-09 |  | 0.9409 | 1.4929 | 2.62 |  | 8.6819e-09 | 1.7623e-08 | 2.3248e-08 | 0.064 |  | 17.1812 |  | 16.6245 |  | 14.1447 |  | 1.4145e-08 | E11
| G11 | Yes | Yes | Yes | 243 | 1.8e-08 | 1.06e-08 |  | 0.94826 | 1.1533 | -1.52 |  | 6.341e-09 | 7.5033e-09 | 9.532e-09 | 0.143 | low gain | 8.3708 | Bad contact S+? | 18.9279 | Alive Again | 14.1082 |  | 1.4108e-08 | G11
| F11 | Yes | Yes | No | 244 | 1.77e-08 | 1.04e-08 |  | 0.9639 | 1.3939 | 1.06 mA |  | 8.6669e-09 | 1.6229e-08 | 2.3022e-08 | 0.067 |  | 15.3557 |  | 16.5581 |  | 13.4583 |  | 1.3458e-08 | F11
| E12 | Yes | Yes | No | 245 | 1.82e-08 | 1.08e-08 |  | 0.96358 | 2.3549 |  |  | 9.0621e-09 | 1.8208e-08 | 2.266e-08 | 0.057 |  | 18.5772 |  | 18.2909 |  | 14.7081 |  | 1.4708e-08 | E12
| G12 | Yes | Yes | No | 246 | 1.69e-08 | 1.06e-08 |  | 0.90871 | 2.4644 |  |  | 9.1588e-09 | 1.6978e-08 | 2.1227e-08 | 0.088 |  | 14.9216 |  | 18.2116 |  | 14.7504 |  | 1.475e-08 | G12
| F12 | Yes | Yes | No | 247 | 2.12e-08 | 1.06e-08 |  | 0.95815 | 1.38 |  |  | 8.395e-09 | 1.8734e-08 | 2.4953e-08 | 0.068 |  | 19.8726 |  | 18.9613 |  | 14.5407 |  | 1.4541e-08 | F12
| G13 | Yes | Yes | Yes | 248 | 1.81e-08 | 1.11e-08 |  | 0.96157 | 1.6794 |  |  | 1.1414e-08 | 1.7209e-08 | 2.4411e-08 | 0.083 |  | 16.4276 |  | 18.8147 |  | 14.5855 |  | 1.4586e-08 | G13
| DK2 | Yes | Yes | No | 249 | 1.71e-08 | 1.13e-08 |  | 0.95478 | 1.4749 |  |  | 1.1086e-08 | 1.5399e-08 | 2.1647e-08 | 0.078 |  | 14.4562 |  | 16.0724 |  | 13.0591 |  | 1.3059e-08 | DK2
| E7 | Yes | Yes | No | 250 | 1.72e-08 | 1e-08 |  | 0.95386 | 1.7237 |  |  | 9.559e-09 | 1.6377e-08 | 2.3265e-08 | 0.066 |  | 15.7554 |  | 16.4642 |  | 13.3138 |  | 1.3314e-08 | E7
| D7 | Yes | Yes | No | 251 | 1.78e-08 | 1.17e-08 |  | 0.96287 | 1.6566 |  |  | 1.0659e-08 | 1.704e-08 | 2.2665e-08 | 0.095 |  | 15.7514 |  | 18.0371 |  | 14.0109 |  | 1.4011e-08 | D7
| F7 | Yes | Yes | Yes | 252 | 0 | 1.51e-08 | N/C? | -0.011801 | 1.418 |  |  | 1.3066e-08 | 0 | 0 | ---- | wiring swap | 0 |  | 18.3812 | Alive again | 15.6805 |  | 1.568e-08 | F7
| E8 | Yes | Yes | No | 253 | 1.87e-08 | 1.23e-08 |  | 0.88757 | 0.6136 |  |  | 9.5064e-09 | 2.0575e-08 | 2.8448e-08 | 0.076 |  | 19.7105 |  | 37.652 | Noisy | 32.8247 | Noisy | 3.2825e-08 | E8
| G8 | Yes | Yes | No | 254 | 1.81e-08 | 1.05e-08 |  | 0.9355 | 1.4092 |  |  | 9.421e-09 | 1.7508e-08 | 2.3733e-08 | 0.066 |  | 17.263 |  | 18.312 |  | 14.4491 |  | 1.4449e-08 | G8
| F8 | Yes | Yes | No | 255 | 2.3e-08 | 1.51e-08 |  | 0.96215 | 1.5754 |  |  | 1.0553e-08 | 2.1634e-08 | 3.2519e-08 | 0.058 |  | 20.17 |  | 17.3154 |  | 13.5418 |  | 1.3542e-08 | F8
| E9 | Yes | Yes | No | 256 | 2.2e-08 | 1.76e-08 |  | 0.96367 | 1.6621 |  |  | 9.3269e-09 | 1.7643e-08 | 2.2617e-08 | 0.106 |  | 16.378 |  | 17.6141 |  | 14.7528 |  | 1.4753e-08 | E9
| G9 | Yes | Yes | No | 257 | 7.21e-09 | 7.37e-09 | S/C? | 0.0049754 | 46.846 |  |  | 5.9612e-09 | 1.4458e-06 | 9.7068e-09 | <0.010 | low gain | 7.9295 | Short to ground S+ | 17.5715 | Alive Again | 13.5938 |  | 1.3594e-08 | G9
| D9 | Yes | Yes | No | 258 | 2.3e-08 | 1.55e-08 |  | 0.93039 | 1.6186 |  |  | 1.0627e-08 | 2.4204e-08 | 3.1946e-08 | 0.067 |  | 22.5271 |  | 18.4658 |  | 14.8604 |  | 1.486e-08 | D9
| F9 | Yes | Yes | No | 259 | 1.88e-08 | 1.3e-08 |  | 0.91938 | 1.2865 |  |  | 9.8918e-09 | 1.7272e-08 | 2.4062e-08 | 0.073 |  | 16.4146 |  | 16.9365 |  | 13.9738 |  | 1.3974e-08 | F9
| E10 | Yes | Yes | Yes | 260 | 1.96e-08 | 1.27e-08 |  | 0.96085 | 1.7579 |  |  | 1.0198e-08 | 2.0215e-08 | 2.8649e-08 | 0.053 |  | 18.9666 |  | 17.4365 |  | 13.692 |  | 1.3692e-08 | E10
| G10 | Yes | Yes | No | 261 | 1.67e-08 | 1.05e-08 |  | 0.9159 | 0.7909 |  |  | 9.2874e-09 | 1.7219e-08 | 2.5514e-08 | 0.076 |  | 15.5414 |  | 17.9051 |  | 14.2883 |  | 1.4288e-08 | G10
| C4 | Yes | Yes | No | 262 | 1.86e-08 | 1.2e-08 |  | 0.95968 | 1.5674 | 15/34 | -1.5853V | 1.1072e-08 | 1.6174e-08 | 2.2398e-08 | 0.071 |  | 14.913 |  | 16.8343 |  | 13.0212 |  | 1.3021e-08 | C4
| B3 | Yes | Yes | No | 263 | 1.73e-08 | 1.07e-08 |  | 0.96016 | 1.5788 | 2.48 |  | 1.0322e-08 | 1.5731e-08 | 2.0242e-08 | 0.073 |  | 240.334 | Wrong noise | 17.8053 | O.k. | 14.4187 |  | 1.4419e-08 | B3
| C3 | Yes | Yes | Yes | 264 | 1.78e-08 | 1.07e-08 |  | 0.95992 | 1.8439 | -1.39 |  | 8.8197e-09 | 1.6043e-08 | 2.1969e-08 | 0.07 |  | 14.8236 |  | 17.5154 |  | 14.1902 |  | 1.419e-08 | C3
| B2 | Yes | Yes | No | 265 | 1.84e-08 | 1.1e-08 |  | 0.96063 | 1.8331 | 1.22 mA |  | 1.089e-08 | 1.7192e-08 | 2.2939e-08 | 0.074 |  | 16.0556 |  | 19.6467 |  | 15.907 |  | 1.5907e-08 | B2
| D2 | Yes | Yes | No | 266 | 1.8e-08 | 1.67e-08 |  | 0.95963 | 1.8149 |  |  | 1.2216e-08 | 1.8386e-08 | 2.4384e-08 | 0.079 |  | 17.4474 |  | 19.8122 |  | 15.9309 |  | 1.5931e-08 | D2
| A3 | Yes | Yes | No | 267 | 1.75e-08 | 1.05e-08 |  | 0.96073 | 1.6203 |  |  | 8.8169e-09 | 1.5701e-08 | 2.0826e-08 | 0.058 |  | 14.6918 |  | 17.0153 |  | 13.4344 |  | 1.3434e-08 | A3
| A2 | Yes | Yes | No | 268 | 1.88e-08 | 1.32e-08 |  | 0.96155 | 1.5419 |  |  | 9.9706e-09 | 1.6028e-08 | 2.0281e-08 | 0.079 |  | 14.6621 |  | 18.5301 |  | 14.208 |  | 1.4208e-08 | A2
| C2 | Yes | Yes | No | 269 | 1.72e-08 | 1.05e-08 |  | 0.95986 | 1.6079 |  |  | 8.8532e-09 | 1.6129e-08 | 2.1328e-08 | 0.081 |  | 14.3516 |  | 18.2648 |  | 14.2997 |  | 1.43e-08 | C2
| B1 | Yes | Yes | No | 270 | 2.67e-08 | 2.32e-08 | M. Noisy, JFETs | 0.95943 | 1.4467 |  |  | 1.2489e-08 | 1.8183e-08 | 2.4774e-08 | 0.052 |  | 17.2647 |  | 18.4413 |  | 14.6206 |  | 1.4621e-08 | B1
| A1 | Yes | Yes | No | 271 | 1.76e-08 | 1.15e-08 |  | 0.95876 | 1.4323 |  |  | 1.0416e-08 | 1.6738e-08 | 2.2551e-08 | 0.058 |  | 15.1473 |  | 18.4841 |  | 14.8558 |  | 1.4856e-08 | A1
| DK1 | Yes | Yes | No | 272 | 1.86e-08 | 1.54e-08 |  | 0.96202 | 1.6405 |  |  | 0 | 0 | 0 | ---- | zero | 0 | Broken S- | 17.7759 | Alive Again | 14.4353 |  | 1.4435e-08 | DK1
| C1 | Yes | Yes | No | 273 | 2.07e-08 | 1.51e-08 |  | 0.95652 | 2.0602 |  |  | 1.0108e-08 | 1.7475e-08 | 2.3754e-08 | 0.07 |  | 16.5839 |  | 19.6251 |  | 14.7259 |  | 1.4726e-08 | C1
| A7 | Yes | Yes | No | 274 | 1.91e-08 | 1.25e-08 |  | 0.95637 | 1.6909 |  |  | 1.0116e-08 | 1.7909e-08 | 2.3454e-08 | 0.077 |  | 17.3406 |  | 17.9263 |  | 13.7952 |  | 1.3795e-08 | A7
| A6 | Yes | Yes | No | 275 | 2.42e-08 | 1.81e-08 |  | 0.96054 | 1.6071 |  |  | 1.159e-08 | 1.8105e-08 | 2.4191e-08 | 0.052 |  | 17.0734 |  | 18.4283 |  | 13.6399 |  | 1.364e-08 | A6
| B6 | Yes | Yes | No | 276 | 0 | 2.62e-08 | N/C? | 0.0074658 | 50.876 |  |  | 6.3066e-09 | 9.7192e-07 | 1.0408e-08 | 0.174 | low gain | 8.5039 | Broken S+ | 16.9016 | Alive Again | 13.4833 |  | 1.3483e-08 | B6
| C7 | Yes | Yes | No | 277 | 1.82e-08 | 1.17e-08 |  | 0.96125 | 2.2858 |  |  | 1.0229e-08 | 1.6514e-08 | 2.1796e-08 | 0.063 |  | 15.0691 |  | 17.4559 |  | 13.9361 |  | 1.3936e-08 | C7
| A5 | Yes | Yes | Yes | 278 | 1.91e-08 | 1.34e-08 |  | 0.95657 | 2.6674 |  |  | 8.8413e-09 | 1.6066e-08 | 2.1245e-08 | 0.082 |  | 14.6599 |  | 16.9925 |  | 13.8928 |  | 1.3893e-08 | A5
| B5 | Yes | Yes | Yes | 279 | 1.97e-08 | 1.37e-08 |  | 0.96135 | 1.6957 |  |  | 1.0537e-08 | 1.6822e-08 | 2.4272e-08 | 0.066 |  | 15.3822 |  | 17.0189 |  | 13.9542 |  | 1.3954e-08 | B5
| C6 | Yes | Yes | No | 280 | 1.81e-08 | 1.08e-08 |  | 0.95791 | 2.629 |  |  | 9.1083e-09 | 1.6001e-08 | 2.1949e-08 | 0.064 |  | 14.9011 |  | 17.9187 |  | 14.7414 |  | 1.4741e-08 | C6
| D6 | Yes | Yes | No | 281 | 1.93e-08 | 1.41e-08 |  | 0.96253 | 1.7391 |  |  | 1.5838e-08 | 2.0714e-08 | 2.7848e-08 | 0.045 |  | 18.5308 |  | 25.3917 | Noisy | 23.1774 | Noisy | 2.3177e-08 | D6
| B4 | Yes | Yes | No | 282 | 2.24e-08 | 1.83e-08 |  | 0.95598 | 1.9192 |  |  | 1.0753e-08 | 1.7038e-08 | 2.2126e-08 | 0.072 |  | 15.8219 |  | 17.2591 |  | 13.9126 |  | 1.3913e-08 | B4
| C5 | Yes | Yes | Yes | 283 | 1.75e-08 | 1.01e-08 |  | 0.9607 | 1.8856 |  |  | 8.5754e-09 | 1.5934e-08 | 2.2877e-08 | 0.055 |  | 14.8502 |  | 16.7486 |  | 13.4726 |  | 1.3473e-08 | C5
| D4 | Yes | Yes | No | 284 | 1.8e-08 | 9.95e-09 |  | 0.96124 | 2.699 |  |  | 8.9319e-09 | 1.6151e-08 | 2.1194e-08 | 0.066 |  | 14.8938 |  | 17.5596 |  | 13.5135 |  | 1.3514e-08 | D4
| A4 | Yes | Yes | No | 285 | 1.67e-08 | 1.02e-08 |  | 0.96207 | 2.1612 |  |  | 9.2505e-09 | 1.4864e-08 | 2.1931e-08 | 0.088 |  | 14.3555 |  | 17.6153 |  | 13.3112 |  | 1.3311e-08 | A4

## PSW
PSW | BODAC |  |  |  | PFM3 |  |  |  | PFM4
| Pixel Name | BDA Operability | Operating Bodac | Noisy Bodac? | Channel # | PFM3 Noise | PFM3 WE Noise | JJB Comment | JFET Gain Test | Warm Resistance test | JFET ID | 4K VSS | JFET + LIA Noise (nV/rtHz) | Noise (nV/rtHz) (1-3 Hz) | Noise (nV/rtHz) (0.06 - 0.3 Hz) (20061127) | 1/f Knee (PTC Control) (20061127) | 0.3 K load curve | Noise taken with correct VSS | Failure mode following cryoharness probe | IST Noise | IST Comment | Flight Noise | Flight Operation
| R1 | Yes | Yes | No | 1 | 1.41e-08 | 1.87e-20 |  | 0.96628 | 4.2734 | 19/42 | -1.68V | 2.2355e-08 | 2.1691e-08 | 2.7016e-08 | <0.010 |  | 19.6763 |  | 14.6144 |  | 9.871 |  | 9.871e-09 | R1
| D16 | Yes | Yes | No | 2 | 1.96e-08 | 1.06e-08 |  | 0.96369 | 1.9979 | 2.51 |  | 8.802e-09 | 1.4974e-08 | 2.0335e-08 | 0.075 |  | 14.3129 |  | 17.9191 |  | 14.0504 |  | 1.405e-08 | D16
| T1 | Yes | Yes | No | 3 | 1.73e-08 | 1.03e-08 |  | 0.9645 | 26.3116 | -1.29 |  | 9.3149e-09 | 2.6677e-08 | 3.4908e-08 | 0.132 |  | 19.6375 |  | 18.151 |  | 17.546 |  | 1.7546e-08 | T1
| B16 | Yes | Yes | No | 4 | 1.66e-08 | 1.02e-08 |  | 0.96666 | 1.1536 | 1.11 mA |  | 1.2808e-08 | 1.8311e-08 | 2.5155e-08 | 0.075 |  | 16.6935 |  | 20.4435 |  | 15.8472 |  | 1.5847e-08 | B16
| C15 | Yes | Yes | No | 5 | 1.74e-08 | 1.02e-08 |  | 0.96225 | 1.4976 |  |  | 1.0204e-08 | 1.7129e-08 | 2.3427e-08 | 0.074 |  | 15.4679 |  | 19.0197 |  | 15.266 |  | 1.5266e-08 | C15
| A15 | Yes | Yes | No | 6 | 1.66e-08 | 1.02e-08 |  | 0.96037 | 1.2796 |  |  | 8.5791e-09 | 1.6687e-08 | 2.2794e-08 | 0.057 |  | 14.9365 |  | 18.3817 |  | 14.2611 |  | 1.4261e-08 | A15
| D15 | Yes | Yes | No | 7 | 1.95e-08 | 1.56e-08 |  | 0.96712 | 1.8444 |  |  | 1.0938e-08 | 1.6719e-08 | 2.3111e-08 | 0.063 |  | 15.5601 |  | 0 | A sad tale - now wired backwards | 4.0818e-11 |  | 4.0818e-20 | D15
| B15 | Yes | Yes | No | 8 | 1.7e-08 | 1.14e-08 |  | 0.96727 | 1.4413 |  |  | 9.2591e-09 | 1.5622e-08 | 2.1213e-08 | 0.071 |  | 14.5228 |  | 18.3661 |  | 14.2928 |  | 1.4293e-08 | B15
| C14 | Yes | Yes | No | 9 | 1.77e-08 | 2.06e-08 |  | 0.96641 | 1.2011 |  |  | 1.0442e-08 | 1.6414e-08 | 2.3171e-08 | 0.066 |  | 15.218 |  | 17.7636 |  | 14.8792 |  | 1.4879e-08 | C14
| D14 | Yes | Yes | No | 10 | 4.76e-08 | 1.07e-08 | V. Noisy, OK later | 0.96587 | 1.1809 |  |  | 9.1635e-09 | 3.0981e-08 | 6.1434e-08 | 0.278 |  | 109.4896 |  | 19.4708 | Not noisy now | 15.0538 |  | 1.5054e-08 | D14
| A14 | Yes | Yes | No | 11 | 2.74e-08 | 1.52e-08 | M. Noisy, JFETs | 0.96484 | 1.5693 |  |  | 9.7064e-09 | 1.5432e-08 | 2.142e-08 | 0.07 |  | 14.1953 |  | 17.4492 |  | 13.4397 |  | 1.344e-08 | A14
| A13 | Yes | Yes | No | 12 | 1.49e-08 | 1.01e-08 | Slow, 360 ms | 0.96675 | 1.1015 |  |  | 9.1293e-09 | 1.4925e-08 | 2.0033e-08 | 0.085 |  | 13.7746 |  | 14.8541 | Very slow | 12.3233 | Slow | 1.2323e-08 | A13
| B14 | Yes | Yes | No | 13 | 1.75e-08 | 1.2e-08 |  | 0.96643 | 1.1285 |  |  | 9.709e-09 | 1.5612e-08 | 2.3048e-08 | 0.075 |  | 14.6174 |  | 18.495 |  | 14.2553 |  | 1.4255e-08 | B14
| C13 | Yes | Yes | No | 14 | 1.64e-08 | 1.01e-08 |  | 0.96685 | 1.2453 |  |  | 9.3033e-09 | 1.7471e-08 | 2.1947e-08 | 0.104 |  | 15.4549 |  | 17.2129 |  | 14.0156 |  | 1.4016e-08 | C13
| B13 | Yes | Yes | No | 15 | 1.75e-08 | 1.23e-08 |  | 0.9666 | 1.8279 |  |  | 8.882e-09 | 1.5537e-08 | 2.0933e-08 | 0.069 |  | 14.1864 |  | 17.7721 |  | 13.8306 |  | 1.3831e-08 | B13
| D13 | Yes | Yes | No | 16 | 1.65e-08 | 1.05e-08 |  | 0.96644 | 1.9323 |  |  | 9.7849e-09 | 1.6443e-08 | 2.462e-08 | 0.072 |  | 15.0936 |  | 17.3714 |  | 13.6344 |  | 1.3634e-08 | D13
| A12 | Yes | Yes | No | 17 | 1.66e-08 | 1.03e-08 |  | 0.93476 | 4.6817 |  |  | 9.0475e-09 | 2.1533e-08 | 2.6915e-08 | 0.02 | weird offset at top | 18.7126 |  | 18.7245 |  | 14.3858 |  | 1.4386e-08 | A12
| C12 | No | Yes | No | 18 | 2.29e-21 | 1.87e-20 | Quiet | 0.96736 | 14462.237 |  |  | 1.3168e-20 | 8.0043e-21 | 8.6026e-20 | <0.010 | dead | 0.24155 |  | 0 | Dead,dead,dead | 8.9493e-09 |  | 8.9493e-18 | C12
| D12 | Yes | Yes | No | 19 | 1.77e-08 | 1.18e-08 |  | 0.96366 | 0.9806 |  |  | 9.3506e-09 | 1.5582e-08 | 2.1304e-08 | 0.068 | weird offset at top | 14.2732 |  | 17.5747 |  | 14.4576 |  | 1.4458e-08 | D12
| B12 | Yes | Yes | No | 20 | 3.34e-08 | 2.01e-08 | M. Noisy, unknown | 0.96732 | 76.3052 |  |  | 1.6822e-08 | 1.9301e-08 | 2.5766e-08 | 0.067 |  | 19.3763 |  | 22.5335 |  | 18.7462 |  | 1.8746e-08 | B12
| E11 | Yes | Yes | No | 21 | 1.74e-08 | 1.16e-08 |  | 0.96695 | 2.0893 |  |  | 9.0239e-09 | 1.5613e-08 | 2.1851e-08 | 0.064 |  | 14.704 |  | 17.9606 |  | 13.9159 |  | 1.3916e-08 | E11
| A11 | Yes | Yes | No | 22 | 1.57e-08 | 1.05e-08 |  | 0.96461 | 1.9797 |  |  | 8.769e-09 | 1.5867e-08 | 2.229e-08 | 0.068 |  | 14.734 |  | 14.9763 | Very slow | 12.8087 | Slow | 1.2809e-08 | A11
| C11 | Yes | Yes | No | 23 | 4.23e-08 | 3.45e-08 | V. Noisy, JFETs | 0.96424 | 2.9588 |  |  | 9.578e-09 | 1.7632e-08 | 2.3253e-08 | 0.058 |  | 15.7506 |  | 16.6092 |  | 13.8068 |  | 1.3807e-08 | C11
| B11 | Yes | Yes | No | 24 | 2.25e-08 | 1.88e-08 |  | 0.96546 | 2.534 |  |  | 9.6814e-09 | 1.5528e-08 | 2.0865e-08 | 0.055 |  | 14.6515 |  | 17.4189 |  | 13.291 |  | 1.3291e-08 | B11
| E1 | Yes | Yes | No | 25 | 1.83e-08 | 1.23e-08 | S/N 13/39 | 0.96731 | 1.8426 | 20/51 | -1.5853V | 1.0739e-08 | 1.6428e-08 | 2.1946e-08 | 0.071 |  | 15.2256 |  | 19.1457 |  | 15.1062 |  | 1.5106e-08 | E1
| F1 | Yes | Yes | Dead LIA | 26 | 1.84e-08 | 1.61e-08 |  | 0.96709 | 2.5686 | 2.57 |  | 9.7801e-09 | 1.541e-08 | 2.2476e-08 | 0.075 |  | 15.134 |  | 18.0846 |  | 14.2706 |  | 1.4271e-08 | F1
| T2 | Yes | Yes | No | 27 | 1.78e-08 | 1.16e-08 |  | 0.96476 | 2.1602 | -1.07 |  | 9.438e-09 | 2.5337e-08 | 4.9658e-08 | 0.234 |  | 21.5156 |  | 17.9469 | Not noisy now | 17.3759 |  | 1.7376e-08 | T2
| H1 | Yes | Yes | No | 28 | 1.67e-08 | 1.02e-08 |  | 0.96713 | 2.4792 | 0.88 mA |  | 9.0095e-09 | 1.5571e-08 | 2.2523e-08 | 0.066 |  | 14.344 |  | 18.3874 |  | 14.208 |  | 1.4208e-08 | H1
| G1 | Yes | Yes | No | 29 | 1.75e-08 | 1.14e-08 |  | 0.96183 | 2.0518 |  |  | 8.9263e-09 | 1.6124e-08 | 2.2819e-08 | 0.071 |  | 15.0106 |  | 18.0168 |  | 15.1505 |  | 1.5151e-08 | G1
| J1 | Yes | Yes | No | 30 | 1.64e-08 | 1.02e-08 |  | 0.96645 | 2.6805 |  |  | 9.352e-09 | 1.5635e-08 | 2.2269e-08 | 0.073 |  | 14.3304 |  | 19.2953 |  | 14.4 |  | 1.44e-08 | J1
| H2 | Yes | Yes | No | 31 | 1.71e-08 | 9.82e-09 |  | 0.96193 | 2.3704 |  |  | 1.0728e-08 | 1.7363e-08 | 2.9217e-08 | 0.113 |  | 15.773 |  | 19.4708 |  | 15.5613 |  | 1.5561e-08 | H2
| F2 | Yes | Yes | No | 32 | 1.6e-08 | 1.01e-08 |  | 0.96049 | 2.6179 |  |  | 9.1333e-09 | 1.5554e-08 | 1.9808e-08 | 0.063 |  | 14.529 |  | 17.0425 |  | 13.917 |  | 1.3917e-08 | F2
| J2 | Yes | Yes | No | 33 | 1.73e-08 | 1.05e-08 |  | 0.94962 | 1.201 |  |  | 9.4421e-09 | 1.6888e-08 | 2.2792e-08 | 0.061 |  | 15.0933 |  | 20.7258 |  | 15.663 |  | 1.5663e-08 | J2
| G2 | Yes | Yes | V. Noisy | 34 | 1.7e-08 | 1.04e-08 |  | 0.96325 | 1.3131 |  |  | 9.7484e-09 | 1.565e-08 | 2.1348e-08 | 0.073 |  | 14.8554 |  | 18.7815 |  | 14.9924 |  | 1.4992e-08 | G2
| H3 | Yes | Yes | No | 35 | 1.64e-08 | 1.06e-08 |  | 0.96601 | 1.4928 |  |  | 9.6128e-09 | 1.5021e-08 | 1.9952e-08 | 0.073 |  | 14.4454 |  | 17.8571 |  | 14.0745 |  | 1.4074e-08 | H3
| J3 | Yes | Yes | No | 36 | 1.77e-08 | 1.05e-08 |  | 0.9629 | 1.0274 |  |  | 9.6047e-09 | 1.5556e-08 | 2.3017e-08 | 0.078 |  | 14.6077 |  | 19.4505 |  | 15.7423 |  | 1.5742e-08 | J3
| E2 | Yes | Yes | No | 37 | 1.67e-08 | 9.91e-09 |  | 0.96649 | 1.2099 |  |  | 9.5949e-09 | 1.5962e-08 | 2.144e-08 | 0.069 |  | 14.5639 |  | 18.2199 |  | 14.5154 |  | 1.4515e-08 | E2
| F3 | Yes | Yes | No | 38 | 1.68e-08 | 1.09e-08 |  | 0.9662 | 1.1819 |  |  | 1.0299e-08 | 1.688e-08 | 2.2857e-08 | 0.086 |  | 15.2695 |  | 19.3012 |  | 15.5318 |  | 1.5532e-08 | F3
| G3 | Yes | Yes | No | 39 | 1.6e-08 | 9.93e-09 |  | 0.96364 | 1.4476 |  |  | 1.105e-08 | 1.7512e-08 | 2.547e-08 | 0.056 |  | 16.3976 |  | 19.1372 |  | 15.2431 |  | 1.5243e-08 | G3
| H4 | Yes | Yes | No | 40 | 2.1e-08 | 1.52e-08 |  | 0.96539 | 1.6399 |  |  | 9.4104e-09 | 1.5632e-08 | 2.1911e-08 | 0.071 |  | 14.4449 |  | 18.3719 |  | 14.417 |  | 1.4417e-08 | H4
| J4 | Yes | Yes | No | 41 | 1.69e-08 | 1.1e-08 |  | 0.96656 | 1.0343 |  |  | 1.1072e-08 | 1.7092e-08 | 2.5152e-08 | 0.061 |  | 16.3996 |  | 19.2171 |  | 14.8568 |  | 1.4857e-08 | J4
| E3 | Yes | Yes | No | 42 | 1.69e-08 | 9.98e-09 |  | 0.967 | 2.072 |  |  | 1.0321e-08 | 1.6664e-08 | 2.5067e-08 | 0.072 |  | 15.0021 |  | 18.8215 |  | 14.1681 |  | 1.4168e-08 | E3
| F4 | Yes | Yes | No | 43 | 1.71e-08 | 9.92e-09 |  | 0.96594 | 2.4926 |  |  | 9.4856e-09 | 1.6489e-08 | 2.24e-08 | 0.074 |  | 15.1613 |  | 18.4391 |  | 15.7912 |  | 1.5791e-08 | F4
| G4 | Yes | Yes | No | 44 | 1.76e-08 | 1.11e-08 |  | 0.96686 | 1.0903 |  |  | 9.288e-09 | 1.6364e-08 | 2.2373e-08 | 0.085 |  | 15.0269 |  | 19.0888 |  | 15.0131 |  | 1.5013e-08 | G4
| H5 | Yes | Yes | No | 45 | 1.66e-08 | 9.97e-09 |  | 0.94938 | 2.1984 |  |  | 1.0004e-08 | 1.6009e-08 | 2.0203e-08 | 0.067 |  | 14.6677 |  | 18.116 |  | 13.961 |  | 1.3961e-08 | H5
| E4 | Yes | Yes | No | 46 | 1.72e-08 | 1.03e-08 |  | 0.9644 | 1.757 |  |  | 9.6619e-09 | 1.5946e-08 | 2.1936e-08 | 0.082 |  | 14.8767 |  | 17.7991 |  | 15.7044 |  | 1.5704e-08 | E4
| J5 | Yes | Yes | No | 47 | 1.7e-08 | 1.02e-08 |  | 0.96446 | 1.3677 |  |  | 1.0017e-08 | 1.5642e-08 | 2.0614e-08 | 0.072 |  | 14.478 |  | 19.1424 |  | 14.115 |  | 1.4115e-08 | J5
| F5 | Yes | Yes | No | 48 | 1.69e-08 | 1.05e-08 |  | 0.96649 | 1.2485 |  |  | 9.4205e-09 | 1.539e-08 | 2.1271e-08 | 0.073 |  | 14.5357 |  | 17.4429 |  | 14.0274 |  | 1.4027e-08 | F5
| D6 | Yes | Yes | No | 49 | 1.79e-08 | 1.03e-08 |  | 0.95484 | 4.04 | 19/44 | -1.3726V | 8.6226e-09 | 1.5671e-08 | 2.2396e-08 | 0.075 |  | 14.5357 |  | 17.6008 |  | 14.3665 |  | 1.4366e-08 | D6
| B6 | Yes | Yes | No | 50 | 1.69e-08 | 1.18e-08 |  | 0.96122 | 1.4283 | 2.47 |  | 8.9061e-09 | 1.5037e-08 | 2.1337e-08 | 0.065 |  | 13.9451 |  | 17.1279 |  | 14.2283 |  | 1.4228e-08 | B6
| C5 | Yes | Yes | No | 51 | 2.01e-08 | 1.15e-08 |  | 0.9634 | 1.2038 | -1.08 |  | 9.1158e-09 | 1.5445e-08 | 2.16e-08 | 0.071 |  | 14.5958 |  | 17.8089 |  | 14.0899 |  | 1.409e-08 | C5
| A5 | Yes | Yes | No | 52 | 1.82e-08 | 1.04e-08 |  | 0.96051 | 1.3513 | 1.25 mA |  | 9.5394e-09 | 1.6362e-08 | 2.1089e-08 | 0.065 |  | 14.5625 |  | 18.4363 |  | 14.1057 |  | 1.4106e-08 | A5
| E5 | Yes | Yes | No | 53 | 1.76e-08 | 1.04e-08 |  | 0.95991 | 1.2709 |  |  | 9.6089e-09 | 1.5442e-08 | 2.1118e-08 | 0.081 |  | 14.446 |  | 18.5156 |  | 15.279 |  | 1.5279e-08 | E5
| B5 | Yes | Yes | M. Noisy | 54 | 1.94e-06 | 1.12e-08 | V. Noisy, unknown | 0.96405 | 1.3291 |  |  | 1.0516e-08 | 2.747e-08 | 8.2203e-08 | 0.213 |  | 24.4631 |  | 19.1427 | Not noisy now | 15.7917 |  | 1.5792e-08 | B5
| D5 | Yes | Yes | No | 55 | 1.72e-08 | 1.03e-08 |  | 0.96134 | 1.186 |  |  | 9.0205e-09 | 1.6834e-08 | 2.3423e-08 | 0.081 |  | 15.7569 |  | 17.8036 |  | 14.1181 |  | 1.4118e-08 | D5
| C4 | Yes | Yes | No | 56 | 1.7e-08 | 1.03e-08 |  | 0.95508 | 1.3305 |  |  | 9.0042e-09 | 1.5466e-08 | 1.7985e-08 | 0.07 |  | 14.0113 |  | 17.1591 |  | 13.0878 |  | 1.3088e-08 | C4
| A4 | Yes | Yes | No | 57 | 1.81e-08 | 1.12e-08 |  | 0.96259 | 1.3379 |  |  | 9.266e-09 | 1.855e-08 | 2.2735e-08 | 0.075 |  | 16.4653 |  | 19.6326 |  | 16.0636 |  | 1.6064e-08 | A4
| D4 | Yes | Yes | No | 58 | 1.66e-08 | 1.07e-08 |  | 0.9626 | 1.4966 |  |  | 9.7631e-09 | 1.6217e-08 | 2.216e-08 | 0.074 |  | 15.2863 |  | 16.7747 |  | 14.237 |  | 1.4237e-08 | D4
| B4 | Yes | Yes | No | 59 | 1.69e-08 | 1.01e-08 |  | 0.95048 | 1.2909 |  |  | 9.0909e-09 | 1.7332e-08 | 2.1998e-08 | 0.069 |  | 19.9641 |  | 17.9742 |  | 15.0233 |  | 1.5023e-08 | B4
| C3 | Yes | Yes | No | 60 | 1.82e-08 | 1.05e-08 |  | 0.96126 | 1.1728 |  |  | 9.5836e-09 | 1.6948e-08 | 2.505e-08 | 0.063 |  | 15.9146 |  | 18.5961 |  | 16.1337 |  | 1.6134e-08 | C3
| B3 | Yes | Yes | No | 61 | 1.75e-08 | 1.07e-08 |  | 0.9588 | 1.1397 |  |  | 6.6368e-09 | 7.8236e-09 | 1.0786e-08 | <0.010 | low gain | 8.6739 |  | 18.8566 | Normal gain now | 14.3546 |  | 1.4355e-08 | B3
| A3 | Yes | Yes | No | 62 | 1.69e-08 | 1.01e-08 |  | 0.96457 | 1.2387 |  |  | 9.1791e-09 | 1.5511e-08 | 2.2964e-08 | 0.077 |  | 14.5857 |  | 19.0415 |  | 14.1925 |  | 1.4192e-08 | A3
| A2 | Yes | Yes | M. Noisy | 63 | 1.65e-08 | 1.16e-08 |  | 0.96205 | 1.3251 |  |  | 8.8755e-09 | 1.494e-08 | 2.0683e-08 | 0.071 |  | 14.1055 |  | 18.008 |  | 14.6639 |  | 1.4664e-08 | A2
| D3 | Yes | Yes | No | 64 | 1.68e-08 | 1.01e-08 |  | 0.96018 | 1.4927 |  |  | 9.5667e-09 | 1.5619e-08 | 2.148e-08 | 0.057 |  | 14.515 |  | 18.1106 |  | 14.1097 |  | 1.411e-08 | D3
| C2 | Yes | Yes | No | 65 | 1.79e-08 | 1.21e-08 |  | 0.96183 | 1.2353 |  |  | 9.6247e-09 | 1.5928e-08 | 2.203e-08 | 0.063 |  | 14.5655 |  | 20.9006 |  | 16.8903 |  | 1.689e-08 | C2
| B2 | Yes | Yes | No | 66 | 1.65e-08 | 1.01e-08 |  | 0.9619 | 1.3027 |  |  | 9.1787e-09 | 1.6721e-08 | 2.2996e-08 | 0.069 |  | 15.379 |  | 17.8198 |  | 14.1797 |  | 1.418e-08 | B2
| D2 | Yes | Yes | No | 67 | 1.8e-08 | 1.09e-08 |  | 0.96317 | 1.2623 |  |  | 9.1509e-09 | 1.5537e-08 | 2.3249e-08 | 0.078 |  | 14.6928 |  | 18.9186 |  | 14.8236 |  | 1.4824e-08 | D2
| A1 | Yes | Yes | No | 68 | 1.73e-08 | 1.05e-08 |  | 0.96328 | 1.1702 |  |  | 9.741e-09 | 1.5762e-08 | 2.0956e-08 | 0.066 |  | 14.9344 |  | 19.9083 |  | 15.1267 |  | 1.5127e-08 | A1
| C1 | Yes | Yes | No | 69 | 1.75e-08 | 1.13e-08 |  | 0.96453 | 1.273 |  |  | 9.1833e-09 | 1.572e-08 | 2.1299e-08 | 0.068 |  | 14.3486 |  | 18.0392 |  | 15.0994 |  | 1.5099e-08 | C1
| B1 | Yes | Yes | No | 70 | 1.7e-08 | 1.09e-08 |  | 0.96292 | 1.2661 |  |  | 9.8035e-09 | 1.5841e-08 | 2.1936e-08 | 0.06 |  | 14.8204 |  | 18.2327 |  | 15.151 |  | 1.5151e-08 | B1
| DK1 | Yes | Yes | No | 71 | 1.68e-08 | 1.05e-08 |  | 0.96268 | 1.2513 |  |  | 9.2173e-09 | 1.6125e-08 | 2.0309e-08 | 0.057 |  | 14.5352 |  | 16.9146 |  | 12.4679 |  | 1.2468e-08 | DK1
| D1 | Yes | Yes | No | 72 | 1.8e-08 | 1.04e-08 |  | 0.95734 | 1.3084 |  |  | 9.7363e-09 | 1.5828e-08 | 1.9838e-08 | 0.061 |  | 14.8828 |  | 17.9938 |  | 14.54 |  | 1.454e-08 | D1
| F12 | Yes | Yes | M. Noisy | 73 | 1.84e-08 | 1.21e-08 | S/N 13/33 | 0.9653 | 1.5123 | 20/50 | -1.5853V | 1.0074e-08 | 1.6403e-08 | 2.0972e-08 | 0.077 |  | 14.9071 |  | 18.1035 |  | 14.9466 |  | 1.4947e-08 | F12
| J11 | Yes | Yes | No | 74 | 1.84e-08 | 1.5e-08 |  | 0.96486 | 1.3215 | 2.53 |  | 1.0658e-08 | 1.6133e-08 | 2.1619e-08 | 0.066 |  | 14.6705 |  | 19.8 | Wrong result from EMC test - value taken from 0xb00016f1 on NHSC Wiki | 14.086 |  | 1.4086e-08 | J11
| E12 | Yes | Yes | No | 75 | 1.69e-08 | 1.01e-08 |  | 0.96385 | 1.3164 | -1.24 |  | 9.9713e-09 | 1.6646e-08 | 2.2313e-08 | 0.068 |  | 14.9552 |  | 18.7223 |  | 14.4954 |  | 1.4495e-08 | E12
| H12 | Yes | Yes | No | 76 | 1.95e-08 | 1.35e-08 |  | 0.92392 | 1.325 | 1.01 mA |  | 9.6256e-09 | 1.6913e-08 | 2.1727e-08 | 0.064 |  | 15.1145 |  | 17.6688 |  | 14.5265 |  | 1.4526e-08 | H12
| G12 | Yes | Yes | No | 77 | 1.64e-08 | 1.12e-08 |  | 0.92536 | 1.3703 |  |  | 9.3052e-09 | 1.5646e-08 | 2.1198e-08 | 0.067 |  | 14.3523 |  | 17.6959 |  | 15.1073 |  | 1.5107e-08 | G12
| F13 | Yes | Yes | No | 78 | 1.84e-08 | 1.28e-08 |  | 0.94246 | 1.1282 |  |  | 1.0352e-08 | 1.6666e-08 | 2.2633e-08 | 0.067 |  | 15.0357 |  | 17.7227 |  | 14.8629 |  | 1.4863e-08 | F13
| E13 | Yes | Yes | No | 79 | 1.71e-08 | 1.04e-08 |  | 0.96669 | 1.4667 |  |  | 9.6792e-09 | 1.5823e-08 | 2.1859e-08 | 0.074 |  | 14.7226 |  | 17.7157 |  | 14.2083 |  | 1.4208e-08 | E13
| J12 | Yes | Yes | No | 80 | 1.73e-08 | 1.17e-08 |  | 0.95787 | 1.0774 |  |  | 1.0078e-08 | 1.632e-08 | 2.1594e-08 | 0.098 |  | 14.876 |  | 19.2014 |  | 14.9257 |  | 1.4926e-08 | J12
| H13 | Yes | Yes | No | 81 | 2.45e-08 | 2.02e-08 |  | 0.96563 | 1.2217 |  |  | 1.0047e-08 | 1.6595e-08 | 2.1907e-08 | 0.077 |  | 15.5833 |  | 19.1177 |  | 14.9715 |  | 1.4971e-08 | H13
| G13 | Yes | Yes | No | 82 | 2.95e-08 | 2.8e-08 | M. Noisy, JFETs | 0.96377 | 407.2638 |  |  | 9.9944e-09 | 1.6574e-08 | 2.2025e-08 | 0.061 |  | 14.7164 |  | 18.9494 |  | 14.544 |  | 1.4544e-08 | G13
| F14 | Yes | Yes | Dead LIA | 83 | 1.74e-08 | 1.06e-08 |  | 0.9674 | 1.4353 |  |  | 1.0049e-08 | 1.5861e-08 | 2.116e-08 | 0.072 |  | 14.865 |  | 18.813 |  | 14.9892 |  | 1.4989e-08 | F14
| E14 | Yes | Yes | No | 84 | 1.61e-08 | 1.01e-08 |  | 0.96383 | 2.5929 |  |  | 1.1653e-08 | 1.7454e-08 | 2.3506e-08 | 0.067 |  | 15.8945 |  | 19.7015 |  | 15.3973 |  | 1.5397e-08 | E14
| J13 | Yes | Yes | No | 85 | 1.68e-08 | 1.07e-08 |  | 0.96273 | 1.1271 |  |  | 9.7368e-09 | 1.5516e-08 | 2.2512e-08 | 0.076 |  | 14.3376 |  | 19.1798 |  | 14.8457 |  | 1.4846e-08 | J13
| H14 | Yes | Yes | No | 86 | 1.64e-08 | 1.06e-08 |  | 0.96462 | 1.1846 |  |  | 9.7228e-09 | 1.547e-08 | 2.2746e-08 | 0.068 |  | 14.701 |  | 17.8515 |  | 14.2779 |  | 1.4278e-08 | H14
| G14 | Yes | Yes | No | 87 | 2.28e-08 | 2.01e-08 |  | 0.96507 | 1.1998 |  |  | 9.9425e-09 | 1.6e-08 | 2.2391e-08 | 0.068 |  | 14.5396 |  | 18.3442 |  | 14.8846 |  | 1.4885e-08 | G14
| J14 | Yes | Yes | No | 88 | 1.71e-08 | 9.85e-09 |  | 0.9636 | 1.1012 |  |  | 1.1138e-08 | 1.6657e-08 | 2.2698e-08 | 0.063 |  | 15.3936 |  | 19.8004 |  | 15.3501 |  | 1.535e-08 | J14
| F15 | Yes | Yes | No | 89 | 1.8e-08 | 1.19e-08 |  | 0.96488 | 1.2316 |  |  | 9.4965e-09 | 1.5741e-08 | 2.2987e-08 | 0.089 |  | 14.5897 |  | 18.0392 |  | 14.178 |  | 1.4178e-08 | F15
| H15 | Yes | Yes | No | 90 | 1.67e-08 | 1.07e-08 |  | 0.96297 | 1.2508 |  |  | 1.0084e-08 | 1.5998e-08 | 2.2314e-08 | 0.07 |  | 14.8008 |  | 18.0708 |  | 14.467 |  | 1.4467e-08 | H15
| J15 | Yes | Yes | No | 91 | 1.72e-08 | 1.14e-08 |  | 0.96403 | 1.3867 |  |  | 9.858e-09 | 1.5469e-08 | 2.004e-08 | 0.081 |  | 14.8064 |  | 18.9536 |  | 14.359 |  | 1.4359e-08 | J15
| G15 | Yes | Yes | No | 92 | 1.72e-08 | 9.95e-09 |  | 0.96606 | 1.1955 |  |  | 9.5491e-09 | 1.6824e-08 | 2.0611e-08 | 0.073 |  | 15.1878 |  | 18.294 |  | 15.0567 |  | 1.5057e-08 | G15
| H16 | Yes | No | No | 93 | 2.05e-08 | 1.47e-08 |  | 0.96549 | 1.1739 |  |  | 9.8075e-09 | 1.6722e-08 | 2.3967e-08 | 0.067 |  | 15.3384 |  | 18.6452 |  | 14.7538 |  | 1.4754e-08 | H16
| DK2 | Yes | Yes | No | 94 | 2.66e-08 | 2.45e-08 | M. Noisy, JFETs | 0.96527 | 1.1393 |  |  | 9.3687e-09 | 1.5515e-08 | 1.994e-08 | 0.063 |  | 14.1661 |  | 16.9948 |  | 13.0191 |  | 1.3019e-08 | DK2
| F16 | Yes | Yes | M. Noisy | 95 | 1.97e-08 | 1.29e-08 |  | 0.96768 | 1.3611 |  |  | 9.4499e-09 | 1.5897e-08 | 2.5134e-08 | 0.065 |  | 15.1524 |  | 18.7971 |  | 15.0006 |  | 1.5001e-08 | F16
| E15 | Yes | Yes | No | 96 | 2.09e-08 | 1.62e-08 |  | 0.96404 | 1.2889 |  |  | 1.0039e-08 | 1.5554e-08 | 2.0405e-08 | 0.063 |  | 14.2992 |  | 17.4764 |  | 13.9313 |  | 1.3931e-08 | E15
| D11 | Yes | Yes | No | 97 | 1.72e-08 | 1.05e-08 | Slow, 45 ms | 0.96867 | 4.3018 | 16/40 | -1.5853V | 8.7909e-09 | 1.5803e-08 | 2.1955e-08 | 0.067 |  | 15.2092 |  | 17.3652 |  | 14.8301 |  | 1.483e-08 | D11
| A10 | Yes | Yes | No | 98 | 0 | 0 | Cryoharness, slow 41 ms | 0.0063589 | 417.0865 | 2.54 |  | 0 | 0 | 0 | ---- | zero | 0 | Bad contact S-? | 16.3034 | Alive Again - BMS check JFET Gain | 12.4701 | A bit Slow | 1.247e-08 | A10
| E10 | Yes | Yes | No | 99 | 1.73e-08 | 1.1e-08 |  | 0.96912 | 1.35 | -1.28 |  | 9.1395e-09 | 1.6216e-08 | 2.1067e-08 | 0.085 |  | 14.6821 |  | 18.8211 |  | 15.426 |  | 1.5426e-08 | E10
| C10 | Yes | Yes | No | 100 | 1.78e-08 | 1.09e-08 |  | 0.96489 | 1.8662 | 1.00 mA |  | 9.0976e-09 | 1.6443e-08 | 2.3678e-08 | 0.08 |  | 15.6178 |  | 17.7348 |  | 14.25 |  | 1.425e-08 | C10
| B10 | Yes | Yes | No | 101 | 1.75e-08 | 1.17e-08 |  | 0.96821 | 1.4952 |  |  | 1.1632e-08 | 1.7462e-08 | 2.3002e-08 | 0.06 |  | 15.5928 |  | 18.4389 |  | 13.842 |  | 1.3842e-08 | B10
| D10 | Yes | Yes | No | 102 | 1.77e-08 | 1.14e-08 |  | 0.96667 | 1.6933 |  |  | 8.9599e-09 | 1.8694e-08 | 2.4588e-08 | 0.07 |  | 16.8166 |  | 17.666 |  | 14.2306 |  | 1.4231e-08 | D10
| A9 | Yes | Yes | No | 103 | 1.71e-08 | 1.05e-08 |  | 0.96931 | 1.2589 |  |  | 9.2667e-09 | 1.5917e-08 | 2.1159e-08 | 0.067 |  | 14.7729 |  | 18.6481 |  | 13.8866 |  | 1.3887e-08 | A9
| E9 | Yes | No | No | 104 | 1.7e-08 | 1.03e-08 |  | 0.9689 | 1.8499 |  |  | 9.0976e-09 | 1.5793e-08 | 2.4361e-08 | 0.093 |  | 16.3396 |  | 17.4836 |  | 14.1387 |  | 1.4139e-08 | E9
| C9 | Yes | Yes | No | 105 | 1.76e-08 | 1.02e-08 |  | 0.96888 | 1.6557 |  |  | 8.9707e-09 | 1.6961e-08 | 2.3339e-08 | 0.087 |  | 15.508 |  | 18.7682 |  | 15.3285 |  | 1.5329e-08 | C9
| B9 | Yes | Yes | No | 106 | 2.23e-08 | 2.1e-08 |  | 0.96902 | 1.5907 |  |  | 1.5393e-08 | 1.7662e-08 | 2.5064e-08 | 0.054 |  | 16.3263 |  | 18.3283 |  | 14.5105 |  | 1.451e-08 | B9
| D9 | Yes | Yes | No | 107 | 1.81e-08 | 1.04e-08 |  | 0.96897 | 1.7706 |  |  | 9.1761e-09 | 1.6115e-08 | 2.197e-08 | 0.068 |  | 15.2255 |  | 18.6016 |  | 15.0151 |  | 1.5015e-08 | D9
| A8 | Yes | Yes | No | 108 | 1.71e-08 | 1.02e-08 |  | 0.96912 | 2.6431 |  |  | 8.9832e-09 | 2.1732e-08 | 5.7802e-08 | 0.166 |  | 33.3238 |  | 19.1003 |  | 14.0858 |  | 1.4086e-08 | A8
| C8 | Yes | Yes | No | 109 | 1.68e-08 | 1.07e-08 |  | 0.96941 | 1.9143 |  |  | 9.4544e-09 | 1.6089e-08 | 2.2382e-08 | 0.069 |  | 14.7871 |  | 18.0383 |  | 14.2823 |  | 1.4282e-08 | C8
| E8 | Yes | Yes | No | 110 | 1.71e-08 | 1.02e-08 |  | 0.96896 | 1.6242 |  |  | 9.0551e-09 | 1.7205e-08 | 2.201e-08 | 0.064 |  | 15.544 |  | 19.5588 |  | 14.6674 |  | 1.4667e-08 | E8
| D8 | Yes | Yes | No | 111 | 1.78e-08 | 1.1e-08 |  | 0.96765 | 1.6889 |  |  | 9.8434e-09 | 1.6135e-08 | 2.1858e-08 | 0.076 |  | 15.0349 |  | 17.3427 |  | 14.1396 |  | 1.414e-08 | D8
| B8 | Yes | Yes | No | 112 | 1.66e-08 | 1.02e-08 |  | 0.96907 | 1.21 |  |  | 9.3864e-09 | 1.5578e-08 | 2.3105e-08 | 0.07 |  | 14.4984 |  | 18.6553 |  | 14.797 |  | 1.4797e-08 | B8
| C7 | Yes | Yes | No | 113 | 1.79e-08 | 1.11e-08 |  | 0.968 | 1.8712 |  |  | 9.749e-09 | 1.749e-08 | 2.3758e-08 | 0.063 |  | 15.799 |  | 18.3889 |  | 14.6944 |  | 1.4694e-08 | C7
| E7 | Yes | Yes | No | 114 | 1.74e-08 | 1.16e-08 |  | 0.96601 | 2.6144 |  |  | 9.1805e-09 | 1.573e-08 | 2.3615e-08 | 0.065 |  | 14.7914 |  | 19.5222 |  | 14.9765 |  | 1.4976e-08 | E7
| A7 | Yes | Yes | No | 115 | 1.77e-08 | 1.26e-08 |  | 0.96792 | 1.5091 |  |  | 1.1282e-08 | 1.6792e-08 | 2.2016e-08 | 0.072 |  | 15.2709 |  | 20.3506 |  | 15.7854 |  | 1.5785e-08 | A7
| D7 | Yes | Yes | No | 116 | 1.71e-08 | 1.17e-08 |  | 0.9694 | 1.7461 |  |  | 8.9728e-09 | 1.5645e-08 | 2.1382e-08 | 0.071 |  | 14.5133 |  | 17.2611 |  | 14.0755 |  | 1.4075e-08 | D7
| B7 | Yes | Yes | No | 117 | 1.68e-08 | 1.04e-08 |  | 0.96831 | 1.4461 |  |  | 9.0181e-09 | 1.6767e-08 | 2.2668e-08 | 0.066 |  | 15.3726 |  | 18.1927 |  | 15.058 |  | 1.5058e-08 | B7
| C6 | Yes | Yes | No | 118 | 1.76e-08 | 1.07e-08 |  | 0.96881 | 2.462 |  |  | 9.5249e-09 | 1.6229e-08 | 2.1601e-08 | 0.065 |  | 15.0641 |  | 17.3791 |  | 14.5559 |  | 1.4556e-08 | C6
| E6 | Yes | Yes | No | 119 | 1.6e-08 | 1.03e-08 |  | 0.96909 | 2.5784 |  |  | 8.9119e-09 | 1.6087e-08 | 2.1362e-08 | 0.075 |  | 14.4296 |  | 17.8514 |  | 13.9187 |  | 1.3919e-08 | E6
| A6 | Yes | Yes | No | 120 | 1.7e-08 | 1.02e-08 |  | 0.96787 | 2.17 |  |  | 9.3276e-09 | 1.8131e-08 | 2.4194e-08 | 0.052 |  | 16.1137 |  | 18.0598 |  | 15.1374 |  | 1.5137e-08 | A6
| G5 | Yes | Yes | No | 121 | 1.69e-08 | 1.07e-08 |  | 0.969 | 2.1692 | 16/45 | -1.5853V | 9.2901e-09 | 1.5685e-08 | 2.3184e-08 | 0.063 |  | 15.4001 |  | 17.2465 |  | 14.0785 |  | 1.4078e-08 | G5
| H6 | Yes | Yes | No | 122 | 2.18e-08 | 1.27e-08 |  | 0.96756 | 2.1485 | 2.56 |  | 9.6312e-09 | 1.5589e-08 | 2.2587e-08 | 0.073 |  | 16.3148 |  | 18.0421 |  | 13.8452 |  | 1.3845e-08 | H6
| J6 | Yes | Yes | No | 123 | 4.15e-08 | 7.3e-09 | Low gain | 0.72991 | 45.647 | -1.37 |  | 6.7606e-09 | 9.5892e-09 | 1.0074e-08 | <0.010 | low gain | 8.4266 | Short to ground S+ | 18.5429 | Normal gain now - BMS check JFET Gain | 13.789 |  | 1.3789e-08 | J6
| F6 | Yes | Yes | No | 124 | 1.86e-08 | 1.11e-08 |  | 0.96811 | 1.907 | 0.90 mA |  | 9.3599e-09 | 1.5895e-08 | 2.1983e-08 | 0.073 |  | 14.7546 |  | 17.4884 |  | 14.4723 |  | 1.4472e-08 | F6
| G6 | Yes | Yes | Yes | 125 | 1.75e-08 | 1.27e-08 |  | 0.96799 | 1.7185 |  |  | 9.0926e-09 | 1.5255e-08 | 2.2479e-08 | 0.064 |  | 15.5418 |  | 17.5656 |  | 13.931 |  | 1.3931e-08 | G6
| H7 | Yes | Yes | M. Noisy | 126 | 1.66e-08 | 1.04e-08 |  | 0.96971 | 1.43 |  |  | 9.2328e-09 | 1.5305e-08 | 2.064e-08 | 0.066 |  | 14.4134 |  | 17.1018 |  | 14.1703 |  | 1.417e-08 | H7
| F7 | Yes | Yes | Yes | 127 | 1.68e-08 | 1.01e-08 |  | 0.97087 | 2.5174 |  |  | 8.6332e-09 | 1.5313e-08 | 2.2848e-08 | 0.059 |  | 15.5791 |  | 17.7167 |  | 13.8986 |  | 1.3899e-08 | F7
| J7 | No | Yes | No | 128 | 1.7e-08 | 1.06e-08 | BDA Intermittant | 0.9686 | 7.3838 |  |  | 1.1698e-07 | 2.194e-08 | 1.4267e-07 | 0.396 | intermittant | 25.903 |  | 18.2396 | Normal | 13.6589 |  | 1.3659e-08 | J7
| G7 | Yes | Yes | Yes | 129 | 1.67e-08 | 1.04e-08 |  | 0.96865 | 2.312 |  |  | 8.6153e-09 | 1.5529e-08 | 2.1735e-08 | 0.064 |  | 14.6391 |  | 17.5143 |  | 13.5573 |  | 1.3557e-08 | G7
| H8 | Yes | Yes | M. Noisy | 130 | 1.83e-08 | 1.05e-08 |  | 0.96882 | 1.8266 |  |  | 9.7114e-09 | 1.5856e-08 | 2.0745e-08 | 0.057 |  | 14.9255 |  | 17.5192 |  | 13.8281 |  | 1.3828e-08 | H8
| F8 | Yes | Yes | Yes | 131 | 1.71e-08 | 1.04e-08 |  | 0.0067818 | 3.0654 |  |  | 6.6835e-09 | 1.1068e-06 | 1.1706e-08 | 0.113 | low gain | 8.5598 | Broken S+ | 18.1729 | Alive Again - BMS check JFET Gain | 13.4025 |  | 1.3403e-08 | F8
| G8 | Yes | Yes | M. Noisy | 132 | 2.29e-21 | 1.87e-20 | BDA dead post PFM2 | 0.97078 | 15582.6494 |  |  | 1.3929e-20 | 7.976e-21 | 8.6026e-20 | <0.010 | dead | 0.24155 |  | 0 | Dead,dead,dead | 2.4714e-09 |  | 2.4714e-18 | G8
| J8 | Yes | Yes | No | 133 | 1.79e-08 | 1.1e-08 |  | 0.96606 | 1.1327 |  |  | 9.4112e-09 | 1.5832e-08 | 2.2361e-08 | 0.075 |  | 15.5928 |  | 18.5483 |  | 14.7799 |  | 1.478e-08 | J8
| F9 | Yes | Yes | No | 134 | 1.84e-08 | 2.27e-08 |  | 0.96756 | 79.8057 |  |  | 1.6259e-08 | 2.2406e-08 | 3.0177e-08 | 0.072 |  | 23.0292 |  | 28.4109 | Noisy | 23.2738 |  | 2.3274e-08 | F9
| H9 | Yes | Yes | No | 135 | 0 | 0 | Cryoharness | 0.96777 | 1.1513 |  |  | 9.2627e-09 | 1.6088e-08 | 2.0645e-08 | 0.069 |  | 14.6608 |  | 18.294 |  | 14.4462 |  | 1.4446e-08 | H9
| G9 | Yes | No | No | 136 | 1.73e-08 | 1.04e-08 |  | 0.96959 | 1.6666 |  |  | 9.4105e-09 | 1.5319e-08 | 2.2162e-08 | 0.075 |  | 14.6662 |  | 17.2739 |  | 14.1893 |  | 1.4189e-08 | G9
| J9 | Yes | Yes | No | 137 | 1.66e-08 | 1.03e-08 |  | 0.97171 | 1.5285 |  |  | 8.9807e-09 | 1.5275e-08 | 2.0452e-08 | 0.079 |  | 14.6271 |  | 18.3535 |  | 13.9277 |  | 1.3928e-08 | J9
| F10 | Yes | Yes | No | 138 | 1.69e-08 | 1.04e-08 |  | 0.96984 | 2.3218 |  |  | 9.4164e-09 | 1.5691e-08 | 2.0589e-08 | 0.074 |  | 14.6096 |  | 17.2542 |  | 13.4719 |  | 1.3472e-08 | F10
| H10 | Yes | Yes | No | 139 | 2.46e-08 | 1.49e-08 |  | 0.97208 | 1.1368 |  |  | 1.0378e-08 | 1.6576e-08 | 2.1451e-08 | 0.086 |  | 15.546 |  | 19.1996 |  | 14.4837 |  | 1.4484e-08 | H10
| G10 | Yes | Yes | No | 140 | 1.65e-08 | 1.04e-08 |  | 0.96979 | 1.2051 |  |  | 1.0255e-08 | 1.5608e-08 | 2.2372e-08 | 0.068 |  | 15.0477 |  | 18.7967 |  | 14.865 |  | 1.4865e-08 | G10
| F11 | Yes | Yes | No | 141 | 1.84e-08 | 1.42e-08 |  | 0.96901 | 1.8659 |  |  | 2.6841e-07 | 1.2769e-06 | 1.7419e-06 | 0.26 |  | 1253.4316 |  | 17.3717 | Not noisy now | 13.5554 |  | 1.3555e-08 | F11
| J10 | Yes | Yes | No | 142 | 1.64e-08 | 2.62e-08 |  | 0.96772 | 1.1227 |  |  | 9.9317e-09 | 1.6084e-08 | 2.1505e-08 | 0.066 |  | 15.1116 |  | 19.6576 |  | 14.7765 |  | 1.4776e-08 | J10
| H11 | Yes | Yes | No | 143 | 1.79e-08 | 1.21e-08 |  | 0.97148 | 2.3893 |  |  | 9.1246e-09 | 1.5325e-08 | 2.158e-08 | 0.056 |  | 14.8621 |  | 17.0945 |  | 13.5881 |  | 1.3588e-08 | H11
| G11 | Yes | Yes | No | 144 | 4.31e-08 | 1.17e-08 | BDA LR contact | 0.97137 | 1.549 |  |  | 9.3601e-09 | 1.2477e-08 | 1.0671e-07 | 0.22 |  | 45.078 |  | 67.0612 | V. Noisy - consider dead? | 38.2148 |  | 3.8215e-08 | G11

## SLW
SLW | BoDAC |  |  |  | PFM3 |  |  |  | PFM4 | [9 empty] | PFM4 in nV | IST
| Pixel Name | BDA Pixel Operability | BoDAC channel Operability | Noisy BoDAC channel | Channel # | Noise | Warm Electronic Test | Comment | JFET Gain Test | Warm Resistance test | JFET ID | VSS at 4K | JFET + LIA Noise (nV/rtHz) | Noise (nV/rtHz) [0.1 - 1.5 Hz] | Noise (20061117) (0.1 - 1.5 Hz) | Noise (20061130) (.5-3 Hz) | 0.3 K load curve | Failure mode following cryoharness probe |  | Noise IST | Comment IST
| R1 | Yes | Yes | No | 49 | N/M | N/M | S/N 22/55 | 0.93443 | 5.0078 | 14/37 | -1.6626V | 2.1867e-08 | 1.5089e-08 | 1.6345e-08 | 1.3606e-08 |  |  | 13.6057 | 15.1359
| T1 | Yes | Yes | No | 50 |  |  |  | 0.93399 | 1.0395 | 2.44 |  | 1.1632e-08 | 2.4304e-08 | 3.2006e-08 | 2.1103e-08 |  |  | 21.1027 | 19.36
| C1 | Yes | Yes | No | 51 |  |  | M. Noisy | 0.93409 | 10.2401 | -1.35 |  | 8.8955e-09 | 2.5586e-08 | 2.1503e-08 | 1.9381e-08 |  |  | 19.3807 | 18.9659
| DK1 | Yes | Yes | No | 52 |  |  |  | 0.93277 | 226.1867 | 1.36 mA |  | 9.4992e-09 | 2.5623e-08 | 2.4511e-08 | 3.8262e-08 |  |  | 38.2616 | 22.1603
| B1 | Yes | Yes | No | 53 |  |  |  | 0.93369 | 2.5226 |  |  | 9.3623e-09 | 2.6883e-08 | 2.1825e-08 | 2.1774e-08 |  |  | 21.7736 | 19.1221
| D1 | Yes | Yes | No | 54 |  |  |  | 0.93283 | 1.833 |  |  | 1.0287e-08 | 2.7443e-08 | 2.1484e-08 | 2.1633e-08 |  |  | 21.6333 | 20.5129
| E1 | Yes | Yes | No | 55 |  |  |  | 0.93299 | 1.0583 |  |  | 9.614e-09 | 2.7653e-08 | 2.1339e-08 | 2.1135e-08 |  |  | 21.1346 | 20.7005
| A1 | Yes | Yes | No | 56 |  |  |  | 0.93455 | 1.4901 |  |  | 9.592e-09 | 2.536e-08 | 2.1376e-08 | 1.9688e-08 |  |  | 19.688 | 18.8951
| C2 | Yes | Yes | No | 57 |  |  | No signal | -0.0003507 | 555.3944 |  |  | 1.105e-08 | -3.6784e-05 | -5.2479e-05 | 1.3488e-08 | low gain | Broken S+ | 13.4877 | 23.5984 | OK - BMS to check JFET gain in cal file
| D2 | Yes | Yes | No | 58 |  |  | M. Noisy | 0.93415 | 1.3809 |  |  | 9.6241e-09 | 3.265e-08 | 2.4422e-08 | 2.3801e-08 |  |  | 23.8008 | 22.5343
| B2 | Yes | Yes | No | 59 |  |  |  | 0.93207 | 1.1475 |  |  | 1.005e-08 | 2.8753e-08 | 2.2834e-08 | 2.2208e-08 |  |  | 22.208 | 22.9879
| E2 | Yes | Yes | No | 60 |  |  | V. Noisy | 0.93475 | 2.2448 |  |  | 9.7851e-09 | 3.1559e-08 | 2.2918e-08 | 2.3312e-08 |  |  | 23.312 | 22.6304
| A2 | Yes | Yes | No | 61 |  |  |  | 0.9326 | 0.8889 |  |  | 9.1913e-09 | 2.7021e-08 | 2.1828e-08 | 2.0752e-08 |  |  | 20.752 | 21.0527
| C3 | Yes | Yes | No | 62 |  |  |  | 0.93443 | 1.3029 |  |  | 9.4673e-09 | 2.9858e-08 | 2.3522e-08 | 2.2948e-08 |  |  | 22.9475 | 23.1687
| D3 | Yes | Yes | No | 63 |  |  |  | 0.9322 | 1.6838 |  |  | 9.4454e-09 | 2.7462e-08 | 2.1555e-08 | 2.1062e-08 |  |  | 21.0624 | 22.2985
| B3 | Yes | Yes | No | 64 |  |  | No signal | 0.93259 | 1.1608 |  |  | 1.131e-08 | 1.4369e-08 | 1.9782e-08 | 1.351e-08 | low gain | Broken S+ | 13.5096 | 22.3559 | OK
| E3 | Yes | Yes | No | 65 |  |  |  | 0.93272 | 1.2133 |  |  | 1.0423e-08 | 2.7447e-08 | 2.2992e-08 | 2.1001e-08 |  |  | 21.0014 | 20.9782
| C4 | Yes | Yes | No | 66 |  |  |  | 0.93408 | 0.7004 |  |  | 8.5104e-09 | 2.6657e-08 | 2.2262e-08 | 2.0787e-08 |  |  | 20.7867 | 19.8256
| DK2 | No | Yes | No | 67 |  |  | BDA dead | 0.93428 | 0.5263 |  |  | 0 | 0 | 0 | 0 | dead |  | 0 | 20.6335 | Alive again!
| D4 | Yes | Yes | No | 68 |  |  |  | 0.93528 | 1.0907 |  |  | 8.4972e-09 | 3.1221e-08 | 2.5419e-08 | 2.3789e-08 |  |  | 23.789 | 20.2993
| C5 | Yes | Yes | No | 69 |  |  |  | 0.93272 | 0.7943 |  |  | 8.7436e-09 | 2.616e-08 | 2.8734e-08 | 2.0433e-08 |  |  | 20.4328 | 19.7112
| B4 | Yes | Yes | No | 70 |  |  |  | 0.93486 | 1.1426 |  |  | 9.305e-09 | 2.5565e-08 | 2.0699e-08 | 2.0201e-08 |  |  | 20.2013 | 18.3884
| A3 | Yes | Yes | No | 71 |  |  | M. Noisy | 0.93466 | 1.9874 |  |  | 9.0543e-09 | 2.4715e-08 | 2.0485e-08 | 1.9771e-08 |  |  | 19.7713 | 19.4889
| T2 | Yes | Yes | No | 72 |  |  | M. Noisy | 0.93296 | 1.6449 |  |  | 9.7831e-09 | 2.0044e-08 | 1.9777e-08 | 1.6604e-08 |  |  | 16.604 | 15.0093

## SSW
SSW | BoDAC |  |  |  | PFM3 |  |  |  | PFM4 | [10 empty] | IST
| Pixel Name | BDA Pixel Operability | BoDAC channel Operability | Noisy BoDAC channel | Channel # | Noise | Warm Electronic Test | Comment | JFET Gain Test | Warm Resistance test | JFET ID | VSS at 4K | JFET + LIA Noise (nV/rtHz) | Noise (nV/rtHz) [0.1 - 1.5 Hz] | Noise (20061117) (0.1 - 1.5 Hz) | Noise (20061130) (.5-3 Hz) | 0.3 K load curves | Failure mode following cryoharness probe | PFM4 Noise in nV | IST Noise | Ist Comment
| R1 | Yes | Yes | No | 1 | N/M | N/M | S/N 23/57 | 0.96266 | 4.2239 | 24/59 | -2.05V | 1.9444e-08 | 1.9218e-08 | 2.0675e-08 | 2.0251e-08 |  |  | 20.2513 | 14.3858
| A4 | Yes | Yes | No | 2 |  |  |  | 0.96306 | 1.6656 | 2.48 |  | 9.6463e-09 | 1.7756e-08 | 1.8754e-08 | 1.6501e-08 |  |  | 16.5014 | 17.6181
| A3 | Yes | Yes | No | 3 |  |  | M. Noisy | 0.957 | 9.603 | -1.54 |  | 1.0096e-08 | 2.8109e-08 | 2.5463e-08 | 2.2044e-08 |  |  | 22.0443 | 19.6177
| A2 | Yes | Yes | No | 4 |  |  | V. Noisy | 0.96294 | 0.7057 | 1.20 mA |  | 9.1509e-09 | 1.9316e-08 | 2.0196e-08 | 1.637e-08 |  |  | 16.3704 | 16.9281
| A1 | Yes | Yes | No | 5 |  |  |  | 0.96208 | 1.5043 |  |  | 8.892e-09 | 1.9125e-08 | 2.1685e-08 | 1.5939e-08 |  |  | 15.939 | 16.3409
| DK1 | Yes | Yes | Yes at 0.1 | 6 |  |  |  | 0.96208 | 2.4131 |  |  | 8.3298e-09 | 1.8086e-08 | 1.7844e-08 | 1.5392e-08 |  |  | 15.3918 | 15.4154
| B3 | Yes | Yes | Yes at 0.1 | 7 |  |  |  | 0.95707 | 1.3574 |  |  | 1.0205e-08 | 2.3927e-08 | 2.4981e-08 | 1.8893e-08 |  |  | 18.8931 | 18.9726
| B2 | Yes | Yes | No | 8 |  |  |  | 0.96166 | 3.4784 |  |  | 9.2052e-09 | 1.9238e-08 | 1.8748e-08 | 1.696e-08 |  |  | 16.9596 | 17.3885
| B1 | Yes | Yes | No | 9 |  |  |  | 0.96261 | 1.0075 |  |  | 1.0475e-08 | 1.8907e-08 | 2.0196e-08 | 1.8261e-08 |  |  | 18.2609 | 17.41
| C3 | Yes | Yes | No | 10 |  |  |  | 0.96243 | 1.683 |  |  | 8.9084e-09 | 1.9326e-08 | 1.8961e-08 | 1.6161e-08 |  |  | 16.1608 | 17.8953
| C2 | Yes | Yes | No | 11 |  |  |  | 0.96251 | 3.3469 |  |  | 8.6452e-09 | 2.0779e-08 | 1.8888e-08 | 1.6514e-08 |  |  | 16.5143 | 17.7169
| C1 | Yes | Yes | No | 12 |  |  |  | 0.96241 | 1.8285 |  |  | 8.9374e-09 | 1.8287e-08 | 1.9056e-08 | 1.7025e-08 |  |  | 17.0246 | 16.8197
| D3 | Yes | Yes | Yes at 0.1 | 13 |  |  |  | 0.96281 | 1.6672 |  |  | 9.9276e-09 | 1.8591e-08 | 1.9789e-08 | 1.5954e-08 |  |  | 15.9538 | 16.9569
| D2 | Yes | Yes | No | 14 |  |  |  | 0.96223 | 0.9686 |  |  | 1.1482e-08 | 1.9018e-08 | 1.9961e-08 | 1.7912e-08 |  |  | 17.9115 | 18.4584
| D1 | Yes | Yes | Yes at 0.1 | 15 |  |  |  | 0.96072 | 0.8705 |  |  | 8.3388e-09 | 1.7799e-08 | 1.8931e-08 | 1.6043e-08 |  |  | 16.0429 | 16.4283
| E3 | Yes | Yes | No | 16 |  |  | M. Noisy | 0.96207 | 1.7479 |  |  | 9.7619e-09 | 1.8606e-08 | 1.8375e-08 | 1.7007e-08 |  |  | 17.0066 | 19.3871
| E2 | Yes | Yes | No | 17 |  |  |  | 0.96215 | 0.6655 |  |  | 9.2201e-09 | 1.798e-08 | 1.8125e-08 | 1.586e-08 |  |  | 15.8605 | 17.2477
| E1 | Yes | Yes | Yes at 0.1 | 18 |  |  |  | 0.9625 | 1.0706 |  |  | 8.9118e-09 | 2.0675e-08 | 1.9369e-08 | 1.6196e-08 |  |  | 16.1965 | 16.8488
| F3 | Yes | Yes | No | 19 |  |  |  | 0.96214 | 1.5442 |  |  | 8.61e-09 | 1.9956e-08 | 2.2962e-08 | 1.8486e-08 |  |  | 18.4863 | 19.8506
| F2 | Yes | Yes | No | 20 |  |  |  | 0.95269 | 2.3713 |  |  | 8.4126e-09 | 2.0468e-08 | 2.0712e-08 | 1.7513e-08 |  |  | 17.5127 | 17.7232
| F1 | Yes | Yes | No | 21 |  |  |  | 0.96193 | 0.8822 |  |  | 1.0501e-08 | 1.8193e-08 | 2.0131e-08 | 1.7088e-08 |  |  | 17.0879 | 17.8639
| G1 | Yes | Yes | No | 22 |  |  |  | 0.96214 | 0.9129 |  |  | 9.1519e-09 | 2.0267e-08 | 2.0986e-08 | 1.8695e-08 |  |  | 18.6948 | 18.521
| T1 | Yes | Yes | No | 23 |  |  |  | 0.94182 | 1.2904 |  |  | 8.9638e-09 | 1.9218e-08 | 1.9645e-08 | 1.5875e-08 |  |  | 15.8748 | 16.0277
| G2 | Yes | Yes | No | 24 |  |  |  | 0.93725 | 1.9352 |  |  | 8.657e-09 | 1.9205e-08 | 1.935e-08 | 1.6352e-08 |  |  | 16.3524 | 17.2845
| E5 | Yes | Yes | No | 25 |  |  | S/N 23/54, Not functional | 0.96257 | 2.493 | 24/60 | -1.58V | 1.1121e-08 | 1.309e-07 | 9.1986e-08 | 2.6891e-08 |  |  | 26.8911 | 18.7373
| E4 | Yes | Yes | No | 26 |  |  | Not functional | 0.96034 | 3.1124 | 2.47 |  | 8.7482e-09 | 1.4682e-07 | 9.1628e-08 | 2.3942e-08 |  |  | 23.9418 | 16.6343
| D7 | Yes | Yes | No | 27 |  |  | Not functional | 0.96181 | 382.9852 | -1.26 |  | 8.298e-09 | 9.95e-09 | 1.1068e-08 | 9.3393e-09 | low gain | Bad contact S+ | 9.3393 | 17.6863 | OK Operating normalling
| D6 | Yes | No | No | 28 |  |  | Not functional | 0.95986 | 4.8226 | 1.25 mA |  | 9.8618e-09 | 1.5836e-07 | 9.4436e-08 | 2.5007e-08 |  |  | 25.0072 | 17.9194
| D5 | No | Yes | No | 29 |  |  | Not functional | 0.9615 | 10264.2501 |  |  | 4.5688e-08 | 1.6849e-07 | 1.615e-07 | 2.8749e-08 | dead |  | 28.749 | 21.3467 | Confirmed dead
| D4 | Yes | Yes | No | 30 |  |  | Not functional | 0.96075 | 3.0965 |  |  | 8.7837e-09 | 9.6071e-08 | 6.393e-08 | 1.9375e-08 |  |  | 19.3752 | 17.136
| C6 | Yes | Yes | No | 31 |  |  | Not functional | 0.96252 | 27.5108 |  |  | 1.0948e-08 | 9.9634e-08 | 6.417e-08 | 2.0316e-08 |  |  | 20.3164 | 18.118
| C5 | Yes | Yes | No | 32 |  |  | Not functional | 0.96194 | 3.4274 |  |  | 1.0136e-08 | 5.5409e-08 | 4.3035e-08 | 2.1146e-08 |  |  | 21.1456 | 17.8259
| C4 | Yes | Yes | No | 33 |  |  | Not functional | 0.96199 | 4.2184 |  |  | 9.6134e-09 | 6.5905e-08 | 4.5507e-08 | 1.8266e-08 |  |  | 18.2662 | 17.2898
| B5 | Yes | Yes | No | 34 |  |  | Not functional | 0.96201 | 2.607 |  |  | 9.7548e-09 | 3.3056e-08 | 3.1892e-08 | 2.073e-08 |  |  | 20.7301 | 17.2363
| B4 | Yes | Yes | No | 35 |  |  | Not functional | 0.96213 | 1.9568 |  |  | 8.4901e-09 | 3.8872e-08 | 3.1668e-08 | 1.7245e-08 |  |  | 17.2446 | 16.8291
| T2 | Yes | Yes | No | 36 |  |  | Not functional | 0.96213 | 2.5441 |  |  | 9.0072e-09 | 2.588e-08 | 2.3984e-08 | 1.7476e-08 |  |  | 17.4757 | 15.8637
| G3 | Yes | Yes | No | 37 |  |  | Not functional | 0.9623 | 2.4624 |  |  | 9.4832e-09 | 4.0112e-08 | 3.2899e-08 | 1.8276e-08 |  |  | 18.2757 | 15.0757
| G4 | Yes | Yes | No | 38 |  |  | Not functional | 0.9487 | 1.0912 |  |  | 9.6267e-09 | 2.8987e-08 | 2.3414e-08 | 1.6437e-08 |  |  | 16.4373 | 15.5546
| DK2 | Yes | Yes | No | 39 |  |  | Not functional | 0.96147 | 1.2539 |  |  | 8.9074e-09 | 7.2701e-08 | 4.8314e-08 | 2.0241e-08 |  |  | 20.2411 | 16.4179
| F5 | Yes | Yes | No | 40 |  |  | Not functional | 0.96208 | 1.9975 |  |  | 8.9286e-09 | 5.987e-08 | 4.8961e-08 | 2.1681e-08 |  |  | 21.6809 | 17.7783
| F4 | Noisy | Yes | No | 41 |  |  | Not functional | 0.95966 | 2.1675 |  |  | 1.3978e-08 | 9.2429e-08 | 1.2151e-07 | 4.9957e-08 |  |  | 49.9569 | 17.2232 | Intermitently very noisy - assume dead
| E6 | Yes | Yes | No | 42 |  |  | Not functional | 0.96191 | 1.5898 |  |  | 9.8563e-09 | 1.154e-07 | 6.7069e-08 | 2.2176e-08 |  |  | 22.1755 | 16.6353

## PTC
PTC | BoDAC |  |  | PFM3 |  |  |  | PFM4
| Pixel Name | BDA Operability | Noisy Bodac? | Channel # | PFM3 Noise | PFM3 WE Noise | JJB Comment | JFET Gain Test | Warm Resistance test | JFET Gain Test | Warm Resistance test | JFET ID | VSS at 4K | JFET + LIA Noise (nV/rtHz) | Noise (nV/rtHz) (1-3Hz) |  | IST Noise | IST Comment
| PTC1 | N/A | N/A | 286 | ? | ? | S/N 22/56 | 0.95711 | 4.1294 | ? | ? | 14/36 | ? | ? | 1.7182e-08 | 17.1825 | 35.187 | Why so noisy? Spike in frequency plot - temp/EMI issue? Not so noisy at other frequencies.
| PTC2 | N/A | N/A | 287 | ? | ? |  | 0.96005 | 2.7689 | ? | ? | 2.5 |  | ? | 1.5356e-08 | 15.3565 | 34.2282 | Why so noisy? Spike in frequency plot - temp/EMI issue? Not so noisy at other frequencies.
| PTC3 | N/A | N/A | 288 | ? | ? | Wiring swapped | 0.958 | 23.586 | ? | ? | -1 |  | ? | 1.3082e-08 | 13.0821 | 38.59 | Why so noisy? Spike in frequency plot - temp/EMI issue? Not so noisy at other frequencies.
1.14 mA

## Appendix1 TW SoR Summary
Analysis by Tim Waskett
Nominal tau for healthy detectors = 5ms
Tau measured by best fit model to EXTERNAL PCal test 0x3001188B, a test perfromed towards the end of the PFM4 test campaign
Model fits are performed by eye and have an approximate (estimated) error of +/- 5 ms
PSW |  |  |  | PMW |  |  |  | PLW
detector | name | ~tau (ms) | comment | detector | name | ~tau (ms) | comment | detector | name | ~tau (ms) | comment
11 | A14 | 35 | bad model fit | 193 | A13 |  | increadibly slow | 148 | A6 | 5 | noisy
12 | A13 |  | increadibly slow | 215 | C8 |  | dead | 166 | A1 | 20
18 | C12 |  | dead | 221 | D1 | 15 |  | 168 | A2 | 10 | noisy
22 | A11 |  | increadibly slow | 241 | F10 | 13 |  | 182 | C6 | 20 | coaligned centre
41 | J4 | 12 | coaligned at edge | 243 | G11 |  | no response to PCal | 189 | E6 | 10
54 | B5 | 13 |  | 246 | G12 | 15
61 | B3 | 5 | noisy | 252 | F7 |  | dead
98 | A10 |  | dead | 257 | G9 |  | no response to PCal
109 | C8 | 13 | noisy | 258 | D9 | 15 | slightly odd model fit
123 | J6 |  | no response to PCal, and coaligned at edge | 276 | B6 |  | no response to PCal
125 | G6 | 10 |  | 282 | B4 | 13
131 | F8 |  | no response to PCal
132 | G8 |  | dead
141 | F11 |  | extremely noisy
|  |  |  | 139 | 88 | 43 | 270
|  |  |  | PSW | PMW | PLW | total
|  |  | Working but unusable detectors | 5 | 4 | 0 | 9
|  |  | Dead detectors | 3 | 2 | 0 | 5
|  |  | Usable but slightly slow/noisy detectors | 6 | 5 | 5 | 16
|  |  | total unusable | 8 | 6 | 0 | 14
|  |  | total bad | 14 | 11 | 5 | 30
|  |  | yield | 0.94245 | 0.93182 | 1 | 0.94815

## PHOT IST Noise
Noise data taken from EMC testing spreadsheet "0xB00014D8 CM PHOT averaged"
PLWR1 | 16.5533 | R1 | 1
PLWA8 | 19.8551 | A8 | 2
PLWA7 | 19.5299 | A7 | 3
PLWA6 | 146.9956 | A6 | 4
PLWA9 | 20.6958 | A9 | 5
PLWC9 | 34.6323 | C9 | 6
PLWB8 | 18.9861 | B8 | 7
PLWB7 | 18.6787 | B7 | 8
PLWC7 | 19.5578 | C7 | 9
PLWB5 | 20.3428 | B5 | 10
PLWB6 | 18.3606 | B6 | 11
PLWA5 | 19.3644 | A5 | 12
PLWT1 | 16.7244 | T1 | 13
PLWB4 | 19.8324 | B4 | 14
PLWC4 | 18.169 | C4 | 15
PLWB3 | 17.7437 | B3 | 16
PLWC2 | 18.3857 | C2 | 17
PLWB2 | 18.875 | B2 | 18
PLWB1 | 19.3054 | B1 | 19
PLWA3 | 19.3835 | A3 | 20
PLWA4 | 18.9284 | A4 | 21
PLWA1 | 18.9725 | A1 | 22
PLWDP1 | 17.2603 | DK1 | 23
PLWA2 | 19.0737 | A2 | 24
PLWE1 | 18.297 | E1 | 25
PLWE2 | 18.8439 | E2 | 26
PLWE3 | 17.6851 | E3 | 27
PLWE4 | 18.6352 | E4 | 28
PLWD1 | 20.5747 | D1 | 29
PLWD2 | 18.1078 | D2 | 30
PLWD3 | 18.1144 | D3 | 31
PLWD4 | 17.779 | D4 | 32
PLWC1 | 18.618 | C1 | 33
PLWC3 | 19.2416 | C3 | 34
PLWC5 | 16.7837 | C5 | 35
PLWT2 | 17.3871 | T2 | 36
PLWE5 | 19.3927 | E5 | 37
PLWC6 | 16.7036 | C6 | 38
PLWC8 | 18.5199 | C8 | 39
PLWD5 | 17.8096 | D5 | 40
PLWD6 | 17.5551 | D6 | 41
PLWD7 | 18.1095 | D7 | 42
PLWD8 | 18.8735 | D8 | 43
PLWE7 | 19.6567 | E7 | 44
PLWE6 | 18.2031 | E6 | 45
PLWE8 | 20.0183 | E8 | 46
PLWDP2 | 17.3849 | DK2 | 47
PLWE9 | 19.2802 | E9 | 48
PMWA13 | 15.3366 | A13 | 1
PMWT1 | 15.8284 | T1 | 2
PMWB12 | 17.5796 | B12 | 3
PMWC13 | 17.5799 | C13 | 4
PMWA12 | 17.6199 | A12 | 5
PMWD12 | 17.6612 | D12 | 6
PMWC12 | 18.2685 | C12 | 7
PMWB11 | 26.6418 | B11 | 8
PMWA11 | 18.2672 | A11 | 9
PMWE13 | 17.5116 | E13 | 10
PMWD11 | 16.6783 | D11 | 11
PMWC11 | 17.4866 | C11 | 12
PMWB10 | 17.8189 | B10 | 13
PMWA10 | 18.2225 | A10 | 14
PMWD10 | 16.6671 | D10 | 15
PMWB9 | 17.0989 | B9 | 16
PMWC10 | 16.5605 | C10 | 17
PMWC9 | 16.3943 | C9 | 18
PMWA9 | 18.5163 | A9 | 19
PMWB8 | 17.6049 | B8 | 20
PMWA8 | 18.2412 | A8 | 21
PMWD8 | 18.5519 | D8 | 22
PMWC8 | 16.6714 | C8 | 23
PMWB7 | 17.0859 | B7 | 24
PMWR1 | 15.5379 | R1 | 25
PMWG1 | 18.4671 | G1 | 26
PMWT2 | 0 | T2 | 27
PMWE1 | 17.7103 | E1 | 28
PMWD1 | 21.6247 | D1 | 29
PMWF1 | 18.715 | F1 | 30
PMWE2 | 17.8475 | E2 | 31
PMWG2 | 18.1211 | G2 | 32
PMWF2 | 18.9046 | F2 | 33
PMWG3 | 18.4014 | G3 | 34
PMWE3 | 17.8716 | E3 | 35
PMWD3 | 18.2436 | D3 | 36
PMWF3 | 18.7089 | F3 | 37
PMWG4 | 19.2449 | G4 | 38
PMWE4 | 18.8585 | E4 | 39
PMWF4 | 17.4071 | F4 | 40
PMWE5 | 16.3427 | E5 | 41
PMWD5 | 17.2038 | D5 | 42
PMWF5 | 17.2439 | F5 | 43
PMWG5 | 18.8312 | G5 | 44
PMWE6 | 16.0969 | E6 | 45
PMWG6 | 17.2882 | G6 | 46
PMWF6 | 16.6313 | F6 | 47
PMWG7 | 17.7392 | G7 | 48
PMWF10 | 18.4097 | F10 | 49
PMWE11 | 16.6245 | E11 | 50
PMWG11 | 18.9279 | G11 | 51
PMWF11 | 16.5581 | F11 | 52
PMWE12 | 18.2909 | E12 | 53
PMWG12 | 18.2116 | G12 | 54
PMWF12 | 18.9613 | F12 | 55
PMWG13 | 18.8147 | G13 | 56
PMWDP2 | 16.0724 | DK2 | 57
PMWE7 | 16.4642 | E7 | 58
PMWD7 | 18.0371 | D7 | 59
PMWF7 | 18.3812 | F7 | 60
PMWE8 | 37.652 | E8 | 61
PMWG8 | 18.312 | G8 | 62
PMWF8 | 17.3154 | F8 | 63
PMWE9 | 17.6141 | E9 | 64
PMWG9 | 17.5715 | G9 | 65
PMWD9 | 18.4658 | D9 | 66
PMWF9 | 16.9365 | F9 | 67
PMWE10 | 17.4365 | E10 | 68
PMWG10 | 17.9051 | G10 | 69
PMWC4 | 16.8343 | C4 | 70
PMWB3 | 17.8053 | B3 | 71
PMWC3 | 17.5154 | C3 | 72
PMWB2 | 19.6467 | B2 | 73
PMWD2 | 19.8122 | D2 | 74
PMWA3 | 17.0153 | A3 | 75
PMWA2 | 18.5301 | A2 | 76
PMWC2 | 18.2648 | C2 | 77
PMWB1 | 18.4413 | B1 | 78
PMWA1 | 18.4841 | A1 | 79
PMWDP1 | 17.7759 | DK1 | 80
PMWC1 | 19.6251 | C1 | 81
PMWA7 | 17.9263 | A7 | 82
PMWA6 | 18.4283 | A6 | 83
PMWB6 | 16.9016 | B6 | 84
PMWC7 | 17.4559 | C7 | 85
PMWA5 | 16.9925 | A5 | 86
PMWB5 | 17.0189 | B5 | 87
PMWC6 | 17.9187 | C6 | 88
PMWD6 | 25.3917 | D6 | 89
PMWB4 | 17.2591 | B4 | 90
PMWC5 | 16.7486 | C5 | 91
PMWD4 | 17.5596 | D4 | 92
PMWA4 | 17.6153 | A4 | 93
PMWP1 | 35.187
PMWP2 | 34.2282
PMWP3 | 38.59
PSWR1 | 14.6144 | R1 | 1
PSWD16 | 17.9191 | D16 | 2
PSWT1 | 18.151 | T1 | 3
PSWB16 | 20.4435 | B16 | 4
PSWC15 | 19.0197 | C15 | 5
PSWA15 | 18.3817 | A15 | 6
PSWD15 | 0 | D15 | 7
PSWB15 | 18.3661 | B15 | 8
PSWC14 | 17.7636 | C14 | 9
PSWD14 | 19.4708 | D14 | 10
PSWA14 | 17.4492 | A14 | 11
PSWA13 | 14.8541 | A13 | 12
PSWB14 | 18.495 | B14 | 13
PSWC13 | 17.2129 | C13 | 14
PSWB13 | 17.7721 | B13 | 15
PSWD13 | 17.3714 | D13 | 16
PSWA12 | 18.7245 | A12 | 17
PSWC12 | 0 | C12 | 18
PSWD12 | 17.5747 | D12 | 19
PSWB12 | 22.5335 | B12 | 20
PSWE11 | 17.9606 | E11 | 21
PSWA11 | 14.9763 | A11 | 22
PSWC11 | 16.6092 | C11 | 23
PSWB11 | 17.4189 | B11 | 24
PSWE1 | 19.1457 | E1 | 25
PSWF1 | 18.0846 | F1 | 26
PSWT2 | 17.9469 | T2 | 27
PSWH1 | 18.3874 | H1 | 28
PSWG1 | 18.0168 | G1 | 29
PSWJ1 | 19.2953 | J1 | 30
PSWH2 | 19.4708 | H2 | 31
PSWF2 | 17.0425 | F2 | 32
PSWJ2 | 20.7258 | J2 | 33
PSWG2 | 18.7815 | G2 | 34
PSWH3 | 17.8571 | H3 | 35
PSWJ3 | 19.4505 | J3 | 36
PSWE2 | 18.2199 | E2 | 37
PSWF3 | 19.3012 | F3 | 38
PSWG3 | 19.1372 | G3 | 39
PSWH4 | 18.3719 | H4 | 40
PSWJ4 | 19.2171 | J4 | 41
PSWE3 | 18.8215 | E3 | 42
PSWF4 | 18.4391 | F4 | 43
PSWG4 | 19.0888 | G4 | 44
PSWH5 | 18.116 | H5 | 45
PSWE4 | 17.7991 | E4 | 46
PSWJ5 | 19.1424 | J5 | 47
PSWF5 | 17.4429 | F5 | 48
PSWD6 | 17.6008 | D6 | 49
PSWB6 | 17.1279 | B6 | 50
PSWC5 | 17.8089 | C5 | 51
PSWA5 | 18.4363 | A5 | 52
PSWE5 | 18.5156 | E5 | 53
PSWB5 | 19.1427 | B5 | 54
PSWD5 | 17.8036 | D5 | 55
PSWC4 | 17.1591 | C4 | 56
PSWA4 | 19.6326 | A4 | 57
PSWD4 | 16.7747 | D4 | 58
PSWB4 | 17.9742 | B4 | 59
PSWC3 | 18.5961 | C3 | 60
PSWB3 | 18.8566 | B3 | 61
PSWA3 | 19.0415 | A3 | 62
PSWA2 | 18.008 | A2 | 63
PSWD3 | 18.1106 | D3 | 64
PSWC2 | 20.9006 | C2 | 65
PSWB2 | 17.8198 | B2 | 66
PSWD2 | 18.9186 | D2 | 67
PSWA1 | 19.9083 | A1 | 68
PSWC1 | 18.0392 | C1 | 69
PSWB1 | 18.2327 | B1 | 70
PSWDP1 | 16.9146 | DK1 | 71
PSWD1 | 17.9938 | D1 | 72
PSWF12 | 18.1035 | F12 | 73
PSWJ11 | 0 | J11 | 74
PSWE12 | 18.7223 | E12 | 75
PSWH12 | 17.6688 | H12 | 76
PSWG12 | 17.6959 | G12 | 77
PSWF13 | 17.7227 | F13 | 78
PSWE13 | 17.7157 | E13 | 79
PSWJ12 | 19.2014 | J12 | 80
PSWH13 | 19.1177 | H13 | 81
PSWG13 | 18.9494 | G13 | 82
PSWF14 | 18.813 | F14 | 83
PSWE14 | 19.7015 | E14 | 84
PSWJ13 | 19.1798 | J13 | 85
PSWH14 | 17.8515 | H14 | 86
PSWG14 | 18.3442 | G14 | 87
PSWJ14 | 19.8004 | J14 | 88
PSWF15 | 18.0392 | F15 | 89
PSWH15 | 18.0708 | H15 | 90
PSWJ15 | 18.9536 | J15 | 91
PSWG15 | 18.294 | G15 | 92
PSWH16 | 18.6452 | H16 | 93
PSWDP2 | 16.9948 | DK2 | 94
PSWF16 | 18.7971 | F16 | 95
PSWE15 | 17.4764 | E15 | 96
PSWD11 | 17.3652 | D11 | 97
PSWA10 | 16.3034 | A10 | 98
PSWE10 | 18.8211 | E10 | 99
PSWC10 | 17.7348 | C10 | 100
PSWB10 | 18.4389 | B10 | 101
PSWD10 | 17.666 | D10 | 102
PSWA9 | 18.6481 | A9 | 103
PSWE9 | 17.4836 | E9 | 104
PSWC9 | 18.7682 | C9 | 105
PSWB9 | 18.3283 | B9 | 106
PSWD9 | 18.6016 | D9 | 107
PSWA8 | 19.1003 | A8 | 108
PSWC8 | 18.0383 | C8 | 109
PSWE8 | 19.5588 | E8 | 110
PSWD8 | 17.3427 | D8 | 111
PSWB8 | 18.6553 | B8 | 112
PSWC7 | 18.3889 | C7 | 113
PSWE7 | 19.5222 | E7 | 114
PSWA7 | 20.3506 | A7 | 115
PSWD7 | 17.2611 | D7 | 116
PSWB7 | 18.1927 | B7 | 117
PSWC6 | 17.3791 | C6 | 118
PSWE6 | 17.8514 | E6 | 119
PSWA6 | 18.0598 | A6 | 120
PSWG5 | 17.2465 | G5 | 121
PSWH6 | 18.0421 | H6 | 122
PSWJ6 | 18.5429 | J6 | 123
PSWF6 | 17.4884 | F6 | 124
PSWG6 | 17.5656 | G6 | 125
PSWH7 | 17.1018 | H7 | 126
PSWF7 | 17.7167 | F7 | 127
PSWJ7 | 18.2396 | J7 | 128
PSWG7 | 17.5143 | G7 | 129
PSWH8 | 17.5192 | H8 | 130
PSWF8 | 18.1729 | F8 | 131
PSWG8 | 0 | G8 | 132
PSWJ8 | 18.5483 | J8 | 133
PSWF9 | 28.4109 | F9 | 134
PSWH9 | 18.294 | H9 | 135
PSWG9 | 17.2739 | G9 | 136
PSWJ9 | 18.3535 | J9 | 137
PSWF10 | 17.2542 | F10 | 138
PSWH10 | 19.1996 | H10 | 139
PSWG10 | 18.7967 | G10 | 140
PSWF11 | 17.3717 | F11 | 141
PSWJ10 | 19.6576 | J10 | 142
PSWH11 | 17.0945 | H11 | 143
PSWG11 | 67.0612 | G11 | 144

## SPEC IST Noise
Reference Noise data taken from EMC testing spreadsheet "0xB00014C7 DM SPECT averaged"
SSWR1 | 14.3858 | R1
SSWA4 | 17.6181 | A4
SSWA3 | 19.6177 | A3
SSWA2 | 16.9281 | A2
SSWA1 | 16.3409 | A1
SSWDP1 | 15.4154 | DK1
SSWB3 | 18.9726 | B3
SSWB2 | 17.3885 | B2
SSWB1 | 17.41 | B1
SSWC3 | 17.8953 | C3
SSWC2 | 17.7169 | C2
SSWC1 | 16.8197 | C1
SSWD3 | 16.9569 | D3
SSWD2 | 18.4584 | D2
SSWD1 | 16.4283 | D1
SSWE3 | 19.3871 | E3
SSWE2 | 17.2477 | E2
SSWE1 | 16.8488 | E1
SSWF3 | 19.8506 | F3
SSWF2 | 17.7232 | F2
SSWF1 | 17.8639 | F1
SSWG1 | 18.521 | G1
SSWT1 | 16.0277 | T1
SSWG2 | 17.2845 | G2
SSWE5 | 18.7373 | E5
SSWE4 | 16.6343 | E4
SSWD7 | 17.6863 | D7
SSWD6 | 17.9194 | D6
SSWD5 | 21.3467 | D5
SSWD4 | 17.136 | D4
SSWC6 | 18.118 | C6
SSWC5 | 17.8259 | C5
SSWC4 | 17.2898 | C4
SSWB5 | 17.2363 | B5
SSWB4 | 16.8291 | B4
SSWT2 | 15.8637 | T2
SSWG3 | 15.0757 | G3
SSWG4 | 15.5546 | G4
SSWDP2 | 16.4179 | DK2
SSWF5 | 17.7783 | F5
SSWF4 | 17.2232 | F4
SSWE6 | 16.6353 | E6
SSWN1 | 18.7945
SSWN2 | 19.1434
SSWN3 | 19.9976
SSWN4 | 19.0314
SSWN5 | 19.5289
SSWN6 | 18.2419
SLWR1 | 15.1359 | R1 | 1
SLWT1 | 19.36 | T1 | 2
SLWC1 | 18.9659 | C1 | 3
SLWDP1 | 22.1603 | DK1 | 4
SLWB1 | 19.1221 | B1 | 5
SLWD1 | 20.5129 | D1 | 6
SLWE1 | 20.7005 | E1 | 7
SLWA1 | 18.8951 | A1 | 8
SLWC2 | 23.5984 | C2 | 9
SLWD2 | 22.5343 | D2 | 10
SLWB2 | 22.9879 | B2 | 11
SLWE2 | 22.6304 | E2 | 12
SLWA2 | 21.0527 | A2 | 13
SLWC3 | 23.1687 | C3 | 14
SLWD3 | 22.2985 | D3 | 15
SLWB3 | 22.3559 | B3 | 16
SLWE3 | 20.9782 | E3 | 17
SLWC4 | 19.8256 | C4 | 18
SLWDP2 | 20.6335 | DK2 | 19
SLWD4 | 20.2993 | D4 | 20
SLWC5 | 19.7112 | C5 | 21
SLWB4 | 18.3884 | B4 | 22
SLWA3 | 19.4889 | A3 | 23
SLWT2 | 15.0093 | T2 | 24

## Sheet1
# White noise | and 1/f knee | table for obs | id=50000145, | bbid=b6b8000a
\BIASPSW = 0.0 | 3007362
\BIASPMW = 0.0 | 3026472
\BIASPLW = 0.0 | 3039036
\BIASFREQ = 13 | 1.0822 | 67
\SUBKTEMP = 0. | 28990236
channel | chName | plateau | fknee1 | fknee2 | fknee | simpleNoise
None | None | V/Hz^(1/2) | Hz | Hz | Hz | V/Hz^(1/2)
45 | PLWR1 | 1.4942e-08 | 0 | 0 | 0 | 1.173e-08 | R1 | 1
7 | PLWA8 | 2.1483e-08 | 0.37941 | 0.040803 | 0.040803 | 1.613e-08 | A8 | 2
6 | PLWA7 | 1.8557e-08 | 0 | 0 | 0 | 1.4449e-08 | A7 | 3
5 | PLWA6 | 1.1047e-07 | 0.34946 | 0.18289 | 0.18289 | 1.1499e-07 | A6 | 4
8 | PLWA9 | 2.0421e-08 | 0 | 0 | 0 | 1.5644e-08 | A9 | 5
25 | PLWC9 | 2.3747e-08 | 0.49923 | 0.053823 | 0.053823 | 1.9715e-08 | C9 | 6
16 | PLWB8 | 2.0471e-08 | 0.099845 | 0.041167 | 0.041167 | 1.6017e-08 | B8 | 7
15 | PLWB7 | 1.9565e-08 | 0 | 0 | 0 | 1.4886e-08 | B7 | 8
23 | PLWC7 | 2.0527e-08 | 0.16974 | 0.054076 | 0.054076 | 1.5459e-08 | C7 | 9
13 | PLWB5 | 2.2756e-08 | 0 | 0 | 0 | 1.7303e-08 | B5 | 10
14 | PLWB6 | 1.8901e-08 | 0 | 0 | 0 | 1.4381e-08 | B6 | 11
4 | PLWA5 | 2.0206e-08 | 0 | 0 | 0 | 1.5531e-08 | A5 | 12
46 | PLWT1 | 3.1334e-08 | 0 | 0 | 0 | 2.2725e-08 | T1 | 13
12 | PLWB4 | 2.1024e-08 | 0 | 0 | 0 | 1.5387e-08 | B4 | 14
20 | PLWC4 | 2.0178e-08 | 0 | 0 | 0 | 1.4884e-08 | C4 | 15
11 | PLWB3 | 1.9885e-08 | 0.29954 | 0.041414 | 0.041414 | 1.4945e-08 | B3 | 16
18 | PLWC2 | 1.9745e-08 | 0 | 0 | 0 | 1.5598e-08 | C2 | 17
10 | PLWB2 | 1.8458e-08 | 0 | 0 | 0 | 1.4528e-08 | B2 | 18
9 | PLWB1 | 1.9608e-08 | 0 | 0 | 0 | 1.4724e-08 | B1 | 19
2 | PLWA3 | 1.9606e-08 | 0 | 0 | 0 | 1.5191e-08 | A3 | 20
3 | PLWA4 | 2.0062e-08 | 0 | 0 | 0 | 1.5276e-08 | A4 | 21
0 | PLWA1 | 2.0989e-08 | 0 | 0 | 0 | 1.5567e-08 | A1 | 22
34 | PLWDP1 | 1.9566e-08 | 0 | 0 | 0 | 1.4844e-08 | DK1 | 23
1 | PLWA2 | 1.9263e-08 | 0.48924 | 0.041176 | 0.041176 | 1.51e-08 | A2 | 24
36 | PLWE1 | 2.0592e-08 | 0 | 0 | 0 | 1.5345e-08 | E1 | 25
37 | PLWE2 | 2.0826e-08 | 0 | 0 | 0 | 1.5105e-08 | E2 | 26
38 | PLWE3 | 1.8454e-08 | 0 | 0 | 0 | 1.3837e-08 | E3 | 27
39 | PLWE4 | 1.9161e-08 | 0 | 0 | 0 | 1.4418e-08 | E4 | 28
26 | PLWD1 | 2.1309e-08 | 0 | 0 | 0 | 1.6381e-08 | D1 | 29
27 | PLWD2 | 2.053e-08 | 0 | 0 | 0 | 1.5591e-08 | D2 | 30
28 | PLWD3 | 1.9639e-08 | 0 | 0 | 0 | 1.5031e-08 | D3 | 31
29 | PLWD4 | 1.8779e-08 | 0 | 0 | 0 | 1.4213e-08 | D4 | 32
17 | PLWC1 | 1.8612e-08 | 0.45929 | 0.036024 | 0.036024 | 1.4531e-08 | C1 | 33
19 | PLWC3 | 1.9843e-08 | 0 | 0 | 0 | 1.5008e-08 | C3 | 34
21 | PLWC5 | 1.7216e-08 | 0.43932 | 0.041181 | 0.041181 | 1.3464e-08 | C5 | 35
47 | PLWT2 | 2.6609e-08 | 0 | 0 | 0 | 1.8943e-08 | T2 | 36
40 | PLWE5 | 2.0544e-08 | 0 | 0 | 0 | 1.5572e-08 | E5 | 37
22 | PLWC6 | 1.7797e-08 | 0 | 0 | 0 | 1.3684e-08 | C6 | 38
24 | PLWC8 | 1.9323e-08 | 0.37941 | 0.047088 | 0.047088 | 1.4562e-08 | C8 | 39
30 | PLWD5 | 1.9356e-08 | 0 | 0 | 0 | 1.4676e-08 | D5 | 40
31 | PLWD6 | 1.7956e-08 | 0 | 0 | 0 | 1.411e-08 | D6 | 41
32 | PLWD7 | 1.9795e-08 | 0 | 0 | 0 | 1.5101e-08 | D7 | 42
33 | PLWD8 | 1.8184e-08 | 0 | 0 | 0 | 1.438e-08 | D8 | 43
42 | PLWE7 | 2.0307e-08 | 0 | 0 | 0 | 1.5468e-08 | E7 | 44
41 | PLWE6 | 2.02e-08 | 0 | 0 | 0 | 1.5242e-08 | E6 | 45
43 | PLWE8 | 1.8956e-08 | 0.3894 | 0.04295 | 0.04295 | 1.4372e-08 | E8 | 46
35 | PLWDP2 | 1.7529e-08 | 0 | 0 | 0 | 1.3638e-08 | DK2 | 47
44 | PLWE9 | 1.9346e-08 | 0.24961 | 0.058232 | 0.058232 | 1.5282e-08 | E9 | 48
52 | PMWA13 | 1.4448e-08 | 0.41935 | 0.16781 | 0.16781 | 1.1894e-08 | A13 | 1
139 | PMWT1 | 2.0345e-08 | 0.49923 | 0.18173 | 0.18173 | 1.731e-08 | T1 | 2
64 | PMWB12 | 1.7245e-08 | 0.21966 | 0.11495 | 0.11495 | 1.3672e-08 | B12 | 3
77 | PMWC13 | 1.7001e-08 | 0.36943 | 0.053423 | 0.053423 | 1.3582e-08 | C13 | 4
51 | PMWA12 | 1.8025e-08 | 0.30952 | 0.081536 | 0.081536 | 1.4102e-08 | A12 | 5
89 | PMWD12 | 1.7981e-08 | 0 | 0 | 0 | 1.4366e-08 | D12 | 6
76 | PMWC12 | 1.8078e-08 | 0.4493 | 0.094261 | 0.094261 | 1.5155e-08 | C12 | 7
63 | PMWB11 | 1.892e-08 | 0.46927 | 0.10268 | 0.10268 | 1.5589e-08 | B11 | 8
50 | PMWA11 | 1.8612e-08 | 0.34946 | 0.079406 | 0.079406 | 1.4336e-08 | A11 | 9
104 | PMWE13 | 1.8598e-08 | 0.36943 | 0.096314 | 0.096314 | 1.4131e-08 | E13 | 10
88 | PMWD11 | 1.7313e-08 | 0.27957 | 0.083523 | 0.083523 | 1.4218e-08 | D11 | 11
75 | PMWC11 | 1.7945e-08 | 0.43932 | 0.068709 | 0.068709 | 1.4273e-08 | C11 | 12
62 | PMWB10 | 1.8088e-08 | 0.49923 | 0.072385 | 0.072385 | 1.3951e-08 | B10 | 13
49 | PMWA10 | 1.7853e-08 | 0 | 0 | 0 | 1.4e-08 | A10 | 14
87 | PMWD10 | 1.7707e-08 | 0.099845 | 0.05906 | 0.05906 | 1.409e-08 | D10 | 15
72 | PMWB9 | 1.7016e-08 | 0.34946 | 0.083102 | 0.083102 | 1.3587e-08 | B9 | 16
74 | PMWC10 | 1.79e-08 | 0 | 0 | 0 | 1.3931e-08 | C10 | 17
85 | PMWC9 | 1.6449e-08 | 0.42933 | 0.10343 | 0.10343 | 1.3511e-08 | C9 | 18
60 | PMWA9 | 1.7416e-08 | 0.10983 | 0.11504 | 0.10983 | 1.4072e-08 | A9 | 19
71 | PMWB8 | 1.8257e-08 | 0.34946 | 0.073801 | 0.073801 | 1.4231e-08 | B8 | 20
59 | PMWA8 | 1.8518e-08 | 0.40937 | 0.092624 | 0.092624 | 1.4483e-08 | A8 | 21
96 | PMWD8 | 1.8793e-08 | 0.26958 | 0.051894 | 0.051894 | 1.4851e-08 | D8 | 22
84 | PMWC8 | 1.7841e-08 | 0.19969 | 0.10976 | 0.10976 | 1.4041e-08 | C8 | 23
70 | PMWB7 | 1.7638e-08 | 0 | 0 | 0 | 1.35e-08 | B7 | 24
138 | PMWR1 | 1.4167e-08 | 0.13978 | 0.10165 | 0.10165 | 1.1829e-08 | R1 | 25
125 | PMWG1 | 1.7748e-08 | 0.30952 | 0.084228 | 0.084228 | 1.3944e-08 | G1 | 26
140 | PMWT2 | 1.613e-17 | 0.47926 | 0.10532 | 0.10532 | 2.9033e-19 | T2 | 27
100 | PMWE1 | 1.7124e-08 | 0.49923 | 0.07428 | 0.07428 | 1.3584e-08 | E1 | 28
86 | PMWD1 | 5.1987e-16 | 0.47926 | 0.10515 | 0.10515 | 1.9717e-18 | D1 | 29
113 | PMWF1 | 1.8373e-08 | 0.30952 | 0.073994 | 0.073994 | 1.4355e-08 | F1 | 30
105 | PMWE2 | 1.7629e-08 | 0 | 0 | 0 | 1.3999e-08 | E2 | 31
130 | PMWG2 | 1.7273e-08 | 0.34946 | 0.11257 | 0.11257 | 1.3915e-08 | G2 | 32
117 | PMWF2 | 1.8314e-08 | 0.49923 | 0.095541 | 0.095541 | 1.4467e-08 | F2 | 33
131 | PMWG3 | 1.7905e-08 | 0.35944 | 0.10859 | 0.10859 | 1.4043e-08 | G3 | 34
106 | PMWE3 | 1.7197e-08 | 0.49923 | 0.091464 | 0.091464 | 1.3589e-08 | E3 | 35
91 | PMWD3 | 1.9007e-08 | 0.43932 | 0.063081 | 0.063081 | 1.5496e-08 | D3 | 36
118 | PMWF3 | 1.7515e-08 | 0.29954 | 0.092575 | 0.092575 | 1.4047e-08 | F3 | 37
132 | PMWG4 | 2.0656e-08 | 0.46927 | 0.075477 | 0.075477 | 1.6103e-08 | G4 | 38
107 | PMWE4 | 1.8368e-08 | 0 | 0 | 0 | 1.4569e-08 | E4 | 39
119 | PMWF4 | 1.6493e-08 | 0.41935 | 0.10387 | 0.10387 | 1.3361e-08 | F4 | 40
108 | PMWE5 | 1.6163e-08 | 0.4493 | 0.055176 | 0.055176 | 1.2959e-08 | E5 | 41
93 | PMWD5 | 1.8192e-08 | 0.47926 | 0.083519 | 0.083519 | 1.4069e-08 | D5 | 42
120 | PMWF5 | 1.7466e-08 | 0.35944 | 0.098104 | 0.098104 | 1.3798e-08 | F5 | 43
133 | PMWG5 | 1.7517e-08 | 0.3894 | 0.10854 | 0.10854 | 1.3972e-08 | G5 | 44
109 | PMWE6 | 1.7173e-08 | 0.33947 | 0.090951 | 0.090951 | 1.3193e-08 | E6 | 45
134 | PMWG6 | 1.7082e-08 | 0.43932 | 0.083613 | 0.083613 | 1.369e-08 | G6 | 46
121 | PMWF6 | 1.6638e-08 | 0.45929 | 0.091515 | 0.091515 | 1.3103e-08 | F6 | 47
135 | PMWG7 | 1.6815e-08 | 0.37941 | 0.084168 | 0.084168 | 1.3742e-08 | G7 | 48
114 | PMWF10 | 1.7838e-08 | 0.33947 | 0.083495 | 0.083495 | 1.3751e-08 | F10 | 49
102 | PMWE11 | 1.7901e-08 | 0 | 0 | 0 | 1.4145e-08 | E11 | 50
127 | PMWG11 | 1.7264e-08 | 0.22964 | 0.10512 | 0.10512 | 1.4108e-08 | G11 | 51
115 | PMWF11 | 1.7688e-08 | 0.21966 | 0.053766 | 0.053766 | 1.3458e-08 | F11 | 52
103 | PMWE12 | 1.9317e-08 | 0.49923 | 0.087828 | 0.087828 | 1.4708e-08 | E12 | 53
128 | PMWG12 | 1.8134e-08 | 0.099845 | 0.094808 | 0.094808 | 1.475e-08 | G12 | 54
116 | PMWF12 | 1.825e-08 | 0.37941 | 0.050278 | 0.050278 | 1.4541e-08 | F12 | 55
129 | PMWG13 | 1.8268e-08 | 0.3894 | 0.078374 | 0.078374 | 1.4586e-08 | G13 | 56
99 | PMWDP2 | 1.6252e-08 | 0.33947 | 0.050566 | 0.050566 | 1.3059e-08 | DK2 | 57
110 | PMWE7 | 1.6763e-08 | 0.35944 | 0.076987 | 0.076987 | 1.3314e-08 | E7 | 58
95 | PMWD7 | 1.8646e-08 | 0 | 0 | 0 | 1.4011e-08 | D7 | 59
122 | PMWF7 | 1.9508e-08 | 0.45929 | 0.049323 | 0.049323 | 1.568e-08 | F7 | 60
111 | PMWE8 | 4.4258e-08 | 0 | 0 | 0 | 3.2825e-08 | E8 | 61
136 | PMWG8 | 1.7733e-08 | 0.48924 | 0.11559 | 0.11559 | 1.4449e-08 | G8 | 62
123 | PMWF8 | 1.7859e-08 | 0.079876 | 0.053163 | 0.053163 | 1.3542e-08 | F8 | 63
112 | PMWE9 | 1.8784e-08 | 0.24961 | 0.04659 | 0.04659 | 1.4753e-08 | E9 | 64
137 | PMWG9 | 1.7533e-08 | 0.099845 | 0.09216 | 0.09216 | 1.3594e-08 | G9 | 65
97 | PMWD9 | 1.8555e-08 | 0.45929 | 0.07885 | 0.07885 | 1.486e-08 | D9 | 66
124 | PMWF9 | 1.774e-08 | 0.34946 | 0.088356 | 0.088356 | 1.3974e-08 | F9 | 67
101 | PMWE10 | 1.7652e-08 | 0.47926 | 0.08667 | 0.08667 | 1.3692e-08 | E10 | 68
126 | PMWG10 | 1.7458e-08 | 0.49923 | 0.12456 | 0.12456 | 1.4288e-08 | G10 | 69
80 | PMWC4 | 1.624e-08 | 0.41935 | 0.061335 | 0.061335 | 1.3021e-08 | C4 | 70
66 | PMWB3 | 1.7403e-08 | 0.27957 | 0.055477 | 0.055477 | 1.4419e-08 | B3 | 71
79 | PMWC3 | 1.8412e-08 | 0.30952 | 0.070114 | 0.070114 | 1.419e-08 | C3 | 72
65 | PMWB2 | 2.0638e-08 | 0.47926 | 0.094722 | 0.094722 | 1.5907e-08 | B2 | 73
90 | PMWD2 | 2.0338e-08 | 0 | 0 | 0 | 1.5931e-08 | D2 | 74
54 | PMWA3 | 1.6941e-08 | 0.32949 | 0.099071 | 0.099071 | 1.3434e-08 | A3 | 75
53 | PMWA2 | 1.7505e-08 | 0.41935 | 0.085032 | 0.085032 | 1.4208e-08 | A2 | 76
78 | PMWC2 | 1.8111e-08 | 0.47926 | 0.10934 | 0.10934 | 1.43e-08 | C2 | 77
61 | PMWB1 | 1.8389e-08 | 0 | 0 | 0 | 1.4621e-08 | B1 | 78
48 | PMWA1 | 1.8609e-08 | 0.28955 | 0.086446 | 0.086446 | 1.4856e-08 | A1 | 79
98 | PMWDP1 | 1.8306e-08 | 0 | 0 | 0 | 1.4435e-08 | DK1 | 80
73 | PMWC1 | 1.8293e-08 | 0.37941 | 0.09486 | 0.09486 | 1.4726e-08 | C1 | 81
58 | PMWA7 | 1.713e-08 | 0.3894 | 0.081239 | 0.081239 | 1.3795e-08 | A7 | 82
57 | PMWA6 | 1.6882e-08 | 0.47926 | 0.076475 | 0.076475 | 1.364e-08 | A6 | 83
69 | PMWB6 | 1.7422e-08 | 0.35944 | 0.078564 | 0.078564 | 1.3483e-08 | B6 | 84
83 | PMWC7 | 1.7517e-08 | 0.32949 | 0.060309 | 0.060309 | 1.3936e-08 | C7 | 85
56 | PMWA5 | 1.7642e-08 | 0.23963 | 0.099788 | 0.099788 | 1.3893e-08 | A5 | 86
68 | PMWB5 | 1.807e-08 | 0.47926 | 0.059586 | 0.059586 | 1.3954e-08 | B5 | 87
82 | PMWC6 | 1.8185e-08 | 0.43932 | 0.095032 | 0.095032 | 1.4741e-08 | C6 | 88
94 | PMWD6 | 3.0278e-08 | 0 | 0 | 0 | 2.3177e-08 | D6 | 89
67 | PMWB4 | 1.8111e-08 | 0.11981 | 0.069656 | 0.069656 | 1.3913e-08 | B4 | 90
81 | PMWC5 | 1.7491e-08 | 0.27957 | 0.090618 | 0.090618 | 1.3473e-08 | C5 | 91
92 | PMWD4 | 1.66e-08 | 0.36943 | 0.077941 | 0.077941 | 1.3514e-08 | D4 | 92
55 | PMWA4 | 1.6798e-08 | 0.21966 | 0.062706 | 0.062706 | 1.3311e-08 | A4 | 93
282 | PSWR1 | 1.2294e-08 | 0 | 0 | 0 | 9.871e-09 | R1 | 1
194 | PSWD16 | 1.8334e-08 | 0 | 0 | 0 | 1.405e-08 | D16 | 2
283 | PSWT1 | 2.239e-08 | 0.48924 | 0.036796 | 0.036796 | 1.7546e-08 | T1 | 3
163 | PSWB16 | 1.9626e-08 | 0 | 0 | 0 | 1.5847e-08 | B16 | 4
178 | PSWC15 | 1.91e-08 | 0.41935 | 0.045197 | 0.045197 | 1.5266e-08 | C15 | 5
147 | PSWA15 | 1.7698e-08 | 0 | 0 | 0 | 1.4261e-08 | A15 | 6
193 | PSWD15 | 1.6078e-17 | 0.47926 | 0.10521 | 0.10521 | 4.0818e-20 | D15 | 7
162 | PSWB15 | 1.8207e-08 | 0 | 0 | 0 | 1.4293e-08 | B15 | 8
177 | PSWC14 | 1.8453e-08 | 0.099845 | 0.063412 | 0.063412 | 1.4879e-08 | C14 | 9
192 | PSWD14 | 1.9629e-08 | 0 | 0 | 0 | 1.5054e-08 | D14 | 10
146 | PSWA14 | 1.707e-08 | 0 | 0 | 0 | 1.344e-08 | A14 | 11
145 | PSWA13 | 1.4989e-08 | 0.3894 | 0.16747 | 0.16747 | 1.2323e-08 | A13 | 12
161 | PSWB14 | 1.7751e-08 | 0 | 0 | 0 | 1.4255e-08 | B14 | 13
176 | PSWC13 | 1.7568e-08 | 0 | 0 | 0 | 1.4016e-08 | C13 | 14
160 | PSWB13 | 1.7545e-08 | 0 | 0 | 0 | 1.3831e-08 | B13 | 15
191 | PSWD13 | 1.7137e-08 | 0 | 0 | 0 | 1.3634e-08 | D13 | 16
144 | PSWA12 | 1.8232e-08 | 0 | 0 | 0 | 1.4386e-08 | A12 | 17
175 | PSWC12 | 8.9111e-16 | 0.47926 | 0.1052 | 0.1052 | 8.9493e-18 | C12 | 18
190 | PSWD12 | 1.9193e-08 | 0 | 0 | 0 | 1.4458e-08 | D12 | 19
159 | PSWB12 | 2.1614e-08 | 0.34946 | 0.094136 | 0.094136 | 1.8746e-08 | B12 | 20
207 | PSWE11 | 1.7591e-08 | 0 | 0 | 0 | 1.3916e-08 | E11 | 21
143 | PSWA11 | 1.5133e-08 | 0.48924 | 0.13335 | 0.13335 | 1.2809e-08 | A11 | 22
174 | PSWC11 | 1.7392e-08 | 0 | 0 | 0 | 1.3807e-08 | C11 | 23
158 | PSWB11 | 1.7081e-08 | 0 | 0 | 0 | 1.3291e-08 | B11 | 24
205 | PSWE1 | 1.9433e-08 | 0 | 0 | 0 | 1.5106e-08 | E1 | 25
220 | PSWF1 | 1.7943e-08 | 0 | 0 | 0 | 1.4271e-08 | F1 | 26
284 | PSWT2 | 2.1879e-08 | 0 | 0 | 0 | 1.7376e-08 | T2 | 27
251 | PSWH1 | 1.7913e-08 | 0 | 0 | 0 | 1.4208e-08 | H1 | 28
236 | PSWG1 | 1.9293e-08 | 0 | 0 | 0 | 1.5151e-08 | G1 | 29
267 | PSWJ1 | 1.8765e-08 | 0 | 0 | 0 | 1.44e-08 | J1 | 30
259 | PSWH2 | 1.9779e-08 | 0 | 0 | 0 | 1.5561e-08 | H2 | 31
228 | PSWF2 | 1.7673e-08 | 0 | 0 | 0 | 1.3917e-08 | F2 | 32
274 | PSWJ2 | 2.0938e-08 | 0 | 0 | 0 | 1.5663e-08 | J2 | 33
243 | PSWG2 | 1.9002e-08 | 0 | 0 | 0 | 1.4992e-08 | G2 | 34
260 | PSWH3 | 1.8489e-08 | 0 | 0 | 0 | 1.4074e-08 | H3 | 35
275 | PSWJ3 | 2.001e-08 | 0 | 0 | 0 | 1.5742e-08 | J3 | 36
212 | PSWE2 | 1.8518e-08 | 0 | 0 | 0 | 1.4515e-08 | E2 | 37
229 | PSWF3 | 2.0231e-08 | 0 | 0 | 0 | 1.5532e-08 | F3 | 38
244 | PSWG3 | 1.9829e-08 | 0 | 0 | 0 | 1.5243e-08 | G3 | 39
261 | PSWH4 | 1.9102e-08 | 0.40937 | 0.050418 | 0.050418 | 1.4417e-08 | H4 | 40
276 | PSWJ4 | 1.924e-08 | 0 | 0 | 0 | 1.4857e-08 | J4 | 41
213 | PSWE3 | 1.8005e-08 | 0 | 0 | 0 | 1.4168e-08 | E3 | 42
230 | PSWF4 | 1.993e-08 | 0 | 0 | 0 | 1.5791e-08 | F4 | 43
245 | PSWG4 | 1.9243e-08 | 0.29954 | 0.035546 | 0.035546 | 1.5013e-08 | G4 | 44
262 | PSWH5 | 1.7945e-08 | 0 | 0 | 0 | 1.3961e-08 | H5 | 45
214 | PSWE4 | 1.9906e-08 | 0 | 0 | 0 | 1.5704e-08 | E4 | 46
277 | PSWJ5 | 1.7982e-08 | 0 | 0 | 0 | 1.4115e-08 | J5 | 47
231 | PSWF5 | 1.8095e-08 | 0 | 0 | 0 | 1.4027e-08 | F5 | 48
199 | PSWD6 | 1.8832e-08 | 0 | 0 | 0 | 1.4366e-08 | D6 | 49
168 | PSWB6 | 1.7492e-08 | 0.47926 | 0.079833 | 0.079833 | 1.4228e-08 | B6 | 50
182 | PSWC5 | 1.7798e-08 | 0 | 0 | 0 | 1.409e-08 | C5 | 51
151 | PSWA5 | 1.7784e-08 | 0 | 0 | 0 | 1.4106e-08 | A5 | 52
215 | PSWE5 | 1.9321e-08 | 0 | 0 | 0 | 1.5279e-08 | E5 | 53
167 | PSWB5 | 1.999e-08 | 0 | 0 | 0 | 1.5792e-08 | B5 | 54
198 | PSWD5 | 1.7974e-08 | 0.26958 | 0.048175 | 0.048175 | 1.4118e-08 | D5 | 55
181 | PSWC4 | 1.7059e-08 | 0 | 0 | 0 | 1.3088e-08 | C4 | 56
150 | PSWA4 | 2.0286e-08 | 0 | 0 | 0 | 1.6064e-08 | A4 | 57
197 | PSWD4 | 1.8014e-08 | 0 | 0 | 0 | 1.4237e-08 | D4 | 58
166 | PSWB4 | 1.8899e-08 | 0 | 0 | 0 | 1.5023e-08 | B4 | 59
180 | PSWC3 | 2.0807e-08 | 0 | 0 | 0 | 1.6134e-08 | C3 | 60
165 | PSWB3 | 1.8515e-08 | 0 | 0 | 0 | 1.4355e-08 | B3 | 61
149 | PSWA3 | 1.8122e-08 | 0 | 0 | 0 | 1.4192e-08 | A3 | 62
148 | PSWA2 | 1.8672e-08 | 0 | 0 | 0 | 1.4664e-08 | A2 | 63
196 | PSWD3 | 1.7881e-08 | 0 | 0 | 0 | 1.411e-08 | D3 | 64
179 | PSWC2 | 2.1692e-08 | 0 | 0 | 0 | 1.689e-08 | C2 | 65
164 | PSWB2 | 1.8042e-08 | 0 | 0 | 0 | 1.418e-08 | B2 | 66
195 | PSWD2 | 1.9513e-08 | 0 | 0 | 0 | 1.4824e-08 | D2 | 67
141 | PSWA1 | 1.846e-08 | 0 | 0 | 0 | 1.5127e-08 | A1 | 68
172 | PSWC1 | 1.8474e-08 | 0 | 0 | 0 | 1.5099e-08 | C1 | 69
156 | PSWB1 | 1.8859e-08 | 0 | 0 | 0 | 1.5151e-08 | B1 | 70
203 | PSWDP1 | 1.5605e-08 | 0.42933 | 0.049164 | 0.049164 | 1.2468e-08 | DK1 | 71
187 | PSWD1 | 1.8031e-08 | 0.14977 | 0.038104 | 0.038104 | 1.454e-08 | D1 | 72
223 | PSWF12 | 1.8868e-08 | 0 | 0 | 0 | 1.4947e-08 | F12 | 73
269 | PSWJ11 | 1.7299e-08 | 0 | 0 | 0 | 1.4086e-08 | J11 | 74
208 | PSWE12 | 1.8662e-08 | 0 | 0 | 0 | 1.4495e-08 | E12 | 75
254 | PSWH12 | 1.8064e-08 | 0 | 0 | 0 | 1.4526e-08 | H12 | 76
239 | PSWG12 | 1.9604e-08 | 0 | 0 | 0 | 1.5107e-08 | G12 | 77
224 | PSWF13 | 1.8831e-08 | 0.47926 | 0.045614 | 0.045614 | 1.4863e-08 | F13 | 78
209 | PSWE13 | 1.8531e-08 | 0 | 0 | 0 | 1.4208e-08 | E13 | 79
270 | PSWJ12 | 1.9132e-08 | 0 | 0 | 0 | 1.4926e-08 | J12 | 80
255 | PSWH13 | 1.8506e-08 | 0 | 0 | 0 | 1.4971e-08 | H13 | 81
240 | PSWG13 | 1.875e-08 | 0 | 0 | 0 | 1.4544e-08 | G13 | 82
225 | PSWF14 | 1.888e-08 | 0 | 0 | 0 | 1.4989e-08 | F14 | 83
210 | PSWE14 | 1.9864e-08 | 0 | 0 | 0 | 1.5397e-08 | E14 | 84
271 | PSWJ13 | 1.9263e-08 | 0.43932 | 0.036901 | 0.036901 | 1.4846e-08 | J13 | 85
256 | PSWH14 | 1.809e-08 | 0 | 0 | 0 | 1.4278e-08 | H14 | 86
241 | PSWG14 | 1.9398e-08 | 0.11981 | 0.041735 | 0.041735 | 1.4885e-08 | G14 | 87
272 | PSWJ14 | 1.8989e-08 | 0.069892 | 0.046532 | 0.046532 | 1.535e-08 | J14 | 88
226 | PSWF15 | 1.7841e-08 | 0 | 0 | 0 | 1.4178e-08 | F15 | 89
257 | PSWH15 | 1.8635e-08 | 0 | 0 | 0 | 1.4467e-08 | H15 | 90
273 | PSWJ15 | 1.8837e-08 | 0 | 0 | 0 | 1.4359e-08 | J15 | 91
242 | PSWG15 | 1.8902e-08 | 0 | 0 | 0 | 1.5057e-08 | G15 | 92
258 | PSWH16 | 1.8914e-08 | 0 | 0 | 0 | 1.4754e-08 | H16 | 93
204 | PSWDP2 | 1.6294e-08 | 0 | 0 | 0 | 1.3019e-08 | DK2 | 94
227 | PSWF16 | 1.8422e-08 | 0.19969 | 0.045059 | 0.045059 | 1.5001e-08 | F16 | 95
211 | PSWE15 | 1.7877e-08 | 0.23963 | 0.045614 | 0.045614 | 1.3931e-08 | E15 | 96
189 | PSWD11 | 1.8759e-08 | 0 | 0 | 0 | 1.483e-08 | D11 | 97
142 | PSWA10 | 1.5806e-08 | 0 | 0 | 0 | 1.247e-08 | A10 | 98
206 | PSWE10 | 1.9191e-08 | 0 | 0 | 0 | 1.5426e-08 | E10 | 99
173 | PSWC10 | 1.8187e-08 | 0 | 0 | 0 | 1.425e-08 | C10 | 100
157 | PSWB10 | 1.7842e-08 | 0 | 0 | 0 | 1.3842e-08 | B10 | 101
188 | PSWD10 | 1.844e-08 | 0.34946 | 0.046326 | 0.046326 | 1.4231e-08 | D10 | 102
155 | PSWA9 | 1.7725e-08 | 0 | 0 | 0 | 1.3887e-08 | A9 | 103
219 | PSWE9 | 1.7476e-08 | 0.43932 | 0.046859 | 0.046859 | 1.4139e-08 | E9 | 104
186 | PSWC9 | 1.9247e-08 | 0 | 0 | 0 | 1.5329e-08 | C9 | 105
171 | PSWB9 | 1.8526e-08 | 0 | 0 | 0 | 1.451e-08 | B9 | 106
202 | PSWD9 | 1.8889e-08 | 0 | 0 | 0 | 1.5015e-08 | D9 | 107
154 | PSWA8 | 1.7859e-08 | 0 | 0 | 0 | 1.4086e-08 | A8 | 108
185 | PSWC8 | 1.8104e-08 | 0 | 0 | 0 | 1.4282e-08 | C8 | 109
218 | PSWE8 | 1.9395e-08 | 0 | 0 | 0 | 1.4667e-08 | E8 | 110
201 | PSWD8 | 1.8584e-08 | 0 | 0 | 0 | 1.414e-08 | D8 | 111
170 | PSWB8 | 1.949e-08 | 0 | 0 | 0 | 1.4797e-08 | B8 | 112
184 | PSWC7 | 1.8911e-08 | 0 | 0 | 0 | 1.4694e-08 | C7 | 113
217 | PSWE7 | 1.9001e-08 | 0 | 0 | 0 | 1.4976e-08 | E7 | 114
153 | PSWA7 | 2.0628e-08 | 0 | 0 | 0 | 1.5785e-08 | A7 | 115
200 | PSWD7 | 1.8329e-08 | 0 | 0 | 0 | 1.4075e-08 | D7 | 116
169 | PSWB7 | 1.929e-08 | 0 | 0 | 0 | 1.5058e-08 | B7 | 117
183 | PSWC6 | 1.8738e-08 | 0 | 0 | 0 | 1.4556e-08 | C6 | 118
216 | PSWE6 | 1.813e-08 | 0 | 0 | 0 | 1.3919e-08 | E6 | 119
152 | PSWA6 | 1.8913e-08 | 0 | 0 | 0 | 1.5137e-08 | A6 | 120
246 | PSWG5 | 1.783e-08 | 0 | 0 | 0 | 1.4078e-08 | G5 | 121
263 | PSWH6 | 1.7518e-08 | 0 | 0 | 0 | 1.3845e-08 | H6 | 122
278 | PSWJ6 | 1.7065e-08 | 0 | 0 | 0 | 1.3789e-08 | J6 | 123
232 | PSWF6 | 1.8739e-08 | 0 | 0 | 0 | 1.4472e-08 | F6 | 124
247 | PSWG6 | 1.7045e-08 | 0.48924 | 0.052145 | 0.052145 | 1.3931e-08 | G6 | 125
264 | PSWH7 | 1.8054e-08 | 0 | 0 | 0 | 1.417e-08 | H7 | 126
233 | PSWF7 | 1.6601e-08 | 0 | 0 | 0 | 1.3899e-08 | F7 | 127
279 | PSWJ7 | 1.7057e-08 | 0 | 0 | 0 | 1.3659e-08 | J7 | 128
248 | PSWG7 | 1.7681e-08 | 0 | 0 | 0 | 1.3557e-08 | G7 | 129
265 | PSWH8 | 1.8042e-08 | 0 | 0 | 0 | 1.3828e-08 | H8 | 130
234 | PSWF8 | 1.7397e-08 | 0 | 0 | 0 | 1.3403e-08 | F8 | 131
249 | PSWG8 | 8.826e-16 | 0.47926 | 0.10521 | 0.10521 | 2.4714e-18 | G8 | 132
280 | PSWJ8 | 1.8854e-08 | 0.49923 | 0.041677 | 0.041677 | 1.478e-08 | J8 | 133
235 | PSWF9 | 2.6275e-08 | 0.37941 | 0.058815 | 0.058815 | 2.3274e-08 | F9 | 134
266 | PSWH9 | 1.8546e-08 | 0 | 0 | 0 | 1.4446e-08 | H9 | 135
250 | PSWG9 | 1.8447e-08 | 0.069892 | 0.041351 | 0.041351 | 1.4189e-08 | G9 | 136
281 | PSWJ9 | 1.7839e-08 | 0 | 0 | 0 | 1.3928e-08 | J9 | 137
221 | PSWF10 | 1.6835e-08 | 0 | 0 | 0 | 1.3472e-08 | F10 | 138
252 | PSWH10 | 1.8451e-08 | 0.4493 | 0.036291 | 0.036291 | 1.4484e-08 | H10 | 139
237 | PSWG10 | 1.8727e-08 | 0 | 0 | 0 | 1.4865e-08 | G10 | 140
222 | PSWF11 | 1.7138e-08 | 0 | 0 | 0 | 1.3555e-08 | F11 | 141
268 | PSWJ10 | 1.9237e-08 | 0.45929 | 0.036257 | 0.036257 | 1.4776e-08 | J10 | 142
253 | PSWH11 | 1.7491e-08 | 0 | 0 | 0 | 1.3588e-08 | H11 | 143
238 | PSWG11 | 5.2457e-08 | 0.41935 | 0.047549 | 0.047549 | 3.8215e-08 | G11 | 144
285 | PTCP1 | 2.674e-08 | 0.45929 | 0.17005 | 0.17005 | 2.7525e-08
286 | PTCP2 | 2.6057e-08 | 0.45929 | 0.05722 | 0.05722 | 2.6087e-08
287 | PTCP3 | 2.2786e-08 | 0.48924 | 0.09525 | 0.09525 | 2.0837e-08