HS_DCU FLIGHT MODEL

REPORT OF FUNCTIONAL TESTS

# Table of contents

1	Table of contents	2

2	Test Configuration	4

3	POWER ON : CONSUMPTION TESTS	7

3.1	Primary currents	7

3.2	Conclusion	7

4	Frame generation Tests	8

4.1	Procedure	8

4.2	ANALYSIS	8

4.3	Test of the DCU frames parity	10

4.4	Conclusion	13

5	Offets	14

6	HouseKeepings	26

6.1	Housekeeping acquisition	26

6.2	Conclusion	31

7	BIAS generation tests	32

7.1	BIAS setting	32

7.2	ANALYSIS	33

7.2.1	Bias frequency photometer	34

7.2.2	Bias frequency spectrometer	37

7.2.3	Bias amplitudes	39

7.2.4	JFET amplitudes Vss	42

7.2.5	JFET On successives (Vdd)	44

7.2.6	Photo Heaters and Spectro Heaters	47

7.3	Conclusion	50

8	BIAS test in stand alone mode : Monitoring of BIAS	51

9	Demodulation Tests	52

9.1	Procedure	52

9.2	Data analysis	53

9.3	Batch edition	53

9.4	Conclusion	57

10	Gain Tests	58

10.1	Procedure	58

10.2	Data analysis	59

10.2.1	Photometer	59

10.2.2	Spectrometer	62

10.3	Conclusion	64

11	Cross Talk	65

11.1	Procedure	65

11.1.1	Photometer	65

11.2	Data analysis	68

11.3	Conclusion	71

12	Noise	72

12.1	Conclusion	76

# Test Configuration

We use the LTU in its latest software configuration, the fully assembled DCU and FCU, the FPU and the PSU Flight Model (PSU FM).

Before to connect any harness, we place savers on all connectors

Connect all harnesses as described below :

- Between DCU and FPU

- Between DCU and LTU

- Between FCU and LTU

- Between FCU and FPU

- Between DCU and PSU FM

- Between FCU and PSU FM

# POWER ON : CONSUMPTION TESTS

## Primary currents

We report thereafter the consumption on the primary voltage (28 V)

Conso 28 V			Main 			redundant			Expected

Power ON			400 	mA		400 	mA			400 	mA

MCU ON 			810 	mA		810 	mA			810 	mA

MCU ON and booted 		870 	mA		870 	mA			870 	mA

LIAP ON 			2100 	mA		2100	mA			2100	mA

LIAS ON			1230	mA		1230	mA			1230	mA

LIAP ON and MCU OFF	1610 	mA		1610 	mA			1600 	mA

LIAS ON and MCU OFF	750 	mA		740 	mA			7540 	mA

## Conclusion

| Consumption TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

These measurements are within the specified values, we can go through the functional tests procedure.

# Frame generation Tests

## Procedure

( FPU simulator switched ON but not used

( Batch executed by LTU : DAQ-IF-TESTS Main Redundant.txt

( 2 x 5 TM file generated by LTU (DPU) (Main and Redundant)

( Rename the files with the following extensions :

Main Side

| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Frame_Test_Main.tm |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Frame_Number_Main.tm |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Photo_frequency_Main.tm |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_spectro_frequency_Main.tm |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Time_reset_Main.tm |

Redundant side

| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Frame_Test_Redundant.tm |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Frame_Number_Redundant.tm |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Photo_frequency_Redundant.tm |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_spectro_frequency_Main.tm |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_Time_reset_Redundant.tm |

## ANALYSIS

Check consistency with specification : Labview Software (button frame_length, frame_id, frame count, sampling rate photo, sampling rate_spectro)

Open last renamed TM files (with extension ref. Frame_Test.tm  to Time_reset.tm)

Frame ID and Frame length :

( Check green light

|  | As run Main | As run Redundant |
| --- | --- | --- |
| Green light OK | OK | OK |
| Frame ID and Frame length | OK | OK |

Frame counter :

( Open last renamed TM file (ref. _Frame_Number)

( Enter 338 in “number of frames” box (340 is theorical value)

Check for “frame count green light led”

Stop frame count analysis

|  | As run Main | As run Redundant |
| --- | --- | --- |
| Green light OK | OK | OK |
| Frame counter | OK | OK |

Sampling frequencies :

Open last renamed TM file (ref. _ photo_frequency)

(  delta t max = 0,007 s	(Criterion)

- Check the "green light" (sampling rate photometer)

- Stop photosampling rate analysis

|  | As run Main | As run Redundant |
| --- | --- | --- |
| Green lights OK | OK | OK |
| Delta tmax < 0,00001 s | OK | OK |
| No error in parity Photometer sampling rate |  |  |

Open last renamed TM file (ref. _ spectro_frequency)

(  delta t max = 0,007 s	(Criterion)

- Check the "green light" (sampling rate spectrometer)

- Stop spectro sampling rate analysis

|  | As run Main | As run Redundant |
| --- | --- | --- |
| Green lights OK | OK | OK |
| Delta tmax < 0,00002 s | OK | OK |
| No error in parity spectrometer sampling rate |  |  |

Time Stamp Reset :

( Open last renamed TM file (ref. _Time_Reset)

( Enter manually the frequency : 50.86 Hz in “frequency box”

( Check the "green light" and no error in “block time vs block number” box

Stop block time analysis

|  | As run Main | As run Redundant |
| --- | --- | --- |
| Green light OK | OK | OK |
| block time vs block number | OK | OK |

## Test of the DCU frames parity

( Check the 4 "green lights" associated with parity

( Enter manually the frequency : 50.86 Hz

( Check two "green lights" :

for the time value at time reset (Block time vs Block number)

for the parity block time

|  | As run Main | As run Redundant |
| --- | --- | --- |
| 4 Green lights OK | OK | OK |

Cut and paste here after hardcopies of labview (Main and Redundant)

Main side

Redundant side

## Conclusion

| Frame generation TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# Offets

In this test, we use a simulator input file that consist in sinus signal with the same pic pic level of 0,2 mV on all channels but at different magnitude levels (id est with different offsets) within the specified range. The purpose is then to check that the automatic offset function allows reading the signal within its specified magnitude range.

We perform signal acquisition using the FPU simulator script : Script_Spire_ScieNiveaux.spt

Simulator freq = 0.32 (0,08 Hz)

We obtain the following telemetry files

Main

TM files of 09h23 and 09h25

Redundant

TM files of 09h27 and 09h29

We thereafter represent the signal as read from TM files. The value of the offset has been added to the signal so as to differentiate the channels and to re create the signal.

On LIAS boards, the same signal is sent on a group of 3 channels. The value of the offset computed by DCU appears on the line above.

| Offsets | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# HouseKeepings

## Housekeeping acquisition

We start the DCU Monitoring Mode from LTU using the batch DRCU Monitoring.txt

We observe the results over a sufficient time (1 minute)

When LIAP LIAS are both OFF

|  | As run |
| --- | --- |
| Check LIAP & LIAS OFF | OK |

When LIAP ON.

Include in as run the hardcopy off LTU monitoring screen (replace typical view provided here after)

Main side :

Redundant side :

When LIAP OFF and LIAS ON.

Include in as run the hardcopy off LTU monitoring screen (replace typical view provided here after)

Main side :

Redundant side :

Switch LIAS OFF and stop LTU monitoring.

## Conclusion

| Housekeeping TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# BIAS generation tests

## BIAS setting

( The FPU simulator must be in slave mode and generating measurement files; one can use any simulator script file (for instance, Script_Spire_01.spt), only the outcoming measurement file matters.

( The Batch executed by LTU is : BIAS-TESTS Main redundant.txt

( 2 times (Main and Redundant) 7 FPU simulator measurement files (14 in total) generated that contain :

1 -> Bias frequency photometer

2 -> Bias frequency spectrometer

3 -> Bias amplitudes (PhotoBiasAmpl LW,PhotoBiasAmpl MW,PhotoBiasAmpl SW,PhotoBiasAmpl TC, SpectroBiasAmpl LW,SpectroBiasAmpl SW)

4 -> JFET amplitudes Vss

PhLWJFetVSS1 0-255

PhLWJFetVSS2 0-255

PhMWJFetVSS1 0-255

PhMWJFetVSS2 0-255

PhMWJFetVSS3 0-255

PhMWJFetVSS4 0-255

PhSWJFetVSS1 0-255

PhSWJFetVSS2 0-255

PhSWJFetVSS3 0-255

PhSWJFetVSS4 0-255

PhSWJFetVSS5 0-255

PhSWJFetVSS6 0-255

PhMWJFetTC 0-255

SpLWJFetVSS 0-255

SpSWJFetVSS1 0-255

SpSWJFetVSS2 0-255

5 -> JFET On successives (Vdd)

PhSWJFetPwr Vdd1

PhSWJFetPwr Vdd2

PhSWJFetPwr Vdd3

PhSWJFetPwr Vdd4

PhSWJFetPwr Vdd5

PhSWJFetPwr Vdd6

PhMLWJFetPwr Vdd1 MW

PhMLWJFetPwr Vdd2 MW

PhMLWJFetPwr Vdd3 MW

PhMLWJFetPwr Vdd4 MW

PhMLWJFetPwr Vdd1 LW

PhMLWJFetPwr Vdd2 LW

PhMLWJFetPwr Vdd TC

SpSLWJFetPwr Vdd SLW

SpSLWJFetPwr Vdd1 SSW

SpSLWJFetPwr Vdd2 SSW

All ON

6 -> Photo Heaters

PhotoHeater 0-255

7 -> Spectro Heaters

SpHeater 0-255

The measurement files obtained are : 		Script_Spire_01_120606_145420\SPIRE_test_BIAS_1_1.mes

.

.

.

Script_Spire_01_120606_145420\SPIRE_test_BIAS_7_1.mes										Script_Spire_01_120606_145420\SPIRE_test_BIAS_8_1.mes															.

.

.

Script_Spire_01_120606_145420\SPIRE_test_BIAS_14_1.mes.

## ANALYSIS

( Check consistency with specification : use the Labview Software

( 7 measurement files to analyze in Main as well as in redundant

### Bias frequency photometer

Measurement file : SPIRE_test_BIAS_1_1.mes for main (resp. SPIRE_test_BIAS_8_1.mes for redundant)

Click on photometer bias frequency button, and select the last generated file SPIRE_test_BIAS_1_1.mes for main (resp. SPIRE_test_BIAS_8_1.mes for redundant).

Check the four green lights.

Include in as run the hardcopy off Labview analysis screen (replace typical view provided here after)

Main side :

Redundant side :

### Bias frequency spectrometer

Measurement file : SPIRE_test_BIAS_2_1.mes for main (resp. SPIRE_test_BIAS_9_1.mes for redundant)

Click on spectrometer bias frequency button, and select the last generated file SPIRE_test_BIAS_2_1.mes for main (resp. SPIRE_test_BIAS_9_1.mes for redundant).

Check the two green lights.

Include in as run the hardcopy off Labview analysis screen (replace typical view provided here after)

Main side :

Redundant side :

### Bias amplitudes

Measurement file : SPIRE_test_BIAS_3_1.mes for main (resp. SPIRE_test_BIAS_10_1.mes for redundant)

Click on photometer bias amplitude button, and select the last generated file SPIRE_test_BIAS_10_1.mes for main (resp. SPIRE_test_BIAS_8_1.mes for redundant).

Include in as run the hardcopy off Labview analysis screen (replace typical view provided here after)

Main side :

Redundant side :

### JFET amplitudes Vss

Measurement file : SPIRE_test_BIAS_4_1.mes for main (resp. SPIRE_test_BIAS_11_1.mes for redundant)

Click on JFET Vss bias button, and select the last generated file SPIRE_test_BIAS_4_1.mes for main (resp. SPIRE_test_BIAS_11_1.mes for redundant).

Main side :

Redundant side :

### JFET On successives (Vdd)

Measurement file : SPIRE_test_BIAS_5_1.mes for main (resp. SPIRE_test_BIAS_12_1.mes for redundant)

Click on JFET VDD bias button, Enter low level 0,0023V and select the last generated file SPIRE_test_BIAS_5_1.mes for main (resp. SPIRE_test_BIAS_12_1.mes for redundant).

Include in as run the hardcopy off Labview analysis screen (replace typical view provided here after)

Main side :

Redundant side :

### Photo Heaters and Spectro Heaters

Measurement file : SPIRE_test_BIAS_6_1.mes for main (resp. SPIRE_test_BIAS_13_1.mes for redundant)

Measurement file : SPIRE_test_BIAS_7_1.mes for main (resp. SPIRE_test_BIAS_14_1.mes for redundant)

Click on photometer & spectrometer & JFET heater bias button, and select the last generated file SPIRE_test_BIAS_6_1.mes & SPIRE_test_BIAS_7_1.mes for main (resp. SPIRE_test_BIAS_13_1.mes and SPIRE_test_BIAS_14_1.mes for redundant)

Include in as run the hardcopy off Labview analysis screen (replace typical view provided here after)

Main side :

Redundant side :

## Conclusion

| BIAS TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# BIAS test in stand alone mode : Monitoring of BIAS

When in SLAVE mode, we can check all the BIAS on a given working point. We the use the test sequence BIAS-TEST(stand alone).txt. (FPU simul in standalone, automatique, measure at 1Hz (script_spire_BIAS.spt)

Check the measurement file on the FPU simulator : Values of all the BIAS at a given level (refer to the batch)

The measurement file obtained is : Script_Spire_01_130606_092942\SPIRE_test_BIAS_1_1.mes.

We set the DAC to their maximum (test sequence : DCU BIAS-TEST(half level)Main Redundant.txt) and measure the BIAS at simulator level. The results are the following :

Refer to file : SPIRE_test_BIAS_1_1.mes

Check BIAS amplitudes

Check consistency with FPU simulator input files

Note : When the half maximum amplitude level is sent by the BIAS boars, the FPU simulator should measure about  the following BIAS amplitude (to upate) :

Main			___Red_____

PhotoBiasAmpl SW 	: Level 80 (HEXA), 128 (dec)	: 	64,341 mV		64,341 mV

PhotoBiasAmpl MW 	: Level 80 (HEXA), 128 (dec)	:	64,656 mV		64,736 mV

PhotoBiasAmpl LW 	: Level 80 (HEXA), 128 (dec)	: 	65,139 mV		65,139 mV

PhotoBiasAmpl TC	: Level 80 (HEXA), 128 (dec)	: 	128,069 mV		128,069 mV

SpectroBiasAmpl SW: Level 80 (HEXA), 128 (dec)	: 	89,154	mV		89,154 mV

SpectroBiasAmpl LW: Level 80 (HEXA), 128 (dec)	: 	89,001	mV		89,001 mV

These new values are therefore the ones to set in the simulator input files :

GainInputAmplMiMax(05mv_Offset0)FM.in

Cross_Talk_Ampl_MiMax(05mv_Offset0)FM.in

The associated script files are :

Script_Spire_GainInputAmplMiMax(05mv_Offset0)FM.spt

Script_Spire_Cross_Ampl_MiMax(05mv_Offset0)FM.spt

# Demodulation Tests

## Procedure

Check LIAP LIAS are both OFF and

|  | As run |
| --- | --- |
| Check LIAP & LIAS OFF | OK |

( FPU simulator switched ON in stand alone

( Run the Script Script_Spire_GainInputAmplMiMax(05mv_Offset0)FM.spt

( Signal frequency = 1 Hz ( simulator setting = 32 Hz

Note : The test is performed with a BIAS frequency of 203 Hz (photo and spectro)

( Batch executed by LTU : “4 DCU DEMODULATION Main redundant.txt”

( 4 TM file generated by LTU (DPU) (Photometer and spectrometer in Main and Redundant)

Note : If the FCU is not there, we need to switch ON the LIAP and the LIAS using the resistor load box. There need to proceed then in two steps (demodulation for photometer and demodulation for spectrometer).

( Rename the files as follows :

|  | As run (Photometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss) _demod_photo_main |

|  | As run (Spectrometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss) _demod_spectro_main |

|  | As run (Photometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss) _demod_photo_redundant |

|  | As run (Spectrometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss) _demod_spectro_redundant |

Check LIAP and LIAS are OFF at the end of the test sequence

|  | As run |
| --- | --- |
| Check LIAP & LIAS OFF | OK |

## Data analysis

( Analysis of the two TM files (photo or spectro) with the Labview software that computes the optimal phases (button demodulation analysis, with photometer and spectrometer selected); select the two TM files renamed above (spectro first, photo last). Wait for end of data analysis by labview (lights “psw completed”, “pmw completed”, “plw completed”, “tc_completed”, “ssw completed” and “slw completed”)

| Optimal Phase/setting Channel group | Main Side Optimal phase (°) | Optimal setting | Redundant Side Optimal phase (°) | Optimal setting |
| --- | --- | --- | --- | --- |
| PSW | 5,2 | 4 | 5,2 | 4 |
| PMW | 5,2 | 4 | 5,2 | 4 |
| PLW | 5,5 | 4 | 5,5 | 4 |
| PTC | 5,6 | 4 | 5,6 | 4 |
| SSW | 1,6 | 1 | 1,6 | 1 |
| SLW | 1,6 | 1 | 1,6 | 1 |

Check the criterions :

channel dispersion < 0,3°

mean optimal phase in [0°;+20°]

## Batch edition

( Set the optimal photometer and spectrometer demodulation values in the test sequences :

DCU ACQUISITIONS (Offsets a zero) Main redundant.txt

|  | As run |
| --- | --- |
| Renamed batch files | DCU ACQUISITIONS (Offsets a zero) Main redundant.txt | DCU ACQUISITIONS (Offsets a zero) Main redundant.txt |

(nécessité de	Acquisitions_photomètre_ampl_mi-max.txt et Acquisitions_spectromètre_ampl_mi-max.txt)

Example : set the xx, yy, zz, tt values to the computed ones

//----------------------------------------

/ Réglage des paramètres de démodulation

//----------------------------------------

/ SetPhotoDemodSW 17 tempo0

# D 0 41A 00xx 0000

/ SetPhotoDemodMW 17 tempo0

# D 0 41B 00yy 0000

/ SetPhotoDemodLW 17 tempo0

# D 0 41C 00zz 0000

/ SetPhotoDemodTC 17 tempo0

# D 0 41D 00tt 0000

Main side

Redundant side

## Conclusion

| Demodulation TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# Gain Tests

## Procedure

Check LIAP and LIAS are both OFF

|  | As run |
| --- | --- |
| Check LIAP & LIAS OFF | OK |

( FPU simulator switched ON in stand alone mode

( Run the Script 	Script_Spire_GainInput _Ampl_MiMax(05mv_Offset0)FM.spt : sinusoidal signal with magnitude 0,5 mv pic pic. The FPU simulator sends on all channels a sinusoïdal signal with a magnitude of 0,5 mV pic pic modulated at 100 % (Offset = 0)

( We use a signal at 0,1 Hz ( simulator frequency = 3,2 Hz

( Batch executed by LTU : use the batch edited in previous § (5 DCU ACQUISITIONS (Offsets a zero) Main redundant.txt) : the LTU sets all offsets at 0 and performs TM acquisitions.

( 4 TM file generated by LTU (DPU) (Photometer and spectrometer in Main and Redundant)

( Rename the files as follows :

|  | As run (Photometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss) _gain_photo_main |

|  | As run (Spectrometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_gain_spectro_main |

|  | As run (Photometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss) _gain_photo_redundant |

|  | As run (Spectrometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_gain_spectro_redundant |

Check LIAP and LIAS are both OFF at the end of the test sequence

## Data analysis

### Photometer

( Analysis of the TM files  with the Labview software (photometer_gain button and select last renamed TM files)

Enter the input signal level in the box “input(mV)” : 0,5

Check green lights for LIAP and paste the hard copy of labview screen here after

Main side :

Redundant side :

### Spectrometer

( Analysis of the TM files  with the Labview software (photometer_gain button and select last renamed TM files)

Enter the input signal level in the box “input(mV)” : 0,5

Check green lights for LIAS and paste the hard copy of labview screen here after

Main side

Redundant Side

## Conclusion

| Gain TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL Erreur de checksum sur fichier TM main Photo => re executed successfully |  |

# Cross Talk

## Procedure

### Photometer

Check LIAP LIAS are both OFF

|  | As run |
| --- | --- |
| Check LIAP & LIAS OFF | OK |

Switch LIAP ON.

( FPU simulator switched ON in stand alone mode

( Run the Script 	Scrip_Spire_Cross_Ampl_MiMax(05mv_Offset0)FM.spt : sinusoidal signal with magnitude 1 mv (1 mV pic pic 0,3536 rms)

( We use a signal at 1 Hz => simulator frequency = 32 Hz

( The FPU simulator sends on some channels a sinusoïdal signal with a magnitude of 1mV pic pic modulated at 100 % (Offset = 0).  The illuminated channels are indicated below (Y) :

Index	0	1	2	3	4	5	6	7

LIAP 1

0		Y	Y	Y	N	N	N	N	N

8		N	N	N	N	N	N	N	Y

16		Y	Y	N	N	N	N	N	N

24		N	N	N	N	N	N	Y	Y

LIAP 2

32		Y	N	N	N	N	N	N	N

40		N	N	N	N	N	Y	Y	Y

48		Y	Y	Y	N	N	N	N	N

56		N	N	N	N	N	N	N	Y

LIAP 3

64		Y	Y	N	N	N	N	N	N

72		N	N	N	N	N	N	Y	Y

80		Y	N	N	N	N	N	N	N

88		N	N	N	N	N	Y	Y	Y

LIAP 4

96		Y	N	N	N	N	N	N	N

104		N	N	N	N	N	Y	Y	Y

112		Y	Y	Y	N	N	N	N	N

120		N	N	N	N	N	N	N	Y

LIAP 5

128		Y	N	N	N	N	N	N	N

136		N	N	N	N	N	Y	Y	Y

144		Y	Y	Y	N	N	N	N	N

152		N	N	N	N	N	N	N	Y

LIAP 6

160		Y	Y	N	N	N	N	N	N

168		N	N	N	N	N	N	Y	Y

176		Y	N	N	N	N	N	N	N

184		N	N	N	N	N	Y	Y	Y

LIAP 7

192		Y	N	N	N	N	N	N	N

200		N	N	N	N	N	Y	Y	Y

208		Y	Y	Y	N	N	N	N	N

216		N	N	N	N	N	N	N	Y

LIAP8

224		Y	N	N	N	N	N	N	N

232		N	N	N	N	N	Y	Y	Y

240		Y	Y	Y	N	N	N	N	N

248		N	N	N	N	N	N	N	Y

LIAP9

256		Y	Y	N	N	N	N	N	N

264		N	N	N	N	N	N	Y	Y

272		Y	N	N	N	N	N	N	N

280		N	N	N	N	N	Y	Y	Y

Index	0	1	2	3	4	5	6	7

LIAS 1

0		Y	Y	Y	N	N	N	N	N

8		N	N	N	N	N	N	N	Y

16		Y	Y	N	N	N	N	N	N

LIAS 2

24	Y	Y	Y	N	N	N	N	N

32	N	Y	Y	Y	Y	Y	Y	N

40	N	N	N	N	N	N	N	N

LIAS 3

48	Y	Y	Y	N	N	N	N	N

56	N	Y	Y	Y	Y	Y	Y	N

64	N	N	N	N	N	N	N	N

( Batch executed by LTU : use the batch edited in § 9.3 (5 DCU ACQUISITIONS (Offsets a zero) Main redundant.txt): the LTU sets all offsets at 0 and performs TM acquisitions.

( 4 TM file generated by LTU (DPU) (Photometer and spectrometer in Main and Redundant)

( Rename the files as follows :

|  | As run (Photometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_cross_photo_main |

|  | As run (Spectrometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_cross_spectro_main |

|  | As run (Photometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_cross_photo_redundant |

|  | As run (Spectrometer) |
| --- | --- |
| Renamed TM file | DCUTm(yyyy.mm.dd_hh.mm.ss)_cross_spectro_redundant |

Check LIAP LIAS are both OFF at the end of the procedure

## Data analysis

( Analysis of the two TM files (photo and spectro) with the Labview software (button cross talk with photometer and spectrometer selected)

Fill the “Photometer input (V)” box with value=0,5, and Input frequency box” with value=1.

Fill the “Spectrometer input (V)” box with value=0,5, and Input frequency box” with value=1.

( The labview sofware represents the results, check green light (note that the global green light is not functional since some LIAS boars are missing).

Paste the hard copy of labview screen here after :

Main side

Redundant side

## Conclusion

| Cross talk TEST | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# Noise

Place the short circuits on the DCU LIAP and DCU LIAS connectors and perform signal acquisition.

We represent the noise, expresses in ADC steps rms values and in nV/sqrt(Hz). This latter value is computed taking into account the gain value as well as the cut off frequencies associated with photometer and spectrometer channels.

We use the following formula, applicable to the kind of differential signal the DCU generates :

oise(nV/sqrt(Hz))=109.ADC (rms value) / [(ADCsteps range/Voltage range) . Gain . sqrt(.fcut off(Hz))]

ADCsteps range = 65535

Voltage range = 5 V

Gain = 5448 (photometer)

Gain = 3528 (spectrometer)

fcut off = 5 Hz  (photometer)

fcut off = 25 Hz (spectrometer)

Photometer Main side

Photometer Redundant side

Spectrometer Main side

Spectrometer Redundant side

## Conclusion

| Noise | OK | NON OK |
| --- | --- | --- |
| Test responsible Henri TRIOU Technical specialist Quality | Comments : TEST SUCCESSFULL |  |

# Text boxes

PSU FM

DCU FLIGHT MODEL TEST CONFIGURATION

FPU simulator

DPU simulator

DPU simulator

FPU simulator

PSU FM