JFET/DCU gain check carried out 10 October 2006

Photometer JFET gains were measured stand alone during the instrument integration testing before the start of PFM4.  This note reports on the photometer side JFET gains and confirmation of system gain.

Datasets used:

70 Hz 		ILT_PERF_DAB_P_SinglePhase_300000F5_

130 Hz		ILT_PERF_DAB_P_SinglePhase_300000F6_

189 Hz		ILT_PERF_DAB_P_SinglePhase_300000F7_

Set up:

Warm JFETs connected to the FM warm electronics via RAL cryoharness.

No FPU present and bias fed directly to JFET inputs through test connectors and the FM backharness.

Electronics driven through standard instrument GSE and test scripts.

Measurements were taken as “standard loadcurves” (ILT_PERF_DAB_P) with fixed phase

Three bias frequencies were used namely 70, 130 and 189 Hz

At each frequency the phase was optimised:

70 Hz   phase 163.76416

130 Hz phase 183.52880

189 Hz phase 196.23464

Example of processing:

Bias converted to rms by dividing converted housekeeping values by sqrt(2).  Gains applied as per DCU design description.  No phase correction applied.  Residual gain derived by fitted first order (only) fit to data before saturation and extracting slope.  See figure 1 for an example.

Results:

The results for each array are summarised in figures 2,3 and 4.  The are also given in handy ascii format at the end of the note.

As a cross check I looked at the value of the resistor during the warm functional checks (FUNC-DCU-13-PHOT_30011021_)  and during an arbitrary cold load curve (ILT_PERF_DAB_P_SinglePhase_30011394_).  The values returned using the same derivation procedure as for the gains were (in M)

|  | EIDP | Warm | Cold |
| --- | --- | --- | --- |
| PSW | 5.04986 | 4.94752 | 5.31180 |
| PMW | 5.36000 | 4.71816 | 5.06183 |
| PLW | 4.92089 | 4.84059 | 5.19588 |

PMW warm value looks anomalous here (12% deviation from EIDP?) otherwise the values are with 2-6% of expected.  Why there is a variation at all is a good question perhaps?  A quick check with another cold loadcurve showed that there is no variation from day to day.

Conclusions:

Basically the results are as expected for the JFET gain – nominal ~0.96 – with some small variation from membrane to membrane and channel to channel.  There is a slight, no more than 2-3%, overall gain variation with frequency.  This is probably due to the phase not being set absolutely precisely but it is hard to tell.  The comparison with the fixed resistor is a little strange but (apart from PMW warm measurement) the reported value is always within 6% of the expectation.  Final conclusion is that the default gains for the DCU are certainly correct to within 6% absolute and probably better than this.

Figure 1:  Example of slope fitting for gain determination.  This is for PLW.  The black line is the measured voltage versus applied bias after gain correction and addition of offsets; the red line is a first order fit to the data.  The slope of the fit is taken as the residual gain.

Figure 2: PLW JFET gains at 70 (red) 130 (green) and 189 Hz (black) versus channel number.

Figure 3: PMW JFET gains at 70 (red) 130 (green) and 189 Hz (black) versus channel number.

Figure 4: PSW JFET gains at 70 (red) 130 (green) and 189 Hz (black) versus channel number.

PLW JFET Gain versus Channel

Chan	70 Hz	130Hz	189 Hz

R1     0.963823   0.957106     0.946430

A8     0.967006   0.960047     0.948999

A7     0.964439   0.957997     0.947228

A6     0.965074   0.958440     0.947720

A9     0.966671   0.959980     0.948990

C9     0.965987   0.958503     0.947432

B8     0.965280   0.958498     0.947672

B7     0.964449   0.958060     0.947234

C7     0.966583   0.960193     0.949279

B5     0.963405   0.957715     0.947130

B6     0.964272   0.958001     0.947237

A5     0.964346   0.958164     0.947463

T1     0.966633   0.959566     0.948499

B4     0.00551464 0.006087   0.006062953

C4     0.965858   0.959429     0.948467

B3     1.33970   0.955295     0.944495

C2     0.963473   0.957327     0.946543

B2     0.966800   0.959841     0.948786

B1     0.963993   0.957756     0.947025

A3     0.964576   0.957813     0.946931

A4     0.966400   0.959106     0.948141

A1     0.963814   0.957416     0.946670

DK1   0.963415   0.957271     0.9465578

A2     0.005961   0.0065688   0.006498414

E1     0.957806   0.951848     0.940983

E2     0.956032   0.948545     0.937317

E3     0.960357   0.954380     0.943274

E4     0.959328   0.952660     0.941445

D1     0.962919   0.955422     0.943883

D2     0.962182   0.955152     0.943743

D3     0.962672   0.955765     0.944480

D4     0.960373   0.954043     0.942842

C1     0.958948   0.952056     0.940622

C3     0.960445   0.954281     0.943117

C5     0.961304   0.954570     0.943222

T2     0.959724   0.952455     0.941105

E5     0.957652   0.951372     0.940420

C6     0.961450   0.954405     0.942977

C8     0.961451   0.954552     0.943190

D5     0.961009   0.954515     0.943321

D6     0.962938   0.955806     0.944361

D7     0.961991   0.954827     0.943451

D8     0.960134   0.952599     0.941205

E7     0.959375   0.952083     0.940740

E6     0.962739   0.955214     0.943668

E8     0.960714   0.954008     0.942629

DK2   0.959971   0.952812    0.9414675

E9     0.959821   0.953570     0.942371

PMW JFET Gain versus Channel

Chan   70 Hz        130 Hz         189 Hz

A13     0.964795     0.959312     0.950434

T1       0.963550      0.958653    0.949820

B12     0.965344     0.959719     0.950672

C13     0.963876     0.959155     0.950523

A12     0.965037     0.959800     0.950795

D12     0.964528     0.959032     0.950072

C12     0.964389     0.959065     0.950278

B11     0.966732     0.960635     0.951461

A11     0.965829     0.960113     0.951012

E13     0.964547     0.958658     0.949475

D11     0.941484     0.935952     0.927010

C11     0.963684     0.959014     0.950304

B10     0.964430     0.959815     0.951071

A10     0.965426     0.960169     0.951223

D10     0.963864     0.958358     0.949370

B9       0.963240   0.958471   0.949843

C10     0.928729     0.923259     0.914357

C9     0.966133   0.960467   0.951477

A9     0.964187   0.958738   0.949923

B8     0.965205   0.960290   0.951517

A8     0.966679   0.961028   0.951950

D8     0.963504   0.958985   0.950281

C8     0.965400   0.960123   0.951132

B7     0.960831   0.955390   0.946466

R1     0.969395   0.963378   0.890225

G1   0.00708896 0.00822181 0.00856990

T2   2.65234e-009 -2.12504e-009  -2.16947e-009

E1   -0.000276926 -0.000225974   -0.00019518

D1     0.958637   0.952747   0.944016

F1     0.962675   0.957646   0.949146

E2     0.966707   0.960770   0.951955

G2     0.967297   0.962419   0.953727

F2     0.967000   0.962200   0.953586

G3     0.968446   0.962895   0.922382

E3     0.968034   0.962927   0.922507

D3     0.968264   0.963219   0.954569

F3     0.968236   0.962830   0.953716

G4   0.00600036 0.00644764 0.00646342

E4     0.967365   0.962274   0.953581

F4     0.964999   0.959305   0.935137

E5     0.962652   0.957817   0.949128

D5     0.941142   0.935477   0.926765

F5     0.969004   0.963684   0.954963

G5      1.01823   0.963453   0.954824

E6     0.969202   0.963535   0.954706

G6     0.969289   0.963790   0.954977

F6     0.968448   0.963389   0.954745

G7     0.968547   0.963165   0.954464

F10     0.967761     0.962879     0.954395

E11     0.946503     0.940895     0.932465

G11     0.953715     0.948264     0.939769

F11     0.969688     0.963897     0.955127

E12     0.969144     0.963580     0.954860

G12     0.913682     0.908713     0.901373

F12     0.963323     0.958148     0.949566

G13     0.966972     0.961567     0.953085

DK2     0.959696     0.954784     0.946262

E7     0.959064   0.953855   0.945262

D7     0.968461   0.962870   0.954216

F7   -0.0126811 -0.0118012  -0.0116634

E8     0.891568   0.887566   0.881205

G8     0.943106   0.935504   0.927282

F8     0.967900   0.962154   0.953440

E9     0.969442   0.963672   0.954973

G9   0.00469866 0.00497541 0.00484126

D9     0.935344   0.930394   0.922454

F9     0.924749   0.919380   0.911588

E10     0.966084     0.960846     0.952300

G10     0.921479     0.915897     0.677865

C4     0.965211   0.959681   0.950644

B3     0.965701   0.960157   0.951260

C3     0.965968   0.959923   0.950769

B2     0.966059   0.960628   0.951577

D2     0.964783   0.959630   0.950761

A3     0.966296   0.960733   0.951816

A2     0.968014   0.961553   0.952291

C2     0.965183   0.959865   0.950998

B1     0.965588   0.959434   0.950200

A1     0.964626   0.958762   0.949582

DK1     0.967768     0.962022     0.952908

C1     0.963542   0.956521   0.947398

A7     0.961984   0.956375   0.947533

A6     0.966522   0.960537   0.951537

B6   0.00685332 0.00746575 0.00732472

C7     0.966477   0.961246   0.952333

A5     0.962690   0.956571   0.947584

B5     0.967316   0.961354   0.952244

C6     0.962953   0.957907   0.949319

D6     0.967566   0.962528   0.953635

B4     0.961243   0.955976   0.947015

C5     0.966631   0.960699   0.951537

D4     0.967070   0.961239   0.952207

A4     0.967728   0.962066   0.952911

PSW JFET Gain versus Channel

Chan   70 Hz        130 Hz         189 Hz

R1     0.969681      0.966282        0.959973

D16     0.966425     0.963690    0.957631

T1     0.968021   0.964498  0.958267

B16     0.970253     0.966660    0.863046

C15     0.965442     0.962253    0.956230

A15     0.963050     0.960371     0.954252

D15     0.970512     0.967118    0.960882

B15     0.971019     0.967274    0.960950

C14     0.969729     0.966411    0.960152

D14     0.968268     0.965871    0.959853

A14     0.968408     0.964845     0.958559

A13     0.969879     0.966750     0.920487

B14     0.969952     0.966427    0.960119

C13     0.970161     0.966850    0.945039

B13     0.969735     0.966596    0.928444

D13     0.969658     0.966442    0.895592

A12     0.938018     0.934758     0.928582

C12     0.970940     0.967357    0.961069

D12     0.966895     0.963662    0.957465

B12     0.970561     0.967316    0.961076

E11     0.970106     0.966949    0.960720

A11     0.968009     0.964608     0.958394

C11     0.967399     0.964238    0.958043

B11     0.969512     0.965460    0.959050

E1     0.970424   0.967308  0.961294

F1     0.970607   0.967092  0.961046

T2     0.968352   0.964764  0.958628

H1     0.970618   0.967130  0.961089

G1     0.964640    0.961825   0.955858

J1     0.969245   0.966447  0.960607

H2     0.965219   0.961926  0.955871

F2     0.963572   0.960495  0.954506

J2     0.952909   0.949619  0.943636

G2     0.966576    0.963251   0.957186

H3     0.969578   0.966015  0.959941

J3     0.966008   0.962901  0.956873

E2     0.969629   0.966492  0.920447

F3     0.969070   0.966205  0.887530

G3     0.966020    0.963637   0.957921

H4     0.968372   0.965388  0.959429

J4     0.970152   0.966560  0.895831

E3     0.970520   0.967000  0.960911

F4     0.968353   0.965943  0.928249

G4     0.970752    0.966855   0.936925

H5     0.951901   0.949380  0.943557

E4     0.967463   0.964401  0.958466

J5     0.967930   0.964463  0.958299

F5     0.969553   0.966488  0.928598

D6     0.958146   0.954837  0.948358

B6     0.964669   0.961222  0.954935

C5     0.965954   0.963405  0.957319

A5     0.964681    0.960513   0.953997

E5     0.962637   0.959908  0.953766

B5     0.967195   0.964048  0.957784

D5     0.964133   0.961343  0.955283

C4     0.958281   0.955080  0.948759

A4     0.966277    0.962594   0.956248

D4     0.966606   0.962601  0.956086

B4     0.953482   0.950481  0.944243

C3     0.964931   0.961260  0.954857

B3     0.962636   0.958796  0.953722

A3     0.968082    0.964565   0.958193

A2     0.965645    0.962048   0.955633

D3     0.962529   0.960185  0.954110

C2     0.965277   0.961831  0.955665

B2     0.964587   0.961897  0.955835

D2     0.967221   0.963165  0.956801

A1     0.966016    0.963284   0.957171

C1     0.967434   0.964527  0.958321

B1     0.966646   0.962917  0.956529

DK1     0.966536     0.962680     0.956299

D1     0.960302   0.957341  0.951166

F12     0.968427     0.965305    0.959375

J11     0.968588     0.964859    0.958821

E12     0.966081     0.963848    0.958166

H12     0.927559     0.923918    0.918025

G12     0.928742     0.925365     0.919455

F13     0.945218     0.942459    0.936705

E13     0.969902     0.966694    0.960776

J12     0.961052     0.957869    0.951918

H13     0.968849     0.965629    0.959673

G13     0.966951     0.963774     0.957830

F14     0.970466     0.967404    0.961473

E14     0.967121     0.963835    0.957841

J13     0.966093     0.962730    0.956664

H14     0.967951     0.964618    0.958567

G14     0.968287     0.965074     0.959176

J14     0.966655     0.963599    0.957704

F15     0.968046     0.964884    0.958935

H15     0.965595     0.962971    0.957188

J15     0.966987     0.964035    0.958148

G15     0.969037     0.966059     0.960147

H16     0.968719     0.965488    0.959503

DK2     0.968815     0.965265     0.959237

F16     0.971151     0.967681    0.961630

E15     0.967408     0.964040    0.958073

D11     0.972181     0.968669    0.962362

A10   0.0057588   0.00635890   0.006366835

E10     0.972404     0.969119    0.962980

C10     0.968753     0.964888    0.958708

B10     0.970738     0.968212    0.962245

D10     0.970124     0.966671    0.960593

A9     0.971986    0.969313   0.963351

E9     0.972860   0.968903  0.962612

C9     0.972388   0.968877  0.962698

B9     0.971713   0.969016  0.963044

D9     0.972339   0.968969  0.962817

A8     0.972721    0.969121   0.962880

C8     0.972940   0.969411  0.963184

E8     0.972417   0.968961  0.962790

D8     0.970484   0.967653  0.961599

B8     0.972320   0.969067  0.962968

C7     0.971440   0.967998  0.961814

E7     0.967875   0.966006  0.952849

A7     0.971568    0.967921   0.961687

D7     0.972993   0.969398  0.963197

B7     0.971650   0.968309  0.962186

C6     0.971348   0.968811  0.962841

E6     0.972604   0.969090  0.962901

A6     0.971449    0.967868   0.961710

G5     0.972115    0.969004   0.962884

H6     0.970557   0.967557  0.960197

J6   0.00607183   0.729912  0.00680532

F6     0.971682   0.968108  0.962085

G6     0.967976    0.967987   0.958291

H7     0.972668   0.969709  0.963607

F7     0.974254   0.970866  0.963863

J7     0.970491   0.968603  0.962821

G7     0.971355    0.968646   0.962781

H8     0.972223   0.968817  0.962796

F8   0.00608885 0.00678184  0.00681399

G8     0.973908    0.970782   0.964673

J8     0.969631   0.966063  0.959529

F9     0.970890   0.967556  0.961614

H9     0.970626   0.967773  0.961984

G9     0.972868    0.969592   0.963551

J9     0.974554   0.971710  0.965112

F10     0.972347     0.969841    0.963879

H10     0.975106     0.972085    0.965425

G10     0.972714     0.969793     0.963239

F11     0.972827     0.969007    0.961767

J10     0.970592     0.967716    0.961920

H11     0.974378     0.971478    0.964873

G11     0.973726     0.971370     0.964862