| SUBJECT: | SPIRE Channel Gain calibration product technical note |
| --- | --- |
| PREPARED BY: | Bruce Swinyard |
| DOCUMENT No: | SPIRE-RAL-NOT-003208 |
| ISSUE: | Draft1.0 | Date: | 15th September 200829 April 2009 |
| APPROVED BY: |  | Date: |  |

Change Record

| ISSUE | DATE | Changes |
| --- | --- | --- |
| 1 | 29th April 2009 | First issue – incorporates reference to JFET gain tables |

# Introduction

## Scope of the document

This Technical Note describes the content of the SPIRE Channel Gain calibration product for both the Photometer and the Spectrometer, written for the delivery of calibration products.  This note deals with the DC signal part of the gain only.  Other notes and calibration products address the correct of the electronics to time varying signals (see AD-1 and RD-.

## Documents

### Applicable Documents

| AD-1 | The SPIRE Analogue Signal Chain and Photometer Detector Data Processing Pipeline (SPIRE-UCF-DOC-002890), Issue 6, 21/11/2008 |
| --- | --- |
| AD-2 | Karki, James, “Analysis of the Sallen-Key Architecture”, Application Report, Texas Instruments, rev. September 2002. (http://focus.ti.com/lit/an/sloa024b/sloa024b.pdf) |

### Reference Documents

| RD-1 | HERSCHEL/SPIRE Detector Control Unit Design Document (SPIRE-SAP-PRJ-001243), Issue 1.0, 11/07/2005 |
| --- | --- |
| RD-2 | SPIRE Low-pass filter calibration file description (SPIRE-BSS-DOC-003188 draft 0.2 April 2009) |
| RD-3 | Warm Photometer JFET Gain Verification (SPIRE-RAL-NOT-002872 issue 2.0 23 March 2009) |
| RD-4 | Warm Spectrometer JFET Gain Verification (SPIRE-RAL-NOT-00xxxx issue 1.0 24th Feb 2009) |

# SPIRE Channel Gains

The readout electronics for SPIRE detectors is composed by the JFET unit and the Detector Control Unit (DCU). The DCU is composed by a Lock-In Amplifier and an Analogue to Digital Converter (ADC) – see RD-1 for a full description.

## JFET Gain

There is one JFET based per-amplifier per detector.  The nominal gain for these amplifiers is 0.96, however there is some small variation from device to device and their gains were measured using SPIRE flight model electronics by feeding the bias supply directly into the input of the units – see RD-3 and RD-4.  We will use the gain per channel measured at RAL as the best value of the system level variation in gain for the spectrometer and photometer and, because we did not measure the PTC channels, we will use fixed nominal values for the PTC channels.

## Lock-In Amplifier

The signal from the JFETs is demodulated by the Lock-In Amplifier (LIA), which has three stages:

- a pre-amplifier band-pass filter to remove the DC component and amplify the signal;

- a square-wave synchronous demodulator which rectifies the signal;

- a low-pass filter.

The pre-amplifier band-pass filter has the following transfer function:

where

and ωb =2πfb is the bias angular frequency.

The parameters in the above formula are resistances and capacities that compose the circuit. The table 1 reports the values of parameters for the pre-amplifier band-pass filter extracted from RD-1 together with the maximum gain values for reference.

| Band-pass filter parameters | Photometer | TC | Spectrometer |
| --- | --- | --- | --- |
| R1 | 15 kΩ | 15 kΩ | 15 kΩ |
| R2 | 15 kΩ | 15 kΩ | 15 kΩ |
| R | 115 Ω | 316 Ω | 267 Ω |
| RB | 100 kΩ | 100 kΩ | 100 kΩ |
| CB | 47 nF | 47 nF | 47 nF |
| CH | 1.5 nF | 1.5 nF | 1.5 nF |
| S | 3.868e-3 | 3.868e-3 | 3.868e-3 |
| G | 262.8 | 96.9 | 114.4 |

Table 1: Pre-amplifier band-pass filter parameters from RD-1 (pages 36 and 49)

After the pre-amplifier band-pass filter, the signal is rectified with a square-wave synchronous demodulator. The demodulator is assumed to have gain:

when RMS voltage is used as input (see AD-1).  When the phase is correctly set – Δφ = 0 – then Hdemod(0) = 0.900316.

The low-pass filters are implemented, for the photometer, as 4-pole Bessel filters, or 2 Low-Pass Sallen-Key Circuit (AD-2) plus a final RC filter. For the spectrometer, 6-pole Bessel filters are used, or 3 Low-Pass Sallen-Key Circuit plus a final RC filter. The transfer function for the low pass filters is described in RD-2 and the values of the various components used in the transfer function are given in the SCalSpecLpfPar and SCalPhotLpfPar products described in RD-2.  We are only concerned in this document with the DC (0 Hz signal frequency) gain of the low pass filters:

HLPF(0, Phot) = 1.93

HLPF(0, Spec) = 2.86

## Digitalization

The output signal of the LIA is multiplexed, then a offset is subtracted, then the signal is amplified by a factor 12 and finally digitalized with a 16 bit ADC. The ADC range is 5 volts which are divided in 65535 values. The digitalization formula is:

where

is the DCU end to end gain. The value of Gtot for various frequencies extracted from RD-1 are reported in table 3.

| Modulation frequency [Hz] | Photometer DCU gain | TC DCU gain | Spectrometer DCU gain |
| --- | --- | --- | --- |
| 50 | 4620.49 | 1688.22 | 2954.24 |
| 60 | 4868.03 | 1777.23 | 3110.28 |
| 70 | 5037.18 | 1838.16 | 3217.09 |
| 80 | 5155.97 | 1881.15 | 3292.43 |
| 90 | 5241.45 | 1912.34 | 3347.07 |
| 100 | 5304.24 | 1935.55 | 3387.69 |
| 110 | 5351.09 | 1953.18 | 3418.51 |
| 120 | 5386.45 | 1966.83 | 3442.32 |
| 130 | 5413.32 | 1977.55 | 3460.99 |
| 140 | 5433.75 | 1986.09 | 3475.81 |
| 150 | 5449.22 | 1992.96 | 3487.69 |
| 160 | 5460.77 | 1998.53 | 3497.29 |
| 170 | 5469.18 | 2003.09 | 3505.08 |
| 180 | 5475.03 | 2006.82 | 3511.42 |
| 190 | 5478.77 | 2009.90 | 3516.59 |
| 200 | 5480.72 | 2012.43 | 3520.79 |
| 210 | 5481.16 | 2014.50 | 3524.19 |
| 220 | 5480.28 | 2016.20 | 3526.91 |
| 230 | 5478.27 | 2017.58 | 3529.06 |
| 240 | 5475.25 | 2018.69 | 3530.71 |
| 250 | 5471.33 | 2019.56 | 3531.94 |
| 260 | 5466.61 | 2020.22 | 3532.80 |
| 270 | 5461.17 | 2020.70 | 3533.32 |
| 280 | 5455.05 | 2021.02 | 3533.56 |
| 290 | 5448.33 | 2021.21 | 3533.53 |
| 300 | 5441.04 | 2021.26 | 3533.27 |

Table 3: DCU end to end gains (Gtot) from RD-1 (p. 77)

# Content of the Channel Gain Calibration product

The Channel Gain calibration product contain the gains for the JFET module and the DPU end to end gain. The DCU end to end gain is written as:

where  .

The channel gain calibration product will contain the value of the reference frequency, the Gtot at the reference frequency and the value of the parameter A. It will contain also the value of the JFET gain for each channel as given in the tables 5a-c and 6a and b.

| Parameter | Photometer | TC | Spectrometer |
| --- | --- | --- | --- |
| Gtot | 5413.32 | 1977.55 | 3497.29 |
| Ref. Frequency [Hz] | 130 | 130 | 160 |
| A | 5.85e-7 | 5.85e-7 | 3.14e-7 |
| JFET gain | Per channel values at 130 Hz | Fixed values of 0.96 | Per channel values at 160 Hz |

Table 4: Values contained in the channel gain calibration product

Table 5a:  PLW JFET Gains

| PLW | 130 Hz |
| --- | --- |
| R1 | 0.957106 |
| A8 | 0.960047 |
| A7 | 0.957997 |
| A6 | 0.958440 |
| A9 | 0.959980 |
| C9 | 0.958503 |
| B8 | 0.958498 |
| B7 | 0.958060 |
| C7 | 0.960193 |
| B5 | 0.957715 |
| B6 | 0.958001 |
| A5 | 0.958164 |
| T1 | 0.959566 |
| B4 | 0.959 |
| C4 | 0.959429 |
| B3 | 0.955295 |
| C2 | 0.957327 |
| B2 | 0.959841 |
| B1 | 0.957756 |
| A3 | 0.957813 |
| A4 | 0.959106 |
| A1 | 0.957416 |
| DK1 | 0.957271 |
| A2 | 0.957 |
| E1 | 0.951848 |
| E2 | 0.948545 |
| E3 | 0.954380 |
| E4 | 0.952660 |
| D1 | 0.955422 |
| D2 | 0.955152 |
| D3 | 0.955765 |
| D4 | 0.954043 |
| C1 | 0.952056 |
| C3 | 0.954281 |
| C5 | 0.954570 |
| T2 | 0.952455 |
| E5 | 0.951372 |
| C6 | 0.954405 |
| C8 | 0.954552 |
| D5 | 0.954515 |
| D6 | 0.955806 |
| D7 | 0.954827 |
| D8 | 0.952599 |
| E7 | 0.952083 |
| E6 | 0.955214 |
| E8 | 0.954008 |
| DK2 | 0.952812 |
| E9 | 0.953570 |

Table 5b: PMW JFET Gains

| PMW | 130 Hz |
| --- | --- |
| A13 | 0.959312 |
| T1 | 0.958653 |
| B12 | 0.959719 |
| C13 | 0.959155 |
| A12 | 0.959800 |
| D12 | 0.959032 |
| C12 | 0.959065 |
| B11 | 0.960635 |
| A11 | 0.960113 |
| E13 | 0.958658 |
| D11 | 0.935952 |
| C11 | 0.959014 |
| B10 | 0.959815 |
| A10 | 0.960169 |
| D10 | 0.958358 |
| B9 | 0.958471 |
| C10 | 0.923259 |
| C9 | 0.960467 |
| A9 | 0.958738 |
| B8 | 0.960290 |
| A8 | 0.961028 |
| D8 | 0.958985 |
| C8 | 0.960123 |
| B7 | 0.955390 |
| R1 | 0.963378 |
| G1 | 0.963 |
| T2 | 0.963 |
| E1 | 0.963 |
| D1 | 0.952747 |
| F1 | 0.957646 |
| E2 | 0.960770 |
| G2 | 0.962419 |
| F2 | 0.962200 |
| G3 | 0.962895 |
| E3 | 0.962927 |
| D3 | 0.963219 |
| F3 | 0.962830 |
| G4 | 0.963 |
| E4 | 0.962274 |
| F4 | 0.959305 |
| E5 | 0.957817 |
| D5 | 0.935477 |
| F5 | 0.963684 |
| G5 | 0.963453 |
| E6 | 0.963535 |
| G6 | 0.963790 |
| F6 | 0.963389 |
| G7 | 0.963165 |
| F10 | 0.962879 |
| E11 | 0.940895 |
| G11 | 0.948264 |
| F11 | 0.963897 |
| E12 | 0.963580 |
| G12 | 0.908713 |
| F12 | 0.958148 |
| G13 | 0.961567 |
| DK2 | 0.954784 |
| E7 | 0.953855 |
| D7 | 0.962870 |
| F7 | 0.963 |
| E8 | 0.887566 |
| G8 | 0.935504 |
| F8 | 0.962154 |
| E9 | 0.963672 |
| G9 | 0.963 |
| D9 | 0.930394 |
| F9 | 0.919380 |
| E10 | 0.960846 |
| G10 | 0.915897 |
| C4 | 0.959681 |
| B3 | 0.960157 |
| C3 | 0.959923 |
| B2 | 0.960628 |
| D2 | 0.959630 |
| A3 | 0.960733 |
| A2 | 0.961553 |
| C2 | 0.959865 |
| B1 | 0.959434 |
| A1 | 0.958762 |
| DK1 | 0.962022 |
| C1 | 0.956521 |
| A7 | 0.956375 |
| A6 | 0.960537 |
| B6 | 0.961 |
| C7 | 0.961246 |
| A5 | 0.956571 |
| B5 | 0.961354 |
| C6 | 0.957907 |
| D6 | 0.962528 |
| B4 | 0.955976 |
| C5 | 0.960699 |
| D4 | 0.961239 |
| A4 | 0.962066 |

Table 5c: PSW JFET Gains

| PSW | 130Hz |
| --- | --- |
| R1 | 0.966282 |
| D16 | 0.963690 |
| T1 | 0.964498 |
| B16 | 0.966660 |
| C15 | 0.962253 |
| A15 | 0.960371 |
| D15 | 0.967118 |
| B15 | 0.967274 |
| C14 | 0.966411 |
| D14 | 0.965871 |
| A14 | 0.964845 |
| A13 | 0.966750 |
| B14 | 0.966427 |
| C13 | 0.966850 |
| B13 | 0.966596 |
| D13 | 0.966442 |
| A12 | 0.934758 |
| C12 | 0.967357 |
| D12 | 0.963662 |
| B12 | 0.967316 |
| E11 | 0.966949 |
| A11 | 0.964608 |
| C11 | 0.964238 |
| B11 | 0.965460 |
| E1 | 0.967308 |
| F1 | 0.967092 |
| T2 | 0.964764 |
| H1 | 0.967130 |
| G1 | 0.961825 |
| J1 | 0.966447 |
| H2 | 0.961926 |
| F2 | 0.960495 |
| J2 | 0.949619 |
| G2 | 0.963251 |
| H3 | 0.966015 |
| J3 | 0.962901 |
| E2 | 0.966492 |
| F3 | 0.966205 |
| G3 | 0.963637 |
| H4 | 0.965388 |
| J4 | 0.966560 |
| E3 | 0.967000 |
| F4 | 0.965943 |
| G4 | 0.966855 |
| H5 | 0.949380 |
| E4 | 0.964401 |
| J5 | 0.964463 |
| F5 | 0.966488 |
| D6 | 0.954837 |
| B6 | 0.961222 |
| C5 | 0.963405 |
| A5 | 0.960513 |
| E5 | 0.959908 |
| B5 | 0.964048 |
| D5 | 0.961343 |
| C4 | 0.955080 |
| A4 | 0.962594 |
| D4 | 0.962601 |
| B4 | 0.950481 |
| C3 | 0.961260 |
| B3 | 0.958796 |
| A3 | 0.964565 |
| A2 | 0.962048 |
| D3 | 0.960185 |
| C2 | 0.961831 |
| B2 | 0.961897 |
| D2 | 0.963165 |
| A1 | 0.963284 |
| C1 | 0.964527 |
| B1 | 0.962917 |
| DK1 | 0.962680 |
| D1 | 0.957341 |
| F12 | 0.965305 |
| J11 | 0.964859 |
| E12 | 0.963848 |
| H12 | 0.923918 |
| G12 | 0.925365 |
| F13 | 0.942459 |
| E13 | 0.966694 |
| J12 | 0.957869 |
| H13 | 0.965629 |
| G13 | 0.963774 |
| F14 | 0.967404 |
| E14 | 0.963835 |
| J13 | 0.962730 |
| H14 | 0.964618 |
| G14 | 0.965074 |
| J14 | 0.963599 |
| F15 | 0.964884 |
| H15 | 0.962971 |
| J15 | 0.964035 |
| G15 | 0.966059 |
| H16 | 0.965488 |
| DK2 | 0.965265 |
| F16 | 0.967681 |
| E15 | 0.964040 |
| D11 | 0.968669 |
| A10 | 0.968 |
| E10 | 0.969119 |
| C10 | 0.964888 |
| B10 | 0.968212 |
| D10 | 0.966671 |
| A9 | 0.969313 |
| E9 | 0.968903 |
| C9 | 0.968877 |
| B9 | 0.969016 |
| D9 | 0.968969 |
| A8 | 0.969121 |
| C8 | 0.969411 |
| E8 | 0.968961 |
| D8 | 0.967653 |
| B8 | 0.969067 |
| C7 | 0.967998 |
| E7 | 0.966006 |
| A7 | 0.967921 |
| D7 | 0.969398 |
| B7 | 0.968309 |
| C6 | 0.968811 |
| E6 | 0.969090 |
| A6 | 0.967868 |
| G5 | 0.969004 |
| H6 | 0.967557 |
| J6 | 0.967 |
| F6 | 0.968108 |
| G6 | 0.967987 |
| H7 | 0.969709 |
| F7 | 0.970866 |
| J7 | 0.968603 |
| G7 | 0.968646 |
| H8 | 0.968817 |
| F8 | 0.969 |
| G8 | 0.970782 |
| J8 | 0.966063 |
| F9 | 0.967556 |
| H9 | 0.967773 |
| G9 | 0.969592 |
| J9 | 0.971710 |
| F10 | 0.969841 |
| H10 | 0.972085 |
| G10 | 0.969793 |
| F11 | 0.969007 |
| J10 | 0.967716 |
| H11 | 0.971478 |
| G11 | 0.971370 |

Table 6a: SLW JFET Gain

| SLW | 160 Hz |
| --- | --- |
| R1 | 0.969421 |
| T1 | 0.968814 |
| C1 | 0.968967 |
| DK1 | 0.967182 |
| B1 | 0.968715 |
| D1 | 0.967019 |
| E1 | 0.967029 |
| A1 | 0.969438 |
| C2 | 0.968 |
| D2 | 0.968810 |
| B2 | 0.965399 |
| E2 | 0.969509 |
| A2 | 0.967010 |
| C3 | 0.969452 |
| D3 | 0.966752 |
| B3 | 0.967136 |
| E3 | 0.967049 |
| C4 | 0.968994 |
| DK2 | 0.969287 |
| D4 | 0.969575 |
| C5 | 0.966935 |
| B4 | 0.969595 |
| A3 | 0.969589 |
| T2 | 0.967410 |

Table 6b: SSW JFET Gain

| SSW | 160 Hz |
| --- | --- |
| R1 | 0.966656 |
| A4 | 0.967052 |
| A3 | 0.960958 |
| A2 | 0.966977 |
| A1 | 0.966246 |
| DK1 | 0.966134 |
| B3 | 0.960801 |
| B2 | 0.965727 |
| B1 | 0.966622 |
| C3 | 0.966589 |
| C2 | 0.966611 |
| C1 | 0.966588 |
| D3 | 0.966960 |
| D2 | 0.966187 |
| D1 | 0.964819 |
| E3 | 0.966087 |
| E2 | 0.966252 |
| E1 | 0.966553 |
| F3 | 0.966163 |
| F2 | 0.956321 |
| F1 | 0.966000 |
| G1 | 0.966227 |
| T1 | 0.944595 |
| G2 | 0.939484 |
| E5 | 0.965506 |
| E4 | 0.961355 |
| D7 | 0.960 |
| D6 | 0.960370 |
| D5 | 0.962957 |
| D4 | 0.961996 |
| C6 | 0.965616 |
| C5 | 0.964981 |
| C4 | 0.965040 |
| B5 | 0.964331 |
| B4 | 0.964868 |
| T2 | 0.965190 |
| G3 | 0.965484 |
| G4 | 0.932831 |
| DK2 | 0.964353 |
| F5 | 0.965167 |
| F4 | 0.962758 |
| E6 | 0.965049 |