### SPIRE

#### Subject:  BDA Electronic Handling Procedure, SPIRE P/SW-PFM S/N013

Prepared by:  Anthony Turner

Document No:

Issue:   Draft								Date: 2/22/05

Checked by: …………………………………….		Date:……………………

Approved by:……………………………………		Date:……………………

# Distribution

Change Record

Issue				Date

Table of Contents

- Introduction…..………………………………………	page	6

- Handling……………………………………………...   	page	6

- Signal Requirements………………………………...	page	6

- Device Isolation………………………………………	page	7

- Room Temperature Detector Values Check……….	page	7

- Load Resistor-Detector Continuity Check…………	page	14

Glossary

1.  Introduction:

This document provides the Electronic Handling Procedure for the Proto-Flight Model-Photometer Short Wavelength Bolometer Detector Array serial number 013.

2.  Handling:

- BDA is Contamination Sensitive:  Handle BDA with Gloves only in a FED-STD-209 Class 10000 clean room (ISO 14644-1 class 7) or better.

- BDA is ESD Sensitive:  Electronic parts included in the P/SW-PFM S/N013 science instrument are subject to electro-static discharge failures.  Please handle with appropriate ESD hardware handling procedures.  Handle with grounding straps, ESD-safe gloves, ESD smocks at an ESD-safer workstation.

3.  Signal Requirements:

The interface circuit for the BDA contains a series of resistive networks as depicted in figure 1.  Two high resistive load resistors (~ 6-14 M() are coupled to a NTD Ge thermistor (R bolo) through a lithographed metalization circuit and provide the bias circuitry for the device.  The maximum DC input voltage for the bias lines V+ and V- lines is +/- 1 V, and the maximum AC input voltage is 100mV rms.

Figure 1:  Interface circuit of the Bolometer Detector Array

DC Voltage-Current Limits for Room Temperature Detector Values Check:

| Signal | Pin | Nominal Value | Max Value |
| --- | --- | --- | --- |
| V+ | 25 | +50mV | +1 V |
| V- | 50 | -50mV | -1 V |
| I+,I- | 25,50 | 10(A | 25(A |

DC Voltage-Current Limits for room temperature Load Resistor-Detector Continuity Check:

| Signal | Pin | Nominal Value | Max Value |
| --- | --- | --- | --- |
| V+ | 25 | +50mV | +1 V |
| V- | 50 | -50mV | -1 V |
| I+, I- | 25,50 | <0.5(A | 1(A |

## 4. Device Isolation Check:

The 300mK stage Kevlar suspended portion of the detector chassis is grounded directly to the electronic ground on pin 51 of the nanonics 51 pin connectors on each side.  A resistance of less than 200( can be checked from the thermal strap of the 300mK stage to electrical ground pin 51.   The 2K stage is electrically isolated from the electronic ground via the Kevlar supports.  Measuring from pin 51 to any metal section of the 2K stage will yield an open circuit.

## 5. Room Temperature Detector Values Check

The final measured DC resistance value for each of the bolometer detector at room temperature is shown in tables 1 through 4.  The measurements bypass the load resistors in the circuit measuring directly through the output signal pins on the two nanonics 51 pin connectors.  All measurements were performed with a Fluke 87 True RMS Multi-meter set in the 4k( range.  All measurements are in k( unless designated otherwise.  Channels that are out of range are re-measured using the 40M( range to determine their value.  The failure mode of any particular channel is also designated in tables 1 through 6.  The designation for the failure modes are open- Channel open at 300mK, short-channel shorted at 300mK, float- channel floating at 300mK or caution- for cautionary concerning due to detector performance status change during testing at 300mK.

## Table 1:  P/SW-PFM S/N013 Room Temperature DC Detector Measurements J01 connector

## Table 2:  P/SW-PFM S/N013 Room Temperature DC Detector Measurements J02 connector

## Table 3:  P/SW-PFM S/N013 Room Temperature DC Detector Measurements J03 connector

Failure mode notes:  Channel C12 was non functional at room temperature, 4K and 300mK.

## Table 4:  P/SW-PFM S/N013 Room Temperature DC Detector Measurements J04 connector

## Table 5:  P/SW-PFM S/N013 Room Temperature DC Detector Measurements J05 connector

## Table 6:  P/SW-PFM S/N013 Room Temperature DC Detector Measurements J06 connector

Failure mode notes:  Channel J7 has a problematic In bump bond and has been designated as a non-functioning pixel at 300mK.  Channel G11 was functional at 300mK but logged to have an open LR during assembly.

6.  Load Resistor-Detector Continuity Check

A DC continuity check of the load resistors in series with the bolometer detectors will complete the electrical checkout at room temperature.  The test can be performed with a Fluke 87 True RMS multi-meter set on the 40M( scale.  The data set measures from V+ to output signal S+ and V- to output signal S- for each channel.  The nominal value read for the live bolometer channels (room temp detector DC resistance ~ 1.5k() should read approximately 3-6M(.  Channels with open bolometer channels will give values 8M( or higher.  The Data sets for the P/SW -PFM S/N013 for the final test through the entire circuit are shown in tables 7 through 12.

## Table 7:  P/SW-PFM S/N013 Load Resistor- Detector DC Continuity Check J01 connector

## Table 8:  P/SW-PFM S/N013 Load Resistor- Detector DC Continuity Check J02 connector

## Table 9:  P/SW-PFM S/N013 Load Resistor- Detector DC Continuity Check J03 connector

## Table 10:  P/SW-PFM S/N013 Load Resistor- Detector DC Continuity Check J04 connector

## Table 11:  P/SW-PFM S/N013 Load Resistor- Detector DC Continuity Check J05 connector

## Table 12:  P/SW-PFM S/N013 Load Resistor- Detector DC Continuity Check J06 connector

Failure mode note:  Channel G11 was measured as shown to have an open LR during assembly.  Re-measurements after performance testing completed on 2/17/2005 revealed the V+ to S+ and V- to S- to now be in the correct range (4.07Mohms and 4.12Mohms respectively).  Channel is listed as cautionary due to this change in status.

# Text boxes

R bolo

R load

R load

R load

R load

R bolo

Output

Signal S-

Output

Signal S+

GND

Pin 51

V-

V+