Complementary straylight measurements on AMIE EM

Date: 20-21 March 2003

Location: ESTEC, Optical lab clean room

Model under test: AMIE Camera EM

The spot given by the lamp source used (same as for measurements on FM in Dec. 2001) shows a ghost image when projected on a white sheet of paper. Indeed, data obtained during the straylight measurements on the FM show such a ghost in the images.

Pictures of this ghost have been taken with a digital camera. The figure below give the positions of the white sheet when taking pictures of the spot: the camera and the support were removed, a sheet of white paper was placed at positions 1, 2, 3 and 4 and pictures of the spot issued from the light source were taken at the different positions.

Positions of the white sheet when taking pictures of the spot

Picture of the spot at position 1

Picture of the spot at position 2

Picture of the spot at position 3

Picture of the spot at position 4

Test set-up (top view)

For the measurements, the source has rotation capabilities (centre of rotation at the extremity of the teleobjective) and the camera can be tilted.

Tilt of the camera

First set-up for horizontal measurements. In the final set-up, the extremity of the teleobjective is at the centre of rotation of the source, as depicted in the set-up for vertical measurements picture

Measurements for horizontal position:

- No tilt of the camera, images with source at center (0°), then rotation of the source from 1° to +10° (step 1°), then back at center, then rotation of the source from –1° to –10° (step 1°)

- Tilt of the camera +1°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera –1°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera –2°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera +2°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera +3°, images with source at center (0°)

- Tilt of the camera +4°, images with source at center (0°)

- Tilt of the camera –3°, images with source at center (0°)

Set-up for vertical measurements

Measurements for vertical position:

- No tilt of the camera, images with source at center (0°), then rotation of the source from 1° to +6° (step 1°), then back at center, then rotation of the source from –1° to –6° (step 1°)

- Tilt of the camera +1°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera –1°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera –2°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)

- Tilt of the camera +2°, images with source at center (0°), then rotation of the source from 2° to +6° (step 2°), then rotation of the source from –2° to –6° (step 2°)