Compact high-resolution space-borne echelle grating spectrometer with AOTF based order sorting for the infrared domain from 2.2 to 4.3 micrometer.

Dennis Nevejans, Eddy Neefs, Emiel Van Ransbeeck, Sophie Berkenbosch, Roland Clairquin

Belgian Institute for Space Aeronomy (BIRA-IASB) Ringlaan 3 B-1180 Brussels Belgium

Lieve De Vos, Wouter Moelans, Stijn Glorieux, Ann Baeke

OIP Sensor Systems, Westerring 21, B-9700 Oudenaarde Belgium

Oleg Korablev, Imant Vinogradov

Space Research Institute (IKI), 117997, 84/32 Profsoyuznaya Str, Moscow, Russia

Yuri Kalinnikov

AFAR, NPO SPURT, 141570, Zelenograd, Moscow Region, Russia

Benny Bach

Bach Research Corp., 2200 Central Avenue Suite D, Boulder, CO 80301, U.S.A.

Jean-Pierre Dubois, Eric Villard

Service d’Aéronomie du CNRS, BP 3, F-91371 Verrières le Buisson Cédex, France.

Abstract

A new compact space borne high resolution spectrometer developed for the ESA Venus Express spacecraft is described. It operates in the IR wavelength range of 2.2 to 4.3 µm and measures absorption spectra of minor constituents in the Venusian atmosphere. It uses a novel echelle grating with groove density of 4 lines per mm in a Littrow configuration, in combination with an IR Acousto-Optic Tunable Filter for order sorting and an actively cooled HgCdTe Focal Plane Array of 256 by 320 pixels. It is designed to obtain an Instrument Line Profile of 0.2 cm-1. First results on optical and spectral properties are reported.

Copyright

OCIS codes  120.4570, 120.6200, 010.1280, 230.1040, 230.1950, 300.6190, 350.6090

Introduction

Until now, all infra-red (IR) optical spectrometers flown on board of spacecraft sent out to measure absorption spectra in the atmospheres of solar system planets, were limited to spectral resolutions of 1 cm-1 or worse. This has seriously hampered the detection of minor atmospheric species. A typical example is the Planetary Fourier Spectrometer (PFS)1,2, an instrument now operating on-board of the ESA Mars Express and Venus Express spacecrafts. Its unapodized spectral resolution is limited to 1-2 cm-1.  Spectrometers3 on earlier planetary missions, such as Voyager and Mariner 9, were Fourier Transform Spectrometers (FTS) having even lower resolutions, ranging from 2.4 to 5 cm-1. These resolutions were much lower than those of optical spectrometers observing the Earth’s atmosphere from orbit: e.g. an older Earth orbiting spectrometer like ATMOS (Atmospheric Trace Molecule Spectroscopy instrument)4 had an unapodized resolution of 0.125 cm-1 and the recent ACE-FTS5 has one of 0.025 cm-1. Ground based instruments reach easily ten times better resolution. On the other hand, all these high resolution IR spectrometers are bulky instruments, making it difficult to install them on planetary or interplanetary spacecrafts having limited resources.

For the ESA Venus Express mission to planet Venus a compact IR spectrometer having low mass, high throughput, rather high resolution (10 times better than PFS), a rather large free spectral range (about 20 cm-1) and a spectral sampling interval and an instrument line profile of 0.1 cm-1 and 0.2 cm-1 respectively, was envisaged. These requirements can be fulfilled by combining a highly dispersive element operated at high diffraction orders (a coarse-spaced echelle grating) with a diffraction order sorting device. Instead of selecting grating orders by means of interchangeable band pass filters on a rotating filter wheel (e.g. the Cryogenic Echelle Spectrograph of the NASA Infrared Telescope Facility) or by means of an orthogonal cross dispersing element, as in VIRTIS6 (Visible and Infra Red Thermal Imaging Spectrometer on the ESA ROSETTA, Mars Express and Venus Express spacecrafts), order filtering can be implemented by inserting an in-line acousto-optic device, known as AOTF (Acousto-Optical Tunable Filter), that may be considered as an electronically tunable constant bandwidth band pass filter.

A first echelle-AOTF combination for the UV-visible (200 to 425 nm) wavelength range was reported in literature in 1999. This Compact High Resolution Spectrometer (CHRS)7,8,9 yielded a spectral resolution of 0.004 to 0.008 nm and realized approximately a tenfold reduction in size and weight, compared to competing baseline instruments that provide comparable resolution. The authors claim it offers also less complex spectrum detection and interpretation than the configuration of an echelle combined with orthogonal dispersion.

The working principle of the echelle-AOTF combination has since then been extended to the near IR (NIR) range of wavelengths and has been demonstrated in two implementations: one10,11, covering the wavelength range from 1 to 1.65 µm with a resolving power of 30000, manufactured, assembled and tested by a joint French-Russian team (Service d’Aéronomie and IKI), and another which was intended as a precursor for the space borne IR spectrometer described in this paper, built by the Belgian Institute for Space Aeronomy (BIRA-IASB) and its industrial partner OIP Sensor Systems (OIP), Oudenaarde, Belgium.

In this paper we describe an echelle based spectrometer with AOTF order sorting based on the same basic principles, but specifically designed for the 2.2 to 4.3 µm range in the IR. This high resolution spectrometer12-15, called SOIR (Solar Occultation in IR), was accepted as part of the scientific payload for the Venus Express mission of the European Space Agency (ESA). The detailed optical study and design, as well as the manufacturing of the instrument, were performed by OIP in association with BIRA-IASB.

Summary of scientific goals for Venus Express

This spectrometer fits within the defined scientific mission themes16 of the ESA Venus Express (VEX) mission, such as the study of dynamics, structure, composition and chemistry of the Venusian atmosphere, and is highly appropriate for the investigation of the altitude range between 110 and 60 km. The application in the Venusian atmosphere of the solar occultation technique, for which this instrument is particularly designed, is limited because of the existence between altitudes from 45 km to 70 km of thick cloud layers, composed of fine droplets of sulphuric acid. However, above these cloud layers the sounding of the atmosphere and the measurement of the vertical abundance profiles of atmospheric constituents is perfectly feasible up to high altitudes.

One of the key scientific themes of the Venus Express mission is the understanding of the past and present water balance in the Venusian atmosphere. The study of the water escape from Venus by means of the measurement of HDO and H2O profiles is therefore of utmost importance. Absorption lines of HDO and H2O of the dark side of Venus have been observed from Earth in the interval 2.34 to 2.43 µm with a resolving power of 0.23 cm-1, which is comparable to the resolving power offered by the described instrument.  These results17 suggest a pronounced deuterium enrichment in the atmosphere of Venus, expressed by an abundance ratio D/H about 120 times larger than the telluric ratio. The measurement of the HDO escape profile between 60 and 80-90 km altitude by SOIR on VEX could confirm or deny this ground based observation.

Besides the quest for HDO, another important scientific objective for SOIR is the measurement of the water vapor abundance profile up to 100 km, in either the 2560-2570 cm-1 (3.9 µm) band or in the stronger 3760 to 3860 cm-1 (2.56 µm) absorption band. Other minor molecular components such as H2, O2, CO, sulphuric compounds like H2S, COS and SO2 and acid vapors like HCl and HF are certainly present and detectable, as gas chromatographic in-situ abundance measurements by the Venera probes18,19 and observations from ground have shown. Also, the detection of minor constituents of the Venusian atmosphere associated with small absorption features, like C2H2, CH4, NO, N2O, H2CO, HCN will be attempted.

Finally, one can derive the Venusian atmosphere’s density profile by measuring the density of its major component. By recording CO2 absorption lines in the 4.3 µm band one can cover at once a large part of the whole altitude range from 60 to 120 km, and probably beyond.

Spectrometer design objectives

The first opportunity to design and manufacture an echelle based high-resolution spectrometer with AOTF order sorting for the 2.2 to 4.3 µm range in the IR appeared when the ESA Venus Express mission was announced. The payload on-board the Venus Express (VEX) spacecraft comprises mainly upgraded versions of instruments already flown on the Mars Express (MEX) and ROSETTA platforms, with some new instruments added. The SOIR instrument (Figure 1) described in this paper is to be considered as a new extension on top of the existing SPICAM-L experiment flown on MEX. The combination of both was called SPICAV.

In accordance with the scientific objectives the SOIR design was tuned for solar occultation observations of the Venusian atmosphere in the 2 to 5 µm band, a requirement that was relaxed to the wavelength range of 2.325 to 4.25 µm because of limitations imposed by the optics and the detector.

A second requirement was to achieve a spectral resolution that was about 10 times better than could be reached before for equal size and mass. An instrument line profile (ILP), full width at half maximum (FWHM), of 0.2 cm-1 and a sampling interval of 0.1 cm-1 per pixel over the whole spectral range were therefore put forward. This will allow deriving the density profile of already known atmospheric constituents at higher altitudes than feasible before; it will also allow the detection of yet unknown atmospheric constituents.

The application of the Sun as a bright light source for the solar occultation technique (Figure 2) determined the field of view (FOV) of the instrument. It constraints which rays emitted from the solar disk, and being refracted in the Venusian atmosphere, will finally enter the spectrometer and reach the detector. Since, at Venus, the Sun has an apparent diameter of 44', the FOV in the spatial direction was limited to 30', in order to avoid solar limb darkening effects. This choice is also compatible with observing the Sun (apparent diameter 32') during ground based tests. For the spectral direction, perpendicular to the spatial direction, a FOV of 2' was considered allowing for a good height resolution during the inversion process of the Venus occultation data. Hence, a rectangular field of view of 2' by 30' was defined. It is the aim to control the attitude of the VEX spacecraft during an occultation observation such that the boresight of the SOIR instrument is pointing the Sun and that the long side of the rectangular FOV remains parallel to the planet’s limb.

Since SOIR had to be put on top of the already existing SPICAM-L part of the SPICAV experiment, the mass available for SOIR was limited to 6.5 kg and its volume to 414 mm by 254 mm by 210 mm. In addition, for thermal reasons the spectrometer’s entrance aperture, and the corresponding aperture in the satellite wall, had to be on a specific side (+Y) of the spacecraft and were restricted in diameter.  Therefore, a periscope-like device had to be added to the side wall of the instrument (visible on Figures 1b and 10), since the VEX solar viewing boresight was at an angle of 60° with respect to the SOIR optical axis.  One of the most severe design constraints was the thermal environment imposed by the satellite in operational (-20 °C to +40 °C) and non-operational (-30 °C to +50 °C) conditions. Table 1 summarizes the spectrometer design objectives.

Spectrometer design

The main drivers for the SOIR design were: the available solar input, the imposed maximum entrance aperture diameter, the signal-to-noise ratio, the desired FOV, the desired spectral sampling interval, the desired ILP, the pixel size and finally the acceptance angle of the AOTF.

It was decided to construct the spectrometer using an echelle grating as a diffracting device; this grating type is known to yield high reciprocal dispersion and high throughput in a compact design. Since an echelle, like all gratings, is generating overlapping orders, an order sorting filtering was required, in this case a tunable bandpass filter. Compared to orthogonal cross-dispersion solutions using prisms or low dispersion gratings, this technique allows the imaging of longer spectral lines on the detector and hence a larger binning factor along the lines, which results in higher signal-to-noise ratios. In SOIR the bandpass filter is an Acousto Optical Tunable Filter (AOTF). It has both the advantage of avoiding mechanical moving mechanisms and of a quick random access to any grating order by electronic control. Moreover, by deactivating the AOTF, one can stop all solar input to the spectrometer; this allows measuring a thermal background spectrum that can be subtracted from the observed spectrum. Figure 3 represents a simplified 2-D set-up of a compact high-resolution spectrometer using an AOTF-echelle combination.

The spectrometer is composed of three main parts: (a) the front-end part which collects the solar light, defines the FOV and restricts the observed wavelength domain; (b) the spectrometer itself which realizes the free spectral range and the ILP and finally (c) the detector system that records the spectra according to a spectral sampling interval.

The front-end starts at the AOTF entrance optics (1) that reduce the diameter of the incoming light beam so that it becomes compatible with the AOTF acceptance aperture. In the intermediate image plane of the entrance optics a first diaphragm (2) is placed that limits the FOV to slightly larger than the desired FOV. Then the beam enters the AOTF (3) to filter out unwanted orders. When activated with an RF signal, the AOTF directs the small fraction of the beam that corresponds to the wavelength domain to be analyzed along its optical axis. The AOTF exit optics (4) images the beam on the spectrometer entrance slit (5), which is the entrance aperture for the spectrometer part of the instrument. A collimating lens (6) captures the light passing through the spectrometer entrance slit, collimates and transports it as a parallel beam to the dispersing element (7), the echelle grating located between collimator lens and imaging lens. Finally, the light diffracted by the grating is imaged via a camera lens (8) on the detector (9).

In order to obtain a compacted optical scheme for the high-resolution spectrometer part, a quasi-Littrow configuration was selected, in which the function of the collimation lens and of the imaging lens are merged into a single off-axis parabolic mirror. An additional benefit is that no co-alignment of two separate lenses is needed. Figure 4 depicts the detailed optical scheme and the associated ray tracing that will be described hereafter. The main recognizable optical units are: the entrance optics (1), the diaphragm (2), the AOTF (3), the AOTF exit optics (4), the spectrometer slit (5), the off-axis parabolic mirror (6), the echelle grating (7), the folding mirror (8), the detector optics (9) and finally the detector (10).

Imager focal length and spectral sampling interval

The focal length of the imaging mirror in the spectrometer, together with the angular dispersion of the echelle grating and the pixel size of the IR detector, define the spectral sampling interval (SSI) of the system. As the SSI is imposed (Table 1), the required focal length of the imaging lens can be derived.

The angular dispersion20 D of a non-tilted echelle grating at Littrow condition is given by:

(1)

with	D,  λ and θ being the angular dispersion in radians per wavelength unit, the wavelength and the blaze angle of the echelle.

The spectral sampling interval (SSI)20 expressed in wavelength units per pixel is then given by:

(2)

and the derived focal length by

(3)

with	p and fimager being the detector pixel size and the focal length of the imaging mirror.

When the SSI is expressed in wavenumber units per pixel, equation (3) becomes:

(4)

It is clear that the focal length computed for a given grating blaze angle and detector pixel size, and for a desired spectral sampling interval expressed in wavenumbers, depends inversely on the wavelength. Knowing the specified SSI (0.1 cm-1) over the whole wavelength range, the pixel size (30 µm) and the blaze angle of the grating (63.43° or arctan(2)), the derived imager focal length becomes 323 mm @ 2.325 µm and 177 mm @ 4.25 µm. Since the optical design had to fit the available volume, a compromise value of 300 mm was adapted for the focal length, thus fulfilling the SSI requirement for wavelengths between 2.5 and 4.25 µm, with a slight deviation between 2.325 and 2.5 µm.

Echelle free spectral range and AOTF characteristics

As demonstrated by formula (5) the free spectral range20 (FSR) of the echelle grating, the range of wavelengths in which there is no interference or superposition by light from adjacent orders, is independent of wavelength when expressed in wavenumber units:

(5)

where d represents the groove step, θ the blaze angle and φ the tilting angle of the echelle.

The FWHM AOTF bandwidth, expressed in wavenumber units, being almost constant over the complete wavelength range, the FSR of the echelle and the AOTF bandwidth can be matched. For a groove step of 250 µm, a blaze angle ( of arctan 2 and a tilting angle φ of 2.6°, the echelle’s FSR equals 22.38 cm-1, which corresponds to an AOTF bandwidth that can be achieved with state of the art technology.

The AOTF used in this instrument has been custom designed and manufactured in Russia by AFAR, NPO SPURT (Zelenograd, Moscow region) according to the specifications stated in Table 2. The actual IR AOTF device consists in a birefringent TeO2 (paratellurite) tetragonal crystal (group 422) that is transparent in the spectral range of 0.35-5 µm. Acousto-Optical Tunable Filters made of this material as well as their design are extensively described in literature21.

The SOIR AOTF is employed in the classical non-collinear large-aperture configuration22 in which a non-polarized incident light beam produces two beams with different polarizations that are diffracted at a small angle with respect to the undiffracted output beams. In SOIR we use only the ordinary polarization of the incoming radiation, (perpendicular to the drawing plane in Figure 5) in order to keep the design simple. At the AOTF output the polarization of the diffracted beam changes to extraordinary. The deflection angle of the diffracted beam, due to Bragg diffraction by the periodic acoustic wave combined with a rotation of the polarization plane, is only weakly dependent on wavelength. This dependency is compensated for by the inclined output facet of the AOTF (4.2° in Figure 6). This oblique cut also refolds the diffracted output beam so that it remains nearly collinear with the incoming beam. The strong undiffracted beam exits the crystal at an angle of 4.6°, and can therefore in a easy way be separated from the useful output beam.

The central wavelength of the AOTF band pass is inversely proportional to the frequency of the applied acoustic wave interacting with the optical waves. The acoustic wave is injected into the crystal by means of a transducer that is converting an electrical RF wave into acoustic power. The RF generator, which generates the RF wave, is itself driven by a small signal RF carrier produced in the SOIR electronics. Both frequency and amplitude can be electrically tuned at random.

The power consumed by an AOTF operating at long IR wavelengths is a potential problem. It not only affects the power budget and thermal regime of the whole instrument, but internal thermal dissipation directly influences the tuning and bandpass characteristics of the AOTF. In order to minimize required RF power, the acousto-optic interaction length should be made as large as possible. This length is constrained by the availability of big pure TeO2 crystals: 36 mm being the maximum that could be achieved, the AOTF configuration as shown in Figure 6 was imposed.

Orders and wavenumber ranges

As the grating groove step and the focal length of the imaging mirror are determined, the wavelength falling on the SOIR detector central column ((c) as well as the wavelengths falling on the edge columns of the detector in the spectral direction ((e) can be computed20 by means of formulas (6) and (7) for each order m.

(6)

(7)

Here d represents the groove step of the echelle, ( the blaze angle, i the angle between incident ray and the echelle facet normal laying in the plane that is perpendicular to the grooves and contains the echelle normal, ( the angle between the incident ray and this plane, and (φ the angle between the optical axis and an edge of the detector sensitive area. It is clear that the product <order times wavelength> remains constant for the detector central and edge columns; hence there is a linear relationship between the wavenumber associated with a pixel and the order number.

Table 3 summarizes the wavenumbers associated with detector pixels 0 and 319 for the echelle orders 101 to 194, computed using formulas (6) and (7), and corrected to take into account a small deviation from the ideal Littrow condition and a small asymmetric distortion in the spectrometer optics. This leads to 94 small wavenumber domains covering the range from 2256 to 4369 cm-1. For orders 101 to 122 the detector width expressed in wavenumbers is smaller than the echelle’s FSR (22.38 cm-1) and hence the detector will miss part of the spectrum. For orders 123 to 194 the inverse happens: the detector width becomes equal or larger than the FSR and only part of the detector will be covered by the order selected by the AOTF.

Echelle grating specifications and manufacturing

The construction of an echelle grating having a groove density of 4 grooves per mm (groove step 250 µm) and effective surface of at least 60 mm by 134 mm is not evident. Bach Research Corp. from Boulder, Colorado, U.S.A. managed to manufacture an echelle based on the confined specifications collected in Table 4, of which the very low groove density was perhaps the most challenging.

There are two major difficulties in realizing these types of coarse echelle gratings23: (a) the high degree of echelle groove pitch accuracy that is required to get a usable wavefront in a high diffraction order, which is difficult to achieve during the fabrication process because of high tool-load and the accompanying elastic deflections experienced; (b) a high echelle efficiency, which means an optically flat and smooth wide groove facet with respect to wavelength.

Based on its experience in both ruling and machining echelles and echelons and given the required tolerances, which were approximately 2 times tighter than previously attempted, Bach Research opted for machining the echelle on a new diamond-machining engine built around the base of a Moore #3 measuring machine. The substrate in which the echelle was machined is a 150 mm by 60 mm by 25 mm block of aluminum 6061-T6 alloy. On the side and back of the substrate Bach Research diamond-turned a set of 30 mm by 30 mm reference mirrors to facilitate the alignment of the echelle in the spectrometer and during the manufacturing process. Their wavefront quality was about (/4 @ 632.8 nm.

Several corrections were made to the engine and the process of cut-and-try was repeated a number of times until the echelle ghost intensities were reduced to the order of 0.2% of the parent line. Each machined echelle was tested for ghosts and efficiency, at both 1.1 µm and 3.3 µm, using a 0.5 m Czerny-Turner spectrometer. Figure 7 depicts a typical ghost trace of a good echelle, measured in the IR domain in the 130th order of a 3.3 µm He-Ne laser.

The expression (8), derived by Rowland in the 19th century, in which m is the order of the parent line, ε the amplitude of the periodic error and d the grating pitch, relates the first order ghost intensity to that of the parent line for the Littrow condition:

(8)

The measured peak ratio of 0.002 suggests that the amplitude of the periodic error on a groove pitch of 250 µm was reduced to about 27 nm. The SOIR echelle was also tested with visible light using an ADE Phase Shift MiniFIZ Surface mapping Fizeau interferometer and was found to have a wavefront quality of approximately 0.7 waves RMS in the 702nd order of 632.8 nm laser light. The non-periodic groove pitch error24 can be estimated from this wavefront data by equation (9), wherein δx is the error in groove placement, # the inverse of the wavefront error, λ/# the fractional wavefront error and ( the angle of use. In a Littrow mounting condition the equation can be re-expressed in terms of groove pitch d and the order of use m:

(9)

Formula (9) yields an estimated non-periodic placement error of 249 nm.

Instrument FOV, entrance optics and relay optics

Since the spectral FOV of 2' must be imaged on 2 detector pixels (2 x 30 µm) to fulfill the Nyquist sampling criterion, the over-all focal length of the spectrometer must be equal to 103 mm. The specified spatial FOV (30') will then obviously be imaged on 30 detector pixels.

On the other hand the allowed acceptance angle of the AOTF and the available entrance aperture directly determine the magnification of the entrance optics in front of the AOTF. Given the entrance pupil diameter of 20 mm and the AOTF acceptance aperture (10 mm x 4 mm), a magnification factor of 5 is required.

The entrance optics consists in a telescope composed of an objective lens (Si-Ge doublet / focal length 58.4 mm) followed by a collimating lens (ZnSe singlet / focal length 11.6 mm). A refracting lens design was preferred to reflecting optics because of limited space and the specific entrance requirements. The Ge lens element also absorbs wavelengths shorter than 1.8 µm (at 300 K), stopping the UV-visible part of the solar spectrum before the AOTF. In the focal plane of the objective lens a rectangular diaphragm of 175 µm by 1250 µm with an edge thickness of 100 µm is placed that reduces the entrance optics FOV to slightly larger than the desired over-all FOV. It avoids imaging the full Sun on the spectrometer slit and suppresses light scattering and ghost images in the rest of the optics; it also acts as a heat sink, since it stops a large part (70 %) of the solar input flux. The entrance optics exit angle in the spatial direction (five times the desired over-all spatial FOV or 2.5º) is still well within the acceptance angle of the AOTF (Table 2).

The relay optics transports the output beam of the AOTF to the echelle grating. It consists in a telescope composed of the AOTF exit optics and the collimator formed by an off-axis parabolic mirror with a focal length of 300 mm; the latter is used 100 mm or 19° off axis and reflects the light out of the horizontal plane by 1.46°. The combined focal length of entry optics and AOTF exit optics is equal to the over-all focal length (103 mm), the spectrometer itself having no effect on focal length due to its magnification factor of 1.

The magnification factor MAXO of the combination of the AOTF exit optics and the collimating mirror (0.0687) is given by (10):

(10)

in which f, fimager and MAEO are the over-all focal length, the focal length of the imager and the magnification factor of the AOTF entrance optics. The spectrometer entrance slit is positioned right behind the AOTF exit optics at the intermediate image plane. Since the slit is imaged one-to-one on the detector, the slit dimensions have been defined as 2 pixels wide (60 µm) by 30 pixels high (900 µm). The thickness of the slit material is 100 µm.

The maximum entrance aperture diameter DEA (20 mm) is imposed by the VEX spacecraft design and has a direct impact on the Fraunhofer diffraction-limited spot size (Airy disc) on the detector. This spot diameter a, given by (11), increases linearly with wavelength (, but at the largest specified wavelength (4.25 µm) it remains still smaller than the specified ILP.

(11)

IR detector and associated optics

The SOIR detector assembly is located aside of the spectrometer section and therefore an extra folding mirror ((8) in Figure 4) is needed to bend the optical path clockwise by 90°. Further, a lens system is positioned after the focal plane of the spectrometer, imaging the entrance aperture of the spectrometer on the cold shield aperture of the cooled IR detector and projecting the slit image, formed by the collimating-imaging mirror, on the detector with a magnification of 1. The system contains five spherical lenses and compensates for aberrations created by the off-axis parabolic mirror. The first and the fourth lens are made of Ge, the third and the fifth of Si and the second of ZnSe. A mirror between the third and the fourth lens bends the light pad upwards by 90°. The final image of the spectral lines on the detector makes a slight angle with respect to the optical axis of the entrance optics and therefore the whole detector assembly is rotated by 2.45° clockwise around an axis perpendicular to the spectrometer’s base plate.

The detector system itself is a slightly modified standard Integrated Detector Dewar Cooler Assembly (IDDCA) of type ID MM0067 manufactured by SOFRADIR, Chatenay-Malabry, France. The main characteristics of this customized version are summarized in Table 5.  It includes a high-sensitivity focal plane array (FPA) comprised of a two dimensional detector array made of photovoltaic Mercury Cadmium Telluride (HgCdTe / MCT) pixels of  30 µm by 30 µm that are organized as 320 columns by 256 rows. The sensitive area of 9600 µm by 7680 µm has a fill factor greater than 90 %. All pixels are connected by means of indium bumps to a CMOS Readout Integrated Circuit (ROIC) that forms a sandwich with the MCT array. The ROIC reads the charge from each MCT photovoltaic diode via direct injection and integrates it on the pixel input capacitor (selectable 0.7 or 2.1 pF) during a programmable integration time. We have opted for power, commanding and simplicity reasons to read-out only a programmable window (8 rows of 320 pixels) of the FPA in the single output mode at a 1 MHz rate. As a consequence of the direct injection read-out method, the linearity of the detector is not constant over the whole video output range. However, a linearity of 95 % can be obtained when the integrated charge is between 7 and 97 % of the pixel saturation value. In this application, in which we observe solar occultation spectra, the integration time shall be chosen such that the detector signal stays within this range. Absorption peaks not descending below 50 % of saturation will be measured with a linearity of 99 %.

The detector FPA is mounted in an evacuated dewar with a customized optical Si window (relative spectral response in Figure 8) and is surrounded with a cold shield with an f/4 aperture. It is cooled by a closed cycle miniature Stirling cooling machine (type K508, RICOR, Israel) that was adapted for space applications. SOFRADIR manufactured the FPA and integrated it with the dewar and the cooler; BIRA-IASB developed a proprietary space qualified cooler motor control circuit (see below).

Detector read-out and system control electronics

The SOIR spectrometer is controlled by a dedicated space qualified electronics unit, the simplified block diagram of which is shown in Figure 9. The central part of the electronics is the Central Control and Communications board (CCC) that contains a Field Programmable Gate Array (FPGA) which performs five main tasks: (1) reception and analysis of telecommands and production of telemetry data issued by or sent to the spacecraft; (2) commanding of the Housekeeping board (HSK) measuring essential housekeeping data, like temperatures of optical parts and supply voltages;  (3) generation of clock signals required for the read-out of the FPA, the analog video signal of which is amplified, sampled and converted to digital on the Video Sampling (VIS) proximity electronics board; (4) generation by means of a Direct Digital Synthesizer (DDS) of the 32-bit wide digital code representing the tunable radio frequency (RF) sine waveform for the AOTF; (5) control via an embedded Proportional-Integral-Derivative (PID) controller program of the RICOR cooler temperature. The digital RF code generated on the CCC board is used by the RF Generator board (RFG) which produces the RF analog signal for the AOTF. The RICOR Cooler Control board (RCC) receives opto-coupled digital motor control signals from the CCC board and regulates the detector cooler temperature by means of Pulse Width Modulation (PWM).

One of the nice features of the spectrometer control electronics is that, because of its very fast RF on/off switching capabilities, the AOTF order sorting filter can be used to make two successive measurements of the same spectrum: one with signal plus thermal background and one with thermal background alone.

SOIR observation modes

Before a solar occultation observation is started and the detector FPA supply voltages are activated the instrument goes into the Pre-cooling Phase, during which the FPA temperature is decreased to its nominal operating temperature (around 90-100 K); during this phase only housekeeping data is sent through telemetry. The parameters required for the pre-cooling phase are sent in a first type of telecommand packet and contain settings like: the pre-cooling duration, the desired FPA temperature and the cooling control loop coefficients. At the end of pre-cooling the spectrometer expects a second type of TC packet containing the observation settings for the occultation. The observation itself consists in the continuous recording and sending of spectra at a rate of 1 spectrum per sec (rate imposed by spacecraft telemetry resources); this process continues until instrument switch-off.

Different types of observation scenarios, all working at the imposed rate, may be executed. The standard observation scenario during solar occultations comprises the repeated read-out of the spectrum corresponding to a fixed grating order selected by the AOTF frequency setting and hence to a fixed wavenumber range. This scenario may be extended by allowing up to 4 AOTF frequency hops within each sampling period of 1 second, allowing the  recording of a sequence of 1 to 4 different wavelength bands. This is only applicable if the sum of all required detector integration times and AOTF settling times does not exceed the sampling period (1 sec). The essential parameters related to the programming of this scenario are: the exact position of the spectral window on the detector, the number of pixels along the spectral lines and their binning factor (1, 2, 4, 8, 16 or 32), the number of different spectral domains per sampling period (1-4) and the 4 possible parameter sets associated with these domains (FPA integration time, number of times a spectrum can be accumulated during 1 sampling period, the AOTF frequency and power setting). An additional parameter flag allows subtracting a background spectrum measured immediately before each signal spectrum. In the latter case 1 to 4 measurement pairs (signal, background) can be recorded per sampling interval.

A different scenario, called “automatic detector window stepping”, can be used for a complete read-out of the focal plane image at a fixed AOTF frequency. Each sampling period the next detector window (8 rows of 320 pixels) is accessed and read-out. In total 32 sampling periods are required to read out the complete detector array once. This scenario is very useful during the commissioning phase of the spacecraft and of SOIR to derive detector and spectrometer characteristics.

A last built-in scenario is “echelle order stepping”. During each sampling period the AOTF frequency is increased by a constant but programmable increment, thus allowing an AOTF frequency scan from a minimum to a maximum frequency. If the minimum frequency and the frequency increment are programmed such that for each step the AOTF bandpass is positioned in the center of the detector window, then we can organize a perfect spectrum scan of succeeding orders. It is obvious that this scenario allows recording scans over the complete wavelength domain, hence from 2.2 µm (order 194) to 4.3 (order 101). This feature can be used during instrument commissioning in cruise phase and in orbit around Venus outside eclipses to get a reference observation of the solar spectrum, allowing the assessment of the optimum integration time settings for solar observations and a calibration of the wavenumber range.

Mechanical structure and optical alignment

Due to the fact that it had to be installed on an internal panel of the VEX spacecraft and on top of an existing instrument that was originally conceived for the MEX mission (SPICAM-L), the surface of SOIR’s base plate was confined to 414 mm (L) by 254 mm (W). This base plate, which carries the holders of the individual optical parts, the electronic boards and the cabling, forms the mechanical interface with SPICAM-L. Its design is a compromise between low mass and stiffness; its overall thickness is 24.6 mm and it has a flat top surface and a rib structured bottom.

The main mechanical subunits depicted in Figure 10 are: (1) the AOTF bridge carrying the AOTF and its entrance and exit optics, a small folding mirror and the spectrometer’s entrance slit, (2) a holder plus baffle for the parabolic mirror, (3) a box-like grating holder providing the necessary tilting of the echelle grating, (4) a support structure for the detector unit, (5) a detector entrance optics holder and (6) a number of boxes lodging the electronic boards and providing radiation shielding. All structural elements, including the base plate, are made of space qualified aluminum AISI 7075. Some elements, like the folding mirror holder or the AOTF support, are manufactured in stainless steel for a better compatibility between optical element, appropriate space qualified glue and structure, and hence a better resistance to vibrations or shocks. In order to avoid light scattering and for thermal reasons all structural elements in or in the neighborhood of the optical path are surface treated by space qualified optically black anodization.

The height of SOIR’s main cover is 100 mm, which is locally reduced to 70 mm in order to fit underneath a VEX subsystem and extended to 210 mm at the location of the detector. It protects the delicate optical parts inside the instrument, contributes to the overall stiffness, provides additional shielding against radiation and serves as a support for the thermal MLI (multi layer insulation) blanket which covers the whole of SOIR.

The completely integrated SPICAV instrument has been successfully qualified during mechanical vibration tests. Furthermore, all SOIR subassemblies, as well as the fully integrated SOIR and SPICAV instruments, have passed several thermal-vacuum tests in an operational temperature range of -40 °C to +50 °C and a non-operational range of -40 °C to +60 °C.

The optical axis of the entrance optics is situated 45 mm above the base plate. In the plane parallel to the base plate it can be aligned to 1' with respect to the mechanical structure and by design it is parallel within 20' to the base plate. SOIR is a sun pointing instrument designed for observations using the solar occultation technique (Figure 2). However, the optical axis of SOIR and the solar port of VEX do not coincide and therefore a two-mirror periscope ((7) in Figure 10) is added externally at the entrance of the instrument. It folds the solar light beam, which passes through an aperture in the spacecraft's side wall under an angle (60°) with respect to the instrument's base plate, into SOIR. One of the periscope mirrors is adjustable in three axes allowing for perfect alignment between the desired inclined line of sight and the SOIR optical axis; an external cube mounted at the outside of the base plate is used as a reference.

Thermal environment

From an optical design point of view the ideal operating temperature range for SOIR is -5 °C to +20 °C, lower temperatures leading to lower thermal background. The instrument is thermally coupled to neighboring instruments and satellite subsystems in three ways: by radiative heat exchange via a 12 layer thermal blanket (Multi Layer Insulation or MLI), by conduction via its own base plate and SPICAM-L to a thermal reference point (TRP1) on the spacecraft, and finally by conduction via a flexible multilayer cupper thermal strap directly from its base plate to a second thermal reference point (TRP2) that is connected to a cold spacecraft wall. The temperatures of both thermal reference points are independently regulated by the spacecraft. Other heat sources and sinks are: the internally dissipated power (17 W during the detector pre-cooling phase and 16 W in the measuring phase), the solar flux absorbed in the entrance optics during sun pointing periods and heat loss to deep space through the open periscope. In cold non-operational cases additional heating for SOIR’s base plate is provided by means of two redundant series of heating elements that are switched on by the satellite when the base plate temperature drops below -5 °C and off when the temperature rises above 0 °C.

Thermal background, signal and noise

A significant contribution to the signal measured by the detector FPA comes from the thermal background emitted by the optical elements and the surrounding mechanical parts. In order to exclude the thermal radiation emitted by the latter the detector dewar is equipped with a f/4 cold shield that is imaged on the SOIR entrance aperture. The only structure which is directly in view of the FPA is the spectrometer slit mount that is in focus for the detector. It shields the FOV limiting diaphragm mounted in the entrance optics, as well as all other optical elements in front of the slit. The little thermal radiation transmitted by the slit is further dispersed by the echelle grating and therefore each detector pixel will receive only energy emitted in a small spectral interval corresponding to the spectral sampling interval (SSI). The mirrors and lenses placed between the grating and the detector will contribute a lot more, since each detector pixel will receive thermal radiation in the much wider 1.8 to 4.3 µm band.

Table 6 represents the thermal background per pixel and per msec for different uniform environmental temperatures, as calculated assuming worst case emissivities and transmission and reflection efficiencies for lenses, mirrors, parabolic mirror, echelle and slit mount, and a constant FPA quantum efficiency of 70 %. These results indicate that at the maximum expected operating temperature of SOIR (293 K) thermal background will be limited to 1 % of full scale per integration time slice of 1 msec.

The noise component associated with the thermal background is one of the main noise sources for the detector; the other is the shot noise of the solar signal itself. Tables 7 and 8 represent the expected solar signal as well as the thermal and shot noise at 2.2 and 4.3 µm and for the same operating temperatures as mentioned in Table 6. For 2.2 µm the integration time for which the detector becomes just saturated by the sum of signal and thermal background equals 36 msec when the operating temperature is 293 K. For this integration time and operating temperature the estimated signal-to-noise ratio (SNR) varies between 2157 at 2.2 µm and 644 at 4.3 µm. The SNR will slightly improve if the operating temperature approaches the nominal SOIR operating temperature (273 K). It is also clear that for the 4.3 µm case one can double the integration time until the pixel saturates and improve SNR by a factor sqrt(2).

Predicted and observed optical and spectral performances

The Instrument Line Profile (ILP) predicted by OIP comes very close to the specified width of 2 pixels. Figure 11 shows the calculated Instrument Line Profiles for 3 wavelengths: 2.32 µm, 3.16 µm and 4.24 µm. Smile, frown and the displacement of the spatial field have also been verified by calculation. Smile (deviation from a straight line of the slit image on the detector at one particular wavelength) will be unnoticeable since it is always less than 1/10th of a pixel and frown (deviation from a straight line of the different images of one particular slit point for different wavelengths within one diffraction order) will never be larger than half a pixel. Finally, the displacement between the central wavelengths of different orders turns out to be less than 1/6th of a pixel.

The predicted spectral performance of the SOIR spectrometer was verified under normal lab conditions at the 3.39 µm line of a He-Ne laser. The laser output was diffused, collimated and sent into the spectrometer’s optical entrance aperture (no periscope installed). Because of the very small bandwidth of the laser line the detector image should be a one-to-one copy of the slit. Figure 12 (a) shows a grayscale representation of a 2-D rectangular cut-out of the detector’s response in the area where the slit image of 2 by 30 pixels is formed. One also clearly observes the excellent aberration correction and the resulting straight spectral line. Figures 17 (b) shows a cross-section of the image along a pixel row, confirming that the ILP is as expected. The over-all noise (detector + electronics) in the detector area outside the spectral lines can also be derived from the same measurement by looking at non-illuminated pixels and amounts to a peak-to-peak noise of 5 LSB.

A more elaborate test with the Sun as a source, and applying the described echelle order stepping scenario, was also performed. Solar rays undergoing absorption by the Earth’s atmosphere were reflected by a mirror outside the lab and directed along the optical axis of SOIR which was located in a clean-room. A typical read-out, recorded at the same AOTF frequency used to perform the 3.39 µm (2949 cm-1) laser test is shown in Figure 13; it shows 8 spectra obtained by applying 4-to-1 binning on 32 detector rows of 320 pixels each. The quality of the spectral lines observed during the laser test is confirmed: spectral lines remain perfectly perpendicular to the diffraction direction and parallel to the detector columns.

During this quick test the selected detector read-out window was not in accordance with the height of the spectral lines and the relation between the wavenumber scale and AOTF frequency was still uncalibrated. This relation and the diffraction in the different orders can be calibrated in flight during the cruise phase using the solar spectrum and its known absorption lines as a reference. Nevertheless, it was demonstrated that the spectrometer attained the projected high-resolution by comparing parts of the obtained spectrum with a synthetic spectrum (Figure 14) computed by means of the PcLnwin program. This program simulated an observation of the Sun from ground through the Earth’s atmosphere under a 45° zenith angle with CH4 as the absorbing molecule and this for a spectral resolution of 0.1 cm-1 half-width-half-maximum. After comparison with the synthetic spectrum (Figure 14) the absorbing features seen in the right part of Figure 13 (pixels 180-320) can clearly be contributed to CH4 absorbing in the 2938 to 2948 cm-1 range and their line width is as expected.

Most of the spectra recorded in a particular echelle order also show interferences by neighboring diffraction orders, because the applied AOTF frequencies used were not calibrated and because the AOTF band pass is not a square fitting exactly one order, but shows a sinc2 behavior giving rise to side lobes. The noticeable feed-through effect can be corrected for by postprocessing knowing the synthetic spectra of expected atmospheric constituents.

Conclusions

The high-resolution IR spectrometer described in this paper is the first spectrometer based on a combination of a highly dispersive echelle and an Acousto-Optical Tunable Filter to be put in space. On November 9 2005 it was launched on board of the ESA Venus Express spacecraft from the Baikonour launching site in Kazakhstan by means of a Soyuz-Fregat launcher. Insertion into a 24 h orbit around Venus is planned for April 11 2006.

It is obvious that the field of applications for the described spectrometer is not limited to the Venus Express mission only. It could be extremely useful to study the presence of methane in the Martian atmosphere by solar occultation and it could perform this with a 10 times higher spectral resolution than is possible with present instruments in orbit around Mars.

Acknowledgements

The authors like to thank Jean-Loup Bertaux, Principal Investigator of the SPICAV experiment, for his many intensive and scientifically founded interventions at the ESA level, defending and promoting the SPICAV experiment and its extension, the SOIR high resolution IR spectrometer. Furthermore, we like to express our gratitude for the financial support we received from the Belgian Federal Science Policy Office, who completely funded the design and manufacturing of SOIR as well as part of the construction of SPICAM-L (ESA PRODEX contract nr. 90113) during years 2003-2005.

Our special thanks go to the technical people at BIRA-IASB and OIP in Belgium and at Service d’Aéronomie, France, who have made the successful assembling and testing of SOIR and SPICAM-L and their integration into SPICAV possible under severe time pressure.

Table 1: Summary of spectrometer design objectives

|  | Value or range | Units |
| --- | --- | --- |
| Wavelength range | 2.32 to 4.25 | µm |
| Wavenumber range | 2353 to 4310 | cm-1 |
| Spectral sampling interval | 0.1 | cm-1/pixel |
| Instrument Line Profile | 2 (FWHM) | pixels |
| Resolving power | 23200 to 43100 | (/((=(/(( |
| Spectral field of view | 2 | arc minutes |
| Spatial field of view | 30 | arc minutes |
| Entrance aperture diameter | 20 | mm |
| Mass | 6.5 | kg |
| Dimensions (w/o periscope) | 414 x 254 x 210 | mm3 |

Table 2: Specified AOTF characteristics

|  | Value or range | Units |
| --- | --- | --- |
| Material | TeO2 |  |
| Wavelength range | 2.325 to 4.25 | µm |
| Spectral bandwidth (FWHM) | (22 | cm-1 |
| Acceptance aperture in plane of diffraction | >10 | mm |
| Acceptance aperture ( to plane of diffraction | >4 | mm |
| Acceptance angle in plane of diffraction | +/- 2 | degrees |
| Acceptance angle ( to plane of diffraction | +/- 3 | degrees |
| Diffraction angle | 4.6 | degrees |
| AOTF RF frequency | 13.5 – 25 | MHz |

Table 3: Computed wavenumbers (cm-1) associated with detector pixels 0 and 319 for echelle orders 101 to 194

| order | (0 | (319 | order | (0 | (319 | order | (0 | (319 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 101 | 2256.41 | 2274.75 | 133 | 2971.32 | 2995.42 | 165 | 3686.22 | 3716.11 |
| 102 | 2278.75 | 2297.27 | 134 | 2993.67 | 3017.95 | 166 | 3708.56 | 3738.63 |
| 103 | 2301.09 | 2319.79 | 135 | 3016.01 | 3040.47 | 167 | 3730.90 | 3761.15 |
| 104 | 2323.43 | 2342.31 | 136 | 3038.35 | 3062.99 | 168 | 3753.24 | 3783.67 |
| 105 | 2345.78 | 2364.83 | 137 | 3060.69 | 3085.51 | 169 | 3775.58 | 3806.19 |
| 106 | 2368.12 | 2387.35 | 138 | 3083.03 | 3108.03 | 170 | 3797.92 | 3828.71 |
| 107 | 2390.46 | 2409.87 | 139 | 3105.37 | 3130.55 | 171 | 3820.26 | 3851.23 |
| 108 | 2412.80 | 2432.39 | 140 | 3127.71 | 3153.07 | 172 | 3842.60 | 3873.76 |
| 109 | 2435.14 | 2454.92 | 141 | 3150.05 | 3175.59 | 173 | 3864.93 | 3896.28 |
| 110 | 2457.48 | 2477.44 | 142 | 3172.39 | 3198.12 | 174 | 3887.27 | 3918.80 |
| 111 | 2479.82 | 2499.96 | 143 | 3194.73 | 3220.64 | 175 | 3909.61 | 3941.32 |
| 112 | 2502.16 | 2522.48 | 144 | 3217.07 | 3243.16 | 176 | 3931.95 | 3963.84 |
| 113 | 2524.51 | 2545.00 | 145 | 3239.41 | 3265.68 | 177 | 3954.29 | 3986.36 |
| 114 | 2546.85 | 2567.52 | 146 | 3261.75 | 3288.20 | 178 | 3976.63 | 4008.88 |
| 115 | 2569.19 | 2590.04 | 147 | 3284.09 | 3310.72 | 179 | 3998.97 | 4031.41 |
| 116 | 2591.53 | 2612.56 | 148 | 3306.43 | 3333.24 | 180 | 4021.31 | 4053.93 |
| 117 | 2613.87 | 2635.08 | 149 | 3328.77 | 3355.76 | 181 | 4043.65 | 4076.45 |
| 118 | 2636.21 | 2657.61 | 150 | 3351.11 | 3378.29 | 182 | 4065.99 | 4098.97 |
| 119 | 2658.55 | 2680.13 | 151 | 3373.45 | 3400.81 | 183 | 4088.33 | 4121.49 |
| 120 | 2680.89 | 2702.65 | 152 | 3395.79 | 3423.33 | 184 | 4110.67 | 4144.01 |
| 121 | 2703.23 | 2725.17 | 153 | 3418.13 | 3445.85 | 185 | 4133.01 | 4166.54 |
| 122 | 2725.58 | 2747.69 | 154 | 3440.47 | 3468.37 | 186 | 4155.35 | 4189.06 |
| 123 | 2747.92 | 2770.21 | 155 | 3462.82 | 3490.89 | 187 | 4177.69 | 4211.58 |
| 124 | 2770.26 | 2792.73 | 156 | 3485.16 | 3513.41 | 188 | 4200.03 | 4234.10 |
| 125 | 2792.60 | 2815.25 | 157 | 3507.50 | 3535.93 | 189 | 4222.37 | 4256.62 |
| 126 | 2814.94 | 2837.78 | 158 | 3529.84 | 3558.46 | 190 | 4244.71 | 4279.14 |
| 127 | 2837.28 | 2860.30 | 159 | 3552.18 | 3580.98 | 191 | 4267.04 | 4301.67 |
| 128 | 2859.62 | 2882.82 | 160 | 3574.52 | 3603.50 | 192 | 4289.38 | 4324.19 |
| 129 | 2881.96 | 2905.34 | 161 | 3596.86 | 3626.02 | 193 | 4311.72 | 4346.71 |
| 130 | 2904.30 | 2927.86 | 162 | 3619.20 | 3648.54 | 194 | 4334.06 | 4369.23 |
| 131 | 2926.64 | 2950.38 | 163 | 3641.54 | 3671.06 |  |  |  |
| 132 | 2948.98 | 2972.90 | 164 | 3663.88 | 3693.58 |  |  |  |

Table 4: Comparison of specified and obtained echelle parameters

|  | Specified value | Obtained value | Units |
| --- | --- | --- | --- |
| Material | Aluminum 6061-T6 |  |
| Coating material | Au (gold) |  |
| Wavelength range | 2 to 4.2 |  | µm |
| Diffraction orders | 100-200 |  |  |
| Blaze angle | arctan(2) | 63.20 | degrees |
| Groove density | 4 |  | per mm |
| Groove spacing | 250 |  | µm |
| Total number of steps | 560 |  |  |
| Structure width | 150 | 150.03 | mm |
| Groove length | 60 | 59.94 | mm |
| Structure thickness | 25 | 24.99 | mm |
| Temperature range | -50 to +50 |  | °C |

Table 5: Main electro-optical characteristics of the customized SOFRADIR ID MM0067 detector

| Characteristics | Value | Unit |
| --- | --- | --- |
| Number of pixel rows | 256 |  |
| Number of pixels per row | 320 |  |
| Cold shield aperture | f/4 |  |
| Integration capacitor for gain setting 0 / 1 | 0.7 / 2.1 | pF |
| Maximum charge for gain setting 0 / 1 | 12 / 37 | 106 e- |
| Video output range | 1.6 - 4.4 | V |
| Video dynamic range | 2.8 | V |
| Minimum integration time | 3 | µs |
| Standard wavelength range | 3.7 - 4.8 | µm |
| Customized cut-on @ 70 % of maximum | 1.8 | µm |
| Customized cut-off @ 70 % of maximum | 4.2 | µm |
| Typical quantum efficiency @ 3.5 µm | 0.75 | e-/photon |
| Typical relative spectral pixel response | See Figure 8 |
| Read-out speed for single video output | 1 - 6 | MHz |
| Typical focal plane temperature | (110 | K |

Table 6: Estimated thermal background and fraction of pixel saturation

| Temp. [K] | 313 | 303 | 293 | 283 | 273 | 263 | 253 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 105 e- pixel-1 .msec-1 | 2.82 | 1.91 | 1.26 | 0.809 | 0.504 | 0.303 | 0.176 |
| % of pixel saturation .msec-1 | 2.23 | 1.52 | 1 | 0.64 | 0.4 | 0.24 | 0.14 |

Table 7: Estimated signal, noise sources and SNR for 2.2 µm and a 36 msec integration time

| Temp. [K] |  | 313 | 303 | 293 | 283 | 273 | 263 | 253 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Thermal electr. | 106 e- | 10.2 | 6.87 | 4.54 | 2.91 | 1.81 | 1.09 | 0.632 |
| Thermal noise | e- | 3186 | 2622 | 2130 | 1707 | 1347 | 1045 | 795 |
| Expected signal | mW.m-2.pixel-1 | 30 |
| Signal electrons | 106 e- | 7.48 |
| Shot noise | e- | 2734 |
| Total noise | e- | 4199 | 3788 | 3466 | 3223 | 3048 | 2927 | 2848 |
| SNR |  | 1781 | 1974 | 2157 | 2319 | 2453 | 2554 | 2625 |

Table 8: Estimated signal, noise sources and SNR for 4.3 µm and a 36 msec integration time

| Temp. [K] |  | 313 | 303 | 293 | 283 | 273 | 263 | 253 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Thermal electr. | 106 e- | 10.2 | 6.87 | 4.54 | 2.91 | 1.81 | 1.09 | 0.632 |
| Thermal noise | e- | 3186 | 2622 | 2130 | 1707 | 1347 | 1045 | 795 |
| Expected signal | mW.m-2.pixel-1 | 3.2 |
| Signal electrons | 106 e- | 1.60 |
| Shot noise | e- | 1263 |
| Total noise | e- | 3428 | 2910 | 2476 | 2123 | 1847 | 1639 | 1493 |
| SNR |  | 465 | 548 | 644 | 751 | 864 | 973 | 1069 |

Captions for figures:

Fig. 1. Photographs of the SOIR high resolution spectrometer mounted on top of the SPICAM-L part of the SPICAV experiment for Venus Express.

(a) SOIR without protective covers and periscope

(b) SOIR with protective covers and periscope installed

Fig. 2. Composition showing the Sun as a source, the Venusian atmosphere and the SOIR spectrometer in orbit around Venus  pointing the Sun during a solar occultation.

Fig. 3. Simplified 2-D optics scheme of SOIR spectrometer.

Fig. 4. 3-D representation of main SOIR optics elements and ray tracing.

Fig. 5. SOIR AOTF: Incoming ordinary (ko) and extraordinary (ke) beams and the outgoing useful diffracted extraordinary beam (ke*).

Fig. 6. SOIR AOTF configuration with orientation of the crystal axes. Dimensions in mm.

Fig. 7. Echelle grating diffraction pattern measured with a monochromatic source at 3.39 µm. Orders 129, 130 and 131 shown.

Fig. 8. Relative spectral response of the customized SOFRADIR ID MM0067 IDDCA. Curves for 4 different pixel locations.

Fig. 9. Simplified block diagram of the SOIR control electronics. Power converters and power switching not shown.

Fig. 10. SOIR mechanical structure in 3-D view.

Fig. 11. Predicted Instrument Line Profile for 3 wavelengths: 2.32, 3.16 and 4.24 µm.

Fig. 12. Image of spectrometer slit on detector with monochromatic 3.39 µm laser source

(a) Gray-scale 2-D image; horizontal axis: relative pixel number along detector row; vertical axis: relative pixel number along detector column;

(b) Image cross-section along a detector row; horizontal axis: relative pixel number along detector row; vertical axis: detector signal in arbitrary units.

Fig. 13. Typical uncalibrated detector output recorded with AOTF tuned to 3.39 µm band during observation through the Earth’s atmosphere. Eight spectra are shown resulting from 32 detector rows of 320 pixels each; each spectrum is obtained by 4-to-1 binning of 4 detector rows. Horizontal axis: pixel number along detector row; vertical axis: detector signal expressed as accumulated ADC codes.

Fig. 14. CH4 absorption spectrum for the range from 2938 to 2948 cm-1, obtained by simulation assuming an observation of the Sun from ground under a 45° zenith angle and through the Earth’s atmosphere. Spectral resolution: 0.1 cm-1 HWHM.

References

Mars Express - The Scientific Payload, ESA Special Publication, Volume: SP-1240, pp. 1-219, 2004.

V. Formisano, V.I. Moroz, F. Angrilli, G. Bianchini, E. Bussoletti, N. Cafaro, F. Capaccioni, M.T. Capria, P. Cerroni, G. Chionchio, L. Colangeli, A. Coradini, A. Di Lellis, S. Fonti, R. Orfei, E. Palomba, G. Piccioni, B. Saggin, A. Ekonomov, A. Grigoriev, V. Gnedykh, I. Khatuntsev, A. Kiselev, I. Matsygorin, B. Moshkin, V. Nechaev, Y. Nikolsky, D. Patsaev, A. Russakov, D. Titov, L. Zasova, M.I. Blecka, A. Jurewicz, M. Michalska, W. Novosielski, P. Orleanski, G. Arnold, H. Hirsch, H. Driesher, J. Lopez-Moreno, R. Rodrigo, J. Rodriguez-Gomez, G. Michel , “PFS: a Fourier spectrometer for the study of martian atmosphere”,  Adv. Space Res., Volume 19, pp. 1277-1280, 1997.

R. Hanel, B. Schlachman, E. Breihan, R. Bywaters, F. Chapman, M. Rhodes, D. Rodgers, and D. Vanous, “Mariner 9 Michelson interferometer”, Applied Optics, Volume 11, No. 11 , pp. 2625-2634, 1972.

M. R. Gunson, M. M. Abbas, M. C. Abrams, M. Allen, L. R. Brown, T. L. Brown, A. Y. Chang, A. Goldman, F. W. Irion, L. L. Lowes, E. Mahieu, G. L. Manney, H. A. Michelsen, M. J. Newchurch, C. P. Rinsland, R. J. Salawitch, G. P. Stiller, G. C. Toon, Y. L. Yung, R. Zander, “The Atmospheric Trace Molecule Spectroscopy (ATMOS) experiment: Deployment on the ATLAS space shuttle missions”, Geophys. Res. Lett., Volume 23, pp. 2333– 2336, 1996.

P. F. Bernath, C. T. McElroy, M. C. Abrams, C. D. Boone, M. Butler, C. Camy-Peyret, M. Carleer, C. Clerbaux, P.-F. Coheur, R. Colin, P. DeCola, M. De Mazière,  J. R. Drummond, D. Dufour, W. F. J. Evans, H. Fast, D. Fussen, K. Gilbert, D. E. Jennings, E. J. Llewellyn, R. P. Lowe, E. Mahieu, J. C. McConnell, M. McHugh, S. D. McLeod, R. Michaud, C. Midwinter, R. Nassar, F. Nichitiu, C. Nowlan, C. P. Rinsland, Y. J. Rochon, N. Rowlands, K. Semeniuk, P. Simon, R. Skelton, J. J. Sloan, M.-A. Soucy, K. Strong, P. Tremblay, D. Turnbull, K. A. Walker, I. Walkty, D. A. Wardle, V. Wehrle, R. Zander, and J. Zou, “Atmospheric Chemistry Experiment (ACE): Mission overview”, Geophys. Res. Lett., Volume 32, L15S01, doi:10.1029/2005GL022386, 2005.

P. Drossart, A. Semery, M. Bouye, Y. Hello, G. Huntzinger, D. Kouach, J.M. Reess, D. Tiphene, Y. Ghomchi, A. Coradini, F. Capaccioni, , G. Arnold, J. Benkhoff, “VIRTIS-H: a high-spectral-resolution channel for the Rosetta infrared imaging spectrometer”, Proc. SPIE Volume 4131, Infrared Spaceborne Remote Sensing VIII, pp. 78-87.

- D.P. Baldwin, D.S. Zamzow, and G.P. Miller, “AOTF-Echelle Spectrometer for Air-ICP-AES Continuous Emission Monitoring of Heavy Metals and Actinides”, SPIE Proceedings Series, Volume 3534, pp. 478-486, 1999.

- Compact high resolution spectrometer, Innovative Technology Summary Report  OST/TMS ID 1564, U.S. Department of Energy, 2000.

- D. P. Baldwin, D. S. Zamzow, D. K. Ottesen and H. A. Johnsen, “Testing of an Echelle Spectrometer as a LIBS Detector at Sandia”, Ames Laboratory Report IS-5148, April 25, 2001.

- O. Korablev and J.-L. Bertaux, “High-resolution IR spectrometer for the studies of planetary atmospheres”, Geophysical Research Abstracts, Geophysical Instrumentation, Volume 3, 2001.

- O. I. Korablev, J.-L. Bertaux, and I. I. Vinogradov, “Compact high-resolution IR spectrometer for atmospheric studies”, Infrared Spaceborne Remote Sensing, Proceedings SPIE, M. Strojnik and B.F. Andresen (ed.), SPIE, Volume 4818, pp. 272-281, 2002.

- O. Korablev, J.-L. Bertaux, I. I. Vinogradov, Y. K. Kalinnikov, D. Nevejans , E. Neefs, T. Le Barbu, G. Durry, A. Fedorova and A. Grigoriev, “High resolution spectrometers for planetary spacecraft”, EGS - AGU - EUG Joint Assembly, Abstracts from the meeting held in Nice, France, 6 - 11 April 2003, abstract #14785, 04/2003.

- O. Korablev, J.-L. Bertaux, D. Nevejans and the SOIR instrument team, “Compact high-resolution IR spectrometer for atmospheric studies”, Geophysical Research Abstracts, Volume 5, 14785, 2003.

- O. I. Korablev, J.-L. Bertaux, I. I. Vinogradov, Y. K. Kalinnikov, D. Nevejans, E. Neefs, T. Le Barbu, and G. Durry, “Compact high-resolution echelle-AOTF NIR spectrometer for atmospheric measurements”, Proceedings of the 5th International Conference on Space Optics (ICSO 2004), 30 March - 2 April 2004, Toulouse, France, B. Warmbein (ed.), ESA SP-554, pp. 73-80. 2004.

O. Korablev, J.-L. Bertaux, I. I. Vinogradov, Y. K. Kalinnikov, D. Nevejans, E. Neefs, T. Le Barbu, G. Durry, A. Fedorova and A. Grigoriev, “High Resolution Spectrometers For Planetary Spacecraft”, 35th COSPAR Scientific Assembly, C3.2/B0.8 Contributions.

Hakan Svedhem, “Venus Express Science Requirements Document”, ESA document VEX-RSSD-SP-002.

C. de Bergh, B. Bezard, T. Owen, D. Crisp, J.P. Maillard, and B.L. Lutz, “Deuterium on Venus - Observations from Earth”, Science (ISSN 0036-8075), Volume 251, pp. 547-549, Feb. 1 1991.

L. M. Mukhin,  B. G. Gelman,  N. I. Lamonov,  V. V. Melnikov, D. F. Nenarokov,  B. P. Okhotnikov,  V. A. Rotin and  V. N. Khokhlov, “Gas-chromatographic analysis of the chemical composition of the Venusian atmosphere on the descent modules of the Venera-13 and Venera-14 probes”, Kosmicheskie Issledovaniia (ISSN 0023-4206), Volume 21, Mar.-Apr., pp. 225-230, 1983.

V. I. Moroz, D. Spankuch, V. M. Linkin, W. Dohier, I. A. Matsygorin, K. Schafer, L. V. Zasova, D. Oertel, A. V. Dyachkov, R. Schuster, V. V. Kerzhanovich, H. Becker-Ross, E. A. Ustinov, and W. Stadthaus, “Venus spacecraft infrared radiance spectra and some aspects of their interpretation”, Applied Optics, Volume 25, No. 10, 15 May 1986.

C. Palmer, “Diffraction Grating Handbook”, 6th edition, Newport Corporation, 2005.

M.S. Gottlieb, “Acousto-optic tunable filters”, in “Design and fabrication of acousto-optic devices”, A.P. Gotzoulis,  D.R. Pape and S.V. Kulakov (ed.), Marcel Dekker N.Y.-Basel-Hong-Kong , pp. 197-283, 1994.

I. C. Chang, “Non-collinear acousto-optic filter with large angular aperture”, Appl. Phys. Lett., 25, p. 370, 1974.

B. W. Bach, Jr., K. G. Bach, M. Schulze, B. Bach, Jr., “Solar Physics And Space Weather Instrumentation, Modern echelons and echelles for infrared spectroscopy”, Proceedings of SPIE Vol. 5901, S. Fineschi, R. and  A. Viereck (ed.) , pp. 184-190, August 2005.

M.C. Hutley, "Diffraction Gratings", Techniques of Physics: 6, Academic press, p. 76, 1982.

Fig. 1. Photographs of the SOIR high resolution spectrometer mounted on top of the SPICAM-L part of the SPICAV experiment for Venus Express.

(a) SOIR without protective covers and periscope

(b) SOIR with protective covers and periscope installed

Fig. 2. Composition showing the Sun as a source, the Venusian atmosphere and the SOIR spectrometer in orbit around Venus  pointing the Sun during a solar occultation.

Fig. 3. Simplified 2-D optics scheme of SOIR spectrometer.

Fig. 4.  3-D representation of main SOIR optics elements and ray tracing.

Fig. 5. SOIR AOTF: Incoming ordinary (ko) and extraordinary (ke) beams and the outgoing useful diffracted extraordinary beam (ke*).

Fig. 6. SOIR AOTF configuration with orientation of the crystal axes also shown. Dimensions in mm.

Fig. 7. Echelle grating diffraction pattern measured with a monochromatic source at 3.39 µm. Orders 129, 130 and 131 shown.

Fig. 8. Relative spectral response of the customized SOFRADIR ID MM0067 IDDCA. Curves for 4 different pixel locations.

Fig. 9. Simplified block diagram of the SOIR control electronics. Power converters and power switching not shown.

Fig. 10. SOIR mechanical structure in 3-D view.

Fig. 11. Predicted Instrument Line Profile for 3 wavelengths: 2.32, 3.16 and 4.24 µm

(a)

(b)

Fig. 12. Image of spectrometer slit on detector with monochromatic 3.39 µm laser source

(a) Gray-scale 2-D image; horizontal axis: relative pixel number along detector row; vertical axis: relative pixel number along detector column;

(b) Image cross-section along a detector row; horizontal axis: relative pixel number along detector row; vertical axis: detector signal in arbitrary units.

.

Fig. 13. Typical uncalibrated detector output recorded with AOTF tuned to 3.39 µm band during observation through the Earth’s atmosphere. Eight spectra are shown resulting from 32 detector rows of 320 pixels each; each spectrum is obtained by 4-to-1 binning of 4 detector rows. Horizontal axis: pixel number along detector row; vertical axis: detector signal expressed as accumulated ADC codes.

Fig. 14. CH4 absorption spectrum for the range from 2938 to 2948 cm-1, obtained by simulation assuming an observation of the Sun from ground under a 45° zenith angle and through the Earth’s atmosphere. Spectral resolution: 0.1 cm-1 HWHM.