Mars Express Bistatic Radar Experiment

Operations Plan and Report

27 August 2005

Summary

=======

DSN Antenna:                 14

Pass:                        0818

Orbit:                       2078

Orbit Start Time:            2005-08-27T11:07:23 SCET

HGA Pointing:                Inertial

Specular Condition:          2005-08-27T13:10:30 ERT

Target:                      (87.39S, 33.20E; Rp = 3382.5 km)

Incidence/Reflection Angle:  74.96 deg

Slant Range:                 7975 km

Slew Angle:                  34.19 deg

Doppler (carrier, fd):       55663 Hz

Doppler (echo, fr):          47830 Hz

Doppler Difference (fdd):    7833 Hz

Earth-Mars Distance:         1.026E11 m

Experiment Set Up

=================

This experiment was conducted using Goldstone's DSS 14.  The support team at JPL included Gene Goltz, Danny Kahan, and Dick Simpson.  In the NOPE area was Jesse Velasco.

Performance Problems

====================

As was the case on the previous day, there were bad interactions between the X-LCP noise diode and its ambient load.  When the diode was on and the ambient load was added, the noise power went to the expected level but then dropped about 1 dB after 30-60 seconds.  The next directive was to turn the noise diode off; but that had little or no effect on the noise power.  When the ambient load was removed, the noise power dropped below the expected sky level for a few seconds; but then the noise diode turned itself back on.  We experimented and found this pattern to be reproducible.  We also made sure the noise diode was turned off before completing the pre-cal.  The problem was documented as DR G105992 on the previous day.

Also as on the previous day the S-LCP channel seemed less stable than the other three in the sense that ADC amplitudes jumped around when constant levels were expected and it was harder to detect changes in noise diode status.  Whether coincidence or not, this instability was only noticed after the X-band noise diode problems began.

Data Acquisition

================

The set up period began at 09:00.  DSS 14 announced that it was ready to begin the pre-calibrations about 15 minutes later.

Except for the X-LCP noise diode problem and the S-LCP stability problem noted above, the pre-calibration was nominal.

RSRs were configured as in Table 1:

Table 1

---------------------------------------------

RSR  Channel Mode  ATT FGAIN    Operator

dB

----- ------- ---- ---- ----- ---------------

RSR3A  S-LCP   1-W auto   60    Kahan

RSR3B  X-LCP   1-W auto   60    Kahan

RSR1A  S-RCP   1-W auto   65    Goltz

RSR1B  X-RCP   1-W auto   60    Goltz

RSR subchannels (SCHAN) were defined as follows:

Table 2

---------------------------------------------------------------------

Subchannel  Sample Rate  Comments

----------  -----------  ---------------------------------------------

1          1 ksps

2          8 ksps

3         25 ksps   Primary recording bandwidth

4        100 ksps   Backup recording

Table 3 lists ADC amplitude levels read from RSR displays during the experiment.  Where available, signal-to-noise ratios, frequency offsets, and other measurements are also listed. Times are in UTC read from a wall clock and may not reflect actual times when configurations changed.  "Steps" are as defined in the briefing message.  Acronyms and abbreviations are explained after Table 3.

Table 3

------------------------------------------------------------------------------

TIME (ERT)  RSR3A  RSR3B  RSR1A  RSR1B                COMMENTS

2005-08-26  S-LCP  X-LCP  S-RCP  X-RCP

ADC    ADC    ADC    ADC

(dB)   (dB)   (dB)   (dB)

---------- ------ ------ ------ ------ ---------------------------------------

End-to-End Configuration Check

09:22                                 DSS 14 reports Pre-Cal Step 01 complete

Reset all RSR ATTEN using AUTO feature

11.0    8.5    4.5   12.5 ATTEN settings

-9.3  -10.0  -10.5  -10.2 ADC amplitudes (all to sky) (dB)

09:25       -9.6   -0.3  -10.5   -0.1 Step 02: X-band to AMB load

09:27       -0.9   -0.2  -10.4   -0.2 Step 03: S-LCP to AMB load

09:29       -0.9   -0.3   -0.6   -0.1 Step 04: S-RCP to AMB load

Pre-Calibration

09:29                                 Ambient load temperatures:

S1 = 18.31C

S2 = 24.12C

X1 = 15.38C

Weather conditions:

T = 24.4C

P = 903.7 mbar

H = 23.3%

Sky clear, winds calm

09:30                                 Reset all ATTEN using auto feature

21.5   20.5   15.5   25.0 ATTEN settings

09:35:01                              All RSR's recording 25 kHz

09:35       -9.8  -10.1   -9.8  -10.2 Step 05: Configure ND controllers

09:40       -9.9  -10.0   -9.8  -10.1 Step 06: XR ND on

09:45       -9.8  -22.2   -9.9  -20.3 Step 07: X-band to sky

09:50       -9.8  -22.2   -9.9  -22.7 Step 08: XR ND off

09:55       -9.7  -19.9   -9.9  -22.8 Step 09: XL ND on

DSS 14 confirms 12.5K ND selected

10:00       -9.8  -10.7   -9.9  -10.2 Step 10: X-band to AMB (XL is too low)

Ambient load temperatures:

S1 = 18.12C

S2 = 23.69C

X1 = 15.19C

10:05       -9.8  -10.8   -9.9  -10.2 Step 11: XL ND off (no change in XL)

10:10       -9.8  -23.4  -10.0  -22.7 Step 12: X-band to sky (XL is too low)

10:12       -9.8  -19.9   -9.9  -22.8 XL ND spontaneously on; level OK?

10:15                                 DC02 SNT E syntax error at DSS 14

10:15:20    -9.7  -19.8   -9.7  -22.7 Step 13: SR ND on

10:20       -9.8  -20.0  -19.0  -22.8 Step 14: SR to sky

10:25       -9.8  -19.9  -21.5  -22.8 Step 15: SR ND off

DSS 14 confirms 12.5 ND selected

10:30       -9.7  -19.8  -21.5  -22.7 Step 16: SL ND on

10:35      -18.4  -20.0  -21.6  -22.8 Step 17: SL to sky

10:40      -19.8  -19.9  -21.5  -22.8 Step 18: SL ND off

Ambient load temperatures:

S1 = 17.88C

S2 = 23.19C

X1 = 14.94C

10:45      -20.1  -19.9               Repeat Step 08: Configure ND controller

10:46   -18.4 to  -22.2  -21.5  -22.7 Repeat Step 11: XL ND off; SL unstable

-20.4

10:50      -20.4  -19.8  -21.5  -22.7 Repeat Step 09: XL ND on

10:55      -20.1   -9.9  -21.6  -10.1 Repeat Step 10: X-band to AMB

10:56      -19.7  -10.7  -21.6  -10.2 XL dropped 0.8 dB

11:00      -19.7  -10.8  -21.4  -10.1 Repeat Step 11: XL ND off

11:05      -19.9  -23.4  -21.5  -22.7 Repeat Step 12: X-band to sky

11:06      -20.1  -19.9  -21.5  -23.0 XL ND turned itself back on; levels OK?

11:10      -20.3  -22.1  -21.4  -22.7 Repeat Step 11: XL ND off

11:15      -19.2  -22.2  -21.6  -22.7 Repeat Step 13: Reconfigure ND controller

Spacecraft Calibration (BCAL) and Slew #1

11:24                                 DSS 14 is on point

11:52                    -1100        S-band AOS (Hz)

11:53      -2500  -7200  -2500  -7200 Entered FRO values (Hz)

11:54      -19.5  -21.8  -21.5  -22.4 ADC amplitudes

1401          1400        Carrier frequency (Hz)

29            54         Carrier to noise ratio (dB)

11:55                                 DSS 14 reports breaker on subreflector

has tripped; technician going to reset

11:58      -20.2  -21.9  -21.5  -20.7 ADC amplitudes (dB)

1391   5134   1389   5121 Carrier frequency

28     31     54     75  Carrier to noise ratio (dB)

12:00               24            68  C/No after TLM on (dB)

12:41      -20.2  -21.9  -21.6  -20.6 ADC amplitudes (dB)

1164   4302   1161   4290 Carrier frequency (Hz)

27     42     54     76  C/No (dB/Hz)

Mini-Cal #1

12:42:23                              Begin Slew #1

12:43      -19.1  -22.0  -21.4  -22.4 ADC amplitudes (dB)

12:44      -19.5  -21.8  -21.6  -20.6 Step 01: XR ND on

12:46                                 Slew echo in all channels

12:47      -18.5  -21.9  -19.0  -22.5 Step 02: XR ND off, SR ND on

12:49                                 Intermittent X-RCP echoes

12:50      -18.4  -19.7  -21.4  -22.4 Step 03: SR ND off, XL ND on

12:53      -17.6  -21.9  -21.4  -22.4 Step 04: XL ND off, SL ND on

4181   1129   4188 Carrier frequency (Hz)

19            36  Carrier to noise (dB/Hz)

12:56      -19.5  -21.9  -21.5  -22.3 Step 05: SL ND off

Bistatic Radar Activities

12:56                                 Begin RSR SCHAN=4 recording (100 kHz)

12:59                                 Weak SR, XR echoes

13:03      -19.4  -22.0  -21.5  -22.2 ADC amplitudes

-1616        Echo frequency (Hz)

23         Echo amplitude (dB/Hz)

4120 Carrier frequency (Hz)

36  Carrier to noise (dB/Hz)

13:03                                 Double peaked XR echo; separation ~700 Hz

13:09      -20.1  -21.9  -21.5  -22.1 ADC amplitude (dB)

-1731  -6271  -1711  -6259 Echo frequency (Hz)

20     23     29     41  Echo peak (dB/Hz)

13:11                      31     47  Echo peak (dB/Hz)

13:12               33                Echo peak (dB/Hz)

13:18       gone   gone    25   faint Echo peak (dB/Hz)

13:21                                 Stop 100 kHz recording (SCHAN=4)

XR, SR echoes intermittent

13:25      -16.7  -21.9  -21.5  -22.2 ADC amplitudes (dB)

4080 Carrier frequency (Hz)

37  Carrier to noise ratio (dB/Hz)

Mini-Cal #2 and Slew #2

13:30      -18.2  -22.0  -21.4  -20.0 Step 01: XR ND on

13:33      -18.6  -21.9  -19.0  -22.2 Step 02: XR ND off, SR ND on

13:36      -18.6  -19.7  -21.4  -22.2 Step 03: SR ND off, XL ND on

13:39      -16.6  -21.9  -21.4  -22.1 Step 04: XL ND off, SL ND on

13:40:54                              Begin Slew #2

13:42      -18.0  -21.9  -21.4  -22.2 Step 05: SL ND off

S-RCP carrier visible

13:49      -19.5  -21.8  -21.4  -20.4 Slew complete

Post-Calibration

14:20                                 DSS 14 going to zenith

14:25       -9.9  -10.0   -9.9  -10.1 Step 01: All channels to AMB at zenith

Ambient load temperatures:

S1 = 18.25C

S2 = 22.38C

X1 = 15.88C

Weather conditions:

T = 23.6C

P = 904.7 mbar

H = 25.0%

Sky clear, winds calm

14:30       -9.8  -10.1   -9.7  -10.1 Step 02: XR ND on

14:33       -9.9  -10.0   -9.5  -10.1 Step 03: XR ND off, SR ND on

14:36       -9.7   -9.9   -9.7  -10.2 Step 04: SR ND off, XL ND on

14:39       -9.6  -10.0   -9.9  -10.1 Step 05: XL ND off, SL ND on

14:42      -17.6  -22.2  -21.4  -22.9 Step 06: All to sky

14:45      -19.8  -20.0  -21.3  -22.7 Step 07: SL ND off, XL ND on

14:48      -20.5  -22.3  -19.1  -22.8 Step 08: XL ND off, SR ND on

14:51      -19.7  -22.2  -21.5  -20.3 Step 09: SR ND off, XR ND on

14:54      -20.0  -22.3  -21.5  -22.8 Step 10: XR ND off

Ambient load temperatures:

S1 = 19.05C

S2 = 22.44C

X1 = 16.69C

Weather conditions:

T = 24.8C

P = 904.9 mbar

H = 24.2%

Sky clear, winds calm

15:00                                 End of track

Stop RSR recording (SCHAN=3)

AMB = ambient load

BOT = Beginning of Track

CONSCAN = conical scan tracking

CW = continuous wave (carrier only)

EOT = End of Track

FRO = frequency offset

HGA = high-gain antenna

LOS = loss of signal

ND = noise diode

No = noise power

Pc = carrier power

S1 = ambient load for S-RCP

S2 = ambient load for S-LCP

SL = S-LCP

SR = S-RCP

TLM = telemetry

X1 = Ambient load for both X-band channels

XL = X-LCP

XR = X-RCP

Dick Simpson

Original: 2005-08-27

Minor edits: 2005-08-28

Added Rp: 2005-11-26