## Sheet1
DATE | TIME | INTLs | KEYWORD | PURPOSE | ACTIVITY DETAILS
2001-06-20 | 17:30 |  | EU Shake | CTS Failure | Start to remove the EU from plate and transport to 306-221B for trobleshooting.
2001-06-22 | 8:30 | PK/BI |  |  | SBEU and USO unit removed from copper plate and transported to mag mapping. Sensor unit removed.
Friday |  |  |  |  | Configuration: EU and CTS mounted on copper plate. J4 pin 3 shorted to chassis. EU J 9 test board installed with test cable to CTS J 12. Cable for CTS hskp in not connected
2001-06-22 | 9:06 |  | Power on 302 |  | 28v current .46A
Friday
| 9:15 | PK |  |  | cmd CTS dual continuum
Testing CTS.
| 9:19 |  | Power off 302
Remove test board connector screws so that connector can mate better to EU J9.
| 9:27 |  | Power on 303 |  | 28V current = .46A
| 9:32 |  |  |  | cmd CTS dual continuum.
| 9:35 |  | Power off 303
Demate tests cable from test board.
| 9:48 |  | Power on 304 |  | 28V current = .47A
| 9:51 |  |  |  | cmd CTD dual continuum
| 9:52 |  |  |  | data storage interupt
| 9:53 |  | Power off 304
Connect pin 2 on test board which was disconnected.
| 10:15 |  | Power on 305 |  | 28V current = .46A
| 10:17 |  |  |  | cmd CTS dual continuum.
| 10:24 |  |  |  | cmd ENG.
Software change: set delay between clk to 25ms from 30sec.
| 10:26 |  |  |  | cmd CTS dual continuum.
| 10:31 |  |  |  | cmd ENG.
| 10:32 |  | Power off 305
| 10:33 |  | Power on 306 |  | Loading s/w to EEPROM: vxWorks_062201_clock5.bin
| 10:49 |  |  |  | Code is running.
| 10:50 |  |  |  | cmd CTS dual continuum.
| 10:51 |  |  |  | 3.3V_Curr = .6 amps.
| 10:56 |  | Power off 306 |  | failure after CTS Run cmd issued
| 10:57 |  | Power on 307 |  | Iin = .46 amps.
| 11:01 |  |  |  | cmd CTS dual continuum.
3.3 V I = .625 amps. Then .591 amps.
| 11:05 |  | Power off 307 |  | failure after CTS Run cmd issued
remove test cable, install flight cable with test board
| 11:15 |  | Power on 308
| 11:17 |  |  |  | cmd Cts Dual continuum.
3.3 V I= .827 then .625 amps.
| 11:19 |  | Power Off 308
| 11:23 |  | Power on 309 |  | Removed Test board and connect Flt cable between CTS-J12 and EU-J9 . 28V current = .46A
| 11:26 |  | mode change |  | CTS-Dual
| 11:34 |  | Power off 309
| 11:35 |  | Power on 310 |  | 28V current = .46A Loading S/W
Loading eeprom with vxWorks_062001_clock3.bin to see if failure is the same as the one on 06-20 at 15:44.
| 11:51 |  | mode change |  | CTS/Dual
| 12:17 |  | Power off 310
| 12:17 |  | Power on 311 |  | Loading eeprom: vxWorks_062201_clock6.bin
| 12:40 |  |  |  | cmd CTS dual continuum.
| 12:49 |  |  |  | ENG
| 12:49 |  | Power off 311
Install test board between EU-J9 and flt cable P9
| 12:55 |  | Power on 312
| 12:58 |  |  |  | cmd CTS dual continuum.
| 13:10 |  | Power off 312
| 13:12 |  | Power on 313
| 13:19 |  |  |  | cmd CTS dual continuum.
| 13:31 |  | Power off 313
2001-06-23
Saturday
| 14:06 |  | Power on 314 |  | 28.4V @ .46A
| 14:10 |  | mode change |  | CTS/Dual
| 14:22 |  | Power off 314
| 15:30 |  |  |  | Paul arrived.
Data cable disconnected from EU elc. Pauls equip. connected to CTS.
| 16:10 |  |  |  | Now operating from Chris's procedures.
| 16:14 |  | Power on 315 |  | Iin = .48 amps.
| 16:19 |  |  |  | cmd CTS dual continuum.
Iin = .88 amps.
| 18:15 |  | Power off 315 |  | Results of testing showed that instrument performed the same with Paul's lab top computer commanding the CTS.
Reconnected data cable to EU
2001-06-25
Monday | 16:10 |  |  |  | Trying more CTS tests - will add code to try turning on 3 data bits at a time and see if failure occurs
| 16:44 |  | Power on 316 |  | Loading eeprom with vxWorks_062501_clock7.bin
| 17:00 |  | cmd |  | CTS/Dual Continuum
Sent the CTS a clocks off command. Waited 3 seconds. Sent it a clocks on command. Waited 3 more seconds. Sent 3 commands separated by 1/10 second each and all 3 worked, unexpectedly! This has always failed in the past with 30 second waits between commands. Then, after 5 minutes we tried to send another command and got a DSI on the first command we tried to send it.
When looked at 3.3 V data, current never went high when all bits set ON but all commands accepted (didn't cause DSI)
| 17:22 |  | Power off 316
2001-06-26
Tuesday | 10:10 |  | Power on 317 |  | Loading eeprom with vxWorks_062601_clock8.bin. This will allow for 2 variable delays between clocks on/off and also between commands.
| 10:26 |  | cmd |  | CTS/Dual Continuum
Sent the CTS a clocks off command. Waited 3 seconds. Sent it a clocks on command. Waited 3 more seconds. Sent 3 commands separated by 1/10 second each and all 3 worked. This has always failed in the past with 30 second waits between commands. Then, after 3 seconds we tried to send another command and got a DSI on the first command we tried to send it.
| 10:43 |  | Power off 317
|  |  | Config. |  | Complete instrument install on copper plate after Mag .
mapping.
| 15:33 |  | Power on 318
Robert is loading new software into the eeprom: vxWorks_061401_ee.bin. Last working software prior to diagnostic additions.
| 15:52 |  |  |  | Cmd to cts /dual cont.
| 15:57 |  |  |  | smm gunn volt = 4
| 15:59 |  |  |  | smm gunn volt = 5
| 15:59 |  |  |  | smm gunn volt = 3
| 16:03 |  | mode change |  | ENG
| 16:04 |  | Power off 318
| 16:21 |  | power on 319
| 16:39 |  |  |  | CMD to cts/dual cont
Running software to send multiple commands after setting all clocks off and all clocks on. Are able to send multiple commands (for about 1 minute) before the DSI happened.
| 16:59 |  | power off 319
| 18:00 |  | Power on 320
| 17:02 |  | cmd |  | CTS Dual Continuum
Running software to send multiple commands after setting all clocks off and all clocks on. This time was able to send multiple commands for about 22 seconds before the DSI happened.
| 17:09 |  | power off 320
| 17:09 |  | power on 321
| 17:11 |  | cmd |  | CTS Dual Continuum
Running software to send multiple commands after setting all clocks off and all clocks on. This time was able to send multiple commands for about 2 seconds before the DSI happened.
| 17:12 |  | power off 321
| 17:13 |  | power on 322
| 17:17 |  | cmd |  | CTS Dual Continuum
| 17:20 |  | Power off 322
| 17:20 |  | Power on 323
| 17:22 |  | mode change |  | CTS/Dual
Running software to send multiple commands after setting all clocks off and all clocks on. This time was able to send multiple commands for about 23 seconds before the DSI happened.
| 17:24 |  | Power off 323
| 17:24 |  | Power on 324
| 17:25 |  | mode change |  | CTS/Dual
Running software to send multiple commands after setting all clocks off and all clocks on. This time was able to send multiple commands for about 4 seconds before the DSI happened.
| 17:26 |  | power off 324
| 17:27 |  | power on 325
| 17:30 |  | cmd |  | CTS Dual Continuum
Running software to send multiple commands after setting all clocks off and all clocks on. This time was able to send multiple commands for about 16 seconds before the DSI happened.
| 17:40 | RN | Power off 325
| 10:45 |  | Power on 326 |  | Testing new SIS s/w version 2
| 11:08 |  | Power off 326
| 11:14 |  | Power on 327 |  | Test s/w loaded in eeprom: vxWorks_062701_pm12.bin
| 11:34 |  | cmd |  | CTS Warmup low level
| 11:35 |  | mode change |  | CTS/DUAL
| 11:35 |  | Power off 327
| 11:36 |  | Power on 328
| 11:39 |  | cmd |  | CTS Warmup low level
| 11:39 |  | mode change |  | CTS/Dual
Instrument died at time about turning off +/- 12V
| 11:43 |  | Power off 328
| 11:44 |  | Power on 329
| 11:46 |  | cmd |  | CTS Warm low level
| 11:47 |  | mode change |  | CTS Dual -
| 11:57 |  | Power off 329 |  | Instrument died at time ablout turning off +/- 12V
| 12:39 |  | Power on 330 |  | Test s/w loaded in eeprom: vxWorks_062701_pm12_2.bin
| 13:00 |  | cmd |  | CTS Warmup low level
| 13:01 |  | mode change |  | CTS/Dual
| 13:02 |  | Power off 330
| 13:04 |  | Power on 331
| 13:06 |  | cmd |  | CTS Warmup low level
| 13:07 |  | mode change |  | CTS/Dual
| 13:25 |  | mode change |  | ENG
| 13:26 |  | cmd |  | CTS Warmup low level
| 13:27 |  | mode change |  | CTS/Dual
| 13:33 |  | mode change |  | ENG
| 13:34 |  | cmd |  | CTS Warmup low level
| 13:37 |  | mode change |  | CTS/Dual
| 13:39 |  | Power off 331
| 13:39 |  | Power on 332
| 13:58 |  | Power off 332
| 13:59 |  | Power on 333 |  | s/w loaded in eeprom vxWorks_062701_ee.bin for TC/TM packet verification test by Peter. Nominal code.
| 14:17 |  | Power off 333
| 14:24 |  | Power on 334 |  | Start of Peter's Test
| 15:36 |  | cmd |  | CTS internal calibration (in order to check rejection s/w)
| 16:06 |  | cmd |  | CTS Warmup level off
| 16:10 |  | cmd |  | CTS Warmup level off
| 16:16 |  | Power off 334
| 16:22 |  | Power on 335
| 16:28 |  | Power off 335
| 16:30 |  | Power on 336
| 16:51 |  | Power off 336
| 16:53 |  | Power on 337 |  | Loading RAM with vxWorks_062701_obfifo.bin to diagnose state of OBFIFO HW both before and after reset
| 17:28 |  | Power off 337
| 17:33 |  | Power on 338
| 17:35 |  | Power off 338
| 18:10 |  | Power on 339
| 18:12 |  | Power off 339
| 18:13 |  | Power on 340
| 18:15 |  | Power off 340
| 18:15 |  | Power on 341
| 18:17 |  | Power off 341
| 18:18 |  | Power on 342
| 18:52 |  | Power off 342
| 18:53 |  | Power on 343
| 18:56 |  | Power off 343
| 18:57 |  | Power on 344
| 19:02 |  | Power off 344
| 19:03 |  | Power on 345
| 19:24 |  | Power off 345
2001-06-28 | 10:34 | RN | Power on 346 |  | No 0xFEFE observed at SIS from HW
| 10:36 | RN | Power off 346
| 10:36 | RN | Power on 347 |  | 0xFEFE observed
| 10:39 | RN | Power off 347
| 10:41 | RN | Power on 348
| 10:47 | RN | Power off 348
| 11:06 | RN | Power on 349 |  | Advanced time synchronization test
| 11:08 | RN | cmd |  | CTS/Dual Continuum
| 12:28 | RN | cmd |  | CTS/Dual Continuum, intercepted at SIS and not sent to instrument
| 13:56 | RN | Power off 349
| 14:01 | RN | Power on 350
| 14:35 | RN | Power off 350
| 14:37 | RN | Power on 351
| 15:10 | RN | Power off 351
| 15:33 | RN | Power on 352
| 17:47 |  | Power off 352 |  | demate primary power cable and connected redundant at SIS unit.
| 17:49 |  | Power on 353 |  | 28.4V @ .66A
| 18:08 |  | Power off 353
| 18:10 |  |  |  | demate redundant power cable and connected primary at SIS unit.
2001-06-29
Friday |  |  |  |  | Dry run of bench test
| 8:43 |  | Power on 354 |  | 28.4V @ .66A
| 8:49 |  | mode change |  | mm cont.
| 8:52 |  | mode change |  | smm cont.
| 8:56 |  | mode change |  | dual cont.
| 8:59 |  | mode change |  | CTS/smm cont
| 9:05 |  | mode change |  | CTS/dual
| 9:12 |  | cmd |  | smm gunn volt auto disable
| 9:12 |  | cmd |  | smm gunn volt = 4
| 9:15 |  | cmd |  | ifp pwr contl = 0
| 9:18 |  | cmd |  | cal htr = on
| 9:20 |  | cmd |  | cal htr = off
| 9:21 |  | cmd |  | USO = off
| 9:22 |  | cmd |  | USO = on
| 9:23 |  | cmd |  | mir pos = hot
| 9:24 |  | cmd |  | mir pos = cold
| 9:25 |  | cmd |  | mir pos = space
| 9:26 |  | cmd |  | CTS inter cal
| 9:30 |  | cmd |  | cont. subtraction 500 mm & 500 smm
| 9:31 |  | cmd |  | cont. subtraction 0 mm & 0 smm
| 9:46 |  | cmd |  | mm lna off
| 9:50 |  | cmd |  | mm lna on
| 9:51 |  | cmd |  | smm lna off
| 9:52 |  | cmd |  | smm lna on
| 9:53 |  | cmd |  | pll reset off
| 9:54 |  | cmd |  | pll reset on
| 9:55 |  | cmd |  | pll reset off
| 9:56 |  | cmd |  | CTS htr contl 40 Deg C low pwr
| 10:00 |  | cmd |  | disable sci data
| 10:00 |  | cmd |  | disable hskp data
| 10:01 |  | cmd |  | enable hskp
| 10:02 |  | cmd |  | enable sci
| 10:03 |  | cmd |  | connection test request
| 10:05 |  | cmd |  | reset tele output buffer
| 10:05 |  | cmd |  | time update
| 10:36 |  | cmd |  | ENG MODE
| 10:43 |  | Power off 354
| 14:33 |  |  |  | Start to shutdown all suport equipement and getting ready to move to 306-221 B on Monday 7-2-01
2001-07-03 | 13:24 | TCK |  |  | All support equipment down in 306-221.
SU, USO, EU-CTS.digital, SBEU mounted on IHF (instrument handling fixture) on Optical table in front of NFR for smm NFR measurement
|  | Pstek |  |  | Setting up external bias for smmRFE for smm NFR AIDS#216861; Note: only using SU
|  | TCK |  |  | See NFR_log04.txt for NFR testing details.
| 18:00 | TCK |  |  | Completed 4 LSB and 4 USB smmNFR patterns.
2001-07-05 |  |  |  |  | configuration CTS, EU (full up instrument)
| 16:38 |  | Power on 355 |  | 28.4V .66A
Receiving eng data.
| 16:49 | AP | cmd |  | CTS/Dual mode
|  | TCK |  |  | Note: smmRFE IF W13 is flex w/ 15 dB pad cnt=6926
|  | TCK |  |  | Note: mmRFE IF W12 is flex w/ 6 dB pad cnt=7195
| 16:53 | AP | cmd |  | CTS heater: 40C, low
| 16:54 | AP | cmd |  | calibration heater ON
Twload1=20.31C Twload2=20.27C
| 18:18 | TCK |  |  | program attn in mmRFE line no extra pad
set attn to 5 dB, mm cnts=7734
| 18:20 | TCK | cmd |  | CTS internal calib
| 18:24 | TCK | cmd |  | CTS/Dual (instr calib)
Note: overs are on 2nd and 1st mirrors!
| 18:27 | TCK |  |  | mm attn = 0, mm cnts=15481
| 18:29 |  |  |  | mm attn = 1, mm cnts=13534
| 18:32 |  |  |  | mm attn = 2, mm cnts=12023
| 18:33 |  |  |  | mm attn = 3, mm cnts=10352
| 18:34 |  |  |  | mm attn = 4, mm cnts=9058
| 18:34 |  |  |  | mm attn = 5, mm cnts=7733
| 18:35 |  |  |  | mm attn = 6, mm cnts=6783
| 18:36 |  |  |  | mm attn = 7, mm cnts=5795
| 18:37 |  |  |  | mm attn = 8, mm cnts=5045
| 18:38 |  |  |  | mm attn = 9, mm cnts=4357
| 18:39 |  |  |  | mm attn = 10, mm cnts=3888
| 18:40 |  |  |  | mm attn = 5
| 18:40 |  | cmd |  | mm LNA off
mm cnts=1681
| 18:41 |  | cmd |  | mm LNA on
move programmable attn to smmRFE line plus 10 dB
set mmRFE pad to 6 dB
| 18:45 |  |  |  | smm attn 0
smm attn = 0, smm cnts=14273
| 18:46 | TCK | cmd |  | CTS/Dual
| 18:51 |  |  |  | smm attn = 1, smm cnts=12081
| 18:51 |  | cmd |  | CTS/Dual
| 18:54 |  |  |  | smm attn = 2, smm cnts=10306
| 18:55 |  | cmd |  | CTS/Dual
| 18:58 |  |  |  | smm attn = 3, smm cnts=8677
| 18:59 |  | cmd |  | CTS/Dual
| 19:03 |  |  |  | smm attn = 4, smm cnts=
| 19:03 |  | cmd |  | CTS/Dual
| 19:06 |  |  |  | smm attn = 5, smm cnts=7396
| 19:06 |  | cmd |  | CTS/Dual
smm attn = 6, smm cnts=5446
| 19:10 |  | cmd |  | CTS/Dual
smm attn = 7, smm cnts=4696
| 19:13 |  | cmd |  | CTS/Dual
smm attn = 8, smm cnts=4081
| 19:17 |  | cmd |  | CTS/Dual
smm attn = 9, smm cnts=3610
| 19:19 |  | cmd |  | CTS/Dual
smm attn = 10, smm cnts=3253
| 19:23 |  | cmd |  | CTS/Dual
| 19:27 |  | cmd |  | smm LNA off
smm cnts=1755
| 19:30 |  | cmd |  | smm LNA on
| 19:31 | TCK | cmd |  | engineering mode
| 19:47 | TCK | Power OFF 355
2001-07-06 | 9:59 | TCK/AP | Power ON 356 |  | looking normal…
2001-07-06
2001-07-06 |  | TCK |  |  | continuing linearity testing of IFP channels with injected noise sources
2001-07-06
2001-07-06 |  | TCK |  |  | Note: using same components as in T/V linearity testing
2001-07-06
2001-07-06 |  |  | config |  | smmRFE 10+5 dB fixed pad
2001-07-06
2001-07-06 |  |  | config |  | mmRFE=0..11 dB attn=coupler=IFP
2001-07-06 |  |  |  |  | coupler=0..70 dB attn=mm Noise Source
2001-07-06 |  |  |  |  | mm Noise Source is Avantek plus 10+1+1 dB
2001-07-06
2001-07-06 | 10:01 | TCK/AP | cmd |  | CTS/dual mode
2001-07-06
2001-07-06 |  |  |  |  | mm attn= 6, noise attn=70 dB, mm cnts=6675
2001-07-06 |  |  |  |  | mm attn= 6, noise attn= 0 dB, mm cnts=7177
2001-07-06
2001-07-06 | 10:09 |  |  |  | mm attn= 4, noise attn=70 dB, mm cnts=8849
2001-07-06 | 10:11 |  |  |  | mm attn= 4, noise attn= 0 dB, mm cnts=9276
2001-07-06
2001-07-06 | 10:13 |  |  |  | mm attn= 2, noise attn=70 dB, mm cnts=11727
2001-07-06 | 10:15 |  |  |  | mm attn= 2, noise attn= 0 dB, mm cnts=12077
2001-07-06
2001-07-06 | 10:16 |  |  |  | mm attn= 0, noise attn=70 dB, mm cnts=15128
2001-07-06 | 10:17 |  |  |  | mm attn= 0, noise attn= 0 dB, mm cnts=15388
2001-07-06
2001-07-06 | 10:19 |  |  |  | mm attn= 8, noise attn=70 dB, mm cnts=4910
2001-07-06 | 10:20 |  |  |  | mm attn= 8, noise attn= 0 dB, mm cnts=5481
2001-07-06
2001-07-06 | 10:21 |  |  |  | mm attn=10, noise attn=70 dB, mm cnts=3793
2001-07-06 | 10:21 |  |  |  | mm attn=10, noise attn= 0 dB, mm cnts=4399
2001-07-06
2001-07-06 | 10:23 |  |  |  | mm attn= 6, noise attn=70 dB, mm cnts=6567
2001-07-06 | 10:25 |  |  |  | mm attn= 6, noise attn= 0 dB, mm cnts=7083
2001-07-06
2001-07-06 | 10:26 |  |  |  | mm attn= 6, noise attn=70 dB
2001-07-06
2001-07-06 | 10:27 | AP |  |  | mm LNA off
2001-07-06 |  |  |  |  | mm cnts=1679
2001-07-06 | 10:29 | AP |  |  | mm LNA on
2001-07-06
2001-07-06 |  |  | config |  | mmRFE with 6 dB fixed pad
2001-07-06
2001-07-06 |  |  | config |  | smmRFE=0..11 dB attn=coupler=IFP
2001-07-06 |  |  |  |  | coupler=0..70 dB attn=smm Noise Source
2001-07-06 | 10:39 |  |  |  | smm Noise Source is JCA amp plus 3 dB
2001-07-06
2001-07-06 | 10:39 |  |  |  | smm attn= 6, noise attn=70 dB, smm cnts=4917
2001-07-06 | 10:41 |  |  |  | smm attn= 6, noise attn= 0 dB, smm cnts=6411
2001-07-06
2001-07-06 | 10:45 |  | config |  | add 6 dB total 3+6 dB on smm Noise source
2001-07-06
2001-07-06 |  |  |  |  | smm attn= 6, noise attn= 0 dB, smm cnts=5298
2001-07-06 | 10:46 |  |  |  | smm attn= 6, noise attn=70 dB, smm cnts=4912
2001-07-06 | 10:52 |  | cmd |  | CTS/Dual
2001-07-06 | 10:55 |  |  |  | smm attn= 6, noise attn= 0 dB, smm cnts=5292
2001-07-06 | 10:55 |  | cmd |  | CTS/Dual
2001-07-06
2001-07-06 | 10:58 |  | cmd |  | CTS heater 40C, high
2001-07-06 | 10:59 | 0:15 | cmd |  | calibration load heater ON
2001-07-06 | 11:13 | 0:03 | cmd |  | CTS internal calib
2001-07-06
2001-07-06 | 11:17 |  |  |  | smm attn= 4, noise attn=70 dB, smm cnts=6583
2001-07-06 | 11:17 | 0:03 | cmd |  | CTS/Dual
2001-07-06 | 11:20 |  |  |  | smm attn= 4, noise attn= 0 dB, smm cnts=6958
2001-07-06 | 11:21 | 0:03 | cmd |  | CTS/Dual
2001-07-06
2001-07-06 | 11:24 |  |  |  | smm attn= 0, noise attn=70 dB, smm cnts=12713 space
2001-07-06 | 11:24 | 0:03 | cmd |  | CTS/Dual
2001-07-06 | 11:27 |  |  |  | smm attn= 0, noise attn= 0 dB, smm cnts=13035 cold
2001-07-06 | 11:28 | 0:03 | cmd |  | CTS/Dual
2001-07-06
2001-07-06 | 11:31 |  |  |  | smm attn=10, noise attn=70 dB, smm cnts=3024 cold
2001-07-06 | 11:31 | 0:03 | cmd |  | CTS/Dual
2001-07-06 | 11:35 |  |  |  | smm attn=10, noise attn= 0 dB, smm cnts=3434 cold?
2001-07-06 | 11:35 | 0:03 | cmd |  | CTS/Dual
2001-07-06
2001-07-06 | 11:38 |  |  |  | smm attn= 6, noise attn=70 dB, smm cnts=4889
2001-07-06 | 11:38 | 0:03 | cmd |  | CTS/Dual
2001-07-06 | 11:41 |  |  |  | smm attn= 6, noise attn= 0 dB, smm cnts=5277
2001-07-06 | 11:41 | 0:03 | cmd |  | CTS/Dual
2001-07-06
2001-07-06 | 11:44 |  |  |  | smm attn= 6, noise attn=70 dB
2001-07-06
2001-07-06 | 11:45 | 0:01 | cmd |  | smm LNA off
2001-07-06 | 11:45 | 0:03 | cmd |  | CTS/Dual
2001-07-06 |  |  |  |  | smm cnts=1755
2001-07-06 | 11:50 | 0:01 | cmd |  | smm LNA on
2001-07-06 |  |  |  |  | empty telemetry queue…
2001-07-06 | 12:06 | AP/TCK | cmd |  | engineering mode
2001-07-06
2001-07-06 | 12:10 | AP | Power OFF 356 |  | Finish with TK test.
2001-07-06
2001-07-06 |  |  | config
2001-07-06 | 13:08 |  | Power on 357 |  | RN is running new code. EEPROM loaded with vxWorks_070601_ee.bin to test engineering HK packet change.
| 13:44 | RN | Power off 357
| 13:46 | RN | Power on 358 |  | Loading new software into EEPROM. vxWorks_070601_ee_2.bin

## Sheet2
Mirror alignment verification test of sorts
time | degrees | mm | smm
3:31 | 0 | 6420 | 8275
3:33 | -0.54 | 6422 | 8269
3:34 | -1.08 | 6441 | 8274
3:34 | -1.62 | 6448 | 8274
3:34 | -2.16 | 6492 | 8280
3:34 | -2.7 | 6499 | 8308
3:34 | -3.24 | 6517 | 8326
3:38 | -3.78 | 6549 | 8346
3:39 | -4.32 | 6569 | 8355
3:40 | -4.68 | 6599 | 8354
3:40 | -5.4 | 6606 | 8350
3:41 | -5.94 | 6621 | 8348
3:41 | -6.48 | 6631 | 8350
3:42 | -7.02 | 6645 | 8352
3:43 | -7.56 | 6648 | 8355
3:43 | -8.1 | 6652 | 8360
3:44 | -8.64 | 6664 | 8361
3:44 | -9.18 | 6662 | 8363
3:45 | -9.72 | 6667 | 8369
3:45 | -9.18 | 6670 | 8362
3:46 | -8.64 | 6668 | 8361
3:47 | -8.1 | 6663 | 8358
3:47 | -7.56 | 6659 | 8357
3:48 | -7.02 | 6650 | 8353
3:49 | -6.48 | 6642 | 8349
3:49 | -5.94 | 6632 | 8345
3:50 | -5.4 | 6634 | 8341
3:50 | -4.86 | 6632 | 8338
3:51 | -4.32 | 6615 | 8339
3:51 | -3.78 | 6597 | 8341
3:52 | -3.24 | 6564 | 8338
3:52 | -2.7 | 6540 | 8328
3:53 | -2.16 | 6520 | 8315
3:54 | -1.62 | 6490 | 8290
3:54 | -1.08 | 6449 | 8245
3:55 | -0.54 | 6430 | 8245
3:56 | 0 | 6418 | 8238
3:57 | 0.54 | 6391 | 8235
3:57 | 1.08 | 6390 | 8235
|  |  |  | 0.03 | 1.0069 | 138633.377

## Sheet3
6:29:01, 631:52 | Hot load |  | calibration | calibration
6:33 | 0 | 0 | 7784 | 6845
6:34 | -5 | -1.35 | 7782 | 6842
6:35 | -10 | -2.7 | 7782 | 6841
6:36 | -15 | -4.05 | 7780 | 6840
6:37 | -20 | -5.4 | 7776 | 6838
6:38 | -25 | -6.75 | 7779 | 6836
6:39 | -30 | -8.1 | 7777 | 6834
6:39 | -35 | -9.45 | 7774 | 6832
6:40 | -40 | -10.8 | 7776 | 6830
6:41 | -45 | -12.15 | 7771 | 6827
6:42 | -50 | -13.5 | 7750 | 6824
6:43 | -55 | -14.85 | 7739 | 6813
6:44 | -60 | -16.2 | 7737 | 6814
6:44 | -65 | -17.55 | 7716 | 6812
6:45 | -70 | -18.9 | 7708 | 6809
6:46 | -75 | -20.25 | 7702 | 6807
6:47 | -80 | -21.6 | 7689 | 6801
6:48 | -85 | -22.95 | 7680 | 6798
6:48 | -90 | -24.3 | 7677 | 6793
6:49:29, 6:52:57 | Hot load | Warm load=43.1,
Cold load=21.1 | Calibration | Calibration
6:53 | -90 | -24.3 | 7674 | 6784
6:54 | -85 | -22.95 | 7676 | 6783
6:55 | -80 | -21.6 | 7684 | 6785
6:56 | -75 | -20.25 | 7689 | 6785
6:57 | -70 | -18.9 | 7695 | 6785
6:58 | -65 | -17.55 | 7702 | 6784
6:59 | -60 | -16.2 | 7713 | 6782
7:00 | -55 | -14.85 | 7716 | 6779
7:0048 | -50 | -13.5 | 7751 | 6780
7:02 | -45 | -12.15 | 7750 | 6785
7:03 | -40 | -10.8 | 7774 | 6786
7:03 | -35 | -9.45 | 7766 | 6784
7:04 | -30 | -8.1 | 7768 | 6783
7:05 | -25 | -6.75 | 7771 | 6781
7:06 | -20 | -5.4 | 7768 | 6779
7:07 | -15 | -4.05 | 7770 | 6778
7:08 | -10 | -2.7 | 7768 | 6777
7:08 | -5 | -1.35 | 7767 | 6775
7:09 | 0 | 0 | 7760 | 6774
7:09:54, 7:12:10 | Hot load | Hot load=44.3
Cold load =21.3 | calibration | calibration
7:12 | 0 | 0 | 7766 | 6767
7:13 | 5 | 1.35 | 7766 | 6766
7:14 | 10 | 2.7 | 7764 | 6763
7:15 | 15 | 4.05 | 7763 | 6763
7:16 | 20 | 5.4 | 7763 | 6761
7:17 | 25 | 6.75 | 7762 | 6760
7:18 | 30 | 8.1 | 7759 | 6759
7:18 | 35 | 9.45 | 7756 | 6758
7:19 | 40 | 10.8 | 7752 | 6756
7:20 | 45 | 12.15 | 7745 | 6757
7:21 | 50 | 13.5 | 7745 | 6753
7:22 | 55 | 14.85 | 7719 | 6749
7:22 | 60 | 16.2 | 7696 | 6743
7:23 | 65 | 17.55 | 7695 | 6739
7:24 | 70 | 18.9 | 7682 | 6737
7:24;58 | 75 | 20.25 | 7674 | 6735
7:26 | 80 | 21.6 | 7663 | 6728
7:27 | 85 | 22.95 | 7648 | 6723
7:27 | 90 | 24.3 | 7641 | 6722
7:28:14, lost mirror position | Hot load | HOT LOAD =44.82
Cold load=21.4 | calibration | calibration
41347 | 0.89517
31401 | 0.67984
19969 | 0.43233
6515 | 0.14105
46189 | 1
35840 | 0.85535
27702 | 0.66113
17954 | 0.42849
5855 | 0.13973
41901 | 1
9.36 | 3.171 | 6649
9.16 | 2.855 | 6933.9
9.09 | 1.932 | 6926.3
9.02 | 0.534 | 6929.6
8.95 | 0.534 | 6980.1
8.74 | 3.059 | 7034.2
8.53 | 3.061 | 4573.2
8.32 | 3.044 | 4575.2
8.53 | 3.061 | 4576.7
8.74 | 3.059 | 7400.5
8.95 | 0.399 | 7032.2
9.02 | 0.629 | 6914.4
9.09 | 2.046 | 6904.4
9.16 | 2.928 | 6903.2
9.36 | 3.178 | 6651.1
9.09 | 1.977 | 6898.5
Variable | MM | Submm
Chot | 6858 | 6899
Ccold | 6817 | 6891
Coff | 1671 | 1754
Y | 1.008 | 1.0016
Tcold (K) (C+273) | 295.3 | 295.3
Thot (K) (C+273) | 305.1 | 305.1
Tsys | 934.7195 | 5997.525

## Sheet4
9.36 | 3.17 | 6504
9.16 | 2.807 | 6790
9.09 | 1.937 | 6786
9.02 | 0.508 | 6803
8.95 | 0.399 | 6915
8.74 | 3.043 | 7286
8.53 | 3.063 | 4525
8.32 | 3.046 | 4527
8.53 | 3.046 | 4529
8.74 | 3.063 | 7298
8.95 | 0.394 | 6921
9.02 | 0.54 | 6832
9.09 | 2.006 | 6789
9.16 | 2.088 | 6787
9.36 | 3.175 | 6477
9.09 | 2.001 | 6785
41317 | 1
35251 | 0.85318
26762 | 0.64772
16869 | 0.40828
5573 | 0.13488
19990 | 0.48382
20008 | 0.48426
Variable | MM | Submm
Chot | 8069 | 6958
Ccold | 8012 | 6950
Coff | 1670 | 1736
Y | 1.009 | 1.0015
Tcold (K) (C+273) | 295.1 | 295.1
Thot (K) (C+273) | 305.4 | 305.4
Tsys | 850.9105 | 6417.925

## Sheet5
DATE | TIME | INTLs | KEYWORD | PURPOSE | ACTIVITY DETAILS
Sam's MASTER - for current master, see Sheet 6
| do not use these - they have been updated |  |  |  | NOTES:
| START TIME | INTERVAL | COMMAND |  | Test Ports in Nominal Configuration
Noise Sources OFF
Spectrum Analyzer on PLE Test Port (J381)
SEE TEMPERATURE CYCLE TABLE
CYCLE NUMBER 4
INSTRUMENT MAY ALREADY BE POWERED ON: ENTER "NEW POWER ON" NUMBER:RECORD VOLTAGE= ### ;AMP = ###
| 19:30 | 0:00 | POWER ON = (SKIP IF ALREADY ON)
| 19:30 | 0:00 | Warm load heater on |  | If USO IS TURNED OFF, ALLOW 30 MINUTES WARMUP BEFORE PROCEEDING OTHERWISE PROCEED AFTER ISSUING COMMAND
| 19:30 | 0:00 | USO WARMUP/ |  | THIS COMMAND IS CYCLE DEPENDENT, HEATER IS ALWAYS 0(HIGH), T CHANGES WITH LARGE PLATE TEMPERATURE; SEE TABLE; ALLOW 30 MINUTES FOR WARMUP
| 19:30 | 0:00 | CTS WARMUP MODE,HEATER =0(HIGH),T=6(60C)/ |  | SYSTEM WILL AUTOMATICALLY CALIBRATE; IFP WILL WARM/COOL SLOWLY...
| 19:30 | 0:40 | CTS -DUAL CONTINUUM/
| 20:10 |  |  |  | REQUEST LARGE PLATE TEMPERATURE= 32C
| 20:10 | 0:00 |  |  | REQUEST SMALL OB PLATE TEMPERATURE= 32 C
| 20:10 |  |  |  | REQUEST COLD LOAD PLATE TEMPERATURE= COLD
20:10
| 20:10 |  |  |  | Record Spectra as per "4.6 LO lock and Frequency Switching Test" as necessary during transition.
| 20:10 |  |  |  | Filename: TV####.tdx
| 20:10 | 0:00 |  |  | Filename: TV####.tdx
| 20:10 | 0:00 |  |  | Filename: TV####.tdx
| 20:10 | 0:00 |  |  | Filename: TV####.tdx
| 20:10 | 0:00 |  |  | AFTER THREE PLATES HAVE STABILIZED, RECORD TEMPERATURES AND PROCEED TLARGE= ### ;TSMALL = ### ; TCOLD = ###
| 20:10 |  | large and small plates should regulate to less than 1 deg c
| 20:10 |  |  |  | Test Ports as per "4.1 Noise Figure of the IFP"
| 20:10 | 0:05 |  |  | Record Trace B
| 20:15 | 0:01 | smm LNA OFF/ |  | Record Trace A; Filename: TV####.tdx
| 20:16 | 0:01 | smm LNA ON/
| 20:17 |  |  |  | Test Ports as per "4.2 Noise Figure of the CTS"
| 20:17 | 0:05 |  |  | "hot" configuration
| 20:22 | 0:05 | CTS -DUAL CONTINUUM/1 |  | "cold" configuration; Filename: TV####.tdx
| 20:27 | 0:05 | CTS -DUAL CONTINUUM/1
| 20:32 |  |  |  | Test Ports as per "4.3 IF Frequency Sweep Test"
| 20:32 | 0:05 |  |  | SigGen=5.753 GHz
| 20:37 | 0:05 | CTS -DUAL CONTINUUM/1 |  | SigGen=5.877
| 20:42 | 0:05 | CTS -DUAL CONTINUUM/2 |  | SigGen=9.667
| 20:47 | 0:05 | CTS -DUAL CONTINUUM/3 |  | SigGen=9.685
| 20:52 | 0:05 | CTS -DUAL CONTINUUM/4 |  | SigGen=10.791
| 20:57 | 0:05 | CTS -DUAL CONTINUUM/5 |  | SigGen=13.455
| 21:02 | 0:05 | CTS -DUAL CONTINUUM/6 |  | SigGen=15.137
| 21:07 | 0:05 | CTS -DUAL CONTINUUM/7 |  | SigGen=16.338
| 21:12 | 0:05 | CTS -DUAL CONTINUUM/8
| 21:17 |  |  |  | Test Ports as per "4.4 Linearity in CTS/Dual Mode"
| 21:17 | 0:05 |  |  | Turn on Chopping Noise Sources
| 21:22 | 0:01 |  |  | nominal pads
| 21:22 | 0:10 | CTS -DUAL CONTINUUM/1 |  | +2 dB pads
| 21:32 | 0:10 | CTS -DUAL CONTINUUM/2 |  | -2 dB pads
| 21:42 | 0:10 | CTS -DUAL CONTINUUM/3
| 21:52 |  |  |  | Test Ports as per "4.5 Linearity in Dual Continuum Mode"
| 21:52 | 0:05
| 21:57 | 0:05 | DUAL-CONTINUUM/ |  | Noise Sources ON
| 22:02 | 0:01 |  |  | nominal pads
| 22:03 | 0:15 | DUAL-CONTINUUM/1 |  | +2 dB pads
| 22:18 | 0:15 | DUAL-CONTINUUM/2 |  | -2 dB pads
| 22:33 | 0:15 | DUAL-CONTINUUM/3
| 22:48 |  |  |  | Noise Sources OFF
| 22:48 | 0:01
| 22:49 |  |  |  | Test Ports as per "4.6 LO lock and Frequency Switching Test" (Nominal)
| 22:49 | 0:05 | CTS -DUAL CONTINUUM/ |  | Record Spectra per 4.6...
| 22:54 |  |  |  | Filename: TV####.tdx
| 22:54 | 0:03 |  |  | Filename: TV####.tdx
| 22:57 | 0:03 |  |  | Filename: TV####.tdx
| 23:00 | 0:03
| 23:03 |  |  |  | Go to cycle 5
| 23:03 |  |  |  | REMEMBER TO INCREASE THE TEMPERTURE SETTINGS FOR CTS WARMUP MODE AND CTS HEAT MODE
23:03

## Sheet6
START TIME | INTERVAL | COMMAND | Test Ports in Nominal Configuration
Noise Sources OFF
Spectrum Analyzer on PLE Test Port (J381)
SEE TEMPERATURE CYCLE TABLE
CYCLE NUMBER 5
0:00 |  |  | Going to +25 at interface
0:00 |  |  | REQUEST LARGE PLATE TEMPERATURE= 22 C
0:00 | 0:00 |  | REQUEST SMALL OB PLATE TEMPERATURE= 22C
0:00 |  |  | REQUEST COLD LOAD PLATE TEMPERATURE= COLD
0:00 |  |  | Record Spectra as per "4.6 LO lock and Frequency Switching Test" as necessary during transition.
0:00 |  |  | Filename: TV0169.tdx to TV0170.tdx (30 MHz sweep only)
0:00 | 0:00 |  | AFTER THREE PLATES HAVE STABILIZED, RECORD TEMPERATURES AND PROCEED TLARGE= xxx ;TSMALL = xxx ; TCOLD = xxx
0:00 |  | large and small plates should regulate to less than 1 deg c
INSTRUMENT MAY ALREADY BE POWERED ON: ENTER "NEW POWER ON" NUMBER:RECORD VOLTAGE= 28.3 ;AMP = 2.48
0:00 | 0:00 | POWER IS ON
0:00 | 0:00 | USO is on | If USO IS TURNED OFF, ALLOW 30 MINUTES WARMUP BEFORE PROCEEDING OTHERWISE PROCEED AFTER ISSUING COMMAND
0:00 | 0:00 |  | THIS COMMAND IS CYCLE DEPENDENT, HEATER IS ALWAYS 0(HIGH), T CHANGES WITH LARGE PLATE TEMPERATURE; SEE TABLE; ALLOW 30 MINUTES FOR WARMUP
0:00 | 0:30 | COMMAND | CTS Heater Contol = 60 degrees C
0:30 | 0:00 | CTS -DUAL CONTINUUM
Needed barrel connectors to terminate for next test. None available. So, took terminations off input cables, re-connected noise sources to input cables, and turned power off to noise sources.
0:30 |  |  | Test Ports as per "4.1 Noise Figure of the IFP"
0:30 | 0:05 |  | Config: disconnect Spec Anal from and terminate PLE Test Port
0:35 |  |  | Config: disconnect IFP output from CTS input (J304 to J307)
0:35 |  |  | Config: terminate CTS input J307; Spec Anal on IFP output J304
0:35 |  |  | Config: Spec Anal, Recall, State 2
0:35 |  |  | Record Trace B
0:35 | 0:01 | smm LNA OFF/ | Record Trace A; Filename: TV0169.tdx
0:36 | 0:01 | smm LNA ON/
"4.2 Noise Figure of the CTS-Spectrum"
Config: No change
0:37 |  |  | "cold" configuration ("IFP off"), Record Trace A; Filename: TV0170.tdx
0:37 | 0:05 | CTS -DUAL CONTINUUM/1
Config: unterminate CTS input J307; demate Spec Anal from IFP output J304
Config: connect IFP output to CTS input (J304 to J307)
Config: unterminate PLE Test Port and connect Spec Anal to PLE Test Port
0:42 |  |  | "hot" configuration ("IFP on") from Test 4.2 (nominal configuration)
0:42 | 0:05 | CTS -DUAL CONTINUUM/1
0:47 |  |  | Test Ports as per "4.3 IF Frequency Sweep Test"
0:47 |  |  | Config: connect IFP output to CTS input (J304 to J307) [already done]
0:47 |  |  | Config: connect Spec Anal to PLE Test Port (J381) [already done]
0:47 |  |  | Config: Spec Anal, Recall, State 0
0:47 |  |  | Config: connect Sig Gen to smm coupler
0:47 |  | RF On | Turn Sig Gen RF On
0:47 |  | Change Frequency | SigGen=5.753 GHz (-80 dBm)
0:47 | 0:05 | CTS -DUAL CONTINUUM/1
|  | Change Frequency | SigGen=5.877 GHz
0:52 | 0:05 | CTS -DUAL CONTINUUM/2
|  | Change Frequency | SigGen=9.667 GHz
0:57 | 0:05 | CTS -DUAL CONTINUUM/3
|  | Change Frequency | SigGen=9.685 GHz
1:02 | 0:05 | CTS -DUAL CONTINUUM/4
|  | Change Frequency | SigGen=10.791 GHz
1:07 | 0:05 | CTS -DUAL CONTINUUM/5
|  | Change Frequency | SigGen=13.455 GHz
1:12 | 0:05 | CTS -DUAL CONTINUUM/6
|  | Change Frequency | SigGen=15.137 GHz
1:17 | 0:05 | CTS -DUAL CONTINUUM/7
|  | Change Frequency | SigGen=16.338 GHz
1:22 | 0:05 | CTS -DUAL CONTINUUM/8
Config: Sig Gen RF off.
1:27 |  | RF Off | Turn RF Off
Test Ports as per "4.4 Linearity in CTS/Dual Continuum Mode"
1:27 |  |  | Remove submm coupler input line from Sig Gen RF out. Put termination on submm coupler input line.
1:27 |  |  | Config: nominal state.
Ensure "0" dB state at smm and mm couplers (2 dB at each coupler) Note: On smm, this is two 1dB pads. On mm, this is one 2 dB pad.
1:27 | 0:05 | CTS-DUAL-CONTINUUM/1
1:32 |  | Connect Noise Sources | Noise Source ON
1:32 | 0:05 | CTS-DUAL-CONTINUUM/1
1:37 |  |  | Noise Source OFF
1:37 |  |  | Config: "+2 dB" state
Remove 2 dB (total) pads from smm and mm couplers. Insert 4 dB pads in their place.
1:37 | 0:05 | CTS-DUAL-CONTINUUM/1
1:42 |  | Connect Noise Sources | Noise Source ON
1:42 | 0:05 | CTS-DUAL-CONTINUUM/1
1:47 |  |  | Noise Source OFF
1:47 |  |  | Config: "-2 dB" state
Remove 4 dB pads from smm and mm couplers. Connect lines directly to couplers.
1:47 | 0:05 | CTS-DUAL-CONTINUUM/1
1:52 |  | Connect Noise Sources | Noise Source ON
1:52 | 0:05 | CTS-DUAL-CONTINUUM/1
1:57 |  |  | Noise Source OFF
Test Ports as per "4.5 Linearity in Dual Continuum Mode"
Leave pads as they are (Config: ("-2" total))
1:57 |  |  | Config: "-2 dB" state
Ensure "-2" dB state at smm and mm couplers. Lines connected directly to couplers
1:57 | 0:05 | DUAL-CONTINUUM/1
2:02 |  | Connect Noise Sources | Noise Source ON
2:02 | 0:05 | DUAL-CONTINUUM/1
2:07 |  |  | Noise Source OFF
2:07 |  |  | Config: "+2 dB" state
Insert 4 dB pads at couplers.
2:07 | 0:05 | DUAL-CONTINUUM/1
2:12 |  | Connect Noise Sources | Noise Source ON
2:12 | 0:05 | DUAL-CONTINUUM/1
2:17 |  |  | Noise Source OFF
2:17 |  |  | Config: nominal state
Remove 4 dB pads from smm and mm couplers. Go to "0" dB state at smm and mm couplers (2 dB at each coupler) Note: On smm, this is two 1dB pads. On mm, this is one 2 dB pad.
2:17 | 0:05 | DUAL-CONTINUUM/1
2:22 |  | Connect Noise Sources | Noise Source ON
2:22 | 0:05 | DUAL-CONTINUUM/1
2:27 |  |  | Noise Source OFF
2:27 |  |  | Test Ports as per "4.6 LO lock and Frequency Switching Test" (Nominal)
2:27 | 0:05 | CTS -DUAL CONTINUUM/ | Record Spectra per 4.6...
2:32 |  |  | Filename: TV0170.tdx - TV0175.tdx
2:32
2:32 |  |  | Go to cycle 6
2:32 |  |  | REMEMBER TO DECREASE THE TEMPERATURE SETTINGS FOR CTS (USE CTS WARMUP COMMAND IF CTS HAS BEEN OFF, USE CTS HEATER CONTROL COMMAND IF CTS IS ALREADY ON)

## linear
IFP linearity tests TCK 7/5/01
| mm-attn |  |  | mm-cnts
| 0 | 10 | 10 | 15481 | 13800 | 41.3988 | 10
| 1 | 9 | 7.9433 | 13534 | 11853 | 40.7383 | 10
| 2 | 8 | 6.3096 | 12023 | 10342 | 40.146 | 10
| 3 | 7 | 5.0119 | 10352 | 8671 | 39.3807 | 10
| 4 | 6 | 3.9811 | 9058 | 7377 | 38.6788 | 10
| 5 | 5 | 3.1623 | 7733 | 6052 | 37.819 | 10
| 6 | 4 | 2.5119 | 6783 | 5102 | 37.0774 | 10
| 7 | 3 | 1.9953 | 5795 | 4114 | 36.1426 | 10
| 8 | 2 | 1.5849 | 5045 | 3364 | 35.2686 | 10
| 9 | 1 | 1.2589 | 4357 | 2676 | 34.2749 | 10
| 10 | 0 | 1 | 3888 | 2207 | 33.438 | 10
|  |  |  | 1681 | 0