Calibration Information for the Upward-Looking Infrared Spectrometer (ULIS)

and Downward-Looking Infrared Spectrometer (DLIS)

20 February 2006

- General Information

The infrared spectrometer in the Descent Imager/Spectral Radiometer (DISR) instrument contains a small grating spectrometer and a pair of linear array detectors.  Optical fibers are used to feed light from upward-looking and downward-looking foreoptics to two locations along the input slit of the grating spectrometer.  The array detectors contain 150 elements each.  The first two elements are not used.  The next 5 detector elements are covered by an opaque mask to provide some information of the dark current of the instrument at the temperature of operation.  The next two elements form a transition from the opaque mask to a transparent region.  The next 132 elements are the active detector elements of the instrument.  The last two elements are used to determine the zero level of the signal, and the 5 elements before have the opaque mask, with the two before forming a transition region.

A mechanical shutter in front of the entrance slit of the spectrometer is closed to measure the thermally generated dark current from the detector.  The measurements with the shutter closed and with the shutter open are both transmitted to the ground.  The integration time and the shutter period are both controlled by auto-exposure algorithms in the instrument.

In order to interpret the data from the IR spectrometer several parameters from the calibration of the instrument are required.  These include the wavelength calibration of the ULIS and DLIS as functions of the temperature of the optics, the spectral resolution of the ULIS and DLIS spectrometers as functions of optics temperature and wavelength,  the absolute responsivity of the ULIS and DLIS instruments as functions of the temperature of the detector arrays, and the relative sensitivity of the instrument across the fields of view of the ULIS and DLIS.  These parameters are provided in tables and functions given below.

A few other pieces of information are required to reduce the data from raw DN to reduced absolutely calibrated average intensity over the field of view of the instrument.  The data are almost always taken in sets of two, one set with the shutter open and the other with the shutter closed.  The charge from the IR detector pixels are read out by two linear rows of CCD detectors.  At the beginning of an exposure, the wells of the detector pixels are filled with charge from the CCD.  As the exposure continues, charge is drained from the detector wells at a rate that is proportional to the light level and the thermally generated “dark current”.  At the end of the exposure, the exposure level is read by recording the amount of charge from the CCD that is required to refill the wells of the detectors.  The signal is the difference between the “precharge” initially placed in the pixels and the amount of charge remaining.  This value is given by the difference in DN between the pixels that record the precharge (pixels 148 and 149) and the DN in a given active pixel in the array.  Finally, the leads from the detector pixels from the odd numbered and even numbered pixels lead to two separate CCD detectors.  Thus, the signal from an even numbered pixel, i, is given by DN(148)-DN(i), and the signal from an odd numbered pixel is given by DN(149)-DN(i).  The net signal is (DN(precharge pixel)-DN(i)) with shutter open – (DN(precharge pixel)-DN(i) with the shutter closed.  The net count rate is this quantity divided by the average exposure time (equal to the “sample time”) in the raw data.  This is the quantity that is proportional to the radiant power on a detector element.

We planned for the operation of the IR spectrometer to be synchronized with the rotation of the probe as determined by our sun sensor system.  The total integration time would be one minute or an integer number of rotations of the probe.  The data collected by the instrument was divided into 8 bins in azimuth for the DLVS and 4 bins in azimuth for the ULVS.  Unfortunately, as the probe swung under the parachute a several times the maximum rate specified for the mission, the sun sensor system was not able to determine the rotation rate accurately at all times during the mission.  The result is that the boundaries of the azimuth bins actual drifted in azimuth during the integration.  Fortunately, the housekeeping data recorded by the IR spectrometer permits the exact times when the shutter was open and the instrument collected a sample to be reconstructed.  This means that for each cycle of operation of the IR spectrometer we have a table of the times at which data were collected in each of the azimuth bins.  Combined with a model of azimuth as a function of mission time determined by other means, we have produced a table of azimuths at which data were collected in each cycle of the IR spectrometer.  The relative weights of each of these samples are also known, as the integration times sometimes varied during single cycles.  This auxiliary information is necessary of one wants to compare the measurements with the signal levels that the machine would have measured for a radiative transfer model.

One final complication is necessary for the reduction of the ULIS measurements to the spectrum of the average intensity in the field of view.  For the ULIS it has been noted in calibration measurements that a weak pattern of fringes (of a few percent amplitude) is superimposed on the intensity spectrum.  The origin of these fringes is not clear to the DISR team, and in fact two of the three DISR instruments produced for the program did not show these fringes.  The flight model, however suffers from these weak fringes.

We planned to take measurements from calibration lamps at four altitudes during the descent into Titan’s atmosphere.  These altitudes were selected to give a good distribution in temperature of the DISR optics and IR detector.  The measurements from these incandescent lamps show the fringe patterns with excellent signal to noise ratio, and interpolation in these fringe patterns with temperature can be used to remove the fringes from the ULIS spectra.  Unfortunately, the electronics in the radar altimeter used by the probe to measure altitude had an error that prevented it from measuring the altitude selected for the fourth calibration data set.  Because of this, we only have the fringe pattern at three of the four expected temperatures.  Nevertheless, we provide below the raw fringe patterns for the ULIS at these three temperatures.  At still colder temperatures, we simply extrapolated in the patterns measured.

- Wavelength calibration

The wavelength scale of the ULIS and DLIS portions of the instrument are given as functions of pixel number and the read-out value of the temperature of the optics.  The expression for the DLIS is

DLIS wavelength (nm) = a +b*N + c*N2

Where N is the number of the detector element and the count begins with 0 and ends with 149.  The values of a, b, and c depend on the temperature of the optics, with

a = 784.04 – 0.023725*T + 4.569x10-5*T2,

b = 7.1568 + 0.0014118*T – 2.8753x10-6*T2,

c = -0.0030065 -1.4648x10-5*T +3.0378x10-8*T2

where T is the optics temperature.

The wavelength scale for the ULIS portion of the IR spectrometer is given by

ULIS wavelength (nm) = d + e*N +f*N2

Where N is the detector element and the count begins with 0 and ends with 149.  The values of  d, e, and f depend on the temperature of the optics, with

d = 766.39 + 0.025199*T,

e = 7.5138 – 8.1082x10-4 * T,

f = -0.0059698 + 5.991x10-6*T.

Again, T is the temperature of the optics.

- Spectral Resolution

The spectral line shape produced by the instrument is a gaussian function with a full width at half maximum response (FWHM) that depends on the wavelength and the temperature of the optics.  For the DLIS, the result is

DLIS FWHM(nm) = A(T) + B(T)* + C(T)*2 + D(T)*3,		(20.)

where			A(T) =529.84 –5.9278*T + 0.015406*T2,			(21.)

B(T) = -1.2194 +0.015864*T –4.1121x10-5*T2,		(22.)

C(T) = 1.1263x10-3 –1.3879x10-6*T + 3.576x10-8*T2,	(23.)

D(T) = -3.1689x10-7 + 3.9644x10-9*T –1.0152x10-11*T2.	(24.)

For the ULIS, the result is

ULIS FWHM(nm) = a(T) + B(T)* +C(T)*2 +D(T)*3		(25.)

where 			A(T) = -224.34 + 0.90546*T					(26.)

B(T) = 0.68932 – 0.0024122*T				(27.)

C(T)= -6.1296x10-4 + 2.0406x10-6*T				(28.)

D(T) = 1.751x10-7 – 5.5931x10-10*T.				(29.)

- The Absolute responsivity of the IR spectrometers

The DISR sensor head was inserted into a 20-inch diameter integrating sphere.  Light from a tungsten lamp was admitted into the integrating sphere through a small hole.  The incident beam bounced off a small baffle, and then illuminated the entire inside of the sphere.  A calibrated laboratory monochrometer was used to observe a spot near the front of the integrating sphere and measured the absolute brightness as a function of wavelength of the wall of the integrating sphere.  This test was repeated at many temperatures over the operating range of the instrument, and the responsivity of the instrument was measured as a function of detector temperature for the ULIS and the DLIS.

The spectral range of the IR spectrometer was from 850 to 1700 nm.  However, the responsivity of the IR detectors was appreciable shortward of 850 nm.  A filter was used in the optical path of DISR to suppress response to second-order light.  However, the cutoff of this filter between 800 and 850 nm was not sufficiently sharp to prevent some response to light between these wavelengths in second-order at the locations of pixels intended to measure light between 1600 and 1700 nm in first order.  In addition, the response of the long wavelength edge of the response of the detector moved from 1700 nm to 1600 nm as the detector cooled from room temperature to its expected operating temperature of 200 K or less.  Thus, the effective range of the IR spectrometer as used in Titan’s atmosphere is limited to the wavelength interval from 850 to 1600 nm.  In the tables below, we give the response of the instrument to light from 850 to 1700 nm in first order, and the response to second-order light from 800 to 850 nm also.

  - Response of the DLIS

The absolute responsivity of the DLIS as a function of pixel number and detector temperature is given below:

| Pixel | Wavelength | 189 | 223 | 233 | 242 | 252 | 255 | 271 | 286 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 7 | 832.18 | 9693.9 | 9075.7 | 8669.7 | 8428.1 | 8215.8 | 8116.7 | 7830.7 | 7431.8 |
| 8 | 839.43 | 13131.2 | 12503.5 | 12112.5 | 11887.5 | 11689.6 | 11581.9 | 11282.7 | 10913.6 |
| 9 | 846.68 | 15262.1 | 14741.1 | 14438.2 | 14251.6 | 14089.4 | 13976.4 | 13693.8 | 13354.2 |
| 10 | 853.91 | 16809.8 | 16439.4 | 16205.8 | 16070.2 | 15947.6 | 15832.8 | 15553.3 | 15219.0 |
| 11 | 861.13 | 18755.0 | 18464.8 | 18296.9 | 18200.9 | 18107.5 | 18020.6 | 17766.5 | 17468.2 |
| 12 | 868.35 | 19882.5 | 19635.4 | 19504.4 | 19448.5 | 19394.5 | 19344.4 | 19154.9 | 18938.4 |
| 13 | 875.55 | 21820.8 | 21452.6 | 21332.2 | 21275.4 | 21223.8 | 21174.7 | 20969.1 | 20705.5 |
| 14 | 882.75 | 23545.8 | 22747.9 | 22572.5 | 22477.3 | 22422.6 | 22367.9 | 22197.7 | 22020.0 |
| 15 | 889.93 | 27194.9 | 25195.1 | 24850.8 | 24631.2 | 24516.6 | 24407.0 | 24196.4 | 24011.1 |
| 16 | 897.11 | 31833.9 | 28003.9 | 27239.3 | 26688.2 | 26407.7 | 26107.5 | 25726.9 | 25428.1 |
| 17 | 904.28 | 38255.4 | 32712.1 | 31400.7 | 30283.1 | 29673.8 | 28978.4 | 28208.2 | 27662.5 |
| 18 | 911.44 | 44164.8 | 38198.3 | 36490.5 | 34872.9 | 33901.7 | 32667.0 | 31236.2 | 30156.7 |
| 19 | 918.58 | 50306.3 | 45147.8 | 43428.5 | 41636.5 | 40456.9 | 38852.4 | 36754.0 | 35016.5 |
| 20 | 925.72 | 54335.2 | 50620.7 | 49253.1 | 47718.6 | 46654.0 | 45107.0 | 42825.7 | 40717.1 |
| 21 | 932.85 | 58350.0 | 56029.0 | 55125.1 | 54021.3 | 53229.6 | 52046.3 | 50063.1 | 48069.0 |
| 22 | 939.97 | 59896.9 | 58827.7 | 58354.1 | 57691.3 | 57221.9 | 56516.1 | 55164.0 | 53737.7 |
| 23 | 947.08 | 61478.8 | 61224.9 | 61090.1 | 60783.7 | 60571.8 | 60286.5 | 59493.6 | 58621.6 |
| 24 | 954.19 | 61179.6 | 61252.1 | 61256.4 | 61146.2 | 61098.6 | 61070.1 | 60706.0 | 60264.6 |
| 25 | 961.28 | 62160.2 | 62193.1 | 62243.5 | 62175.3 | 62178.8 | 62260.8 | 62060.4 | 61798.4 |
| 26 | 968.36 | 62169.7 | 61947.7 | 61910.4 | 61788.7 | 61764.1 | 61804.9 | 61602.7 | 61339.6 |
| 27 | 975.43 | 64136.6 | 63601.5 | 63481.3 | 63276.8 | 63186.1 | 63156.9 | 62839.0 | 62469.7 |
| 28 | 982.50 | 65536.1 | 64724.4 | 64504.0 | 64205.5 | 64063.0 | 63936.9 | 63563.8 | 63138.9 |
| 29 | 989.55 | 68686.1 | 67724.9 | 67471.9 | 67121.8 | 66932.8 | 66755.7 | 66293.5 | 65777.4 |
| 30 | 996.60 | 70812.8 | 69814.5 | 69517.8 | 69122.7 | 68919.2 | 68668.5 | 68114.1 | 67464.9 |
| 31 | 1003.63 | 74121.5 | 73219.7 | 72969.3 | 72596.4 | 72398.8 | 72189.8 | 71650.5 | 71019.9 |
| 32 | 1010.66 | 75279.0 | 74619.8 | 74415.8 | 74087.8 | 73918.0 | 73702.5 | 73125.7 | 72390.3 |
| 33 | 1017.68 | 77230.1 | 76874.2 | 76771.3 | 76533.4 | 76430.6 | 76348.5 | 75978.4 | 75471.9 |
| 34 | 1024.68 | 76605.8 | 76553.2 | 76519.7 | 76357.8 | 76318.4 | 76313.4 | 76094.5 | 75764.8 |
| 35 | 1031.68 | 77034.8 | 77179.7 | 77232.2 | 77146.1 | 77148.8 | 77242.3 | 77168.3 | 77029.3 |
| 36 | 1038.67 | 75427.1 | 75674.8 | 75731.8 | 75662.1 | 75695.0 | 75801.4 | 75765.0 | 75609.5 |
| 37 | 1045.65 | 75600.9 | 75813.5 | 75888.2 | 75826.9 | 75849.1 | 75982.9 | 75975.1 | 75851.1 |
| 38 | 1052.62 | 74295.2 | 74380.9 | 74391.6 | 74288.0 | 74296.1 | 74368.8 | 74323.7 | 74138.2 |
| 39 | 1059.58 | 74961.4 | 74860.1 | 74848.8 | 74700.0 | 74682.9 | 74761.0 | 74753.4 | 74673.5 |
| 40 | 1066.53 | 74430.1 | 74093.2 | 73984.2 | 73765.9 | 73706.9 | 73702.8 | 73633.1 | 73496.9 |
| 41 | 1073.47 | 76218.9 | 75665.3 | 75512.2 | 75230.3 | 75122.5 | 75073.2 | 74901.1 | 74646.9 |
| 42 | 1080.40 | 76651.2 | 75911.5 | 75671.7 | 75326.6 | 75176.7 | 75045.4 | 74835.5 | 74526.3 |
| 43 | 1087.32 | 79575.8 | 78653.7 | 78380.6 | 77967.1 | 77779.6 | 77610.7 | 77340.7 | 77004.9 |
| 44 | 1094.24 | 80499.4 | 79516.0 | 79178.4 | 78730.6 | 78516.3 | 78305.7 | 78004.8 | 77619.3 |
| 45 | 1101.14 | 83445.9 | 82441.7 | 82121.7 | 81662.0 | 81437.1 | 81220.8 | 80887.4 | 80479.6 |
| 46 | 1108.04 | 84543.8 | 83618.4 | 83285.2 | 82820.4 | 82597.3 | 82372.7 | 82058.3 | 81640.9 |
| 47 | 1114.92 | 86798.3 | 86005.2 | 85720.0 | 85286.0 | 85081.7 | 84894.6 | 84591.1 | 84198.0 |
| 48 | 1121.80 | 86283.5 | 85689.3 | 85434.1 | 85041.1 | 84869.1 | 84711.2 | 84473.8 | 84103.5 |
| 49 | 1128.66 | 87291.3 | 86913.7 | 86733.0 | 86403.7 | 86266.6 | 86182.7 | 86002.3 | 85683.8 |
| 50 | 1135.52 | 85228.1 | 85099.4 | 84985.3 | 84728.7 | 84646.3 | 84627.3 | 84551.7 | 84288.8 |
| 51 | 1142.37 | 86425.2 | 86402.0 | 86328.2 | 86094.8 | 86027.5 | 86057.4 | 86037.1 | 85864.5 |
| 52 | 1149.20 | 84204.3 | 84274.1 | 84207.2 | 84003.8 | 83966.3 | 84020.6 | 84063.1 | 83958.3 |
| 53 | 1156.03 | 84150.1 | 84242.9 | 84206.4 | 84016.8 | 83989.7 | 84076.9 | 84146.8 | 84063.0 |
| 54 | 1162.85 | 81799.9 | 81875.1 | 81821.3 | 81632.9 | 81608.6 | 81685.7 | 81782.0 | 81720.9 |
| 55 | 1169.66 | 81656.8 | 81674.1 | 81625.1 | 81420.5 | 81383.9 | 81465.9 | 81546.2 | 81474.6 |
| 56 | 1176.46 | 79168.8 | 79082.6 | 78988.2 | 78767.8 | 78724.3 | 78759.7 | 78825.7 | 78754.1 |
| 57 | 1183.25 | 79589.1 | 79385.5 | 79288.1 | 79030.7 | 78950.2 | 78967.9 | 79004.4 | 78894.6 |
| 58 | 1190.03 | 78168.2 | 77833.4 | 77659.9 | 77345.3 | 77207.8 | 77066.7 | 76708.6 | 75867.6 |
| 59 | 1196.81 | 78473.8 | 78012.2 | 77843.4 | 77512.7 | 77380.6 | 77303.3 | 77169.0 | 76810.5 |
| 60 | 1203.57 | 77703.8 | 77120.7 | 76889.1 | 76537.0 | 76403.0 | 76284.6 | 76136.6 | 75810.9 |
| 61 | 1210.32 | 79640.5 | 78928.2 | 78669.1 | 78275.4 | 78102.1 | 77974.3 | 77844.3 | 77596.9 |
| 62 | 1217.07 | 78966.3 | 78176.8 | 77885.7 | 77459.7 | 77287.3 | 77124.4 | 76961.5 | 76654.8 |
| 63 | 1223.80 | 79531.8 | 78693.5 | 78402.5 | 77971.2 | 77786.6 | 77611.6 | 77402.2 | 77052.0 |
| 64 | 1230.52 | 80134.1 | 79236.3 | 78901.4 | 78442.5 | 78247.5 | 78042.7 | 77837.9 | 77504.1 |
| 65 | 1237.24 | 81708.5 | 80813.8 | 80491.1 | 80025.2 | 79806.5 | 79610.7 | 79379.7 | 79027.2 |
| 66 | 1243.95 | 81157.2 | 80284.2 | 79949.0 | 79480.9 | 79273.2 | 79062.1 | 78842.2 | 78499.8 |
| 67 | 1250.64 | 82512.2 | 81707.7 | 81405.1 | 80941.4 | 80734.1 | 80550.3 | 80328.5 | 79972.1 |
| 68 | 1257.33 | 81802.0 | 81043.4 | 80736.1 | 80298.6 | 80109.8 | 79922.3 | 79723.8 | 79389.8 |
| 69 | 1264.01 | 82889.8 | 82191.3 | 81899.7 | 81424.3 | 81168.1 | 80844.8 | 80135.3 | 78836.7 |
| 70 | 1270.68 | 81897.3 | 81314.5 | 81046.9 | 80640.4 | 80487.0 | 80337.2 | 80206.2 | 79944.1 |
| 71 | 1277.34 | 82496.7 | 81996.6 | 81766.1 | 81390.9 | 81231.6 | 81122.3 | 80999.5 | 80740.0 |
| 72 | 1283.99 | 80306.5 | 79932.3 | 79722.3 | 79381.7 | 79267.2 | 79173.1 | 79098.5 | 78853.3 |
| 73 | 1290.63 | 80981.1 | 80646.2 | 80462.8 | 80133.5 | 80011.2 | 79932.8 | 79864.7 | 79648.2 |
| 74 | 1297.26 | 79254.8 | 78986.3 | 78813.9 | 78502.0 | 78406.1 | 78340.8 | 78305.3 | 78107.6 |
| 75 | 1303.88 | 79427.4 | 79181.7 | 79030.5 | 78728.5 | 78626.2 | 78584.2 | 78548.3 | 78369.7 |
| 76 | 1310.49 | 77634.8 | 77418.1 | 77253.4 | 76961.4 | 76881.4 | 76826.1 | 76811.9 | 76640.5 |
| 77 | 1317.09 | 77850.0 | 77618.9 | 77468.1 | 77178.2 | 77086.5 | 77047.6 | 77038.4 | 76864.1 |
| 78 | 1323.69 | 75994.2 | 75766.8 | 75591.2 | 75303.5 | 75227.1 | 75171.9 | 75155.1 | 74977.7 |
| 79 | 1330.27 | 76122.9 | 75850.3 | 75664.8 | 75376.2 | 75279.5 | 75233.5 | 75211.3 | 75035.8 |
| 80 | 1336.84 | 74389.2 | 74113.7 | 73906.4 | 73616.5 | 73530.3 | 73461.4 | 73430.8 | 73227.3 |
| 81 | 1343.41 | 74122.4 | 73776.5 | 73545.8 | 73245.1 | 73129.2 | 73061.4 | 72999.5 | 72769.0 |
| 82 | 1349.96 | 71332.9 | 70944.8 | 70668.3 | 70366.9 | 70259.8 | 70154.8 | 70089.4 | 69856.0 |
| 83 | 1356.51 | 69986.0 | 69497.0 | 69195.5 | 68872.5 | 68735.9 | 68627.0 | 68532.7 | 68282.8 |
| 84 | 1363.05 | 68237.7 | 67702.3 | 67356.7 | 67053.8 | 66920.2 | 66789.1 | 66667.5 | 66393.7 |
| 85 | 1369.57 | 69008.0 | 68468.5 | 68123.3 | 67815.7 | 67674.1 | 67547.7 | 67398.6 | 67104.5 |
| 86 | 1376.09 | 67933.9 | 67380.9 | 66941.2 | 66635.5 | 66475.6 | 66319.4 | 66111.8 | 65750.9 |
| 87 | 1382.60 | 67797.0 | 67168.1 | 66660.7 | 66327.5 | 66125.2 | 65952.7 | 65694.1 | 65251.5 |
| 88 | 1389.10 | 65902.1 | 65233.4 | 64706.4 | 64399.9 | 64217.4 | 64052.0 | 63786.3 | 63347.5 |
| 89 | 1395.59 | 65703.1 | 64985.1 | 64445.6 | 64150.4 | 63943.0 | 63817.5 | 63535.0 | 63087.3 |
| 90 | 1402.07 | 64468.5 | 63714.6 | 63127.9 | 62827.7 | 62626.1 | 62509.1 | 62182.6 | 61670.4 |
| 91 | 1408.54 | 65945.0 | 65093.4 | 64461.2 | 64138.5 | 63906.8 | 63763.9 | 63426.4 | 62895.1 |
| 92 | 1415.00 | 65500.0 | 64676.3 | 64018.6 | 63677.3 | 63442.6 | 63297.1 | 62943.6 | 62386.9 |
| 93 | 1421.45 | 67245.3 | 66404.2 | 65755.9 | 65405.1 | 65168.2 | 65006.8 | 64635.8 | 64058.9 |
| 94 | 1427.90 | 67168.2 | 66416.4 | 65784.8 | 65460.4 | 65248.2 | 65059.3 | 64682.5 | 64116.9 |
| 95 | 1434.33 | 68745.8 | 68054.8 | 67493.6 | 67175.8 | 66948.9 | 66755.2 | 66391.5 | 65851.5 |
| 96 | 1440.75 | 68712.6 | 68114.8 | 67575.4 | 67251.5 | 67052.8 | 66840.8 | 66479.1 | 65932.4 |
| 97 | 1447.17 | 70167.9 | 69626.1 | 69124.9 | 68814.1 | 68611.9 | 68357.4 | 67989.4 | 67423.7 |
| 98 | 1453.57 | 69617.8 | 69163.4 | 68685.8 | 68376.1 | 68180.0 | 67901.7 | 67556.0 | 67017.5 |
| 99 | 1459.97 | 70501.0 | 70111.0 | 69667.8 | 69354.8 | 69151.7 | 68870.6 | 68551.2 | 68047.0 |
| 100 | 1466.35 | 69287.2 | 68997.6 | 68549.3 | 68265.6 | 68090.8 | 67798.1 | 67470.9 | 66975.7 |
| 101 | 1472.73 | 69039.9 | 68817.4 | 68449.3 | 68180.4 | 68029.9 | 67731.4 | 67402.5 | 66829.1 |
| 102 | 1479.10 | 67949.9 | 67835.3 | 67472.4 | 67223.9 | 67073.3 | 66795.1 | 66501.3 | 66023.2 |
| 103 | 1485.46 | 67234.8 | 67225.6 | 66933.9 | 66706.2 | 66565.8 | 66293.1 | 66019.5 | 65486.7 |
| 104 | 1491.80 | 64598.6 | 64807.9 | 64546.4 | 64379.2 | 64288.2 | 64055.9 | 63783.0 | 63207.3 |
| 105 | 1498.14 | 66054.1 | 66138.0 | 65868.4 | 65696.8 | 65586.8 | 65350.6 | 65142.7 | 64751.1 |
| 106 | 1504.47 | 64359.2 | 64513.4 | 64259.0 | 64098.0 | 63998.0 | 63777.3 | 63570.3 | 63209.9 |
| 107 | 1510.79 | 63992.6 | 64181.9 | 63966.2 | 63795.6 | 63691.3 | 63481.5 | 63310.4 | 62943.9 |
| 108 | 1517.10 | 60599.5 | 61036.4 | 60849.1 | 60747.4 | 60710.9 | 60544.8 | 60393.4 | 59929.4 |
| 109 | 1523.40 | 61390.7 | 61899.8 | 61727.5 | 61635.5 | 61569.9 | 61413.7 | 61281.8 | 60925.2 |
| 110 | 1529.70 | 60062.1 | 60452.0 | 60325.2 | 60213.9 | 60152.9 | 59986.9 | 59855.1 | 59537.7 |
| 111 | 1535.98 | 59422.2 | 59883.2 | 59748.9 | 59657.5 | 59658.8 | 59522.0 | 59424.6 | 59127.5 |
| 112 | 1542.25 | 57542.0 | 58067.9 | 57915.4 | 57833.3 | 57811.1 | 57706.5 | 57677.6 | 57403.1 |
| 113 | 1548.52 | 56973.5 | 57588.9 | 57479.5 | 57442.9 | 57420.1 | 57321.2 | 57259.2 | 56983.1 |
| 114 | 1554.77 | 54956.7 | 55642.5 | 55550.3 | 55543.0 | 55550.7 | 55470.1 | 55423.4 | 55192.6 |
| 115 | 1561.02 | 54234.5 | 54941.5 | 54887.7 | 54890.4 | 54899.7 | 54847.0 | 54859.4 | 54649.9 |
| 116 | 1567.25 | 52221.7 | 53050.5 | 52965.6 | 52973.6 | 53007.7 | 52974.5 | 53015.7 | 52853.6 |
| 117 | 1573.48 | 51662.3 | 52576.6 | 52500.2 | 52508.3 | 52532.6 | 52504.2 | 52556.5 | 52381.3 |
| 118 | 1579.69 | 49756.9 | 50849.2 | 50779.1 | 50827.8 | 50851.9 | 50823.1 | 50850.4 | 50694.2 |
| 119 | 1585.90 | 48524.7 | 50010.2 | 49987.7 | 50023.3 | 50038.1 | 49979.1 | 49968.7 | 49602.4 |
| 120 | 1592.10 | 45986.3 | 48319.8 | 48304.4 | 48393.5 | 48441.6 | 48445.4 | 48539.1 | 48410.0 |
| 121 | 1598.29 | 43120.2 | 47641.4 | 47684.0 | 47818.3 | 47895.5 | 47915.2 | 48079.8 | 47952.7 |
| 122 | 1604.47 | 36586.6 | 45857.9 | 45982.6 | 46170.1 | 46289.5 | 46355.4 | 46540.2 | 46455.2 |
| 123 | 1610.64 | 27910.0 | 44707.6 | 45073.8 | 45390.2 | 45552.7 | 45671.0 | 45912.4 | 45834.0 |
| 124 | 1616.80 | 18380.1 | 42304.5 | 43071.4 | 43609.6 | 43839.4 | 44017.5 | 44365.0 | 44402.3 |
| 125 | 1622.95 | 12252.2 | 39789.4 | 41748.8 | 42699.9 | 43066.3 | 43347.3 | 43813.5 | 43897.2 |
| 126 | 1629.09 | 7803.2 | 34017.5 | 38373.7 | 40285.7 | 40916.2 | 41437.5 | 42024.8 | 42216.4 |
| 127 | 1635.22 | 5887.0 | 27473.4 | 34698.5 | 38624.8 | 39791.5 | 40696.9 | 41530.3 | 41798.4 |
| 128 | 1641.34 | 4279.4 | 19843.7 | 28173.4 | 34756.1 | 37069.6 | 38707.6 | 39863.8 | 40317.1 |
| 129 | 1647.46 | 3736.1 | 15081.2 | 22107.3 | 29857.3 | 33616.4 | 36618.1 | 38399.6 | 38922.6 |
| 130 | 1653.56 | 2733.0 | 10117.0 | 15289.9 | 22933.7 | 28174.8 | 33633.1 | 36931.7 | 38090.0 |
| 131 | 1659.65 | 2446.3 | 7911.2 | 11629.4 | 17678.5 | 22792.3 | 29927.3 | 35532.4 | 37478.1 |
| 132 | 1665.74 | 1743.5 | 5367.4 | 7821.3 | 12144.2 | 16286.7 | 23609.7 | 31722.0 | 35372.3 |
| 133 | 1671.81 | 1954.9 | 4721.6 | 6551.3 | 9638.8 | 12621.9 | 18392.0 | 27110.8 | 32620.6 |
| 134 | 1677.88 | 1214.7 | 3156.6 | 4383.1 | 6564.1 | 8700.8 | 13042.2 | 21317.0 | 28971.8 |
| 135 | 1683.94 | 1287.7 | 2612.7 | 3475.6 | 5025.2 | 6561.4 | 9758.5 | 16482.0 | 24430.8 |
| 136 | 1689.99 | 940.1 | 1853.9 | 2426.4 | 3520.1 | 4621.9 | 6850.2 | 11705.8 | 18404.0 |
| 137 | 1696.02 | 1117.0 | 1810.3 | 2239.8 | 3024.0 | 3843.2 | 5505.7 | 9162.2 | 14283.1 |
| 138 | 1702.05 | 802.7 | 1328.4 | 1595.3 | 2165.3 | 2728.2 | 3824.5 | 6316.8 | 10092.7 |
| 139 | 1708.07 | 1270.8 | 1715.9 | 1925.7 | 2353.1 | 2825.4 | 3695.8 | 5676.6 | 8518.7 |
| 140 | 1714.08 | 693.8 | 1050.7 | 1142.2 | 1457.4 | 1795.0 | 2399.1 | 3759.5 | 5713.6 |
| 141 | 1720.08 | 901.6 | 1176.4 | 1284.2 | 1494.8 | 1727.5 | 2188.9 | 3241.6 | 4671.1 |
| 142 | 1726.07 | 619.4 | 836.3 | 866.4 | 1026.7 | 1215.7 | 1540.0 | 2262.7 | 3261.0 |
|  | Second-Order Response |  |  |  |  |  |  |  |  |
| 105 | 749.07 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 106 | 752.24 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 107 | 755.40 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 108 | 758.55 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 109 | 761.70 | 15.2 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 110 | 764.85 | 18.2 | 15.1 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 111 | 767.99 | 22.4 | 18.0 | 16.8 | 15.5 | 0.0 | 0.0 | 0.0 | 0.0 |
| 112 | 771.13 | 29.7 | 22.0 | 19.8 | 18.5 | 17.3 | 16.6 | 0.0 | 0.0 |
| 113 | 774.26 | 46.6 | 29.2 | 25.9 | 22.9 | 20.6 | 19.6 | 17.2 | 14.9 |
| 114 | 777.39 | 63.8 | 45.6 | 38.8 | 31.1 | 27.3 | 25.3 | 20.3 | 17.7 |
| 115 | 780.51 | 80.9 | 62.8 | 55.9 | 48.2 | 41.6 | 37.5 | 26.8 | 21.4 |
| 116 | 783.63 | 98.1 | 79.9 | 73.1 | 65.4 | 58.8 | 54.7 | 40.7 | 28.3 |
| 117 | 786.74 | 129.8 | 97.1 | 90.2 | 82.5 | 75.9 | 71.9 | 57.8 | 43.7 |
| 118 | 789.85 | 169.0 | 127.9 | 114.4 | 99.7 | 93.1 | 89.0 | 75.0 | 60.9 |
| 119 | 792.95 | 220.5 | 166.2 | 148.0 | 133.0 | 120.0 | 112.1 | 92.2 | 78.1 |
| 120 | 796.05 | 285.6 | 217.1 | 194.8 | 173.4 | 155.1 | 145.6 | 118.2 | 95.2 |
| 121 | 799.14 | 370.1 | 281.5 | 252.9 | 226.0 | 203.3 | 191.4 | 152.5 | 124.2 |
| 122 | 802.23 | 475.4 | 364.5 | 328.8 | 292.3 | 264.8 | 248.2 | 200.2 | 161.0 |
| 123 | 805.32 | 601.3 | 468.7 | 424.5 | 379.0 | 342.9 | 322.8 | 260.9 | 210.7 |
| 124 | 808.40 | 749.0 | 593.7 | 542.4 | 486.3 | 442.2 | 417.0 | 338.4 | 273.7 |
| 125 | 811.47 | 913.4 | 739.9 | 678.6 | 614.6 | 563.7 | 533.3 | 436.5 | 354.0 |
| 126 | 814.54 | 1088.9 | 903.6 | 836.7 | 763.8 | 702.7 | 668.5 | 556.8 | 456.0 |
| 127 | 817.61 | 1275.7 | 1078.4 | 1005.5 | 929.2 | 864.5 | 824.9 | 694.7 | 579.6 |
| 128 | 820.67 | 1461.3 | 1265.0 | 1190.6 | 1106.1 | 1035.8 | 993.1 | 855.5 | 722.4 |
| 129 | 823.73 | 1647.2 | 1450.2 | 1374.6 | 1293.0 | 1221.6 | 1177.3 | 1026.0 | 885.2 |
| 130 | 826.78 | 1819.1 | 1637.1 | 1564.2 | 1479.3 | 1405.2 | 1361.7 | 1211.7 | 1058.4 |
| 131 | 829.83 | 1986.9 | 1809.0 | 1740.7 | 1663.5 | 1595.6 | 1550.9 | 1395.1 | 1244.7 |
| 132 | 832.87 | 2137.1 | 1977.6 | 1913.0 | 1835.5 | 1768.9 | 1728.7 | 1585.5 | 1429.1 |
| 133 | 835.91 | 2276.2 | 2128.5 | 2068.7 | 2001.8 | 1939.8 | 1901.5 | 1759.8 | 1617.8 |
| 134 | 838.94 | 2401.1 | 2268.6 | 2215.8 | 2151.1 | 2093.6 | 2058.1 | 1931.1 | 1790.0 |
| 135 | 841.97 | 2516.4 | 2394.0 | 2344.5 | 2288.4 | 2237.7 | 2206.5 | 2085.5 | 1959.9 |
| 136 | 844.99 | 2607.5 | 2511.1 | 2466.3 | 2412.3 | 2365.0 | 2335.7 | 2230.7 | 2112.1 |
| 137 | 848.01 | 2692.3 | 2602.7 | 2566.1 | 2525.0 | 2486.4 | 2457.8 | 2358.4 | 2254.1 |
| 138 | 851.03 | 2761.0 | 2687.4 | 2653.4 | 2615.4 | 2581.4 | 2559.6 | 2479.9 | 2380.4 |
| 139 | 854.04 | 2820.6 | 2757.1 | 2730.2 | 2700.1 | 2667.5 | 2647.4 | 2576.4 | 2501.0 |
| 140 | 857.04 | 2869.9 | 2817.8 | 2797.3 | 2767.3 | 2741.4 | 2725.5 | 2663.0 | 2592.8 |
| 141 | 860.04 | 2910.3 | 2867.1 | 2847.3 | 2825.2 | 2806.2 | 2792.6 | 2737.8 | 2678.1 |
| 142 | 863.04 | 2936.9 | 2908.8 | 2896.6 | 2874.5 | 2855.5 | 2843.8 | 2803.6 | 2749.7 |

The response of the ULIS instrument is given below:

| Pixel | Wavelength | 189 | 223 | 233 | 243 | 252 | 255 | 271 | 286 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 7 | 822.54 | 507.99 | 487.36 | 470.53 | 450.65 | 442.27 | 436.64 | 428.16 | 401.97 |
| 8 | 829.82 | 814.19 | 790.21 | 768.52 | 749.67 | 736.40 | 732.23 | 718.44 | 692.60 |
| 9 | 837.10 | 1006.48 | 984.25 | 975.38 | 952.17 | 942.83 | 936.17 | 920.20 | 873.96 |
| 10 | 844.36 | 1128.62 | 1118.40 | 1112.11 | 1100.29 | 1091.69 | 1090.05 | 1066.53 | 1012.60 |
| 11 | 851.61 | 1276.78 | 1275.12 | 1284.38 | 1269.63 | 1266.83 | 1266.66 | 1254.58 | 1228.89 |
| 12 | 858.86 | 1390.79 | 1401.65 | 1413.20 | 1411.41 | 1409.72 | 1417.17 | 1404.31 | 1383.42 |
| 13 | 866.09 | 1535.60 | 1546.51 | 1570.68 | 1566.54 | 1573.11 | 1581.27 | 1575.52 | 1555.25 |
| 14 | 873.32 | 1677.42 | 1673.94 | 1690.80 | 1693.37 | 1695.77 | 1708.80 | 1699.16 | 1681.22 |
| 15 | 880.53 | 1889.02 | 1827.13 | 1844.84 | 1833.33 | 1837.13 | 1844.11 | 1835.01 | 1814.70 |
| 16 | 887.74 | 2193.54 | 2012.82 | 2003.23 | 1984.10 | 1976.42 | 1977.54 | 1957.49 | 1934.67 |
| 17 | 894.93 | 2698.70 | 2364.15 | 2326.17 | 2266.31 | 2243.04 | 2219.22 | 2175.82 | 2121.64 |
| 18 | 902.12 | 3217.29 | 2776.42 | 2689.49 | 2594.47 | 2537.65 | 2477.33 | 2384.52 | 2312.75 |
| 19 | 909.29 | 3713.19 | 3276.45 | 3175.10 | 3044.73 | 2964.01 | 2863.61 | 2706.09 | 2561.80 |
| 20 | 916.46 | 4082.47 | 3741.58 | 3644.83 | 3523.59 | 3437.10 | 3318.97 | 3127.39 | 2957.97 |
| 21 | 923.61 | 4439.93 | 4214.27 | 4165.22 | 4068.25 | 3999.49 | 3898.68 | 3686.71 | 3456.20 |
| 22 | 930.76 | 4815.29 | 4676.23 | 4648.28 | 4586.74 | 4535.39 | 4452.10 | 4236.82 | 3971.30 |
| 23 | 937.90 | 5121.96 | 5049.74 | 5052.20 | 5006.03 | 4973.03 | 4923.67 | 4774.65 | 4602.82 |
| 24 | 945.02 | 5183.63 | 5166.09 | 5171.49 | 5150.58 | 5131.00 | 5105.65 | 4994.64 | 4875.86 |
| 25 | 952.14 | 5213.11 | 5240.88 | 5276.46 | 5267.26 | 5262.22 | 5251.65 | 5121.52 | 4918.98 |
| 26 | 959.25 | 5235.43 | 5282.19 | 5318.15 | 5330.04 | 5335.12 | 5344.86 | 5281.36 | 5197.34 |
| 27 | 966.35 | 5332.70 | 5377.08 | 5430.01 | 5437.61 | 5447.06 | 5466.69 | 5426.32 | 5364.73 |
| 28 | 973.43 | 5422.90 | 5437.21 | 5466.33 | 5469.73 | 5470.00 | 5479.34 | 5425.64 | 5367.54 |
| 29 | 980.51 | 5506.42 | 5485.70 | 5511.28 | 5495.83 | 5488.97 | 5484.54 | 5394.33 | 5242.62 |
| 30 | 987.58 | 5638.80 | 5592.24 | 5600.26 | 5585.50 | 5573.90 | 5566.05 | 5493.67 | 5412.15 |
| 31 | 994.64 | 5899.59 | 5850.88 | 5878.21 | 5859.44 | 5851.20 | 5844.14 | 5778.82 | 5693.61 |
| 32 | 1001.69 | 6174.41 | 6129.13 | 6150.33 | 6140.84 | 6132.59 | 6122.91 | 6047.98 | 5976.67 |
| 33 | 1008.73 | 6529.87 | 6481.63 | 6513.96 | 6495.63 | 6486.48 | 6476.73 | 6399.33 | 6301.69 |
| 34 | 1015.76 | 6680.93 | 6632.78 | 6643.50 | 6625.40 | 6610.89 | 6597.97 | 6513.66 | 6426.44 |
| 35 | 1022.77 | 6734.03 | 6702.54 | 6724.38 | 6701.61 | 6690.00 | 6681.37 | 6613.01 | 6516.01 |
| 36 | 1029.78 | 6631.04 | 6634.64 | 6657.48 | 6654.14 | 6650.41 | 6651.51 | 6589.62 | 6524.29 |
| 37 | 1036.78 | 6633.66 | 6675.47 | 6732.78 | 6735.68 | 6744.53 | 6757.36 | 6716.07 | 6643.75 |
| 38 | 1043.77 | 6646.42 | 6705.96 | 6759.35 | 6776.47 | 6786.35 | 6801.61 | 6756.70 | 6703.71 |
| 39 | 1050.75 | 6711.29 | 6764.41 | 6824.77 | 6828.13 | 6832.47 | 6834.93 | 6740.43 | 6567.48 |
| 40 | 1057.73 | 6724.49 | 6748.13 | 6777.29 | 6775.07 | 6772.40 | 6778.13 | 6725.60 | 6663.35 |
| 41 | 1064.69 | 6685.88 | 6679.86 | 6707.79 | 6689.06 | 6680.65 | 6681.23 | 6627.71 | 6529.56 |
| 42 | 1071.64 | 6602.61 | 6586.76 | 6601.13 | 6588.41 | 6580.36 | 6576.18 | 6524.52 | 6456.29 |
| 43 | 1078.58 | 6714.52 | 6699.93 | 6738.93 | 6726.01 | 6723.61 | 6721.77 | 6678.61 | 6600.90 |
| 44 | 1085.51 | 6905.29 | 6887.28 | 6920.39 | 6915.53 | 6911.63 | 6905.87 | 6850.04 | 6786.46 |
| 45 | 1092.43 | 7263.63 | 7225.25 | 7269.55 | 7251.79 | 7247.37 | 7239.11 | 7182.63 | 7097.68 |
| 46 | 1099.34 | 7519.34 | 7456.05 | 7467.58 | 7446.61 | 7429.24 | 7409.98 | 7330.59 | 7241.27 |
| 47 | 1106.25 | 7768.21 | 7678.50 | 7687.63 | 7644.65 | 7621.46 | 7594.35 | 7517.29 | 7401.44 |
| 48 | 1113.14 | 7747.42 | 7653.30 | 7638.06 | 7602.19 | 7575.62 | 7548.67 | 7463.49 | 7365.54 |
| 49 | 1120.02 | 7717.81 | 7652.27 | 7665.62 | 7631.43 | 7613.82 | 7594.30 | 7528.31 | 7426.25 |
| 50 | 1126.89 | 7676.78 | 7653.31 | 7675.06 | 7661.19 | 7654.22 | 7644.83 | 7581.55 | 7504.63 |
| 51 | 1133.76 | 7833.78 | 7849.04 | 7908.06 | 7901.28 | 7902.60 | 7902.04 | 7850.45 | 7746.36 |
| 52 | 1140.61 | 7916.09 | 7949.65 | 8000.39 | 8005.79 | 8007.06 | 8010.22 | 7947.53 | 7837.91 |
| 53 | 1147.45 | 8124.23 | 8154.03 | 8209.96 | 8202.26 | 8201.88 | 8206.72 | 8161.48 | 8066.84 |
| 54 | 1154.29 | 8090.02 | 8104.94 | 8125.60 | 8114.43 | 8103.95 | 8103.28 | 8041.65 | 7951.46 |
| 55 | 1161.11 | 7902.12 | 7911.72 | 7937.79 | 7914.10 | 7899.46 | 7884.64 | 7759.24 | 7504.71 |
| 56 | 1167.93 | 7701.64 | 7714.79 | 7729.07 | 7717.10 | 7707.34 | 7707.76 | 7663.14 | 7593.07 |
| 57 | 1174.73 | 7440.95 | 7478.75 | 7525.18 | 7517.84 | 7519.00 | 7526.41 | 7473.15 | 7326.71 |
| 58 | 1181.53 | 7349.07 | 7406.69 | 7455.20 | 7464.63 | 7470.29 | 7477.70 | 7427.65 | 7295.47 |
| 59 | 1188.31 | 7514.50 | 7575.54 | 7648.88 | 7654.22 | 7667.14 | 7682.46 | 7660.85 | 7593.96 |
| 60 | 1195.09 | 7651.88 | 7702.35 | 7756.82 | 7765.28 | 7769.83 | 7778.08 | 7743.80 | 7683.54 |
| 61 | 1201.85 | 7869.26 | 7894.64 | 7953.57 | 7942.05 | 7939.57 | 7945.01 | 7904.91 | 7818.68 |
| 62 | 1208.61 | 7917.89 | 7909.59 | 7932.06 | 7916.04 | 7902.23 | 7894.54 | 7839.01 | 7757.61 |
| 63 | 1215.35 | 7954.40 | 7919.70 | 7941.83 | 7908.16 | 7889.93 | 7881.59 | 7823.26 | 7711.14 |
| 64 | 1222.09 | 7859.41 | 7814.47 | 7820.05 | 7793.69 | 7777.74 | 7763.57 | 7705.36 | 7618.37 |
| 65 | 1228.82 | 7845.46 | 7809.59 | 7842.12 | 7814.88 | 7804.54 | 7799.22 | 7749.47 | 7653.37 |
| 66 | 1235.53 | 7826.38 | 7807.98 | 7838.12 | 7828.03 | 7822.04 | 7815.99 | 7771.51 | 7704.15 |
| 67 | 1242.24 | 7966.49 | 7969.64 | 8032.19 | 8022.00 | 8021.23 | 8022.32 | 7981.99 | 7898.29 |
| 68 | 1248.94 | 8119.37 | 8137.19 | 8187.79 | 8188.03 | 8189.24 | 8185.72 | 8132.45 | 8053.21 |
| 69 | 1255.63 | 8351.53 | 8366.00 | 8433.79 | 8424.82 | 8424.84 | 8426.02 | 8375.11 | 8276.79 |
| 70 | 1262.30 | 8502.88 | 8506.61 | 8545.94 | 8538.07 | 8529.23 | 8522.97 | 8464.80 | 8380.24 |
| 71 | 1268.97 | 8629.16 | 8616.05 | 8655.77 | 8626.13 | 8615.16 | 8609.52 | 8548.73 | 8432.83 |
| 72 | 1275.63 | 8566.81 | 8540.52 | 8547.87 | 8524.94 | 8505.74 | 8492.55 | 8429.27 | 8332.99 |
| 73 | 1282.28 | 8466.86 | 8451.80 | 8478.49 | 8446.45 | 8434.66 | 8431.02 | 8378.25 | 8268.81 |
| 74 | 1288.92 | 8263.04 | 8274.46 | 8299.23 | 8286.42 | 8277.28 | 8278.45 | 8233.60 | 8158.70 |
| 75 | 1295.55 | 8162.13 | 8198.88 | 8252.40 | 8239.88 | 8239.47 | 8247.57 | 8215.28 | 8127.84 |
| 76 | 1302.16 | 8046.64 | 8103.90 | 8152.24 | 8157.93 | 8160.18 | 8173.30 | 8143.24 | 8079.27 |
| 77 | 1308.77 | 8059.08 | 8138.90 | 8222.14 | 8224.32 | 8237.47 | 8259.52 | 8239.10 | 8164.59 |
| 78 | 1315.37 | 8083.30 | 8179.04 | 8251.90 | 8269.58 | 8279.94 | 8302.36 | 8280.76 | 8225.45 |
| 79 | 1321.96 | 8110.96 | 8211.21 | 8304.97 | 8313.10 | 8329.57 | 8356.16 | 8322.72 | 8196.06 |
| 80 | 1328.54 | 8244.73 | 8335.57 | 8402.88 | 8411.99 | 8418.81 | 8437.50 | 8406.41 | 8336.28 |
| 81 | 1335.11 | 8299.49 | 8371.64 | 8438.87 | 8429.24 | 8427.46 | 8445.78 | 8413.90 | 8320.77 |
| 82 | 1341.67 | 8177.01 | 8222.98 | 8257.65 | 8248.51 | 8237.51 | 8247.58 | 8213.95 | 8133.29 |
| 83 | 1348.22 | 7937.64 | 7969.77 | 8013.00 | 7988.32 | 7980.39 | 7992.22 | 7964.72 | 7872.43 |
| 84 | 1354.77 | 7592.50 | 7622.17 | 7652.00 | 7647.23 | 7641.05 | 7654.72 | 7630.51 | 7567.40 |
| 85 | 1361.30 | 7359.39 | 7397.70 | 7457.67 | 7440.17 | 7443.11 | 7463.80 | 7451.03 | 7380.43 |
| 86 | 1367.82 | 7335.20 | 7380.33 | 7426.84 | 7417.35 | 7411.65 | 7423.77 | 7376.82 | 7288.31 |
| 87 | 1374.33 | 7344.10 | 7394.29 | 7464.14 | 7449.56 | 7453.12 | 7479.12 | 7460.36 | 7378.69 |
| 88 | 1380.83 | 7407.48 | 7465.83 | 7519.76 | 7518.24 | 7515.41 | 7530.86 | 7474.16 | 7355.25 |
| 89 | 1387.32 | 7498.47 | 7553.65 | 7644.11 | 7627.64 | 7642.92 | 7669.73 | 7665.28 | 7596.04 |
| 90 | 1393.81 | 7737.24 | 7792.02 | 7861.86 | 7874.31 | 7880.77 | 7924.17 | 7910.28 | 7851.92 |
| 91 | 1400.28 | 7916.14 | 7953.13 | 8045.40 | 8032.66 | 8044.37 | 8080.65 | 8057.97 | 7962.40 |
| 92 | 1406.74 | 8114.30 | 8143.83 | 8197.92 | 8196.55 | 8191.13 | 8229.15 | 8187.74 | 8083.07 |
| 93 | 1413.19 | 8269.38 | 8279.78 | 8346.90 | 8312.37 | 8309.36 | 8339.32 | 8307.39 | 8209.32 |
| 94 | 1419.64 | 8347.82 | 8348.05 | 8374.27 | 8358.16 | 8343.77 | 8377.91 | 8330.39 | 8217.66 |
| 95 | 1426.07 | 8188.05 | 8188.44 | 8250.65 | 8213.28 | 8212.19 | 8234.23 | 8190.50 | 8005.78 |
| 96 | 1432.50 | 8290.12 | 8305.82 | 8346.61 | 8333.21 | 8310.40 | 8289.27 | 8072.17 | 7655.01 |
| 97 | 1438.91 | 8273.56 | 8304.77 | 8402.31 | 8369.19 | 8380.89 | 8408.26 | 8394.97 | 8300.51 |
| 98 | 1445.31 | 8255.16 | 8318.14 | 8386.29 | 8383.99 | 8383.18 | 8405.40 | 8340.75 | 8173.53 |
| 99 | 1451.71 | 8269.36 | 8327.58 | 8438.88 | 8406.31 | 8418.58 | 8438.32 | 8420.46 | 8315.72 |
| 100 | 1458.09 | 8172.05 | 8255.50 | 8332.91 | 8335.16 | 8338.82 | 8364.62 | 8334.95 | 8233.60 |
| 101 | 1464.47 | 8145.10 | 8227.48 | 8356.74 | 8326.71 | 8348.02 | 8366.63 | 8355.94 | 8264.35 |
| 102 | 1470.83 | 7727.94 | 7864.24 | 7977.53 | 7993.12 | 8012.94 | 8049.25 | 8006.12 | 7769.86 |
| 103 | 1477.19 | 8110.15 | 8239.68 | 8389.73 | 8375.84 | 8407.49 | 8423.19 | 8416.94 | 8314.93 |
| 104 | 1483.53 | 8089.76 | 8240.97 | 8367.12 | 8385.52 | 8403.86 | 8441.54 | 8425.01 | 8336.44 |
| 105 | 1489.87 | 8150.85 | 8289.77 | 8447.69 | 8432.38 | 8460.91 | 8472.14 | 8432.30 | 8256.27 |
| 106 | 1496.20 | 8115.92 | 8265.87 | 8377.86 | 8388.70 | 8396.71 | 8426.82 | 8398.81 | 8294.89 |
| 107 | 1502.51 | 8079.41 | 8199.93 | 8333.25 | 8300.54 | 8318.57 | 8332.55 | 8314.75 | 8193.86 |
| 108 | 1508.82 | 7936.66 | 8056.82 | 8146.32 | 8139.81 | 8138.99 | 8160.14 | 8137.96 | 8031.86 |
| 109 | 1515.12 | 7775.32 | 7894.30 | 8015.72 | 7984.95 | 8000.19 | 8018.52 | 8006.70 | 7871.95 |
| 110 | 1521.40 | 7615.90 | 7751.99 | 7846.07 | 7858.81 | 7868.37 | 7902.63 | 7895.34 | 7798.29 |
| 111 | 1527.68 | 7474.17 | 7626.52 | 7771.06 | 7753.53 | 7792.05 | 7820.63 | 7812.19 | 7712.50 |
| 112 | 1533.95 | 7351.11 | 7533.05 | 7653.43 | 7675.21 | 7699.44 | 7738.16 | 7724.20 | 7607.54 |
| 113 | 1540.21 | 7235.16 | 7409.90 | 7561.26 | 7557.78 | 7591.26 | 7618.73 | 7629.90 | 7535.30 |
| 114 | 1546.45 | 7098.32 | 7287.16 | 7399.24 | 7422.13 | 7444.27 | 7492.44 | 7495.69 | 7437.37 |
| 115 | 1552.69 | 6939.49 | 7126.90 | 7268.81 | 7259.72 | 7288.79 | 7318.49 | 7336.83 | 7251.33 |
| 116 | 1558.92 | 6728.64 | 6930.48 | 7040.73 | 7066.90 | 7093.07 | 7134.04 | 7148.72 | 7085.11 |
| 117 | 1565.14 | 6592.38 | 6804.31 | 6947.09 | 6952.20 | 6994.61 | 7032.28 | 7070.78 | 7009.13 |
| 118 | 1571.35 | 6417.67 | 6661.69 | 6783.09 | 6822.30 | 6861.19 | 6921.90 | 6959.96 | 6935.05 |
| 119 | 1577.55 | 6367.69 | 6624.36 | 6775.55 | 6789.45 | 6837.41 | 6885.92 | 6937.91 | 6888.88 |
| 120 | 1583.74 | 6257.26 | 6561.33 | 6684.41 | 6725.58 | 6762.92 | 6823.01 | 6862.40 | 6821.90 |
| 121 | 1589.92 | 6069.01 | 6461.65 | 6607.82 | 6619.10 | 6670.50 | 6712.07 | 6764.10 | 6721.27 |
| 122 | 1596.09 | 5733.24 | 6343.93 | 6450.84 | 6491.69 | 6518.34 | 6574.72 | 6604.72 | 6532.84 |
| 123 | 1602.25 | 5103.79 | 6210.91 | 6339.63 | 6345.20 | 6382.67 | 6420.88 | 6458.98 | 6391.93 |
| 124 | 1608.40 | 4064.24 | 5983.05 | 6098.16 | 6133.28 | 6157.59 | 6213.28 | 6244.98 | 6209.55 |
| 125 | 1614.54 | 2999.50 | 5715.46 | 5885.70 | 5910.96 | 5956.18 | 6000.94 | 6056.26 | 6015.81 |
| 126 | 1620.67 | 2012.98 | 5296.07 | 5550.84 | 5654.93 | 5701.81 | 5773.71 | 5842.16 | 5819.95 |
| 127 | 1626.79 | 1392.62 | 4728.47 | 5245.99 | 5414.36 | 5508.48 | 5586.72 | 5668.54 | 5629.01 |
| 128 | 1632.90 | 983.78 | 3948.54 | 4748.36 | 5129.39 | 5275.38 | 5414.74 | 5521.77 | 5524.54 |
| 129 | 1639.00 | 768.35 | 3109.86 | 4124.82 | 4758.16 | 5016.71 | 5209.74 | 5348.22 | 5301.57 |
| 130 | 1645.09 | 696.82 | 2446.01 | 3341.75 | 4119.49 | 4498.00 | 4812.28 | 4964.98 | 4846.12 |
| 131 | 1651.18 | 511.89 | 1714.20 | 2550.22 | 3434.33 | 4014.40 | 4559.13 | 4888.31 | 4903.56 |
| 132 | 1657.25 | 396.92 | 1224.27 | 1815.95 | 2603.30 | 3214.69 | 3990.47 | 4528.97 | 4659.53 |
| 133 | 1663.31 | 338.53 | 899.55 | 1313.23 | 1887.94 | 2439.53 | 3285.79 | 4129.80 | 4399.99 |
| 134 | 1669.36 | 285.26 | 703.02 | 997.43 | 1412.34 | 1828.99 | 2605.38 | 3637.16 | 4141.40 |
| 135 | 1675.41 | 258.80 | 556.41 | 778.65 | 1067.67 | 1381.10 | 1991.60 | 3060.54 | 3803.74 |
| 136 | 1681.44 | 238.50 | 460.39 | 621.33 | 853.98 | 1096.34 | 1582.70 | 2512.15 | 3361.82 |
| 137 | 1687.46 | 234.90 | 391.17 | 517.11 | 667.60 | 841.54 | 1188.13 | 1937.66 | 2743.98 |
| 138 | 1693.48 | 215.41 | 331.16 | 424.23 | 545.94 | 672.55 | 939.64 | 1520.18 | 2164.26 |
| 139 | 1699.48 | 206.79 | 310.55 | 383.23 | 463.56 | 561.85 | 750.70 | 1194.82 | 1737.35 |
| 140 | 1705.48 | 177.54 | 258.02 | 318.12 | 395.50 | 464.20 | 613.03 | 936.25 | 1353.58 |
| 141 | 1711.46 | 189.35 | 253.06 | 310.01 | 350.67 | 415.08 | 528.01 | 775.37 | 1049.02 |
| 142 | 1717.43 | 24.60 | 89.66 | 134.89 | 176.26 | 215.49 | 303.87 | 490.65 | 686.33 |
| 90 | 696.90 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 |
| 91 | 700.14 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 |
| 92 | 703.37 | 0.14 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 |
| 93 | 706.60 | 0.41 | 0.12 | 0.01 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 |
| 94 | 709.82 | 0.68 | 0.38 | 0.27 | 0.16 | 0.06 | 0.00 | 0.00 | 0.00 |
| 95 | 713.04 | 0.97 | 0.66 | 0.55 | 0.43 | 0.32 | 0.26 | 0.04 | 0.00 |
| 96 | 716.25 | 1.27 | 0.95 | 0.83 | 0.71 | 0.60 | 0.53 | 0.30 | 0.09 |
| 97 | 719.45 | 1.57 | 1.24 | 1.12 | 0.99 | 0.88 | 0.81 | 0.58 | 0.36 |
| 98 | 722.66 | 1.88 | 1.54 | 1.42 | 1.29 | 1.17 | 1.10 | 0.86 | 0.64 |
| 99 | 725.85 | 2.20 | 1.86 | 1.73 | 1.59 | 1.47 | 1.40 | 1.15 | 0.92 |
| 100 | 729.05 | 2.54 | 2.18 | 2.05 | 1.91 | 1.78 | 1.71 | 1.45 | 1.22 |
| 101 | 732.23 | 2.88 | 2.51 | 2.37 | 2.23 | 2.10 | 2.03 | 1.76 | 1.52 |
| 102 | 735.42 | 3.23 | 2.85 | 2.71 | 2.56 | 2.43 | 2.35 | 2.08 | 1.83 |
| 103 | 738.59 | 3.60 | 3.20 | 3.06 | 2.91 | 2.77 | 2.69 | 2.41 | 2.15 |
| 104 | 741.77 | 3.98 | 3.57 | 3.42 | 3.26 | 3.12 | 3.04 | 2.75 | 2.48 |
| 105 | 744.94 | 4.31 | 3.95 | 3.79 | 3.63 | 3.48 | 3.39 | 3.09 | 2.82 |
| 106 | 748.10 | 4.64 | 4.28 | 4.15 | 4.01 | 3.86 | 3.77 | 3.46 | 3.17 |
| 107 | 751.26 | 5.03 | 4.62 | 4.48 | 4.33 | 4.20 | 4.13 | 3.83 | 3.54 |
| 108 | 754.41 | 5.48 | 4.99 | 4.83 | 4.67 | 4.54 | 4.45 | 4.18 | 3.91 |
| 109 | 757.56 | 6.04 | 5.44 | 5.25 | 5.06 | 4.90 | 4.81 | 4.51 | 4.25 |
| 110 | 760.70 | 6.62 | 5.99 | 5.76 | 5.53 | 5.33 | 5.22 | 4.87 | 4.59 |
| 111 | 763.84 | 7.36 | 6.57 | 6.32 | 6.08 | 5.86 | 5.72 | 5.30 | 4.96 |
| 112 | 766.97 | 8.29 | 7.29 | 6.98 | 6.67 | 6.42 | 6.28 | 5.82 | 5.40 |
| 113 | 770.10 | 9.48 | 8.22 | 7.84 | 7.43 | 7.11 | 6.93 | 6.38 | 5.94 |
| 114 | 773.23 | 11.43 | 9.35 | 8.83 | 8.37 | 8.02 | 7.78 | 7.05 | 6.51 |
| 115 | 776.35 | 14.16 | 11.24 | 10.38 | 9.61 | 9.04 | 8.75 | 7.95 | 7.22 |
| 116 | 779.46 | 17.28 | 13.91 | 12.76 | 11.63 | 10.74 | 10.26 | 8.95 | 8.14 |
| 117 | 782.57 | 21.02 | 17.00 | 15.73 | 14.41 | 13.25 | 12.58 | 10.59 | 9.23 |
| 118 | 785.67 | 25.59 | 20.69 | 19.12 | 17.56 | 16.27 | 15.53 | 13.04 | 11.05 |
| 119 | 788.77 | 31.13 | 25.18 | 23.25 | 21.36 | 19.80 | 18.88 | 16.04 | 13.66 |
| 120 | 791.87 | 37.55 | 30.64 | 28.31 | 26.00 | 24.07 | 22.96 | 19.52 | 16.73 |
| 121 | 794.96 | 44.08 | 37.01 | 34.39 | 31.63 | 29.30 | 27.94 | 23.73 | 20.36 |
| 122 | 798.04 | 50.95 | 43.54 | 40.94 | 38.09 | 35.52 | 33.96 | 28.88 | 24.77 |
| 123 | 801.12 | 60.21 | 50.29 | 47.41 | 44.62 | 42.07 | 40.52 | 35.05 | 30.13 |
| 124 | 804.20 | 70.20 | 59.39 | 55.51 | 51.63 | 48.60 | 46.98 | 41.60 | 36.45 |
| 125 | 807.27 | 81.15 | 69.38 | 65.45 | 61.04 | 57.16 | 54.91 | 48.09 | 42.99 |
| 126 | 810.33 | 94.07 | 80.17 | 75.60 | 71.03 | 67.15 | 64.80 | 56.47 | 49.64 |
| 127 | 813.39 | 106.63 | 93.00 | 87.68 | 82.14 | 77.54 | 74.89 | 66.44 | 58.55 |

Use of the Absolute Responsivities

The net count rate from the IR spectrometer is related to the average brightness over the field of view of the instrument according to:

Net DN/sec(i) = Rmeas(pixel i, Detector Temp. T) * <I()>

Where <I(l)> denotes the average intensity over the field of view of the detector in watts/sq.m-micron-sr.  The units of Rmeas is (DN/sec)/(Watt/sq.m-micron-sr).  To evaluate the average over the field of view of the instrument, the relative response across the field of view of the instrument is required.

- Relative response over the fields of view of the ULIS and DLIS

5.1 Field of view of the ULIS

For the ULIS, the field of view covers approximately 170 degrees in azimuth and extends from zenith angles of 5 degrees to 90 degrees.  The relative response is given in the table below.

Relative Spatial Response of ULIS (times 10,000)

| AZ | ZA = 1 | ZA = 2 | ZA = 3 | ZA = 4 | ZA = 5 | ZA = 6 | ZA = 7 | ZA = 8 | ZA = 9 | ZA = 10 | ZA = 11 | ZA = 12 | ZA = 13 | ZA = 14 | ZA = 15 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| -89 | 0 | 0 | 0 | 1 | 1 | 2 | 2 | 2 | 2 | 5 | 5 | 6 | 6 | 7 | 1 |
| -88 | 0 | 1 | 1 | 1 | 2 | 3 | 3 | 4 | 4 | 8 | 11 | 12 | 13 | 14 | 15 |
| -87 | 0 | 1 | 1 | 2 | 3 | 4 | 5 | 5 | 6 | 17 | 32 | 32 | 33 | 30 | 17 |
| -86 | 1 | 1 | 2 | 2 | 4 | 5 | 6 | 14 | 14 | 21 | 36 | 38 | 40 | 35 | 24 |
| -85 | 1 | 2 | 2 | 3 | 4 | 13 | 14 | 15 | 16 | 24 | 41 | 58 | 58 | 52 | 26 |
| -84 | 1 | 2 | 3 | 4 | 10 | 14 | 16 | 17 | 18 | 38 | 55 | 64 | 64 | 54 | 33 |
| -83 | 1 | 2 | 3 | 4 | 11 | 15 | 17 | 19 | 29 | 41 | 59 | 69 | 81 | 61 | 35 |
| -82 | 1 | 2 | 4 | 9 | 11 | 16 | 19 | 27 | 31 | 44 | 73 | 90 | 84 | 63 | 56 |
| -81 | 1 | 3 | 4 | 10 | 12 | 17 | 27 | 29 | 34 | 58 | 76 | 95 | 94 | 92 | 85 |
| -80 | 2 | 3 | 9 | 11 | 13 | 23 | 28 | 30 | 42 | 61 | 80 | 165 | 143 | 123 | 176 |
| -79 | 2 | 3 | 10 | 11 | 14 | 24 | 29 | 39 | 45 | 64 | 107 | 208 | 328 | 247 | 224 |
| -78 | 2 | 4 | 10 | 12 | 19 | 25 | 31 | 40 | 48 | 89 | 131 | 291 | 407 | 311 | 381 |
| -77 | 2 | 4 | 10 | 12 | 20 | 26 | 37 | 42 | 66 | 110 | 218 | 349 | 486 | 539 | 449 |
| -76 | 2 | 4 | 11 | 13 | 21 | 27 | 38 | 44 | 85 | 132 | 242 | 407 | 730 | 625 | 685 |
| -75 | 2 | 9 | 11 | 18 | 21 | 33 | 39 | 69 | 104 | 216 | 265 | 629 | 809 | 875 | 825 |
| -74 | 2 | 9 | 12 | 19 | 22 | 34 | 40 | 75 | 123 | 238 | 366 | 684 | 1102 | 1009 | 1081 |
| -73 | 3 | 10 | 12 | 19 | 23 | 34 | 45 | 81 | 206 | 259 | 403 | 798 | 1181 | 1259 | 1173 |
| -72 | 3 | 10 | 13 | 20 | 29 | 35 | 45 | 86 | 224 | 343 | 395 | 851 | 1425 | 1344 | 1434 |
| -71 | 3 | 10 | 13 | 21 | 29 | 36 | 44 | 68 | 243 | 360 | 533 | 903 | 1503 | 1594 | 1527 |
| -70 | 3 | 11 | 18 | 21 | 30 | 41 | 43 | 78 | 321 | 381 | 568 | 1119 | 1747 | 1679 | 1790 |
| -69 | 3 | 11 | 18 | 22 | 31 | 41 | 41 | 98 | 340 | 465 | 602 | 1169 | 1826 | 1931 | 1933 |
| -68 | 3 | 11 | 19 | 27 | 32 | 40 | 39 | 95 | 358 | 485 | 684 | 1382 | 1904 | 2062 | 2224 |
| -67 | 3 | 11 | 19 | 27 | 32 | 40 | 39 | 104 | 440 | 505 | 728 | 1430 | 2027 | 2322 | 2338 |
| -66 | 4 | 12 | 20 | 28 | 37 | 38 | 38 | 113 | 458 | 589 | 935 | 1371 | 2077 | 2417 | 2670 |
| -65 | 4 | 12 | 20 | 28 | 37 | 36 | 37 | 122 | 475 | 608 | 871 | 1580 | 2291 | 2694 | 2789 |
| -64 | 4 | 12 | 21 | 29 | 36 | 35 | 35 | 193 | 493 | 624 | 911 | 1624 | 2338 | 2792 | 3140 |
| -63 | 4 | 13 | 21 | 30 | 35 | 35 | 34 | 201 | 574 | 707 | 1114 | 1831 | 2384 | 2910 | 3232 |
| -62 | 4 | 13 | 21 | 30 | 35 | 34 | 33 | 208 | 517 | 726 | 1152 | 1873 | 2494 | 3170 | 3497 |
| -61 | 4 | 13 | 22 | 31 | 34 | 34 | 33 | 148 | 524 | 745 | 1189 | 1807 | 2543 | 3250 | 3582 |
| -60 | 9 | 18 | 27 | 36 | 32 | 31 | 35 | 218 | 595 | 846 | 1389 | 2013 | 2771 | 3509 | 3855 |
| -59 | 9 | 18 | 27 | 35 | 31 | 30 | 28 | 224 | 601 | 871 | 1318 | 2055 | 2817 | 3474 | 3939 |
| -58 | 9 | 18 | 27 | 35 | 31 | 30 | 28 | 229 | 607 | 860 | 1350 | 2096 | 2756 | 3702 | 4202 |
| -57 | 9 | 19 | 28 | 35 | 31 | 29 | 27 | 234 | 613 | 878 | 1548 | 2316 | 2979 | 3749 | 4284 |
| -56 | 9 | 19 | 28 | 34 | 30 | 28 | 27 | 238 | 615 | 1001 | 1584 | 2249 | 3021 | 3793 | 4553 |
| -55 | 10 | 19 | 29 | 34 | 30 | 27 | 41 | 306 | 619 | 1023 | 1512 | 2286 | 3241 | 3911 | 4633 |
| -54 | 10 | 19 | 29 | 34 | 29 | 25 | 42 | 243 | 623 | 1044 | 1703 | 2502 | 3174 | 3953 | 4712 |
| -53 | 10 | 20 | 29 | 34 | 27 | 24 | 43 | 246 | 627 | 1042 | 1751 | 2536 | 3211 | 4172 | 4849 |
| -52 | 10 | 20 | 30 | 33 | 26 | 24 | 44 | 249 | 693 | 1084 | 1797 | 2465 | 3419 | 4089 | 4890 |
| -51 | 10 | 20 | 30 | 31 | 26 | 23 | 21 | 315 | 629 | 1125 | 1842 | 2772 | 3472 | 4146 | 5107 |
| -50 | 10 | 20 | 30 | 31 | 26 | 23 | 42 | 316 | 631 | 1164 | 2122 | 2837 | 3509 | 4466 | 5140 |
| -49 | 10 | 21 | 31 | 31 | 25 | 22 | 42 | 318 | 633 | 1383 | 2180 | 2901 | 3817 | 4540 | 5219 |
| -48 | 10 | 21 | 35 | 30 | 25 | 19 | 43 | 252 | 711 | 1369 | 2237 | 3157 | 3884 | 4567 | 5483 |
| -47 | 10 | 21 | 35 | 30 | 25 | 18 | 42 | 252 | 742 | 1404 | 2533 | 3218 | 3904 | 4877 | 5450 |
| -46 | 11 | 21 | 35 | 30 | 25 | 18 | 42 | 316 | 706 | 1438 | 2464 | 3277 | 4208 | 4946 | 5539 |
| -45 | 11 | 21 | 35 | 30 | 22 | 17 | 14 | 315 | 735 | 1654 | 2500 | 3529 | 4271 | 4968 | 5551 |
| -44 | 11 | 22 | 35 | 29 | 22 | 17 | 38 | 314 | 763 | 1633 | 2534 | 3585 | 4286 | 5218 | 5516 |
| -43 | 11 | 22 | 35 | 29 | 21 | 16 | 37 | 246 | 988 | 1663 | 2697 | 3640 | 4584 | 5182 | 5602 |
| -42 | 11 | 22 | 34 | 29 | 21 | 16 | 36 | 244 | 1014 | 1691 | 2728 | 3584 | 4642 | 5190 | 5566 |
| -41 | 11 | 27 | 34 | 27 | 21 | 13 | 35 | 320 | 973 | 1900 | 2757 | 3855 | 4652 | 5274 | 5574 |
| -40 | 11 | 27 | 34 | 26 | 21 | 13 | 8 | 342 | 997 | 1875 | 2785 | 3885 | 4936 | 5236 | 5658 |
| -39 | 11 | 27 | 34 | 26 | 20 | 13 | 8 | 362 | 1217 | 1900 | 2941 | 3912 | 4942 | 5242 | 5619 |
| -38 | 11 | 27 | 34 | 26 | 20 | 12 | 27 | 315 | 1239 | 1923 | 2965 | 4069 | 4812 | 5324 | 5625 |
| -37 | 12 | 27 | 34 | 26 | 20 | 12 | 25 | 334 | 1193 | 2075 | 2987 | 4093 | 4839 | 5283 | 5661 |
| -36 | 12 | 28 | 34 | 26 | 18 | 12 | 22 | 351 | 1213 | 2095 | 2899 | 4116 | 5131 | 5287 | 5665 |
| -35 | 12 | 28 | 33 | 25 | 17 | 11 | 20 | 566 | 1232 | 2114 | 3157 | 4027 | 5000 | 5405 | 5668 |
| -34 | 12 | 28 | 33 | 25 | 17 | 11 | 2 | 514 | 1249 | 2132 | 3174 | 4045 | 5024 | 5440 | 5702 |
| -33 | 12 | 28 | 31 | 25 | 17 | 7 | 3 | 529 | 1397 | 2067 | 3080 | 4297 | 5046 | 5474 | 5703 |
| -32 | 12 | 28 | 31 | 25 | 17 | 8 | 4 | 543 | 1412 | 2293 | 3094 | 4199 | 5178 | 5542 | 5758 |
| -31 | 12 | 28 | 31 | 25 | 16 | 9 | 16 | 556 | 1426 | 2307 | 3106 | 4207 | 5198 | 5653 | 5793 |
| -30 | 12 | 29 | 31 | 24 | 15 | 18 | 49 | 575 | 1373 | 2245 | 3247 | 4105 | 5146 | 5720 | 5940 |
| -29 | 12 | 29 | 30 | 24 | 16 | 20 | 74 | 801 | 1385 | 2258 | 3255 | 4330 | 5152 | 5753 | 5975 |
| -28 | 12 | 29 | 30 | 20 | 16 | 21 | 98 | 829 | 1395 | 2270 | 3261 | 4333 | 5363 | 5785 | 6043 |
| -27 | 12 | 29 | 30 | 20 | 17 | 23 | 122 | 856 | 1603 | 2281 | 3157 | 4225 | 5267 | 5928 | 6077 |
| -26 | 12 | 29 | 30 | 21 | 18 | 25 | 145 | 889 | 1628 | 2421 | 3160 | 4225 | 5267 | 5959 | 6187 |
| -25 | 12 | 29 | 30 | 21 | 26 | 27 | 174 | 914 | 1656 | 2429 | 3161 | 4223 | 5155 | 6024 | 6255 |
| -24 | 13 | 29 | 30 | 22 | 28 | 36 | 195 | 938 | 1684 | 2436 | 3289 | 4110 | 5151 | 6055 | 6287 |
| -23 | 13 | 30 | 30 | 22 | 36 | 42 | 216 | 962 | 1710 | 2465 | 3284 | 4324 | 5144 | 5954 | 6354 |
| -22 | 13 | 30 | 30 | 31 | 38 | 44 | 236 | 992 | 1742 | 2493 | 3277 | 4208 | 5247 | 6225 | 6385 |
| -21 | 13 | 30 | 30 | 32 | 39 | 46 | 452 | 1013 | 1766 | 2520 | 3288 | 4198 | 5237 | 6084 | 6518 |
| -20 | 13 | 30 | 30 | 34 | 41 | 48 | 478 | 1231 | 1790 | 2553 | 3329 | 4187 | 5116 | 6092 | 6530 |
| -19 | 13 | 30 | 28 | 35 | 42 | 56 | 495 | 1251 | 2016 | 2577 | 3368 | 4255 | 5102 | 6098 | 6578 |
| -18 | 13 | 30 | 29 | 36 | 43 | 58 | 512 | 1278 | 2038 | 2799 | 3426 | 4301 | 5135 | 6013 | 6591 |
| -17 | 13 | 30 | 29 | 37 | 52 | 60 | 528 | 1296 | 2058 | 2828 | 3706 | 4390 | 5183 | 5993 | 6638 |
| -16 | 13 | 30 | 30 | 38 | 53 | 61 | 550 | 1313 | 2084 | 2849 | 3741 | 4653 | 5450 | 6071 | 6652 |
| -15 | 13 | 30 | 30 | 39 | 55 | 63 | 565 | 1329 | 2102 | 2870 | 3780 | 4695 | 5539 | 6339 | 6680 |
| -14 | 13 | 30 | 30 | 47 | 56 | 72 | 578 | 1352 | 2120 | 2896 | 3807 | 4779 | 5582 | 6430 | 6696 |
| -13 | 13 | 30 | 31 | 49 | 58 | 74 | 591 | 1367 | 2136 | 2914 | 3838 | 4818 | 5669 | 6475 | 6760 |
| -12 | 13 | 31 | 31 | 50 | 59 | 75 | 610 | 1380 | 2158 | 2931 | 3862 | 4852 | 5709 | 6563 | 6778 |
| -11 | 13 | 31 | 32 | 51 | 67 | 77 | 621 | 1393 | 2172 | 2953 | 3884 | 4882 | 5792 | 6604 | 6840 |
| -10 | 13 | 31 | 32 | 52 | 69 | 79 | 632 | 1412 | 2185 | 2968 | 3910 | 4910 | 5829 | 6688 | 6859 |
| -9 | 13 | 31 | 43 | 53 | 70 | 87 | 641 | 1423 | 2205 | 2981 | 3929 | 4951 | 5865 | 6727 | 6924 |
| -8 | 13 | 31 | 44 | 54 | 71 | 89 | 657 | 1433 | 2216 | 3000 | 3946 | 4969 | 5943 | 6807 | 6944 |
| -7 | 13 | 31 | 45 | 62 | 73 | 90 | 665 | 1449 | 2226 | 3011 | 3967 | 4993 | 5975 | 6842 | 7009 |
| -6 | 13 | 31 | 45 | 63 | 74 | 92 | 672 | 1457 | 2242 | 3021 | 3981 | 5007 | 6004 | 6919 | 7030 |
| -5 | 13 | 31 | 46 | 64 | 82 | 94 | 678 | 1464 | 2250 | 3036 | 3999 | 5020 | 6026 | 6950 | 7095 |
| -4 | 13 | 31 | 47 | 65 | 84 | 102 | 691 | 1470 | 2257 | 3044 | 4009 | 5038 | 6066 | 7024 | 7117 |
| -3 | 13 | 31 | 48 | 66 | 85 | 104 | 696 | 1483 | 2270 | 3051 | 4018 | 5047 | 6076 | 7052 | 7183 |
| -2 | 13 | 31 | 48 | 67 | 86 | 105 | 699 | 1487 | 2275 | 3063 | 4031 | 5061 | 6090 | 7083 | 7205 |
| -1 | 13 | 31 | 49 | 68 | 88 | 107 | 702 | 1491 | 2279 | 3067 | 4036 | 5066 | 6096 | 7099 | 7272 |
| 0 | 13 | 31 | 55 | 69 | 89 | 108 | 902 | 1690 | 2479 | 3267 | 4297 | 5070 | 6358 | 7388 | 7264 |
| 1 | 13 | 31 | 58 | 77 | 97 | 117 | 713 | 1502 | 2290 | 3093 | 4109 | 5141 | 6173 | 7205 | 7348 |
| 2 | 13 | 31 | 58 | 78 | 98 | 118 | 713 | 1502 | 2291 | 3107 | 4125 | 5159 | 6192 | 7226 | 7417 |
| 3 | 13 | 31 | 59 | 79 | 100 | 120 | 713 | 1502 | 2290 | 3250 | 4232 | 5226 | 6261 | 7295 | 7440 |
| 4 | 13 | 31 | 60 | 80 | 101 | 121 | 712 | 1507 | 2296 | 3274 | 4253 | 5240 | 6276 | 7312 | 7507 |
| 5 | 13 | 31 | 61 | 81 | 102 | 130 | 717 | 1505 | 2293 | 3297 | 4272 | 5303 | 6341 | 7378 | 7535 |
| 6 | 13 | 31 | 61 | 82 | 110 | 131 | 714 | 1502 | 2290 | 3407 | 4367 | 5314 | 6352 | 7391 | 7614 |
| 7 | 13 | 31 | 62 | 83 | 111 | 132 | 710 | 1498 | 2294 | 3427 | 4391 | 5323 | 6413 | 7452 | 7643 |
| 8 | 13 | 31 | 63 | 91 | 112 | 134 | 706 | 1500 | 2299 | 3445 | 4502 | 5429 | 6420 | 7460 | 7722 |
| 9 | 13 | 31 | 63 | 92 | 113 | 135 | 708 | 1493 | 2303 | 3549 | 4523 | 5443 | 6477 | 7517 | 7751 |
| 10 | 13 | 31 | 64 | 93 | 115 | 143 | 701 | 1486 | 2428 | 3564 | 4543 | 5534 | 6480 | 7521 | 7832 |
| 11 | 13 | 31 | 64 | 94 | 116 | 145 | 694 | 1486 | 2440 | 3576 | 4649 | 5554 | 6482 | 7573 | 7861 |
| 12 | 13 | 31 | 64 | 94 | 124 | 146 | 687 | 1477 | 2450 | 3676 | 4666 | 5572 | 6587 | 7573 | 7942 |
| 13 | 13 | 31 | 64 | 95 | 125 | 147 | 685 | 1467 | 2546 | 3685 | 4681 | 5677 | 6595 | 7621 | 7973 |
| 14 | 13 | 30 | 64 | 96 | 126 | 148 | 675 | 1456 | 2553 | 3693 | 4782 | 5691 | 6689 | 7617 | 8054 |
| 15 | 13 | 30 | 64 | 104 | 127 | 157 | 665 | 1435 | 2558 | 3787 | 4794 | 5793 | 6703 | 7719 | 8086 |
| 16 | 13 | 30 | 65 | 105 | 128 | 158 | 654 | 1430 | 2562 | 3791 | 4893 | 5804 | 6804 | 7496 | 8206 |
| 17 | 13 | 30 | 65 | 105 | 129 | 159 | 649 | 1424 | 2652 | 3793 | 4902 | 5594 | 6595 | 7546 | 8231 |
| 18 | 13 | 30 | 65 | 106 | 137 | 160 | 637 | 1417 | 2652 | 3882 | 4689 | 5690 | 6604 | 7535 | 8298 |
| 19 | 13 | 30 | 65 | 107 | 137 | 161 | 623 | 1510 | 2651 | 3602 | 4783 | 5697 | 6637 | 7567 | 8330 |
| 20 | 13 | 30 | 67 | 108 | 138 | 169 | 609 | 1508 | 2459 | 3599 | 4787 | 5739 | 6623 | 7552 | 8406 |
| 21 | 13 | 30 | 67 | 108 | 139 | 170 | 601 | 1199 | 2454 | 3683 | 4790 | 5724 | 6652 | 7536 | 8438 |
| 22 | 13 | 30 | 66 | 109 | 140 | 171 | 388 | 1187 | 2448 | 3676 | 5146 | 5708 | 6635 | 7561 | 8268 |
| 23 | 13 | 30 | 66 | 107 | 141 | 172 | 371 | 1302 | 2529 | 3668 | 5187 | 5735 | 6440 | 7321 | 8247 |
| 24 | 13 | 30 | 65 | 107 | 149 | 168 | 354 | 1294 | 2520 | 3952 | 5228 | 5811 | 6419 | 7343 | 8268 |
| 25 | 13 | 29 | 65 | 107 | 150 | 176 | 335 | 1284 | 2509 | 4000 | 5290 | 6045 | 6397 | 7320 | 8242 |
| 26 | 12 | 29 | 64 | 107 | 137 | 177 | 323 | 1273 | 2497 | 4045 | 5340 | 6151 | 6749 | 7338 | 8259 |
| 27 | 12 | 29 | 64 | 107 | 137 | 178 | 303 | 1349 | 2786 | 4085 | 5387 | 6255 | 6793 | 7311 | 8010 |
| 28 | 12 | 29 | 63 | 107 | 137 | 179 | 283 | 1336 | 2544 | 4126 | 5707 | 6635 | 7120 | 7617 | 7979 |
| 29 | 12 | 29 | 62 | 95 | 137 | 180 | 262 | 1320 | 2579 | 4164 | 5750 | 6738 | 7244 | 7737 | 8341 |
| 30 | 12 | 29 | 62 | 94 | 136 | 180 | 240 | 1026 | 2611 | 4200 | 5790 | 6852 | 7354 | 7856 | 8387 |
| 31 | 12 | 29 | 61 | 94 | 138 | 178 | 192 | 1324 | 2917 | 4509 | 5824 | 7228 | 7739 | 8250 | 8771 |
| 32 | 12 | 28 | 61 | 93 | 137 | 178 | 208 | 1349 | 2945 | 4541 | 5859 | 7329 | 7848 | 8378 | 8896 |
| 33 | 12 | 28 | 60 | 92 | 136 | 178 | 208 | 1373 | 2971 | 4292 | 5891 | 7703 | 8243 | 8770 | 8987 |
| 34 | 12 | 28 | 47 | 91 | 135 | 166 | 207 | 1394 | 2908 | 4594 | 6196 | 7520 | 8349 | 8885 | 8999 |
| 35 | 12 | 28 | 47 | 90 | 133 | 165 | 164 | 1566 | 2913 | 4619 | 6224 | 7550 | 8455 | 8998 | 9041 |
| 36 | 12 | 28 | 46 | 89 | 132 | 164 | 169 | 1233 | 2915 | 4641 | 5970 | 7852 | 8835 | 9052 | 9053 |
| 37 | 12 | 28 | 45 | 88 | 119 | 162 | 172 | 1217 | 2915 | 4661 | 6164 | 7877 | 8949 | 9063 | 9065 |
| 38 | 12 | 27 | 44 | 87 | 118 | 161 | 176 | 1199 | 3115 | 4477 | 6162 | 7898 | 9052 | 9093 | 9107 |
| 39 | 11 | 27 | 44 | 86 | 117 | 159 | 195 | 1374 | 3093 | 4453 | 6157 | 7729 | 9162 | 9169 | 9118 |
| 40 | 11 | 27 | 43 | 85 | 115 | 157 | 194 | 1352 | 2709 | 4426 | 5979 | 7697 | 9185 | 9194 | 9210 |
| 41 | 11 | 27 | 42 | 84 | 114 | 156 | 173 | 1329 | 2683 | 4591 | 5947 | 7663 | 8757 | 9278 | 9283 |
| 42 | 11 | 22 | 41 | 71 | 113 | 145 | 173 | 1112 | 2848 | 4200 | 5912 | 7266 | 8732 | 9369 | 9325 |
| 43 | 11 | 22 | 40 | 70 | 111 | 144 | 174 | 1114 | 2817 | 4166 | 5874 | 7245 | 8643 | 9408 | 9427 |
| 44 | 11 | 22 | 40 | 69 | 110 | 143 | 173 | 1309 | 2425 | 4129 | 5669 | 7161 | 8193 | 9508 | 9521 |
| 45 | 11 | 22 | 39 | 68 | 108 | 142 | 178 | 1308 | 2390 | 4284 | 5626 | 7074 | 8160 | 9187 | 9624 |
| 46 | 11 | 21 | 38 | 66 | 95 | 140 | 161 | 1305 | 2352 | 3884 | 5581 | 6687 | 8064 | 9144 | 9665 |
| 47 | 11 | 21 | 37 | 65 | 93 | 139 | 159 | 1099 | 2507 | 3840 | 5588 | 6596 | 7606 | 9038 | 9759 |
| 48 | 10 | 21 | 36 | 64 | 92 | 138 | 156 | 1094 | 2465 | 3794 | 5143 | 6503 | 7566 | 8571 | 9862 |
| 49 | 10 | 21 | 31 | 63 | 90 | 127 | 152 | 1282 | 2240 | 3939 | 5056 | 6109 | 7463 | 8520 | 9519 |
| 50 | 10 | 21 | 31 | 61 | 89 | 126 | 148 | 1274 | 2231 | 3529 | 4966 | 6012 | 6998 | 8407 | 9398 |
| 51 | 10 | 20 | 30 | 60 | 87 | 124 | 159 | 1265 | 2221 | 3476 | 4641 | 5912 | 7031 | 8019 | 9428 |
| 52 | 10 | 20 | 30 | 47 | 86 | 123 | 129 | 1052 | 2403 | 3420 | 4564 | 5542 | 6941 | 7922 | 9020 |
| 53 | 10 | 20 | 30 | 46 | 84 | 121 | 123 | 1040 | 2187 | 3362 | 4486 | 5564 | 6546 | 7950 | 8936 |
| 54 | 10 | 20 | 29 | 44 | 74 | 120 | 116 | 1027 | 2172 | 3097 | 4405 | 5494 | 6471 | 7571 | 8558 |
| 55 | 10 | 19 | 29 | 43 | 72 | 106 | 109 | 1206 | 2156 | 3035 | 4033 | 5124 | 6515 | 7487 | 8439 |
| 56 | 10 | 19 | 29 | 42 | 71 | 105 | 141 | 987 | 2138 | 2972 | 4087 | 5050 | 6137 | 7225 | 8318 |
| 57 | 9 | 19 | 28 | 40 | 70 | 106 | 139 | 970 | 2110 | 2721 | 4018 | 5096 | 6055 | 7137 | 7920 |
| 58 | 9 | 19 | 28 | 39 | 68 | 104 | 137 | 951 | 2089 | 2662 | 3618 | 4720 | 5672 | 7046 | 7795 |
| 59 | 9 | 18 | 27 | 37 | 67 | 102 | 135 | 932 | 2067 | 2574 | 3537 | 4641 | 5708 | 6655 | 7370 |
| 60 | 9 | 18 | 27 | 36 | 65 | 101 | 54 | 910 | 2043 | 2506 | 3575 | 4559 | 5619 | 6565 | 7242 |
| 61 | 4 | 13 | 22 | 31 | 52 | 87 | 122 | 685 | 1815 | 2279 | 3168 | 4147 | 5230 | 6146 | 6838 |
| 62 | 4 | 13 | 22 | 30 | 50 | 85 | 120 | 855 | 1788 | 2217 | 3081 | 4174 | 5138 | 6023 | 6706 |
| 63 | 4 | 13 | 21 | 30 | 49 | 83 | 118 | 830 | 1757 | 2154 | 2993 | 4077 | 4841 | 5456 | 6180 |
| 64 | 4 | 12 | 21 | 29 | 47 | 81 | 115 | 803 | 1497 | 1888 | 2580 | 3656 | 4734 | 5223 | 6021 |
| 65 | 4 | 12 | 20 | 29 | 48 | 79 | 101 | 572 | 1439 | 1826 | 2488 | 3556 | 4625 | 5021 | 5245 |
| 66 | 4 | 12 | 20 | 28 | 47 | 65 | 99 | 543 | 1381 | 1761 | 2516 | 3131 | 4192 | 4443 | 4974 |
| 67 | 4 | 12 | 20 | 28 | 39 | 66 | 96 | 513 | 1323 | 1494 | 2097 | 3148 | 4079 | 4250 | 4240 |
| 68 | 3 | 11 | 19 | 27 | 39 | 64 | 94 | 481 | 1061 | 1428 | 2000 | 3042 | 3507 | 3755 | 4002 |
| 69 | 3 | 11 | 19 | 22 | 38 | 62 | 95 | 179 | 1002 | 1362 | 1550 | 2612 | 3350 | 3466 | 3407 |
| 70 | 3 | 11 | 18 | 21 | 37 | 60 | 80 | 405 | 942 | 1093 | 1463 | 2502 | 3193 | 2979 | 3115 |
| 71 | 3 | 10 | 13 | 21 | 36 | 52 | 78 | 370 | 681 | 1026 | 1375 | 2068 | 2713 | 2809 | 2586 |
| 72 | 3 | 10 | 13 | 20 | 35 | 51 | 75 | 113 | 621 | 960 | 1096 | 1955 | 2556 | 2317 | 2401 |
| 73 | 3 | 10 | 12 | 20 | 29 | 50 | 73 | 92 | 560 | 690 | 995 | 1840 | 2076 | 2147 | 1897 |
| 74 | 2 | 9 | 12 | 19 | 28 | 49 | 65 | 71 | 296 | 622 | 905 | 1523 | 1919 | 1655 | 1714 |
| 75 | 2 | 9 | 12 | 18 | 27 | 48 | 64 | 49 | 234 | 553 | 670 | 1405 | 1324 | 1370 | 1210 |
| 76 | 2 | 4 | 11 | 13 | 26 | 41 | 62 | 72 | 172 | 281 | 594 | 963 | 1167 | 878 | 912 |
| 77 | 2 | 4 | 11 | 13 | 25 | 40 | 61 | 69 | 110 | 212 | 518 | 842 | 687 | 708 | 584 |
| 78 | 2 | 4 | 10 | 12 | 24 | 39 | 49 | 67 | 78 | 142 | 238 | 720 | 531 | 395 | 468 |
| 79 | 2 | 3 | 10 | 11 | 17 | 38 | 48 | 64 | 74 | 87 | 161 | 262 | 375 | 287 | 254 |
| 80 | 2 | 3 | 9 | 11 | 16 | 37 | 46 | 51 | 70 | 83 | 96 | 157 | 224 | 150 | 191 |
| 81 | 1 | 3 | 4 | 10 | 15 | 27 | 44 | 48 | 52 | 79 | 91 | 101 | 149 | 126 | 111 |
| 82 | 1 | 3 | 4 | 10 | 14 | 26 | 32 | 45 | 49 | 61 | 87 | 95 | 96 | 87 | 95 |
| 83 | 1 | 2 | 3 | 4 | 13 | 24 | 29 | 32 | 45 | 57 | 69 | 75 | 90 | 82 | 73 |
| 84 | 1 | 2 | 3 | 4 | 12 | 23 | 26 | 29 | 31 | 53 | 64 | 68 | 71 | 65 | 67 |
| 85 | 1 | 2 | 2 | 3 | 5 | 22 | 24 | 26 | 28 | 35 | 43 | 62 | 64 | 60 | 51 |
| 86 | 1 | 1 | 2 | 3 | 4 | 8 | 11 | 23 | 24 | 31 | 38 | 40 | 43 | 44 | 45 |
| 87 | 0 | 1 | 1 | 2 | 3 | 6 | 8 | 9 | 10 | 26 | 32 | 34 | 36 | 37 | 28 |
| 88 | 0 | 1 | 1 | 1 | 2 | 4 | 5 | 6 | 7 | 8 | 12 | 13 | 14 | 15 | 22 |
| 89 | 0 | 0 | 0 | 1 | 1 | 2 | 3 | 3 | 3 | 4 | 6 | 6 | 7 | 7 | 7 |
| 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Table 17b.

Relative Spatial Response of ULIS (times 10,000)

| AZ | ZA = 16 | ZA = 17 | ZA = 18 | ZA = 19 | ZA = 20 | ZA = 21 | ZA = 22 | ZA = 23 | ZA = 24 | ZA = 25 | ZA = 26 | ZA = 27 | ZA = 28 | ZA = 29 | ZA = 30 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| -89 | 0 | 0 | 0 | 1 | 6 | 4 | 4 | 4 | 5 | 5 | 4 | 9 | 9 | 10 | 10 |
| -88 | 2 | 2 | 2 | 5 | 15 | 15 | 13 | 14 | 14 | 15 | 14 | 18 | 19 | 19 | 20 |
| -87 | 2 | 2 | 2 | 13 | 27 | 28 | 25 | 23 | 24 | 25 | 24 | 28 | 28 | 29 | 29 |
| -86 | 2 | 3 | 8 | 18 | 32 | 41 | 38 | 35 | 34 | 34 | 34 | 43 | 42 | 57 | 65 |
| -85 | 4 | 6 | 16 | 30 | 45 | 47 | 55 | 71 | 69 | 87 | 101 | 185 | 175 | 157 | 167 |
| -84 | 8 | 9 | 20 | 56 | 92 | 115 | 135 | 254 | 244 | 223 | 240 | 319 | 275 | 256 | 269 |
| -83 | 14 | 29 | 57 | 107 | 249 | 290 | 313 | 337 | 380 | 359 | 379 | 454 | 375 | 356 | 415 |
| -82 | 45 | 62 | 159 | 234 | 320 | 464 | 492 | 501 | 516 | 495 | 517 | 552 | 476 | 456 | 512 |
| -81 | 128 | 162 | 224 | 407 | 489 | 539 | 670 | 674 | 582 | 630 | 644 | 651 | 577 | 555 | 585 |
| -80 | 163 | 273 | 379 | 472 | 657 | 714 | 844 | 807 | 719 | 766 | 775 | 750 | 678 | 646 | 672 |
| -79 | 284 | 337 | 548 | 637 | 727 | 887 | 913 | 941 | 856 | 898 | 895 | 872 | 803 | 782 | 782 |
| -78 | 348 | 508 | 609 | 801 | 894 | 1061 | 1082 | 1105 | 1022 | 1025 | 992 | 980 | 962 | 1035 | 992 |
| -77 | 525 | 566 | 769 | 854 | 1050 | 1136 | 1248 | 1168 | 1159 | 1081 | 1138 | 1266 | 1243 | 1297 | 1207 |
| -76 | 598 | 727 | 818 | 1017 | 1217 | 1309 | 1379 | 1302 | 1322 | 1285 | 1418 | 1552 | 1529 | 1558 | 1481 |
| -75 | 898 | 1010 | 1000 | 1178 | 1284 | 1469 | 1440 | 1482 | 1613 | 1559 | 1697 | 1836 | 1829 | 1863 | 1787 |
| -74 | 1162 | 1121 | 1298 | 1285 | 1475 | 1649 | 1631 | 1770 | 1747 | 1876 | 1999 | 2003 | 2086 | 2126 | 2176 |
| -73 | 1268 | 1419 | 1427 | 1616 | 1825 | 1769 | 1942 | 2086 | 2066 | 2153 | 2277 | 2273 | 2342 | 2388 | 2442 |
| -72 | 1561 | 1552 | 1773 | 1995 | 2031 | 2123 | 2222 | 2217 | 2354 | 2430 | 2554 | 2549 | 2599 | 2692 | 2729 |
| -71 | 1679 | 1914 | 2182 | 2199 | 2388 | 2407 | 2347 | 2503 | 2624 | 2707 | 2854 | 2843 | 2898 | 2909 | 2824 |
| -70 | 1998 | 2078 | 2336 | 2585 | 2700 | 2536 | 2626 | 2785 | 2938 | 3027 | 2988 | 3094 | 3093 | 2977 | 2920 |
| -69 | 2152 | 2442 | 2722 | 2937 | 2823 | 2821 | 2934 | 3096 | 3209 | 3161 | 3252 | 3275 | 3198 | 3089 | 3039 |
| -68 | 2503 | 2824 | 2867 | 3059 | 3174 | 3179 | 3213 | 3380 | 3337 | 3437 | 3499 | 3379 | 3261 | 3157 | 3130 |
| -67 | 2635 | 2960 | 3237 | 3443 | 3452 | 3463 | 3335 | 3502 | 3650 | 3678 | 3559 | 3441 | 3324 | 3246 | 3249 |
| -66 | 2999 | 3338 | 3612 | 3559 | 3575 | 3592 | 3612 | 3809 | 3856 | 3738 | 3620 | 3503 | 3431 | 3363 | 3340 |
| -65 | 3144 | 3456 | 3741 | 3831 | 3853 | 3950 | 3978 | 4031 | 3914 | 3797 | 3724 | 3608 | 3500 | 3452 | 3459 |
| -64 | 3458 | 3776 | 4074 | 4101 | 4204 | 4227 | 4215 | 4088 | 4022 | 3900 | 3786 | 3671 | 3611 | 3569 | 3529 |
| -63 | 3555 | 3880 | 4183 | 4280 | 4311 | 4314 | 4290 | 4136 | 4048 | 3937 | 3867 | 3778 | 3676 | 3646 | 3547 |
| -62 | 3825 | 4163 | 4439 | 4523 | 4512 | 4494 | 4325 | 4163 | 4074 | 4014 | 3928 | 3818 | 3789 | 3664 | 3538 |
| -61 | 3924 | 4260 | 4528 | 4601 | 4695 | 4542 | 4360 | 4242 | 4132 | 4044 | 3956 | 3898 | 3792 | 3655 | 3557 |
| -60 | 4195 | 4524 | 4734 | 4794 | 4724 | 4564 | 4437 | 4270 | 4161 | 4103 | 4016 | 3916 | 3812 | 3686 | 3550 |
| -59 | 4284 | 4609 | 4936 | 4919 | 4768 | 4645 | 4473 | 4307 | 4210 | 4133 | 4042 | 3937 | 3805 | 3679 | 3570 |
| -58 | 4553 | 4878 | 4989 | 4947 | 4844 | 4670 | 4510 | 4396 | 4239 | 4154 | 4076 | 3929 | 3827 | 3699 | 3563 |
| -57 | 4648 | 4958 | 5111 | 5035 | 4883 | 4752 | 4589 | 4434 | 4326 | 4202 | 4068 | 3951 | 3821 | 3721 | 3597 |
| -56 | 4910 | 5167 | 5179 | 5059 | 4964 | 4789 | 4626 | 4515 | 4349 | 4195 | 4090 | 3957 | 3844 | 3716 | 3619 |
| -55 | 4994 | 5238 | 5221 | 5130 | 4987 | 4816 | 4704 | 4530 | 4391 | 4230 | 4084 | 3980 | 3851 | 3740 | 3613 |
| -54 | 5252 | 5362 | 5283 | 5147 | 5012 | 4886 | 4716 | 4552 | 4393 | 4263 | 4119 | 3974 | 3875 | 3735 | 3632 |
| -53 | 5318 | 5385 | 5303 | 5164 | 5058 | 4893 | 4747 | 4566 | 4443 | 4262 | 4143 | 3998 | 3871 | 3769 | 3627 |
| -52 | 5541 | 5453 | 5315 | 5209 | 5062 | 4942 | 4750 | 4609 | 4443 | 4301 | 4138 | 4035 | 3895 | 3789 | 3648 |
| -51 | 5601 | 5465 | 5352 | 5213 | 5108 | 4921 | 4776 | 4621 | 4482 | 4315 | 4174 | 4031 | 3915 | 3781 | 3680 |
| -50 | 5614 | 5501 | 5363 | 5259 | 5108 | 4965 | 4782 | 4662 | 4495 | 4356 | 4188 | 4056 | 3922 | 3802 | 3677 |
| -49 | 5615 | 5511 | 5401 | 5263 | 5152 | 4950 | 4835 | 4663 | 4536 | 4352 | 4218 | 4063 | 3943 | 3838 | 3699 |
| -48 | 5664 | 5515 | 5412 | 5275 | 5147 | 5004 | 4841 | 4708 | 4519 | 4385 | 4197 | 4081 | 3965 | 3832 | 3721 |
| -47 | 5669 | 5561 | 5424 | 5315 | 5193 | 5011 | 4879 | 4686 | 4552 | 4363 | 4233 | 4099 | 3973 | 3854 | 3756 |
| -46 | 5719 | 5572 | 5464 | 5328 | 5181 | 5060 | 4857 | 4719 | 4530 | 4397 | 4252 | 4106 | 3993 | 3877 | 3754 |
| -45 | 5731 | 5615 | 5476 | 5363 | 5236 | 5043 | 4897 | 4698 | 4564 | 4409 | 4240 | 4126 | 4012 | 3888 | 3776 |
| -44 | 5736 | 5628 | 5516 | 5372 | 5231 | 5083 | 4876 | 4736 | 4574 | 4398 | 4279 | 4145 | 4021 | 3908 | 3789 |
| -43 | 5787 | 5635 | 5529 | 5409 | 5271 | 5062 | 4916 | 4751 | 4555 | 4436 | 4299 | 4154 | 4041 | 3928 | 3817 |
| -42 | 5793 | 5687 | 5574 | 5419 | 5250 | 5102 | 4897 | 4734 | 4598 | 4457 | 4308 | 4174 | 4062 | 3959 | 3803 |
| -41 | 5845 | 5701 | 5589 | 5438 | 5291 | 5082 | 4938 | 4774 | 4623 | 4466 | 4329 | 4195 | 4099 | 3942 | 3812 |
| -40 | 5859 | 5747 | 5643 | 5444 | 5271 | 5124 | 4954 | 4758 | 4634 | 4491 | 4370 | 4237 | 4092 | 3951 | 3839 |
| -39 | 5904 | 5762 | 5644 | 5488 | 5313 | 5105 | 4962 | 4802 | 4660 | 4537 | 4362 | 4228 | 4124 | 3984 | 3848 |
| -38 | 5918 | 5808 | 5690 | 5470 | 5294 | 5148 | 4979 | 4847 | 4672 | 4531 | 4389 | 4271 | 4134 | 3994 | 3881 |
| -37 | 5933 | 5824 | 5670 | 5515 | 5339 | 5165 | 4988 | 4840 | 4699 | 4577 | 4436 | 4287 | 4145 | 4028 | 3865 |
| -36 | 5978 | 5878 | 5718 | 5498 | 5349 | 5174 | 5008 | 4886 | 4746 | 4597 | 4448 | 4324 | 4153 | 4014 | 3876 |
| -35 | 5994 | 5884 | 5727 | 5544 | 5369 | 5221 | 5054 | 4906 | 4757 | 4605 | 4456 | 4311 | 4168 | 4049 | 3912 |
| -34 | 5941 | 5932 | 5747 | 5592 | 5380 | 5206 | 5038 | 4914 | 4766 | 4617 | 4469 | 4349 | 4207 | 4065 | 3925 |
| -33 | 6056 | 5941 | 5757 | 5574 | 5428 | 5254 | 5073 | 4926 | 4802 | 4654 | 4507 | 4365 | 4223 | 4104 | 3962 |
| -32 | 5990 | 5962 | 5806 | 5624 | 5449 | 5276 | 5119 | 4962 | 4815 | 4667 | 4546 | 4404 | 4263 | 4114 | 3965 |
| -31 | 5998 | 5972 | 5789 | 5636 | 5461 | 5286 | 5093 | 4971 | 4824 | 4677 | 4534 | 4393 | 4266 | 4144 | 3999 |
| -30 | 6083 | 6022 | 5840 | 5658 | 5511 | 5334 | 5142 | 4984 | 4838 | 4716 | 4573 | 4423 | 4274 | 4124 | 4011 |
| -29 | 6180 | 6033 | 5852 | 5671 | 5497 | 5312 | 5186 | 5021 | 4876 | 4730 | 4575 | 4426 | 4304 | 4156 | 4022 |
| -28 | 6226 | 6055 | 5874 | 5722 | 5547 | 5361 | 5168 | 5008 | 4887 | 4764 | 4606 | 4457 | 4309 | 4161 | 4035 |
| -27 | 6237 | 6067 | 5888 | 5746 | 5598 | 5410 | 5215 | 5046 | 4901 | 4743 | 4638 | 4489 | 4341 | 4196 | 4069 |
| -26 | 6282 | 6118 | 5939 | 5761 | 5585 | 5390 | 5266 | 5085 | 4941 | 4780 | 4620 | 4495 | 4347 | 4230 | 4095 |
| -25 | 6269 | 6132 | 5954 | 5814 | 5634 | 5441 | 5247 | 5073 | 4981 | 4819 | 4653 | 4505 | 4381 | 4229 | 4096 |
| -24 | 6315 | 6155 | 5978 | 5829 | 5687 | 5492 | 5299 | 5112 | 4968 | 4826 | 4660 | 4539 | 4392 | 4261 | 4122 |
| -23 | 6328 | 6169 | 6031 | 5854 | 5657 | 5507 | 5351 | 5165 | 5007 | 4838 | 4699 | 4547 | 4399 | 4263 | 4148 |
| -22 | 6375 | 6217 | 6046 | 5871 | 5710 | 5517 | 5334 | 5213 | 5046 | 4878 | 4738 | 4582 | 4433 | 4295 | 4149 |
| -21 | 6388 | 6231 | 6100 | 5925 | 5723 | 5561 | 5380 | 5199 | 5055 | 4887 | 4751 | 4618 | 4468 | 4307 | 4177 |
| -20 | 6411 | 6255 | 6116 | 5981 | 5775 | 5574 | 5425 | 5244 | 5091 | 4928 | 4762 | 4627 | 4473 | 4312 | 4211 |
| -19 | 6424 | 6305 | 6143 | 5969 | 5829 | 5619 | 5439 | 5290 | 5101 | 4965 | 4804 | 4668 | 4509 | 4349 | 4215 |
| -18 | 6473 | 6320 | 6198 | 6023 | 5805 | 5631 | 5452 | 5337 | 5146 | 4976 | 4841 | 4683 | 4546 | 4386 | 4250 |
| -17 | 6487 | 6370 | 6215 | 6077 | 5860 | 5646 | 5498 | 5318 | 5151 | 4980 | 4845 | 4718 | 4576 | 4424 | 4288 |
| -16 | 6536 | 6385 | 6267 | 6094 | 5916 | 5703 | 5545 | 5365 | 5196 | 5016 | 4881 | 4722 | 4572 | 4422 | 4290 |
| -15 | 6551 | 6435 | 6284 | 6149 | 5934 | 5758 | 5562 | 5412 | 5212 | 5052 | 4894 | 4758 | 4606 | 4449 | 4318 |
| -14 | 6580 | 6427 | 6337 | 6167 | 5991 | 5815 | 5610 | 5460 | 5261 | 5098 | 4931 | 4795 | 4618 | 4478 | 4346 |
| -13 | 6597 | 6483 | 6368 | 6183 | 6008 | 5796 | 5667 | 5477 | 5310 | 5143 | 4969 | 4832 | 4654 | 4507 | 4376 |
| -12 | 6653 | 6539 | 6388 | 6228 | 6028 | 5854 | 5684 | 5487 | 5321 | 5124 | 4976 | 4870 | 4692 | 4543 | 4406 |
| -11 | 6671 | 6558 | 6446 | 6241 | 6071 | 5902 | 5705 | 5537 | 5371 | 5174 | 5015 | 4871 | 4695 | 4544 | 4405 |
| -10 | 6727 | 6616 | 6466 | 6297 | 6114 | 5917 | 5749 | 5576 | 5391 | 5226 | 5065 | 4887 | 4734 | 4582 | 4436 |
| -9 | 6791 | 6635 | 6525 | 6350 | 6128 | 5960 | 5793 | 5588 | 5426 | 5265 | 5104 | 4926 | 4774 | 4622 | 4474 |
| -8 | 6811 | 6693 | 6584 | 6370 | 6171 | 6004 | 5837 | 5632 | 5462 | 5301 | 5140 | 4953 | 4803 | 4654 | 4504 |
| -7 | 6874 | 6713 | 6605 | 6428 | 6227 | 6021 | 5854 | 5675 | 5498 | 5337 | 5155 | 4988 | 4830 | 4681 | 4532 |
| -6 | 6897 | 6772 | 6665 | 6487 | 6246 | 6065 | 5900 | 5721 | 5541 | 5353 | 5193 | 5023 | 4857 | 4708 | 4560 |
| -5 | 6964 | 6837 | 6687 | 6509 | 6303 | 6111 | 5946 | 5768 | 5558 | 5390 | 5231 | 5062 | 4893 | 4737 | 4588 |
| -4 | 6987 | 6860 | 6747 | 6569 | 6364 | 6168 | 5993 | 5787 | 5604 | 5428 | 5270 | 5101 | 4928 | 4765 | 4617 |
| -3 | 7054 | 6925 | 6808 | 6630 | 6388 | 6190 | 6012 | 5835 | 5652 | 5474 | 5309 | 5141 | 4968 | 4801 | 4647 |
| -2 | 7078 | 6950 | 6830 | 6653 | 6451 | 6248 | 6060 | 5883 | 5702 | 5520 | 5349 | 5181 | 5009 | 4837 | 4677 |
| -1 | 7146 | 7020 | 6899 | 6715 | 6514 | 6312 | 6119 | 5932 | 5751 | 5571 | 5396 | 5222 | 5051 | 4879 | 4714 |
| 0 | 7139 | 7015 | 6890 | 6690 | 6489 | 6289 | 6088 | 5909 | 5729 | 5549 | 5369 | 5199 | 5028 | 4858 | 4687 |
| 1 | 7225 | 7101 | 6978 | 6788 | 6589 | 6389 | 6190 | 5991 | 5812 | 5633 | 5453 | 5270 | 5100 | 4930 | 4760 |
| 2 | 7279 | 7157 | 7035 | 6844 | 6639 | 6440 | 6242 | 6039 | 5850 | 5672 | 5493 | 5307 | 5129 | 4960 | 4790 |
| 3 | 7299 | 7178 | 7057 | 6866 | 6689 | 6492 | 6294 | 6087 | 5889 | 5711 | 5533 | 5345 | 5160 | 4990 | 4821 |
| 4 | 7354 | 7235 | 7115 | 6923 | 6711 | 6515 | 6318 | 6141 | 5938 | 5751 | 5574 | 5387 | 5199 | 5022 | 4853 |
| 5 | 7375 | 7256 | 7173 | 6983 | 6770 | 6567 | 6372 | 6167 | 5987 | 5792 | 5616 | 5429 | 5239 | 5054 | 4886 |
| 6 | 7450 | 7313 | 7196 | 7009 | 6828 | 6621 | 6426 | 6223 | 6015 | 5843 | 5658 | 5473 | 5283 | 5095 | 4919 |
| 7 | 7475 | 7370 | 7254 | 7071 | 6856 | 6681 | 6481 | 6279 | 6072 | 5869 | 5701 | 5516 | 5327 | 5138 | 4952 |
| 8 | 7550 | 7393 | 7278 | 7133 | 6919 | 6708 | 6507 | 6336 | 6130 | 5924 | 5723 | 5561 | 5372 | 5184 | 4997 |
| 9 | 7581 | 7471 | 7337 | 7160 | 6983 | 6771 | 6563 | 6393 | 6189 | 5984 | 5782 | 5594 | 5403 | 5213 | 5022 |
| 10 | 7663 | 7499 | 7396 | 7223 | 7013 | 6837 | 6629 | 6422 | 6230 | 6022 | 5813 | 5624 | 5430 | 5239 | 5049 |
| 11 | 7745 | 7578 | 7421 | 7287 | 7078 | 6868 | 6678 | 6456 | 6248 | 6054 | 5846 | 5642 | 5457 | 5267 | 5077 |
| 12 | 7778 | 7613 | 7505 | 7316 | 7128 | 6915 | 6701 | 6500 | 6280 | 6086 | 5879 | 5727 | 5540 | 5344 | 5154 |
| 13 | 7861 | 7698 | 7536 | 7362 | 7150 | 6959 | 6803 | 6606 | 6377 | 6160 | 5966 | 5761 | 5575 | 5377 | 5184 |
| 14 | 7893 | 7757 | 7632 | 7472 | 7250 | 7039 | 6849 | 6633 | 6420 | 6193 | 6001 | 5796 | 5610 | 5411 | 5214 |
| 15 | 7996 | 7825 | 7682 | 7499 | 7301 | 7084 | 6895 | 6682 | 6470 | 6240 | 6036 | 5831 | 5633 | 5448 | 5250 |
| 16 | 8017 | 7875 | 7705 | 7553 | 7327 | 7131 | 6921 | 6731 | 6499 | 6288 | 6059 | 5867 | 5670 | 5486 | 5288 |
| 17 | 8065 | 7897 | 7756 | 7581 | 7381 | 7185 | 6969 | 6781 | 6550 | 6341 | 6110 | 5967 | 5759 | 5562 | 5359 |
| 18 | 8087 | 7947 | 7780 | 7636 | 7438 | 7214 | 7018 | 6867 | 6664 | 6455 | 6225 | 6002 | 5780 | 5580 | 5378 |
| 19 | 8161 | 7970 | 7832 | 7692 | 7529 | 7332 | 7135 | 6917 | 6717 | 6488 | 6262 | 6026 | 5809 | 5599 | 5397 |
| 20 | 8190 | 8021 | 7930 | 7766 | 7587 | 7392 | 7169 | 6976 | 6749 | 6522 | 6295 | 6059 | 5839 | 5629 | 5469 |
| 21 | 8329 | 8156 | 7956 | 7819 | 7646 | 7424 | 7230 | 7014 | 6781 | 6555 | 6391 | 6155 | 5932 | 5719 | 5500 |
| 22 | 8364 | 8193 | 8032 | 7873 | 7678 | 7485 | 7270 | 7098 | 6872 | 6649 | 6416 | 6191 | 5967 | 5746 | 5533 |
| 23 | 8444 | 8275 | 8065 | 7901 | 7737 | 7586 | 7354 | 7128 | 6904 | 6682 | 6452 | 6227 | 6003 | 5782 | 5629 |
| 24 | 8480 | 8313 | 8146 | 8042 | 7832 | 7608 | 7384 | 7160 | 6936 | 6707 | 6486 | 6324 | 6101 | 5878 | 5667 |
| 25 | 8560 | 8459 | 8294 | 8084 | 7871 | 7638 | 7415 | 7192 | 7022 | 6801 | 6581 | 6360 | 6131 | 5916 | 5690 |
| 26 | 8661 | 8498 | 8335 | 8130 | 7911 | 7679 | 7501 | 7278 | 7056 | 6835 | 6616 | 6389 | 6169 | 5992 | 5766 |
| 27 | 8742 | 8582 | 8387 | 8156 | 8003 | 7774 | 7543 | 7312 | 7090 | 6870 | 6705 | 6486 | 6247 | 6007 | 5782 |
| 28 | 8779 | 8621 | 8413 | 8261 | 8046 | 7819 | 7586 | 7356 | 7194 | 6959 | 6741 | 6503 | 6262 | 6021 | 5797 |
| 29 | 8861 | 8731 | 8517 | 8303 | 8089 | 7865 | 7687 | 7455 | 7226 | 6995 | 6759 | 6519 | 6338 | 6097 | 5868 |
| 30 | 8990 | 8756 | 8544 | 8331 | 8182 | 7965 | 7735 | 7505 | 7274 | 7088 | 6840 | 6596 | 6355 | 6115 | 5889 |
| 31 | 8941 | 8804 | 8585 | 8438 | 8226 | 8012 | 7784 | 7623 | 7374 | 7112 | 6861 | 6618 | 6378 | 6198 | 5961 |
| 32 | 8953 | 8832 | 8685 | 8467 | 8257 | 8046 | 7887 | 7659 | 7400 | 7140 | 6951 | 6706 | 6463 | 6218 | 5979 |
| 33 | 8985 | 8951 | 8734 | 8519 | 8373 | 8157 | 7937 | 7685 | 7494 | 7234 | 6976 | 6733 | 6489 | 6308 | 6064 |
| 34 | 9009 | 8982 | 8767 | 8621 | 8425 | 8212 | 7969 | 7779 | 7520 | 7262 | 7002 | 6823 | 6580 | 6336 | 6072 |
| 35 | 9041 | 9034 | 8889 | 8676 | 8462 | 8317 | 8059 | 7807 | 7549 | 7358 | 7100 | 6851 | 6609 | 6412 | 6140 |
| 36 | 9053 | 9065 | 8924 | 8732 | 8588 | 8351 | 8081 | 7895 | 7642 | 7389 | 7132 | 6945 | 6688 | 6417 | 6146 |
| 37 | 9096 | 9098 | 8978 | 8837 | 8627 | 8380 | 8180 | 7917 | 7666 | 7480 | 7227 | 6975 | 6710 | 6503 | 6216 |
| 38 | 9109 | 9111 | 9082 | 8894 | 8660 | 8479 | 8211 | 8010 | 7756 | 7505 | 7254 | 7045 | 6773 | 6510 | 6303 |
| 39 | 9121 | 9155 | 9137 | 9000 | 8758 | 8510 | 8314 | 8034 | 7785 | 7601 | 7331 | 7045 | 6830 | 6569 | 6305 |
| 40 | 9219 | 9168 | 9182 | 9036 | 8789 | 8610 | 8347 | 8059 | 7877 | 7628 | 7417 | 7132 | 6829 | 6567 | 6364 |
| 41 | 9247 | 9299 | 9274 | 9065 | 8888 | 8642 | 8451 | 8169 | 7908 | 7703 | 7419 | 7135 | 6915 | 6626 | 6364 |
| 42 | 9340 | 9346 | 9306 | 9156 | 8920 | 8744 | 8481 | 8201 | 8001 | 7725 | 7487 | 7204 | 6920 | 6706 | 6424 |
| 43 | 9437 | 9454 | 9391 | 9183 | 9020 | 8778 | 8592 | 8313 | 8005 | 7789 | 7512 | 7271 | 6991 | 6707 | 6506 |
| 44 | 9542 | 9554 | 9416 | 9288 | 9053 | 8881 | 8624 | 8421 | 8108 | 7789 | 7574 | 7295 | 7056 | 6775 | 6506 |
| 45 | 9639 | 9601 | 9500 | 9322 | 9139 | 8916 | 8737 | 8435 | 8189 | 7873 | 7573 | 7358 | 7080 | 6830 | 6536 |
| 46 | 9683 | 9692 | 9526 | 9416 | 9164 | 9011 | 8770 | 8539 | 8195 | 7949 | 7659 | 7421 | 7144 | 6850 | 6580 |
| 47 | 9788 | 9717 | 9611 | 9445 | 9278 | 9041 | 8885 | 8545 | 8301 | 7958 | 7723 | 7447 | 7199 | 6900 | 6663 |
| 48 | 9885 | 9748 | 9637 | 9532 | 9317 | 9128 | 8881 | 8643 | 8310 | 8066 | 7725 | 7512 | 7220 | 6952 | 6653 |
| 49 | 9935 | 9826 | 9664 | 9561 | 9431 | 9154 | 8962 | 8637 | 8401 | 8068 | 7830 | 7562 | 7273 | 7038 | 6707 |
| 50 | 9984 | 9852 | 9750 | 9642 | 9472 | 9276 | 8951 | 8726 | 8461 | 8160 | 7894 | 7581 | 7310 | 7018 | 6746 |
| 51 | 10000 | 9923 | 9778 | 9678 | 9564 | 9283 | 9025 | 8783 | 8485 | 8219 | 7920 | 7627 | 7325 | 7052 | 6818 |
| 52 | 9989 | 9944 | 9850 | 9760 | 9600 | 9391 | 9101 | 8807 | 8543 | 8245 | 7969 | 7715 | 7358 | 7125 | 6797 |
| 53 | 9687 | 9939 | 9874 | 9780 | 9690 | 9383 | 9138 | 8856 | 8633 | 8305 | 8057 | 7700 | 7407 | 7160 | 6832 |
| 54 | 9581 | 9931 | 9902 | 9835 | 9648 | 9459 | 9195 | 8878 | 8629 | 8398 | 8107 | 7751 | 7498 | 7141 | 6908 |
| 55 | 9184 | 9781 | 9890 | 9849 | 9668 | 9419 | 9218 | 8951 | 8720 | 8449 | 8094 | 7842 | 7540 | 7233 | 6930 |
| 56 | 9054 | 9679 | 9870 | 9834 | 9657 | 9401 | 9162 | 8973 | 8722 | 8422 | 8171 | 7873 | 7562 | 7260 | 7001 |
| 57 | 8625 | 9343 | 9840 | 9825 | 9609 | 9394 | 9155 | 8905 | 8719 | 8499 | 8198 | 7890 | 7587 | 7333 | 7022 |
| 58 | 8515 | 9206 | 9628 | 9741 | 9607 | 9328 | 9079 | 8898 | 8623 | 8423 | 8270 | 7914 | 7661 | 7355 | 6989 |
| 59 | 8082 | 8794 | 9547 | 9713 | 9531 | 9313 | 9073 | 8833 | 8604 | 8382 | 8240 | 7989 | 7688 | 7379 | 7065 |
| 60 | 7945 | 8649 | 9298 | 9486 | 9489 | 9263 | 9069 | 8827 | 8490 | 8316 | 8189 | 8016 | 7764 | 7455 | 7090 |
| 61 | 7508 | 8204 | 8892 | 9106 | 9461 | 9256 | 8994 | 8811 | 8473 | 8202 | 8073 | 7993 | 7776 | 7409 | 7149 |
| 62 | 7391 | 8054 | 8735 | 8965 | 9158 | 9232 | 8991 | 8713 | 8373 | 8128 | 8006 | 7876 | 7791 | 7467 | 7141 |
| 63 | 6862 | 7548 | 8283 | 8512 | 8716 | 8884 | 8987 | 8700 | 8359 | 8021 | 7940 | 7833 | 7573 | 7458 | 7195 |
| 64 | 6693 | 7367 | 7938 | 8156 | 8517 | 8726 | 8916 | 8592 | 8335 | 7988 | 7850 | 7649 | 7410 | 7323 | 7189 |
| 65 | 6027 | 6733 | 7304 | 7549 | 7758 | 8094 | 8294 | 8581 | 8238 | 7883 | 7785 | 7557 | 7198 | 7098 | 7068 |
| 66 | 5705 | 6485 | 7030 | 6914 | 7108 | 7306 | 7558 | 8034 | 8227 | 7884 | 7671 | 7360 | 7112 | 6935 | 6839 |
| 67 | 4887 | 5600 | 6213 | 6648 | 6828 | 7011 | 6893 | 7270 | 7736 | 7873 | 7607 | 7271 | 6913 | 6710 | 6673 |
| 68 | 4584 | 5289 | 5436 | 5876 | 6176 | 6345 | 6597 | 6960 | 6957 | 7328 | 7531 | 7184 | 6823 | 6547 | 6444 |
| 69 | 3837 | 4423 | 5114 | 5603 | 5391 | 5545 | 5930 | 6278 | 6631 | 6645 | 6975 | 6991 | 6735 | 6379 | 6279 |
| 70 | 3504 | 3638 | 4223 | 4865 | 5109 | 4877 | 5176 | 5510 | 5933 | 6315 | 6282 | 6508 | 6541 | 6169 | 6013 |
| 71 | 2816 | 3285 | 3875 | 3985 | 4460 | 4580 | 4507 | 4855 | 5209 | 5582 | 5936 | 5799 | 6015 | 6088 | 5834 |
| 72 | 2573 | 2575 | 3033 | 3553 | 3673 | 3912 | 4206 | 4199 | 4555 | 4914 | 5133 | 4982 | 5183 | 5496 | 5656 |
| 73 | 2012 | 2304 | 2369 | 2787 | 3254 | 3188 | 3572 | 3906 | 3899 | 4244 | 4437 | 4270 | 4533 | 4808 | 4971 |
| 74 | 1803 | 1711 | 2115 | 2249 | 2682 | 2910 | 2875 | 3200 | 3179 | 3573 | 3739 | 3636 | 3898 | 4198 | 4333 |
| 75 | 1261 | 1497 | 1653 | 2045 | 2269 | 2560 | 2422 | 2556 | 2870 | 2786 | 3012 | 3257 | 3299 | 3587 | 3402 |
| 76 | 791 | 1080 | 1274 | 1711 | 2150 | 2232 | 2268 | 2123 | 2211 | 2175 | 2400 | 2669 | 2649 | 2869 | 2664 |
| 77 | 666 | 776 | 1206 | 1414 | 1846 | 1904 | 1933 | 1780 | 1811 | 1746 | 1824 | 2080 | 2048 | 2243 | 2115 |
| 78 | 355 | 675 | 914 | 1339 | 1582 | 1696 | 1597 | 1619 | 1564 | 1628 | 1491 | 1491 | 1446 | 1616 | 1592 |
| 79 | 257 | 413 | 804 | 1040 | 1277 | 1368 | 1199 | 1302 | 1318 | 1341 | 1249 | 1227 | 1087 | 1126 | 1121 |
| 80 | 175 | 316 | 511 | 740 | 1153 | 1040 | 1046 | 1047 | 1095 | 1104 | 957 | 929 | 910 | 940 | 923 |
| 81 | 147 | 157 | 233 | 622 | 845 | 713 | 811 | 828 | 871 | 868 | 715 | 761 | 768 | 796 | 800 |
| 82 | 92 | 93 | 138 | 320 | 537 | 568 | 571 | 608 | 766 | 632 | 541 | 601 | 626 | 651 | 704 |
| 83 | 76 | 75 | 75 | 112 | 410 | 345 | 356 | 389 | 541 | 396 | 371 | 462 | 483 | 505 | 602 |
| 84 | 61 | 61 | 54 | 58 | 97 | 113 | 141 | 288 | 316 | 239 | 235 | 322 | 341 | 359 | 378 |
| 85 | 46 | 57 | 50 | 43 | 36 | 36 | 45 | 68 | 91 | 83 | 98 | 183 | 198 | 213 | 229 |
| 86 | 41 | 42 | 41 | 34 | 27 | 32 | 31 | 30 | 31 | 32 | 32 | 43 | 55 | 67 | 80 |
| 87 | 26 | 27 | 28 | 30 | 23 | 22 | 21 | 21 | 22 | 23 | 23 | 28 | 29 | 30 | 30 |
| 88 | 20 | 21 | 22 | 21 | 13 | 12 | 12 | 13 | 13 | 13 | 14 | 19 | 19 | 20 | 20 |
| 89 | 5 | 6 | 6 | 6 | 5 | 4 | 4 | 4 | 4 | 4 | 5 | 9 | 10 | 10 | 10 |
| 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Table 17c.

Relative Spatial Response of ULIS (times 10,000)

| AZ | ZA = 31 | ZA = 32 | ZA = 33 | ZA = 34 | ZA = 35 | ZA = 36 | ZA = 37 | ZA = 38 | ZA = 39 | ZA = 40 | ZA = 41 | ZA = 42 | ZA = 43 | ZA = 44 | ZA = 45 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| -89 | 10 | 10 | 10 | 11 | 11 | 11 | 11 | 12 | 14 | 20 | 13 | 14 | 14 | 14 | 15 |
| -88 | 20 | 20 | 21 | 21 | 22 | 22 | 28 | 34 | 42 | 50 | 39 | 33 | 33 | 34 | 34 |
| -87 | 30 | 30 | 31 | 37 | 40 | 43 | 46 | 55 | 63 | 81 | 70 | 52 | 52 | 53 | 54 |
| -86 | 74 | 70 | 71 | 80 | 83 | 87 | 87 | 84 | 93 | 102 | 93 | 83 | 71 | 72 | 74 |
| -85 | 179 | 194 | 144 | 122 | 127 | 150 | 133 | 105 | 123 | 132 | 124 | 107 | 98 | 92 | 93 |
| -84 | 318 | 269 | 221 | 205 | 189 | 195 | 198 | 172 | 145 | 162 | 156 | 139 | 131 | 114 | 109 |
| -83 | 392 | 345 | 331 | 284 | 235 | 257 | 245 | 220 | 213 | 186 | 192 | 198 | 204 | 264 | 182 |
| -82 | 466 | 421 | 408 | 363 | 351 | 304 | 292 | 310 | 313 | 387 | 352 | 424 | 434 | 486 | 415 |
| -81 | 541 | 530 | 486 | 478 | 465 | 467 | 510 | 516 | 601 | 575 | 574 | 645 | 659 | 692 | 608 |
| -80 | 630 | 645 | 634 | 722 | 713 | 781 | 704 | 792 | 869 | 795 | 791 | 807 | 878 | 882 | 803 |
| -79 | 871 | 864 | 955 | 950 | 1036 | 976 | 975 | 989 | 1027 | 1015 | 949 | 1027 | 1103 | 1054 | 1025 |
| -78 | 1086 | 1083 | 1178 | 1178 | 1260 | 1250 | 1167 | 1253 | 1271 | 1204 | 1196 | 1251 | 1258 | 1244 | 1219 |
| -77 | 1299 | 1398 | 1401 | 1500 | 1486 | 1481 | 1476 | 1435 | 1429 | 1423 | 1420 | 1411 | 1459 | 1435 | 1412 |
| -76 | 1566 | 1622 | 1638 | 1732 | 1756 | 1676 | 1664 | 1648 | 1645 | 1585 | 1584 | 1634 | 1675 | 1654 | 1634 |
| -75 | 1839 | 1892 | 1952 | 1947 | 1938 | 1908 | 1879 | 1833 | 1831 | 1832 | 1834 | 1846 | 1862 | 1842 | 1748 |
| -74 | 2111 | 2170 | 2209 | 2174 | 2121 | 2095 | 2069 | 2048 | 2050 | 2054 | 1999 | 2002 | 1997 | 1902 | 1817 |
| -73 | 2383 | 2429 | 2355 | 2328 | 2318 | 2247 | 2225 | 2205 | 2210 | 2244 | 2222 | 2124 | 2037 | 1942 | 1884 |
| -72 | 2648 | 2528 | 2462 | 2454 | 2447 | 2469 | 2464 | 2458 | 2450 | 2347 | 2244 | 2160 | 2104 | 2017 | 1982 |
| -71 | 2745 | 2628 | 2613 | 2608 | 2604 | 2600 | 2598 | 2569 | 2469 | 2394 | 2294 | 2213 | 2138 | 2059 | 2079 |
| -70 | 2883 | 2812 | 2789 | 2734 | 2733 | 2732 | 2707 | 2599 | 2516 | 2415 | 2324 | 2273 | 2201 | 2144 | 2085 |
| -69 | 2981 | 2914 | 2913 | 2939 | 2930 | 2816 | 2702 | 2607 | 2556 | 2463 | 2377 | 2304 | 2222 | 2164 | 2105 |
| -68 | 3079 | 3062 | 3063 | 3045 | 2931 | 2817 | 2730 | 2641 | 2572 | 2487 | 2436 | 2334 | 2237 | 2170 | 2125 |
| -67 | 3217 | 3183 | 3159 | 3073 | 2959 | 2846 | 2736 | 2656 | 2609 | 2535 | 2433 | 2331 | 2261 | 2194 | 2132 |
| -66 | 3320 | 3301 | 3188 | 3074 | 2961 | 2875 | 2762 | 2693 | 2621 | 2551 | 2454 | 2347 | 2270 | 2220 | 2152 |
| -65 | 3414 | 3302 | 3190 | 3104 | 2991 | 2872 | 2778 | 2724 | 2634 | 2540 | 2469 | 2371 | 2295 | 2226 | 2164 |
| -64 | 3405 | 3319 | 3212 | 3106 | 2988 | 2903 | 2808 | 2718 | 2641 | 2555 | 2468 | 2388 | 2319 | 2238 | 2170 |
| -63 | 3424 | 3314 | 3207 | 3114 | 3006 | 2900 | 2822 | 2725 | 2637 | 2570 | 2477 | 2395 | 2331 | 2243 | 2168 |
| -62 | 3419 | 3333 | 3225 | 3109 | 3026 | 2922 | 2818 | 2741 | 2652 | 2560 | 2493 | 2408 | 2320 | 2257 | 2176 |
| -61 | 3439 | 3329 | 3244 | 3129 | 3025 | 2944 | 2832 | 2737 | 2669 | 2576 | 2486 | 2413 | 2333 | 2263 | 2190 |
| -60 | 3460 | 3350 | 3240 | 3152 | 3047 | 2943 | 2856 | 2754 | 2660 | 2588 | 2499 | 2426 | 2347 | 2278 | 2195 |
| -59 | 3456 | 3347 | 3261 | 3146 | 3067 | 2956 | 2855 | 2776 | 2672 | 2601 | 2502 | 2440 | 2354 | 2276 | 2205 |
| -58 | 3478 | 3367 | 3254 | 3162 | 3061 | 2972 | 2871 | 2793 | 2685 | 2595 | 2516 | 2430 | 2367 | 2279 | 2217 |
| -57 | 3475 | 3385 | 3272 | 3180 | 3078 | 2967 | 2889 | 2780 | 2702 | 2599 | 2531 | 2444 | 2377 | 2289 | 2220 |
| -56 | 3493 | 3378 | 3290 | 3176 | 3095 | 2985 | 2907 | 2799 | 2721 | 2613 | 2546 | 2454 | 2378 | 2300 | 2232 |
| -55 | 3511 | 3397 | 3310 | 3195 | 3091 | 3004 | 2905 | 2818 | 2719 | 2632 | 2532 | 2465 | 2389 | 2312 | 2243 |
| -54 | 3517 | 3416 | 3307 | 3215 | 3110 | 3017 | 2920 | 2832 | 2735 | 2647 | 2543 | 2476 | 2401 | 2315 | 2255 |
| -53 | 3537 | 3411 | 3327 | 3212 | 3123 | 3031 | 2935 | 2847 | 2750 | 2665 | 2561 | 2488 | 2396 | 2328 | 2256 |
| -52 | 3557 | 3432 | 3348 | 3226 | 3137 | 3026 | 2950 | 2842 | 2766 | 2683 | 2579 | 2495 | 2408 | 2338 | 2267 |
| -51 | 3554 | 3464 | 3344 | 3240 | 3152 | 3041 | 2965 | 2858 | 2775 | 2701 | 2598 | 2511 | 2420 | 2343 | 2268 |
| -50 | 3575 | 3485 | 3359 | 3269 | 3143 | 3057 | 2981 | 2872 | 2798 | 2712 | 2608 | 2522 | 2447 | 2344 | 2268 |
| -49 | 3598 | 3481 | 3374 | 3284 | 3159 | 3083 | 2974 | 2880 | 2808 | 2705 | 2619 | 2532 | 2458 | 2354 | 2269 |
| -48 | 3607 | 3493 | 3401 | 3277 | 3175 | 3100 | 2983 | 2889 | 2818 | 2716 | 2630 | 2544 | 2470 | 2366 | 2279 |
| -47 | 3627 | 3505 | 3413 | 3293 | 3192 | 3114 | 2992 | 2898 | 2827 | 2728 | 2642 | 2549 | 2473 | 2366 | 2276 |
| -46 | 3639 | 3534 | 3406 | 3307 | 3224 | 3103 | 3001 | 2919 | 2837 | 2736 | 2637 | 2544 | 2468 | 2375 | 2280 |
| -45 | 3653 | 3524 | 3419 | 3335 | 3236 | 3113 | 3011 | 2929 | 2847 | 2735 | 2636 | 2556 | 2463 | 2373 | 2285 |
| -44 | 3684 | 3538 | 3433 | 3349 | 3242 | 3124 | 3037 | 2940 | 2856 | 2736 | 2637 | 2556 | 2459 | 2379 | 2304 |
| -43 | 3674 | 3553 | 3462 | 3362 | 3227 | 3146 | 3049 | 2951 | 2856 | 2737 | 2654 | 2557 | 2462 | 2384 | 2310 |
| -42 | 3689 | 3585 | 3477 | 3369 | 3254 | 3153 | 3057 | 2961 | 2858 | 2773 | 2656 | 2559 | 2462 | 2386 | 2316 |
| -41 | 3717 | 3601 | 3491 | 3375 | 3262 | 3161 | 3065 | 2961 | 2878 | 2775 | 2674 | 2561 | 2463 | 2386 | 2318 |
| -40 | 3726 | 3613 | 3518 | 3383 | 3271 | 3170 | 3089 | 2981 | 2879 | 2778 | 2677 | 2579 | 2465 | 2398 | 2331 |
| -39 | 3735 | 3641 | 3506 | 3392 | 3280 | 3194 | 3095 | 2982 | 2879 | 2780 | 2680 | 2599 | 2500 | 2410 | 2342 |
| -38 | 3763 | 3628 | 3510 | 3420 | 3307 | 3203 | 3096 | 2983 | 2880 | 2798 | 2700 | 2602 | 2503 | 2411 | 2342 |
| -37 | 3753 | 3633 | 3533 | 3426 | 3318 | 3212 | 3098 | 3004 | 2899 | 2800 | 2700 | 2603 | 2507 | 2413 | 2338 |
| -36 | 3782 | 3661 | 3538 | 3431 | 3324 | 3234 | 3121 | 3007 | 2901 | 2802 | 2701 | 2604 | 2509 | 2435 | 2350 |
| -35 | 3792 | 3667 | 3543 | 3455 | 3349 | 3234 | 3124 | 3010 | 2903 | 2803 | 2721 | 2623 | 2528 | 2446 | 2357 |
| -34 | 3820 | 3695 | 3572 | 3461 | 3354 | 3232 | 3144 | 3034 | 2924 | 2824 | 2723 | 2639 | 2545 | 2446 | 2354 |
| -33 | 3827 | 3703 | 3579 | 3486 | 3377 | 3255 | 3143 | 3035 | 2927 | 2841 | 2741 | 2641 | 2547 | 2445 | 2368 |
| -32 | 3837 | 3732 | 3608 | 3493 | 3377 | 3256 | 3143 | 3035 | 2961 | 2859 | 2763 | 2663 | 2567 | 2465 | 2366 |
| -31 | 3871 | 3743 | 3617 | 3501 | 3401 | 3279 | 3164 | 3071 | 2959 | 2855 | 2761 | 2667 | 2567 | 2480 | 2380 |
| -30 | 3905 | 3778 | 3650 | 3525 | 3403 | 3281 | 3165 | 3072 | 2957 | 2851 | 2776 | 2683 | 2595 | 2500 | 2399 |
| -29 | 3918 | 3792 | 3658 | 3529 | 3428 | 3307 | 3205 | 3093 | 2979 | 2867 | 2774 | 2690 | 2583 | 2493 | 2402 |
| -28 | 3953 | 3819 | 3684 | 3556 | 3433 | 3310 | 3208 | 3093 | 2978 | 2864 | 2783 | 2704 | 2596 | 2501 | 2418 |
| -27 | 3934 | 3819 | 3685 | 3562 | 3462 | 3358 | 3235 | 3115 | 3001 | 2902 | 2800 | 2696 | 2586 | 2498 | 2408 |
| -26 | 3961 | 3846 | 3713 | 3590 | 3468 | 3356 | 3221 | 3111 | 3001 | 2903 | 2798 | 2711 | 2612 | 2507 | 2417 |
| -25 | 3962 | 3847 | 3742 | 3620 | 3495 | 3370 | 3239 | 3129 | 3033 | 2923 | 2813 | 2715 | 2605 | 2497 | 2416 |
| -24 | 3989 | 3875 | 3749 | 3614 | 3492 | 3357 | 3235 | 3126 | 3052 | 2942 | 2845 | 2733 | 2622 | 2524 | 2427 |
| -23 | 4016 | 3902 | 3765 | 3630 | 3508 | 3372 | 3254 | 3158 | 3048 | 2939 | 2841 | 2730 | 2631 | 2520 | 2419 |
| -22 | 4017 | 3903 | 3757 | 3617 | 3524 | 3389 | 3273 | 3177 | 3068 | 2971 | 2861 | 2750 | 2651 | 2536 | 2439 |
| -21 | 4045 | 3930 | 3781 | 3640 | 3516 | 3379 | 3283 | 3174 | 3065 | 2969 | 2858 | 2758 | 2642 | 2538 | 2431 |
| -20 | 4073 | 3957 | 3807 | 3665 | 3541 | 3404 | 3298 | 3194 | 3099 | 2990 | 2889 | 2774 | 2670 | 2555 | 2459 |
| -19 | 4076 | 3953 | 3804 | 3677 | 3537 | 3426 | 3310 | 3208 | 3114 | 3010 | 2905 | 2790 | 2686 | 2573 | 2477 |
| -18 | 4105 | 3981 | 3832 | 3703 | 3563 | 3438 | 3301 | 3207 | 3105 | 3006 | 2894 | 2791 | 2679 | 2579 | 2473 |
| -17 | 4135 | 4010 | 3861 | 3731 | 3590 | 3465 | 3325 | 3220 | 3118 | 3019 | 2907 | 2804 | 2693 | 2595 | 2500 |
| -16 | 4139 | 4039 | 3890 | 3760 | 3617 | 3492 | 3350 | 3234 | 3141 | 3032 | 2929 | 2818 | 2715 | 2609 | 2515 |
| -15 | 4168 | 4031 | 3890 | 3760 | 3615 | 3490 | 3365 | 3226 | 3133 | 3021 | 2918 | 2816 | 2705 | 2611 | 2507 |
| -14 | 4195 | 4050 | 3910 | 3783 | 3656 | 3518 | 3393 | 3253 | 3147 | 3044 | 2933 | 2830 | 2729 | 2626 | 2528 |
| -13 | 4223 | 4076 | 3930 | 3802 | 3675 | 3538 | 3415 | 3293 | 3162 | 3059 | 2957 | 2845 | 2744 | 2647 | 2538 |
| -12 | 4253 | 4102 | 3950 | 3823 | 3696 | 3557 | 3434 | 3312 | 3180 | 3071 | 2969 | 2868 | 2760 | 2657 | 2554 |
| -11 | 4246 | 4097 | 3952 | 3813 | 3686 | 3590 | 3454 | 3332 | 3210 | 3081 | 2980 | 2878 | 2777 | 2667 | 2565 |
| -10 | 4277 | 4128 | 3981 | 3834 | 3708 | 3581 | 3456 | 3323 | 3201 | 3081 | 2966 | 2865 | 2762 | 2660 | 2551 |
| -9 | 4308 | 4159 | 4010 | 3856 | 3730 | 3604 | 3478 | 3343 | 3222 | 3100 | 2984 | 2876 | 2774 | 2671 | 2569 |
| -8 | 4340 | 4191 | 4043 | 3898 | 3753 | 3627 | 3501 | 3377 | 3244 | 3122 | 3003 | 2895 | 2793 | 2684 | 2581 |
| -7 | 4363 | 4216 | 4069 | 3922 | 3776 | 3651 | 3525 | 3400 | 3278 | 3157 | 3024 | 2908 | 2806 | 2703 | 2601 |
| -6 | 4390 | 4239 | 4093 | 3946 | 3814 | 3676 | 3550 | 3424 | 3294 | 3174 | 3054 | 2929 | 2819 | 2716 | 2614 |
| -5 | 4418 | 4263 | 4117 | 3970 | 3839 | 3714 | 3576 | 3436 | 3308 | 3188 | 3068 | 2944 | 2830 | 2728 | 2622 |
| -4 | 4447 | 4291 | 4141 | 3995 | 3865 | 3740 | 3604 | 3457 | 3322 | 3203 | 3082 | 2953 | 2839 | 2735 | 2633 |
| -3 | 4477 | 4320 | 4167 | 4021 | 3891 | 3766 | 3618 | 3471 | 3345 | 3225 | 3097 | 2963 | 2848 | 2743 | 2641 |
| -2 | 4508 | 4351 | 4196 | 4047 | 3918 | 3779 | 3633 | 3486 | 3360 | 3235 | 3107 | 2979 | 2863 | 2755 | 2653 |
| -1 | 4539 | 4383 | 4227 | 4074 | 3943 | 3794 | 3648 | 3502 | 3375 | 3246 | 3118 | 2990 | 2874 | 2765 | 2662 |
| 0 | 4532 | 4376 | 4221 | 4065 | 3919 | 3773 | 3627 | 3481 | 3353 | 3225 | 3097 | 2969 | 2858 | 2747 | 2646 |
| 1 | 4607 | 4448 | 4293 | 4138 | 3990 | 3844 | 3698 | 3548 | 3406 | 3274 | 3147 | 3019 | 2903 | 2792 | 2684 |
| 2 | 4634 | 4471 | 4316 | 4161 | 4014 | 3868 | 3722 | 3569 | 3424 | 3287 | 3159 | 3031 | 2916 | 2805 | 2696 |
| 3 | 4666 | 4499 | 4340 | 4186 | 4038 | 3893 | 3743 | 3590 | 3445 | 3305 | 3172 | 3050 | 2934 | 2824 | 2713 |
| 4 | 4698 | 4528 | 4365 | 4210 | 4063 | 3918 | 3764 | 3611 | 3475 | 3334 | 3195 | 3064 | 2948 | 2838 | 2727 |
| 5 | 4731 | 4562 | 4395 | 4236 | 4089 | 3939 | 3795 | 3642 | 3498 | 3356 | 3215 | 3078 | 2962 | 2852 | 2746 |
| 6 | 4765 | 4596 | 4427 | 4262 | 4114 | 3970 | 3818 | 3665 | 3521 | 3380 | 3239 | 3100 | 2981 | 2870 | 2760 |
| 7 | 4799 | 4631 | 4462 | 4295 | 4139 | 3986 | 3832 | 3679 | 3533 | 3391 | 3252 | 3113 | 2994 | 2883 | 2773 |
| 8 | 4821 | 4651 | 4480 | 4310 | 4151 | 3996 | 3843 | 3689 | 3546 | 3404 | 3262 | 3127 | 3007 | 2901 | 2782 |
| 9 | 4844 | 4665 | 4495 | 4329 | 4165 | 4007 | 3854 | 3703 | 3557 | 3415 | 3277 | 3146 | 3025 | 2915 | 2784 |
| 10 | 4868 | 4680 | 4510 | 4345 | 4180 | 4021 | 3865 | 3715 | 3568 | 3427 | 3296 | 3161 | 3039 | 2919 | 2786 |
| 11 | 4897 | 4707 | 4526 | 4361 | 4196 | 4077 | 3923 | 3766 | 3616 | 3480 | 3345 | 3210 | 3081 | 2954 | 2823 |
| 12 | 4975 | 4784 | 4603 | 4421 | 4253 | 4099 | 3940 | 3782 | 3631 | 3496 | 3361 | 3227 | 3091 | 2959 | 2827 |
| 13 | 5005 | 4814 | 4634 | 4447 | 4271 | 4116 | 3958 | 3799 | 3650 | 3505 | 3369 | 3233 | 3097 | 2964 | 2831 |
| 14 | 5035 | 4845 | 4666 | 4475 | 4289 | 4137 | 3979 | 3826 | 3668 | 3512 | 3376 | 3240 | 3104 | 2970 | 2836 |
| 15 | 5067 | 4877 | 4691 | 4499 | 4313 | 4159 | 4001 | 3849 | 3691 | 3539 | 3384 | 3248 | 3111 | 2977 | 2843 |
| 16 | 5088 | 4949 | 4762 | 4570 | 4383 | 4218 | 4069 | 3913 | 3755 | 3603 | 3445 | 3297 | 3153 | 3019 | 2885 |
| 17 | 5159 | 4968 | 4784 | 4593 | 4406 | 4239 | 4091 | 3937 | 3785 | 3627 | 3475 | 3318 | 3170 | 3027 | 2894 |
| 18 | 5178 | 4987 | 4804 | 4616 | 4430 | 4260 | 4112 | 3959 | 3810 | 3653 | 3501 | 3343 | 3174 | 3025 | 2881 |
| 19 | 5198 | 5006 | 4825 | 4641 | 4454 | 4326 | 4174 | 4026 | 3873 | 3710 | 3546 | 3375 | 3211 | 3045 | 2908 |
| 20 | 5270 | 5072 | 4892 | 4711 | 4526 | 4349 | 4197 | 4034 | 3875 | 3709 | 3543 | 3372 | 3208 | 3040 | 2903 |
| 21 | 5292 | 5093 | 4912 | 4732 | 4548 | 4361 | 4199 | 4034 | 3868 | 3709 | 3536 | 3370 | 3199 | 3035 | 2890 |
| 22 | 5324 | 5114 | 4934 | 4746 | 4606 | 4414 | 4237 | 4073 | 3908 | 3749 | 3578 | 3405 | 3239 | 3066 | 2921 |
| 23 | 5417 | 5198 | 4993 | 4802 | 4615 | 4424 | 4245 | 4079 | 3909 | 3743 | 3580 | 3407 | 3240 | 3067 | 2921 |
| 24 | 5442 | 5214 | 5005 | 4812 | 4625 | 4434 | 4254 | 4082 | 3957 | 3787 | 3623 | 3450 | 3281 | 3117 | 2958 |
| 25 | 5460 | 5232 | 5068 | 4874 | 4681 | 4493 | 4305 | 4134 | 3966 | 3795 | 3621 | 3460 | 3293 | 3126 | 2965 |
| 26 | 5536 | 5309 | 5084 | 4890 | 4696 | 4505 | 4315 | 4144 | 3971 | 3803 | 3629 | 3506 | 3347 | 3179 | 3011 |
| 27 | 5556 | 5329 | 5100 | 4907 | 4761 | 4570 | 4376 | 4195 | 4022 | 3853 | 3679 | 3513 | 3351 | 3192 | 3025 |
| 28 | 5633 | 5408 | 5178 | 4973 | 4779 | 4577 | 4378 | 4195 | 4031 | 3857 | 3689 | 3559 | 3397 | 3235 | 3077 |
| 29 | 5650 | 5425 | 5201 | 4991 | 4836 | 4627 | 4428 | 4228 | 4066 | 3902 | 3736 | 3568 | 3406 | 3244 | 3075 |
| 30 | 5722 | 5498 | 5275 | 5045 | 4838 | 4629 | 4420 | 4224 | 4062 | 3901 | 3775 | 3608 | 3454 | 3283 | 3110 |
| 31 | 5739 | 5516 | 5282 | 5046 | 4887 | 4680 | 4472 | 4274 | 4095 | 3934 | 3772 | 3598 | 3429 | 3280 | 3108 |
| 32 | 5757 | 5579 | 5343 | 5106 | 4888 | 4682 | 4475 | 4267 | 4129 | 3967 | 3791 | 3613 | 3438 | 3284 | 3123 |
| 33 | 5821 | 5585 | 5349 | 5168 | 4938 | 4732 | 4526 | 4319 | 4111 | 3948 | 3770 | 3591 | 3413 | 3259 | 3134 |
| 34 | 5874 | 5643 | 5412 | 5176 | 4942 | 4736 | 4577 | 4364 | 4148 | 3953 | 3771 | 3608 | 3428 | 3264 | 3109 |
| 35 | 5872 | 5644 | 5414 | 5236 | 5004 | 4788 | 4575 | 4359 | 4143 | 3943 | 3797 | 3612 | 3429 | 3270 | 3115 |
| 36 | 5937 | 5696 | 5467 | 5239 | 5008 | 4832 | 4619 | 4403 | 4187 | 3973 | 3788 | 3603 | 3417 | 3279 | 3133 |
| 37 | 5945 | 5697 | 5520 | 5292 | 5059 | 4825 | 4613 | 4445 | 4232 | 4015 | 3819 | 3634 | 3459 | 3287 | 3135 |
| 38 | 6032 | 5750 | 5523 | 5342 | 5108 | 4872 | 4652 | 4440 | 4227 | 4059 | 3851 | 3678 | 3507 | 3336 | 3168 |
| 39 | 6042 | 5835 | 5574 | 5340 | 5106 | 4923 | 4691 | 4480 | 4267 | 4054 | 3833 | 3708 | 3556 | 3385 | 3213 |
| 40 | 6102 | 5838 | 5637 | 5386 | 5157 | 4922 | 4739 | 4520 | 4308 | 4088 | 3866 | 3687 | 3519 | 3398 | 3240 |
| 41 | 6161 | 5894 | 5629 | 5426 | 5199 | 4969 | 4740 | 4517 | 4343 | 4121 | 3899 | 3662 | 3516 | 3370 | 3223 |
| 42 | 6160 | 5951 | 5683 | 5440 | 5237 | 5009 | 4781 | 4549 | 4332 | 4156 | 3911 | 3689 | 3514 | 3367 | 3223 |
| 43 | 6217 | 5950 | 5741 | 5479 | 5229 | 5049 | 4817 | 4583 | 4359 | 4116 | 3941 | 3721 | 3513 | 3368 | 3224 |
| 44 | 6298 | 6009 | 5742 | 5533 | 5293 | 5039 | 4853 | 4619 | 4386 | 4143 | 3926 | 3750 | 3538 | 3372 | 3226 |
| 45 | 6282 | 6047 | 5789 | 5582 | 5334 | 5076 | 4843 | 4656 | 4421 | 4172 | 3954 | 3785 | 3576 | 3377 | 3231 |
| 46 | 6310 | 6107 | 5821 | 5567 | 5362 | 5096 | 4870 | 4685 | 4460 | 4202 | 3990 | 3781 | 3615 | 3405 | 3217 |
| 47 | 6338 | 6086 | 5882 | 5596 | 5341 | 5141 | 4885 | 4662 | 4478 | 4223 | 4015 | 3822 | 3655 | 3425 | 3224 |
| 48 | 6384 | 6116 | 5912 | 5625 | 5370 | 5164 | 4898 | 4675 | 4491 | 4241 | 4024 | 3830 | 3668 | 3443 | 3243 |
| 49 | 6408 | 6146 | 5893 | 5689 | 5401 | 5195 | 4912 | 4689 | 4509 | 4259 | 4040 | 3834 | 3673 | 3447 | 3263 |
| 50 | 6453 | 6217 | 5916 | 5712 | 5425 | 5179 | 4972 | 4707 | 4529 | 4310 | 4060 | 3840 | 3679 | 3463 | 3284 |
| 51 | 6488 | 6253 | 5930 | 5683 | 5479 | 5193 | 4988 | 4716 | 4502 | 4324 | 4075 | 3861 | 3651 | 3484 | 3306 |
| 52 | 6562 | 6233 | 5998 | 5695 | 5491 | 5205 | 5007 | 4733 | 4544 | 4331 | 4082 | 3868 | 3657 | 3476 | 3309 |
| 53 | 6599 | 6270 | 6036 | 5708 | 5504 | 5257 | 5028 | 4789 | 4558 | 4338 | 4089 | 3911 | 3666 | 3457 | 3292 |
| 54 | 6636 | 6348 | 6075 | 5784 | 5519 | 5280 | 5050 | 4812 | 4581 | 4346 | 4098 | 3922 | 3709 | 3463 | 3296 |
| 55 | 6659 | 6369 | 6135 | 5806 | 5530 | 5286 | 5050 | 4819 | 4582 | 4350 | 4101 | 3925 | 3720 | 3502 | 3295 |
| 56 | 6658 | 6373 | 6138 | 5808 | 5583 | 5275 | 5081 | 4808 | 4613 | 4340 | 4160 | 3918 | 3711 | 3495 | 3292 |
| 57 | 6658 | 6423 | 6137 | 5859 | 5592 | 5282 | 5067 | 4795 | 4599 | 4327 | 4145 | 3909 | 3732 | 3487 | 3308 |
| 58 | 6729 | 6424 | 6138 | 5869 | 5604 | 5339 | 5072 | 4860 | 4588 | 4348 | 4132 | 3896 | 3725 | 3498 | 3358 |
| 59 | 6754 | 6442 | 6192 | 5882 | 5661 | 5352 | 5076 | 4854 | 4578 | 4383 | 4121 | 3951 | 3715 | 3518 | 3358 |
| 60 | 6831 | 6504 | 6196 | 5929 | 5657 | 5382 | 5123 | 4857 | 4570 | 4375 | 4137 | 3941 | 3720 | 3525 | 3327 |
| 61 | 6825 | 6498 | 6254 | 5922 | 5648 | 5426 | 5114 | 4861 | 4656 | 4369 | 4129 | 3930 | 3697 | 3497 | 3311 |
| 62 | 6815 | 6551 | 6261 | 5928 | 5693 | 5419 | 5153 | 4911 | 4661 | 4373 | 4169 | 3939 | 3675 | 3488 | 3285 |
| 63 | 6870 | 6558 | 6270 | 5989 | 5701 | 5415 | 5204 | 4917 | 4653 | 4445 | 4153 | 3919 | 3730 | 3465 | 3278 |
| 64 | 6863 | 6620 | 6332 | 6021 | 5683 | 5440 | 5184 | 4920 | 4683 | 4419 | 4149 | 3929 | 3712 | 3475 | 3293 |
| 65 | 6921 | 6603 | 6295 | 6047 | 5728 | 5384 | 5123 | 4947 | 4683 | 4390 | 4174 | 3903 | 3675 | 3459 | 3289 |
| 66 | 6697 | 6628 | 6308 | 5988 | 5668 | 5419 | 5107 | 4887 | 4657 | 4446 | 4170 | 3867 | 3652 | 3486 | 3273 |
| 67 | 6516 | 6384 | 6251 | 6002 | 5682 | 5361 | 5044 | 4823 | 4650 | 4417 | 4141 | 3865 | 3661 | 3448 | 3271 |
| 68 | 6231 | 6149 | 6069 | 5947 | 5626 | 5304 | 5055 | 4809 | 4583 | 4318 | 4172 | 3896 | 3641 | 3441 | 3289 |
| 69 | 6045 | 5843 | 5757 | 5745 | 5643 | 5322 | 4999 | 4741 | 4568 | 4291 | 4066 | 3858 | 3622 | 3460 | 3290 |
| 70 | 5861 | 5610 | 5518 | 5333 | 5246 | 5162 | 5043 | 4732 | 4486 | 4200 | 3963 | 3812 | 3604 | 3455 | 3281 |
| 71 | 5579 | 5257 | 5182 | 5089 | 4997 | 4907 | 4820 | 4670 | 4398 | 4188 | 3913 | 3705 | 3485 | 3302 | 3296 |
| 72 | 5397 | 5071 | 4872 | 4773 | 4676 | 4653 | 4554 | 4434 | 4381 | 4102 | 3819 | 3604 | 3436 | 3194 | 3118 |
| 73 | 4810 | 4886 | 4629 | 4529 | 4424 | 4176 | 4035 | 3933 | 3888 | 3916 | 3813 | 3557 | 3318 | 3076 | 2936 |
| 74 | 4158 | 4279 | 4331 | 4182 | 4016 | 3866 | 3719 | 3620 | 3569 | 3520 | 3371 | 3333 | 3267 | 3024 | 2787 |
| 75 | 3506 | 3613 | 3848 | 3734 | 3621 | 3464 | 3315 | 3190 | 3135 | 3097 | 3063 | 3024 | 2970 | 2912 | 2657 |
| 76 | 2852 | 3080 | 3110 | 3359 | 3251 | 2998 | 2888 | 2828 | 2782 | 2638 | 2602 | 2611 | 2550 | 2492 | 2438 |
| 77 | 2306 | 2554 | 2585 | 2855 | 2706 | 2600 | 2526 | 2437 | 2390 | 2345 | 2303 | 2132 | 2065 | 2000 | 2055 |
| 78 | 1827 | 1845 | 2086 | 2110 | 2223 | 2109 | 1969 | 2099 | 2104 | 1948 | 1899 | 1824 | 1645 | 1666 | 1765 |
| 79 | 1347 | 1355 | 1585 | 1599 | 1827 | 1556 | 1587 | 1642 | 1655 | 1598 | 1428 | 1345 | 1340 | 1362 | 1473 |
| 80 | 838 | 864 | 864 | 1086 | 1092 | 1160 | 1050 | 1252 | 1369 | 1200 | 1127 | 973 | 1039 | 1144 | 1113 |
| 81 | 713 | 668 | 581 | 546 | 541 | 559 | 666 | 705 | 901 | 803 | 722 | 735 | 755 | 855 | 803 |
| 82 | 614 | 525 | 479 | 390 | 345 | 304 | 269 | 302 | 335 | 454 | 351 | 449 | 469 | 565 | 494 |
| 83 | 516 | 425 | 376 | 285 | 242 | 257 | 251 | 216 | 189 | 154 | 149 | 163 | 178 | 269 | 189 |
| 84 | 419 | 325 | 231 | 187 | 196 | 214 | 231 | 198 | 163 | 135 | 117 | 102 | 93 | 82 | 76 |
| 85 | 245 | 226 | 132 | 132 | 156 | 183 | 187 | 174 | 145 | 110 | 92 | 76 | 67 | 66 | 68 |
| 86 | 92 | 94 | 88 | 105 | 127 | 139 | 154 | 156 | 120 | 85 | 67 | 58 | 51 | 52 | 54 |
| 87 | 30 | 31 | 42 | 77 | 99 | 107 | 109 | 112 | 97 | 67 | 49 | 38 | 37 | 38 | 40 |
| 88 | 20 | 20 | 22 | 36 | 58 | 63 | 77 | 79 | 79 | 42 | 24 | 23 | 24 | 24 | 25 |
| 89 | 10 | 10 | 10 | 12 | 29 | 32 | 32 | 33 | 34 | 19 | 10 | 10 | 10 | 10 | 11 |
| 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Table 17d.

Relative Spatial Response of ULIS (times 10,000)

| AZ | ZA = 46 | ZA = 47 | ZA = 48 | ZA = 49 | ZA = 50 | ZA = 51 | ZA = 52 | ZA = 53 | ZA = 54 | ZA = 55 | ZA = 56 | ZA = 57 | ZA = 58 | ZA = 59 | ZA = 60 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| -89 | 19 | 19 | 19 | 18 | 13 | 7 | 8 | 8 | 8 | 12 | 23 | 26 | 20 | 13 | 12 |
| -88 | 39 | 39 | 37 | 32 | 23 | 16 | 17 | 17 | 20 | 30 | 37 | 30 | 22 | 21 | 24 |
| -87 | 59 | 58 | 49 | 41 | 37 | 28 | 24 | 25 | 34 | 37 | 36 | 31 | 30 | 33 | 38 |
| -86 | 78 | 70 | 57 | 51 | 45 | 39 | 34 | 35 | 38 | 37 | 36 | 38 | 41 | 47 | 50 |
| -85 | 90 | 78 | 68 | 59 | 54 | 55 | 53 | 41 | 42 | 43 | 43 | 47 | 53 | 58 | 64 |
| -84 | 105 | 102 | 119 | 120 | 145 | 145 | 134 | 98 | 74 | 50 | 51 | 55 | 61 | 67 | 85 |
| -83 | 191 | 189 | 209 | 211 | 239 | 238 | 202 | 179 | 156 | 121 | 105 | 113 | 160 | 185 | 197 |
| -82 | 373 | 282 | 301 | 324 | 317 | 294 | 274 | 248 | 247 | 247 | 248 | 253 | 301 | 317 | 372 |
| -81 | 570 | 532 | 448 | 413 | 409 | 384 | 383 | 384 | 385 | 387 | 399 | 403 | 444 | 505 | 507 |
| -80 | 767 | 733 | 648 | 576 | 579 | 576 | 576 | 544 | 547 | 551 | 557 | 570 | 637 | 610 | 584 |
| -79 | 994 | 963 | 871 | 786 | 746 | 731 | 734 | 739 | 745 | 753 | 781 | 797 | 765 | 749 | 719 |
| -78 | 1191 | 1159 | 1077 | 958 | 913 | 922 | 911 | 915 | 924 | 970 | 982 | 932 | 884 | 893 | 865 |
| -77 | 1418 | 1364 | 1285 | 1207 | 1139 | 1129 | 1121 | 1115 | 1121 | 1141 | 1104 | 1080 | 1036 | 1017 | 992 |
| -76 | 1613 | 1531 | 1454 | 1379 | 1306 | 1321 | 1316 | 1312 | 1288 | 1259 | 1226 | 1209 | 1188 | 1140 | 1108 |
| -75 | 1704 | 1682 | 1640 | 1595 | 1520 | 1490 | 1464 | 1407 | 1361 | 1344 | 1311 | 1291 | 1259 | 1200 | 1156 |
| -74 | 1784 | 1752 | 1731 | 1691 | 1666 | 1589 | 1510 | 1438 | 1424 | 1422 | 1402 | 1337 | 1273 | 1236 | 1196 |
| -73 | 1883 | 1864 | 1824 | 1760 | 1677 | 1613 | 1558 | 1511 | 1498 | 1471 | 1427 | 1369 | 1305 | 1258 | 1223 |
| -72 | 1968 | 1939 | 1876 | 1790 | 1716 | 1652 | 1585 | 1578 | 1543 | 1489 | 1425 | 1387 | 1335 | 1276 | 1228 |
| -71 | 2019 | 1960 | 1886 | 1801 | 1755 | 1692 | 1626 | 1586 | 1552 | 1499 | 1441 | 1395 | 1335 | 1279 | 1231 |
| -70 | 2024 | 1980 | 1917 | 1838 | 1774 | 1711 | 1643 | 1596 | 1560 | 1502 | 1447 | 1389 | 1338 | 1278 | 1230 |
| -69 | 2046 | 1986 | 1942 | 1872 | 1785 | 1706 | 1662 | 1607 | 1555 | 1508 | 1454 | 1393 | 1342 | 1282 | 1234 |
| -68 | 2066 | 2008 | 1938 | 1882 | 1796 | 1726 | 1681 | 1623 | 1561 | 1515 | 1456 | 1397 | 1347 | 1287 | 1239 |
| -67 | 2087 | 2017 | 1948 | 1876 | 1813 | 1745 | 1697 | 1634 | 1569 | 1523 | 1461 | 1398 | 1347 | 1288 | 1244 |
| -66 | 2082 | 2021 | 1953 | 1881 | 1828 | 1760 | 1694 | 1647 | 1581 | 1530 | 1467 | 1404 | 1353 | 1294 | 1256 |
| -65 | 2086 | 2017 | 1958 | 1892 | 1834 | 1765 | 1704 | 1649 | 1583 | 1532 | 1472 | 1423 | 1360 | 1315 | 1258 |
| -64 | 2097 | 2022 | 1970 | 1898 | 1827 | 1776 | 1706 | 1650 | 1581 | 1532 | 1474 | 1426 | 1366 | 1318 | 1260 |
| -63 | 2104 | 2034 | 1976 | 1904 | 1839 | 1783 | 1714 | 1657 | 1590 | 1542 | 1482 | 1431 | 1370 | 1316 | 1259 |
| -62 | 2118 | 2040 | 1968 | 1916 | 1846 | 1791 | 1718 | 1662 | 1597 | 1549 | 1482 | 1428 | 1378 | 1314 | 1267 |
| -61 | 2122 | 2052 | 1981 | 1923 | 1854 | 1802 | 1724 | 1664 | 1605 | 1550 | 1493 | 1427 | 1377 | 1319 | 1266 |
| -60 | 2117 | 2057 | 1988 | 1937 | 1864 | 1799 | 1722 | 1672 | 1618 | 1547 | 1492 | 1426 | 1376 | 1328 | 1272 |
| -59 | 2119 | 2070 | 2000 | 1939 | 1862 | 1806 | 1745 | 1683 | 1624 | 1554 | 1499 | 1444 | 1382 | 1329 | 1283 |
| -58 | 2132 | 2076 | 2009 | 1942 | 1869 | 1805 | 1751 | 1681 | 1622 | 1555 | 1500 | 1445 | 1392 | 1335 | 1284 |
| -57 | 2145 | 2087 | 2003 | 1933 | 1881 | 1815 | 1757 | 1687 | 1629 | 1574 | 1508 | 1453 | 1393 | 1347 | 1292 |
| -56 | 2152 | 2092 | 2008 | 1937 | 1869 | 1814 | 1756 | 1698 | 1630 | 1576 | 1521 | 1456 | 1402 | 1352 | 1297 |
| -55 | 2158 | 2088 | 2028 | 1944 | 1876 | 1808 | 1764 | 1706 | 1639 | 1585 | 1530 | 1468 | 1394 | 1347 | 1299 |
| -54 | 2167 | 2096 | 2026 | 1943 | 1874 | 1826 | 1764 | 1714 | 1657 | 1577 | 1515 | 1460 | 1405 | 1358 | 1304 |
| -53 | 2178 | 2106 | 2035 | 1951 | 1883 | 1827 | 1766 | 1707 | 1643 | 1582 | 1526 | 1461 | 1407 | 1353 | 1307 |
| -52 | 2180 | 2117 | 2042 | 1975 | 1892 | 1835 | 1783 | 1707 | 1641 | 1583 | 1528 | 1473 | 1411 | 1364 | 1305 |
| -51 | 2192 | 2104 | 2048 | 1983 | 1902 | 1844 | 1783 | 1718 | 1654 | 1585 | 1522 | 1467 | 1414 | 1366 | 1312 |
| -50 | 2193 | 2111 | 2045 | 1990 | 1924 | 1850 | 1785 | 1721 | 1657 | 1587 | 1533 | 1480 | 1417 | 1366 | 1301 |
| -49 | 2195 | 2118 | 2053 | 1986 | 1919 | 1852 | 1790 | 1715 | 1651 | 1589 | 1535 | 1482 | 1428 | 1361 | 1306 |
| -48 | 2201 | 2135 | 2056 | 1990 | 1920 | 1852 | 1792 | 1730 | 1666 | 1603 | 1538 | 1485 | 1429 | 1362 | 1304 |
| -47 | 2205 | 2139 | 2072 | 1994 | 1921 | 1854 | 1793 | 1733 | 1671 | 1608 | 1550 | 1496 | 1429 | 1372 | 1313 |
| -46 | 2209 | 2143 | 2077 | 2007 | 1935 | 1865 | 1795 | 1735 | 1674 | 1614 | 1553 | 1498 | 1430 | 1371 | 1312 |
| -45 | 2214 | 2148 | 2082 | 2010 | 1937 | 1867 | 1807 | 1747 | 1687 | 1627 | 1564 | 1505 | 1436 | 1377 | 1321 |
| -44 | 2231 | 2154 | 2080 | 2009 | 1940 | 1869 | 1810 | 1750 | 1687 | 1623 | 1560 | 1502 | 1435 | 1377 | 1327 |
| -43 | 2237 | 2170 | 2091 | 2017 | 1950 | 1872 | 1809 | 1747 | 1691 | 1628 | 1564 | 1505 | 1449 | 1380 | 1321 |
| -42 | 2243 | 2170 | 2091 | 2017 | 1949 | 1880 | 1814 | 1752 | 1689 | 1625 | 1562 | 1502 | 1446 | 1381 | 1322 |
| -41 | 2249 | 2184 | 2105 | 2028 | 1961 | 1890 | 1811 | 1749 | 1686 | 1623 | 1568 | 1506 | 1449 | 1387 | 1316 |
| -40 | 2260 | 2185 | 2106 | 2029 | 1970 | 1898 | 1829 | 1766 | 1692 | 1629 | 1566 | 1504 | 1444 | 1389 | 1329 |
| -39 | 2265 | 2189 | 2116 | 2042 | 1969 | 1897 | 1827 | 1765 | 1702 | 1639 | 1573 | 1508 | 1447 | 1385 | 1324 |
| -38 | 2261 | 2184 | 2111 | 2038 | 1968 | 1896 | 1825 | 1771 | 1709 | 1646 | 1581 | 1504 | 1449 | 1388 | 1327 |
| -37 | 2257 | 2180 | 2107 | 2034 | 1972 | 1904 | 1835 | 1770 | 1705 | 1640 | 1581 | 1518 | 1457 | 1384 | 1330 |
| -36 | 2270 | 2200 | 2126 | 2054 | 1982 | 1914 | 1845 | 1771 | 1706 | 1641 | 1576 | 1515 | 1460 | 1399 | 1339 |
| -35 | 2277 | 2197 | 2123 | 2051 | 1979 | 1919 | 1850 | 1778 | 1713 | 1649 | 1579 | 1518 | 1456 | 1402 | 1344 |
| -34 | 2274 | 2195 | 2120 | 2057 | 1985 | 1916 | 1845 | 1772 | 1715 | 1650 | 1585 | 1532 | 1472 | 1400 | 1346 |
| -33 | 2288 | 2218 | 2139 | 2069 | 1997 | 1914 | 1849 | 1776 | 1710 | 1652 | 1586 | 1525 | 1467 | 1412 | 1349 |
| -32 | 2296 | 2216 | 2132 | 2064 | 2002 | 1933 | 1860 | 1787 | 1724 | 1660 | 1580 | 1524 | 1465 | 1403 | 1343 |
| -31 | 2295 | 2212 | 2126 | 2065 | 1997 | 1925 | 1862 | 1792 | 1720 | 1661 | 1595 | 1530 | 1477 | 1413 | 1338 |
| -30 | 2310 | 2236 | 2150 | 2073 | 2005 | 1939 | 1867 | 1799 | 1722 | 1657 | 1597 | 1531 | 1470 | 1409 | 1343 |
| -29 | 2319 | 2233 | 2147 | 2075 | 2007 | 1931 | 1864 | 1796 | 1728 | 1666 | 1603 | 1545 | 1481 | 1401 | 1339 |
| -28 | 2330 | 2248 | 2172 | 2089 | 2015 | 1947 | 1872 | 1788 | 1725 | 1657 | 1599 | 1536 | 1474 | 1409 | 1349 |
| -27 | 2323 | 2249 | 2167 | 2086 | 2015 | 1940 | 1871 | 1800 | 1732 | 1669 | 1604 | 1546 | 1478 | 1400 | 1337 |
| -26 | 2341 | 2260 | 2178 | 2101 | 2021 | 1954 | 1865 | 1798 | 1730 | 1662 | 1600 | 1537 | 1472 | 1402 | 1343 |
| -25 | 2335 | 2253 | 2175 | 2090 | 2016 | 1944 | 1877 | 1805 | 1742 | 1675 | 1608 | 1547 | 1477 | 1412 | 1335 |
| -24 | 2346 | 2267 | 2183 | 2105 | 2022 | 1955 | 1883 | 1815 | 1731 | 1670 | 1604 | 1540 | 1472 | 1406 | 1338 |
| -23 | 2341 | 2257 | 2179 | 2094 | 2017 | 1945 | 1878 | 1804 | 1740 | 1674 | 1613 | 1544 | 1479 | 1411 | 1346 |
| -22 | 2350 | 2272 | 2188 | 2110 | 2025 | 1957 | 1885 | 1816 | 1744 | 1683 | 1618 | 1552 | 1468 | 1403 | 1339 |
| -21 | 2347 | 2263 | 2185 | 2100 | 2022 | 1948 | 1881 | 1806 | 1739 | 1674 | 1612 | 1545 | 1476 | 1412 | 1343 |
| -20 | 2356 | 2279 | 2195 | 2116 | 2031 | 1960 | 1888 | 1818 | 1744 | 1683 | 1621 | 1550 | 1484 | 1416 | 1352 |
| -19 | 2372 | 2288 | 2204 | 2126 | 2047 | 1968 | 1901 | 1826 | 1756 | 1693 | 1626 | 1543 | 1474 | 1410 | 1345 |
| -18 | 2377 | 2280 | 2202 | 2124 | 2039 | 1965 | 1892 | 1822 | 1752 | 1684 | 1620 | 1549 | 1483 | 1419 | 1350 |
| -17 | 2396 | 2303 | 2213 | 2135 | 2056 | 1973 | 1905 | 1830 | 1760 | 1695 | 1626 | 1559 | 1493 | 1424 | 1360 |
| -16 | 2421 | 2317 | 2227 | 2146 | 2068 | 1991 | 1913 | 1844 | 1774 | 1701 | 1636 | 1570 | 1499 | 1434 | 1353 |
| -15 | 2410 | 2307 | 2213 | 2136 | 2059 | 1978 | 1907 | 1838 | 1766 | 1697 | 1631 | 1560 | 1494 | 1428 | 1363 |
| -14 | 2419 | 2317 | 2222 | 2141 | 2064 | 1987 | 1913 | 1841 | 1773 | 1704 | 1634 | 1568 | 1503 | 1437 | 1370 |
| -13 | 2435 | 2333 | 2231 | 2150 | 2074 | 1993 | 1916 | 1847 | 1779 | 1707 | 1638 | 1572 | 1507 | 1438 | 1372 |
| -12 | 2446 | 2344 | 2246 | 2156 | 2080 | 2003 | 1925 | 1851 | 1782 | 1713 | 1643 | 1574 | 1509 | 1442 | 1375 |
| -11 | 2463 | 2355 | 2258 | 2167 | 2090 | 2009 | 1931 | 1858 | 1789 | 1720 | 1628 | 1563 | 1497 | 1430 | 1360 |
| -10 | 2449 | 2349 | 2252 | 2154 | 2078 | 2000 | 1922 | 1848 | 1777 | 1707 | 1635 | 1567 | 1502 | 1433 | 1364 |
| -9 | 2467 | 2361 | 2264 | 2169 | 2089 | 2011 | 1929 | 1852 | 1784 | 1714 | 1643 | 1570 | 1505 | 1438 | 1369 |
| -8 | 2479 | 2380 | 2284 | 2181 | 2096 | 2019 | 1941 | 1863 | 1792 | 1719 | 1648 | 1577 | 1511 | 1443 | 1374 |
| -7 | 2492 | 2393 | 2297 | 2200 | 2108 | 2031 | 1949 | 1871 | 1797 | 1727 | 1656 | 1584 | 1516 | 1449 | 1378 |
| -6 | 2512 | 2413 | 2317 | 2214 | 2120 | 2039 | 1961 | 1883 | 1807 | 1735 | 1664 | 1590 | 1520 | 1453 | 1384 |
| -5 | 2520 | 2420 | 2324 | 2228 | 2132 | 2048 | 1970 | 1893 | 1814 | 1741 | 1670 | 1599 | 1528 | 1459 | 1390 |
| -4 | 2532 | 2428 | 2331 | 2233 | 2137 | 2050 | 1973 | 1895 | 1818 | 1745 | 1673 | 1602 | 1532 | 1463 | 1394 |
| -3 | 2540 | 2435 | 2335 | 2240 | 2144 | 2056 | 1976 | 1899 | 1822 | 1748 | 1677 | 1605 | 1534 | 1465 | 1397 |
| -2 | 2552 | 2446 | 2343 | 2245 | 2149 | 2061 | 1979 | 1902 | 1825 | 1751 | 1680 | 1608 | 1538 | 1469 | 1400 |
| -1 | 2560 | 2455 | 2350 | 2253 | 2157 | 2069 | 1985 | 1906 | 1829 | 1756 | 1684 | 1613 | 1542 | 1473 | 1405 |
| 0 | 2544 | 2437 | 2330 | 2235 | 2139 | 2053 | 1968 | 1891 | 1813 | 1742 | 1670 | 1599 | 1527 | 1460 | 1392 |
| 1 | 2583 | 2478 | 2371 | 2267 | 2172 | 2084 | 1998 | 1915 | 1838 | 1765 | 1693 | 1622 | 1548 | 1476 | 1408 |
| 2 | 2593 | 2488 | 2381 | 2279 | 2181 | 2093 | 2007 | 1923 | 1843 | 1770 | 1699 | 1626 | 1551 | 1479 | 1408 |
| 3 | 2608 | 2502 | 2396 | 2289 | 2188 | 2100 | 2015 | 1929 | 1848 | 1775 | 1703 | 1629 | 1554 | 1482 | 1410 |
| 4 | 2619 | 2514 | 2407 | 2304 | 2200 | 2110 | 2025 | 1939 | 1855 | 1781 | 1706 | 1631 | 1557 | 1484 | 1413 |
| 5 | 2637 | 2529 | 2422 | 2316 | 2210 | 2118 | 2032 | 1947 | 1861 | 1783 | 1708 | 1633 | 1559 | 1486 | 1414 |
| 6 | 2648 | 2534 | 2427 | 2319 | 2212 | 2119 | 2033 | 1948 | 1862 | 1783 | 1708 | 1633 | 1558 | 1485 | 1413 |
| 7 | 2648 | 2535 | 2428 | 2321 | 2214 | 2120 | 2034 | 1949 | 1863 | 1783 | 1708 | 1633 | 1558 | 1485 | 1413 |
| 8 | 2650 | 2536 | 2431 | 2323 | 2216 | 2122 | 2036 | 1950 | 1864 | 1784 | 1708 | 1633 | 1558 | 1485 | 1413 |
| 9 | 2652 | 2537 | 2434 | 2327 | 2219 | 2124 | 2038 | 1952 | 1866 | 1786 | 1709 | 1633 | 1559 | 1485 | 1413 |
| 10 | 2655 | 2539 | 2436 | 2330 | 2223 | 2127 | 2040 | 1954 | 1868 | 1787 | 1711 | 1635 | 1560 | 1486 | 1414 |
| 11 | 2691 | 2565 | 2463 | 2360 | 2253 | 2151 | 2064 | 1978 | 1892 | 1809 | 1713 | 1636 | 1561 | 1487 | 1414 |
| 12 | 2694 | 2568 | 2466 | 2363 | 2258 | 2154 | 2068 | 1981 | 1895 | 1811 | 1734 | 1657 | 1581 | 1506 | 1434 |
| 13 | 2699 | 2571 | 2468 | 2366 | 2263 | 2158 | 2071 | 1985 | 1898 | 1813 | 1737 | 1660 | 1583 | 1507 | 1435 |
| 14 | 2704 | 2574 | 2472 | 2368 | 2266 | 2162 | 2075 | 1989 | 1902 | 1817 | 1739 | 1663 | 1586 | 1510 | 1434 |
| 15 | 2710 | 2578 | 2475 | 2372 | 2270 | 2166 | 2080 | 1993 | 1903 | 1816 | 1738 | 1658 | 1580 | 1503 | 1430 |
| 16 | 2751 | 2617 | 2505 | 2399 | 2296 | 2193 | 2098 | 2011 | 1923 | 1833 | 1752 | 1674 | 1593 | 1515 | 1423 |
| 17 | 2751 | 2614 | 2503 | 2399 | 2296 | 2189 | 2097 | 2007 | 1919 | 1832 | 1747 | 1668 | 1590 | 1509 | 1434 |
| 18 | 2744 | 2611 | 2502 | 2400 | 2293 | 2190 | 2093 | 2007 | 1919 | 1828 | 1745 | 1663 | 1585 | 1507 | 1428 |
| 19 | 2764 | 2642 | 2522 | 2422 | 2319 | 2213 | 2113 | 2024 | 1937 | 1849 | 1761 | 1661 | 1579 | 1501 | 1425 |
| 20 | 2765 | 2644 | 2521 | 2422 | 2318 | 2214 | 2110 | 2024 | 1934 | 1847 | 1760 | 1677 | 1597 | 1515 | 1438 |
| 21 | 2767 | 2642 | 2521 | 2419 | 2320 | 2213 | 2109 | 2020 | 1934 | 1844 | 1757 | 1677 | 1594 | 1514 | 1433 |
| 22 | 2797 | 2675 | 2550 | 2444 | 2341 | 2240 | 2133 | 2042 | 1953 | 1866 | 1776 | 1693 | 1591 | 1510 | 1431 |
| 23 | 2800 | 2675 | 2553 | 2441 | 2342 | 2238 | 2136 | 2039 | 1953 | 1864 | 1777 | 1691 | 1610 | 1527 | 1447 |
| 24 | 2831 | 2709 | 2584 | 2466 | 2363 | 2264 | 2160 | 2062 | 1951 | 1865 | 1775 | 1691 | 1608 | 1527 | 1443 |
| 25 | 2838 | 2711 | 2588 | 2465 | 2365 | 2262 | 2162 | 2060 | 1974 | 1882 | 1790 | 1705 | 1623 | 1543 | 1442 |
| 26 | 2876 | 2749 | 2621 | 2498 | 2388 | 2289 | 2160 | 2060 | 1968 | 1876 | 1788 | 1697 | 1620 | 1538 | 1459 |
| 27 | 2881 | 2757 | 2630 | 2501 | 2390 | 2286 | 2186 | 2081 | 1984 | 1896 | 1804 | 1716 | 1631 | 1536 | 1454 |
| 28 | 2921 | 2793 | 2668 | 2538 | 2414 | 2313 | 2207 | 2078 | 1984 | 1892 | 1803 | 1711 | 1629 | 1547 | 1470 |
| 29 | 2917 | 2800 | 2671 | 2546 | 2418 | 2313 | 2211 | 2105 | 2001 | 1913 | 1820 | 1731 | 1639 | 1538 | 1459 |
| 30 | 2937 | 2817 | 2700 | 2581 | 2451 | 2342 | 2236 | 2129 | 2002 | 1908 | 1818 | 1723 | 1629 | 1551 | 1467 |
| 31 | 2935 | 2794 | 2681 | 2577 | 2460 | 2343 | 2241 | 2128 | 2014 | 1923 | 1832 | 1737 | 1647 | 1560 | 1464 |
| 32 | 2971 | 2806 | 2693 | 2579 | 2477 | 2361 | 2253 | 2140 | 2034 | 1926 | 1818 | 1736 | 1640 | 1551 | 1474 |
| 33 | 2979 | 2835 | 2706 | 2591 | 2476 | 2337 | 2243 | 2136 | 2022 | 1917 | 1828 | 1738 | 1648 | 1566 | 1484 |
| 34 | 2954 | 2811 | 2685 | 2579 | 2463 | 2348 | 2238 | 2133 | 2037 | 1928 | 1830 | 1748 | 1660 | 1553 | 1482 |
| 35 | 2976 | 2840 | 2705 | 2585 | 2466 | 2360 | 2245 | 2139 | 2033 | 1932 | 1824 | 1736 | 1647 | 1565 | 1483 |
| 36 | 2996 | 2869 | 2731 | 2606 | 2486 | 2365 | 2246 | 2126 | 2028 | 1930 | 1833 | 1743 | 1660 | 1571 | 1484 |
| 37 | 2997 | 2859 | 2722 | 2593 | 2482 | 2361 | 2232 | 2133 | 2042 | 1944 | 1852 | 1759 | 1669 | 1559 | 1477 |
| 38 | 3030 | 2892 | 2753 | 2616 | 2489 | 2359 | 2234 | 2137 | 2049 | 1958 | 1866 | 1750 | 1666 | 1575 | 1477 |
| 39 | 3064 | 2925 | 2787 | 2638 | 2501 | 2371 | 2247 | 2136 | 2041 | 1952 | 1857 | 1767 | 1669 | 1570 | 1477 |
| 40 | 3089 | 2945 | 2796 | 2646 | 2514 | 2383 | 2261 | 2150 | 2028 | 1939 | 1850 | 1760 | 1665 | 1576 | 1485 |
| 41 | 3076 | 2960 | 2819 | 2669 | 2518 | 2382 | 2243 | 2132 | 2021 | 1932 | 1853 | 1763 | 1671 | 1572 | 1469 |
| 42 | 3078 | 2938 | 2800 | 2662 | 2512 | 2370 | 2259 | 2147 | 2035 | 1936 | 1846 | 1756 | 1665 | 1561 | 1479 |
| 43 | 3079 | 2935 | 2798 | 2661 | 2523 | 2373 | 2264 | 2152 | 2051 | 1941 | 1851 | 1761 | 1670 | 1563 | 1480 |
| 44 | 3081 | 2915 | 2774 | 2637 | 2505 | 2382 | 2274 | 2166 | 2057 | 1944 | 1846 | 1754 | 1646 | 1560 | 1490 |
| 45 | 3064 | 2918 | 2773 | 2642 | 2513 | 2391 | 2283 | 2175 | 2067 | 1957 | 1852 | 1760 | 1650 | 1564 | 1484 |
| 46 | 3070 | 2924 | 2773 | 2651 | 2522 | 2401 | 2280 | 2171 | 2063 | 1956 | 1850 | 1754 | 1644 | 1560 | 1478 |
| 47 | 3077 | 2927 | 2776 | 2645 | 2516 | 2399 | 2290 | 2182 | 2076 | 1970 | 1864 | 1760 | 1650 | 1566 | 1483 |
| 48 | 3082 | 2930 | 2764 | 2650 | 2527 | 2410 | 2302 | 2196 | 2090 | 1984 | 1866 | 1760 | 1662 | 1563 | 1475 |
| 49 | 3087 | 2912 | 2769 | 2654 | 2539 | 2423 | 2317 | 2200 | 2095 | 1988 | 1882 | 1774 | 1670 | 1564 | 1471 |
| 50 | 3105 | 2917 | 2775 | 2680 | 2564 | 2426 | 2319 | 2212 | 2105 | 1982 | 1875 | 1767 | 1657 | 1563 | 1453 |
| 51 | 3126 | 2924 | 2800 | 2685 | 2536 | 2418 | 2309 | 2201 | 2094 | 1966 | 1849 | 1753 | 1657 | 1561 | 1467 |
| 52 | 3109 | 2964 | 2806 | 2681 | 2530 | 2412 | 2299 | 2171 | 2063 | 1963 | 1867 | 1771 | 1659 | 1562 | 1452 |
| 53 | 3116 | 2953 | 2806 | 2654 | 2525 | 2406 | 2273 | 2162 | 2063 | 1967 | 1870 | 1756 | 1659 | 1549 | 1452 |
| 54 | 3099 | 2948 | 2804 | 2651 | 2522 | 2420 | 2286 | 2170 | 2071 | 1941 | 1843 | 1746 | 1648 | 1550 | 1443 |
| 55 | 3094 | 2949 | 2827 | 2671 | 2539 | 2397 | 2281 | 2163 | 2039 | 1941 | 1843 | 1745 | 1614 | 1516 | 1428 |
| 56 | 3096 | 2973 | 2803 | 2671 | 2539 | 2393 | 2275 | 2157 | 2027 | 1929 | 1831 | 1713 | 1614 | 1515 | 1419 |
| 57 | 3143 | 2976 | 2806 | 2673 | 2561 | 2396 | 2288 | 2148 | 2031 | 1931 | 1813 | 1714 | 1602 | 1502 | 1406 |
| 58 | 3163 | 2993 | 2833 | 2697 | 2528 | 2384 | 2285 | 2145 | 2026 | 1902 | 1803 | 1703 | 1604 | 1490 | 1398 |
| 59 | 3164 | 3025 | 2830 | 2694 | 2523 | 2399 | 2266 | 2160 | 2041 | 1907 | 1806 | 1707 | 1587 | 1483 | 1402 |
| 60 | 3150 | 2995 | 2811 | 2697 | 2540 | 2395 | 2223 | 2140 | 2039 | 1900 | 1798 | 1679 | 1578 | 1489 | 1392 |
| 61 | 3152 | 2974 | 2805 | 2674 | 2525 | 2412 | 2241 | 2121 | 2019 | 1915 | 1806 | 1684 | 1584 | 1480 | 1388 |
| 62 | 3137 | 2951 | 2785 | 2666 | 2517 | 2396 | 2240 | 2115 | 2001 | 1920 | 1795 | 1691 | 1590 | 1476 | 1395 |
| 63 | 3111 | 2951 | 2813 | 2646 | 2509 | 2388 | 2238 | 2115 | 1983 | 1902 | 1803 | 1702 | 1579 | 1485 | 1388 |
| 64 | 3119 | 2942 | 2814 | 2644 | 2490 | 2381 | 2231 | 2109 | 1966 | 1887 | 1785 | 1691 | 1573 | 1494 | 1395 |
| 65 | 3111 | 2953 | 2806 | 2647 | 2513 | 2368 | 2242 | 2124 | 1984 | 1892 | 1779 | 1685 | 1569 | 1489 | 1391 |
| 66 | 3111 | 2968 | 2804 | 2626 | 2512 | 2369 | 2230 | 2129 | 1989 | 1888 | 1773 | 1658 | 1565 | 1467 | 1391 |
| 67 | 3120 | 2956 | 2791 | 2621 | 2513 | 2364 | 2247 | 2119 | 1983 | 1879 | 1769 | 1655 | 1563 | 1467 | 1393 |
| 68 | 3119 | 2948 | 2766 | 2651 | 2502 | 2366 | 2244 | 2110 | 1981 | 1877 | 1762 | 1660 | 1569 | 1479 | 1405 |
| 69 | 3116 | 2928 | 2791 | 2638 | 2487 | 2344 | 2226 | 2094 | 1980 | 1876 | 1764 | 1653 | 1563 | 1479 | 1409 |
| 70 | 3093 | 2956 | 2788 | 2628 | 2457 | 2347 | 2196 | 2077 | 1986 | 1876 | 1758 | 1645 | 1553 | 1464 | 1398 |
| 71 | 3121 | 2947 | 2786 | 2612 | 2486 | 2326 | 2176 | 2061 | 1966 | 1862 | 1751 | 1660 | 1544 | 1456 | 1387 |
| 72 | 3043 | 2931 | 2803 | 2627 | 2454 | 2279 | 2154 | 2070 | 1947 | 1839 | 1728 | 1647 | 1534 | 1444 | 1374 |
| 73 | 2898 | 2812 | 2715 | 2573 | 2397 | 2239 | 2152 | 2042 | 1933 | 1816 | 1733 | 1614 | 1494 | 1429 | 1368 |
| 74 | 2714 | 2617 | 2550 | 2452 | 2386 | 2239 | 2106 | 1988 | 1860 | 1770 | 1698 | 1580 | 1459 | 1403 | 1335 |
| 75 | 2558 | 2498 | 2396 | 2251 | 2114 | 2058 | 1990 | 1896 | 1772 | 1659 | 1578 | 1513 | 1455 | 1357 | 1279 |
| 76 | 2366 | 2217 | 2064 | 1912 | 1790 | 1787 | 1744 | 1703 | 1628 | 1550 | 1458 | 1405 | 1339 | 1275 | 1212 |
| 77 | 2083 | 1951 | 1793 | 1636 | 1523 | 1472 | 1424 | 1392 | 1389 | 1404 | 1315 | 1243 | 1150 | 1121 | 1076 |
| 78 | 1713 | 1616 | 1453 | 1233 | 1153 | 1142 | 1103 | 1100 | 1106 | 1164 | 1125 | 1042 | 965 | 969 | 922 |
| 79 | 1397 | 1318 | 1125 | 959 | 886 | 853 | 845 | 844 | 845 | 849 | 879 | 850 | 772 | 757 | 713 |
| 80 | 1027 | 943 | 787 | 643 | 636 | 625 | 618 | 563 | 560 | 559 | 560 | 552 | 593 | 560 | 523 |
| 81 | 712 | 622 | 460 | 427 | 409 | 372 | 365 | 360 | 355 | 351 | 362 | 361 | 408 | 463 | 456 |
| 82 | 397 | 300 | 298 | 329 | 311 | 281 | 254 | 223 | 216 | 210 | 205 | 219 | 275 | 281 | 324 |
| 83 | 170 | 172 | 198 | 206 | 240 | 227 | 185 | 155 | 125 | 84 | 61 | 78 | 127 | 154 | 154 |
| 84 | 76 | 77 | 101 | 108 | 139 | 148 | 117 | 74 | 43 | 27 | 26 | 29 | 35 | 38 | 56 |
| 85 | 64 | 55 | 47 | 44 | 41 | 45 | 51 | 38 | 33 | 32 | 31 | 35 | 39 | 40 | 40 |
| 86 | 61 | 52 | 43 | 40 | 38 | 32 | 31 | 35 | 35 | 35 | 35 | 34 | 33 | 33 | 32 |
| 87 | 46 | 49 | 40 | 37 | 32 | 23 | 24 | 29 | 35 | 35 | 35 | 27 | 26 | 24 | 24 |
| 88 | 32 | 35 | 37 | 28 | 20 | 15 | 20 | 24 | 28 | 33 | 35 | 26 | 18 | 17 | 15 |
| 89 | 17 | 19 | 19 | 19 | 10 | 9 | 13 | 18 | 19 | 21 | 21 | 26 | 17 | 9 | 8 |
| 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Table 17e.

Relative Spatial Response of ULIS (times 10,000)

| AZ | ZA = 61 | ZA = 62 | ZA = 63 | ZA = 64 | ZA = 65 | ZA = 66 | ZA = 67 | ZA = 68 | ZA = 69 | ZA= 70 | ZA = 71 | ZA= 72 | ZA= 73 | ZA= 74 | ZA= 75 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| -89 | 7 | 6 | 7 | 8 | 8 | 11 | 20 | 32 | 24 | 9 | 5 | 6 | 8 | 13 | 18 |
| -88 | 20 | 14 | 14 | 15 | 16 | 20 | 25 | 32 | 26 | 14 | 14 | 19 | 23 | 28 | 33 |
| -87 | 33 | 26 | 22 | 23 | 25 | 30 | 35 | 39 | 36 | 28 | 20 | 24 | 29 | 33 | 27 |
| -86 | 48 | 40 | 33 | 31 | 34 | 38 | 43 | 48 | 47 | 41 | 35 | 29 | 34 | 32 | 34 |
| -85 | 61 | 55 | 49 | 44 | 42 | 47 | 51 | 56 | 59 | 53 | 49 | 44 | 40 | 40 | 42 |
| -84 | 97 | 84 | 78 | 75 | 72 | 70 | 69 | 68 | 65 | 67 | 59 | 56 | 53 | 62 | 46 |
| -83 | 206 | 216 | 214 | 225 | 198 | 174 | 178 | 182 | 189 | 183 | 169 | 173 | 119 | 106 | 99 |
| -82 | 370 | 374 | 348 | 322 | 308 | 287 | 291 | 298 | 282 | 277 | 224 | 187 | 171 | 159 | 181 |
| -81 | 480 | 454 | 441 | 417 | 405 | 398 | 402 | 398 | 347 | 268 | 242 | 238 | 254 | 283 | 253 |
| -80 | 572 | 557 | 536 | 524 | 504 | 496 | 473 | 405 | 329 | 307 | 308 | 342 | 371 | 344 | 285 |
| -79 | 691 | 684 | 678 | 674 | 645 | 559 | 450 | 373 | 371 | 394 | 410 | 437 | 435 | 377 | 341 |
| -78 | 858 | 834 | 811 | 769 | 719 | 570 | 434 | 446 | 468 | 493 | 520 | 526 | 470 | 419 | 387 |
| -77 | 996 | 945 | 902 | 837 | 706 | 661 | 611 | 563 | 563 | 590 | 602 | 564 | 513 | 481 | 458 |
| -76 | 1066 | 1035 | 975 | 896 | 859 | 859 | 811 | 750 | 733 | 690 | 642 | 605 | 557 | 522 | 474 |
| -75 | 1110 | 1053 | 1012 | 959 | 920 | 888 | 842 | 792 | 752 | 704 | 668 | 621 | 579 | 523 | 484 |
| -74 | 1142 | 1087 | 1049 | 995 | 945 | 905 | 854 | 801 | 765 | 718 | 678 | 625 | 579 | 526 | 489 |
| -73 | 1167 | 1111 | 1062 | 1004 | 953 | 914 | 861 | 827 | 777 | 734 | 682 | 639 | 584 | 539 | 503 |
| -72 | 1169 | 1109 | 1061 | 1012 | 961 | 923 | 871 | 833 | 781 | 739 | 687 | 642 | 599 | 544 | 502 |
| -71 | 1171 | 1112 | 1068 | 1020 | 981 | 921 | 882 | 832 | 790 | 739 | 697 | 642 | 599 | 548 | 507 |
| -70 | 1171 | 1116 | 1079 | 1027 | 978 | 918 | 880 | 830 | 789 | 738 | 696 | 654 | 602 | 546 | 509 |
| -69 | 1176 | 1128 | 1087 | 1027 | 975 | 921 | 879 | 838 | 788 | 746 | 704 | 649 | 597 | 555 | 516 |
| -68 | 1185 | 1146 | 1087 | 1033 | 978 | 929 | 883 | 842 | 800 | 751 | 705 | 653 | 606 | 555 | 519 |
| -67 | 1196 | 1147 | 1088 | 1037 | 981 | 933 | 892 | 842 | 802 | 755 | 709 | 654 | 608 | 566 | 523 |
| -66 | 1197 | 1149 | 1096 | 1041 | 991 | 937 | 896 | 856 | 802 | 756 | 710 | 664 | 614 | 575 | 528 |
| -65 | 1210 | 1151 | 1100 | 1045 | 996 | 948 | 902 | 857 | 812 | 766 | 721 | 667 | 616 | 574 | 526 |
| -64 | 1213 | 1149 | 1105 | 1061 | 1002 | 953 | 912 | 866 | 813 | 767 | 721 | 674 | 616 | 576 | 531 |
| -63 | 1211 | 1164 | 1102 | 1058 | 1008 | 958 | 910 | 864 | 819 | 773 | 725 | 672 | 616 | 577 | 537 |
| -62 | 1210 | 1163 | 1111 | 1055 | 1005 | 956 | 915 | 870 | 818 | 769 | 718 | 663 | 616 | 573 | 534 |
| -61 | 1220 | 1164 | 1111 | 1065 | 1012 | 960 | 914 | 868 | 820 | 771 | 715 | 663 | 617 | 579 | 536 |
| -60 | 1220 | 1174 | 1118 | 1066 | 1020 | 968 | 918 | 871 | 829 | 775 | 720 | 670 | 619 | 581 | 543 |
| -59 | 1227 | 1176 | 1130 | 1077 | 1028 | 961 | 914 | 873 | 821 | 766 | 712 | 666 | 621 | 588 | 544 |
| -58 | 1239 | 1193 | 1129 | 1080 | 1018 | 970 | 923 | 873 | 819 | 765 | 720 | 674 | 623 | 586 | 550 |
| -57 | 1245 | 1189 | 1140 | 1079 | 1024 | 967 | 919 | 872 | 818 | 765 | 722 | 677 | 633 | 593 | 551 |
| -56 | 1248 | 1192 | 1132 | 1077 | 1017 | 977 | 925 | 871 | 818 | 772 | 725 | 680 | 641 | 600 | 553 |
| -55 | 1251 | 1193 | 1139 | 1075 | 1025 | 974 | 931 | 878 | 825 | 773 | 729 | 690 | 645 | 601 | 555 |
| -54 | 1246 | 1192 | 1128 | 1076 | 1018 | 974 | 924 | 878 | 826 | 782 | 739 | 689 | 639 | 598 | 556 |
| -53 | 1252 | 1189 | 1135 | 1086 | 1028 | 973 | 930 | 878 | 834 | 786 | 736 | 690 | 641 | 592 | 558 |
| -52 | 1242 | 1188 | 1138 | 1081 | 1031 | 974 | 935 | 883 | 831 | 787 | 737 | 688 | 643 | 594 | 556 |
| -51 | 1248 | 1192 | 1133 | 1084 | 1035 | 978 | 925 | 880 | 826 | 788 | 739 | 689 | 641 | 597 | 557 |
| -50 | 1245 | 1190 | 1130 | 1082 | 1031 | 980 | 927 | 883 | 828 | 786 | 741 | 690 | 642 | 605 | 558 |
| -49 | 1250 | 1196 | 1134 | 1086 | 1028 | 978 | 925 | 886 | 830 | 788 | 739 | 692 | 643 | 602 | 557 |
| -48 | 1249 | 1185 | 1139 | 1090 | 1031 | 981 | 929 | 883 | 832 | 791 | 741 | 694 | 651 | 603 | 557 |
| -47 | 1256 | 1192 | 1138 | 1089 | 1029 | 985 | 933 | 885 | 830 | 794 | 743 | 692 | 654 | 604 | 558 |
| -46 | 1265 | 1201 | 1146 | 1093 | 1034 | 984 | 935 | 888 | 833 | 792 | 746 | 703 | 653 | 603 | 563 |
| -45 | 1266 | 1199 | 1140 | 1094 | 1037 | 989 | 933 | 891 | 836 | 795 | 747 | 702 | 648 | 607 | 560 |
| -44 | 1269 | 1199 | 1140 | 1089 | 1033 | 990 | 943 | 889 | 848 | 796 | 751 | 700 | 655 | 605 | 557 |
| -43 | 1263 | 1205 | 1141 | 1090 | 1043 | 983 | 941 | 889 | 844 | 793 | 749 | 699 | 654 | 604 | 562 |
| -42 | 1264 | 1205 | 1136 | 1092 | 1045 | 982 | 936 | 896 | 840 | 798 | 746 | 702 | 652 | 608 | 559 |
| -41 | 1259 | 1206 | 1148 | 1088 | 1045 | 991 | 936 | 896 | 837 | 792 | 752 | 697 | 656 | 604 | 564 |
| -40 | 1260 | 1202 | 1150 | 1090 | 1040 | 991 | 935 | 891 | 844 | 790 | 749 | 693 | 652 | 610 | 560 |
| -39 | 1262 | 1204 | 1147 | 1102 | 1040 | 986 | 944 | 889 | 842 | 796 | 744 | 699 | 646 | 606 | 566 |
| -38 | 1269 | 1218 | 1149 | 1099 | 1050 | 987 | 940 | 896 | 836 | 794 | 750 | 696 | 651 | 603 | 563 |
| -37 | 1272 | 1214 | 1163 | 1099 | 1046 | 998 | 940 | 894 | 844 | 800 | 747 | 703 | 648 | 609 | 567 |
| -36 | 1277 | 1219 | 1161 | 1108 | 1047 | 999 | 949 | 899 | 843 | 796 | 753 | 700 | 654 | 615 | 564 |
| -35 | 1287 | 1217 | 1154 | 1099 | 1053 | 1003 | 945 | 898 | 852 | 794 | 751 | 704 | 651 | 612 | 570 |
| -34 | 1283 | 1220 | 1162 | 1095 | 1048 | 998 | 950 | 892 | 847 | 800 | 756 | 702 | 659 | 616 | 562 |
| -33 | 1277 | 1215 | 1154 | 1100 | 1051 | 994 | 941 | 896 | 849 | 793 | 750 | 705 | 663 | 607 | 564 |
| -32 | 1281 | 1223 | 1150 | 1096 | 1047 | 1000 | 947 | 890 | 844 | 798 | 751 | 699 | 656 | 609 | 567 |
| -31 | 1276 | 1214 | 1159 | 1101 | 1054 | 992 | 942 | 893 | 849 | 803 | 745 | 703 | 658 | 602 | 559 |
| -30 | 1284 | 1223 | 1156 | 1097 | 1050 | 998 | 948 | 898 | 844 | 794 | 748 | 707 | 651 | 605 | 562 |
| -29 | 1275 | 1220 | 1161 | 1107 | 1054 | 994 | 940 | 893 | 849 | 800 | 743 | 701 | 655 | 609 | 565 |
| -28 | 1285 | 1229 | 1158 | 1104 | 1050 | 1001 | 948 | 899 | 843 | 794 | 748 | 703 | 659 | 602 | 558 |
| -27 | 1278 | 1221 | 1167 | 1110 | 1058 | 1005 | 944 | 895 | 846 | 800 | 754 | 707 | 652 | 607 | 560 |
| -26 | 1281 | 1227 | 1174 | 1104 | 1049 | 999 | 950 | 901 | 852 | 806 | 749 | 702 | 657 | 611 | 564 |
| -25 | 1275 | 1222 | 1164 | 1109 | 1052 | 1002 | 953 | 902 | 843 | 798 | 751 | 706 | 660 | 615 | 558 |
| -24 | 1282 | 1225 | 1169 | 1114 | 1057 | 993 | 945 | 894 | 845 | 800 | 753 | 707 | 661 | 605 | 560 |
| -23 | 1287 | 1230 | 1162 | 1104 | 1049 | 997 | 947 | 898 | 849 | 801 | 756 | 698 | 652 | 606 | 562 |
| -22 | 1276 | 1222 | 1167 | 1110 | 1055 | 999 | 951 | 902 | 841 | 792 | 747 | 701 | 654 | 608 | 564 |
| -21 | 1280 | 1227 | 1171 | 1116 | 1061 | 991 | 943 | 894 | 843 | 795 | 750 | 704 | 655 | 610 | 566 |
| -20 | 1287 | 1229 | 1162 | 1109 | 1052 | 998 | 947 | 899 | 847 | 799 | 753 | 707 | 658 | 601 | 557 |
| -19 | 1276 | 1220 | 1167 | 1112 | 1058 | 1004 | 952 | 901 | 852 | 803 | 756 | 697 | 650 | 603 | 559 |
| -18 | 1285 | 1225 | 1172 | 1117 | 1064 | 1011 | 956 | 906 | 845 | 796 | 748 | 700 | 653 | 607 | 562 |
| -17 | 1294 | 1231 | 1175 | 1122 | 1056 | 1003 | 947 | 898 | 850 | 801 | 750 | 704 | 657 | 610 | 564 |
| -16 | 1288 | 1220 | 1167 | 1114 | 1062 | 1005 | 951 | 903 | 854 | 803 | 753 | 706 | 659 | 613 | 565 |
| -15 | 1295 | 1230 | 1172 | 1120 | 1064 | 1009 | 956 | 907 | 857 | 805 | 755 | 709 | 662 | 614 | 565 |
| -14 | 1305 | 1240 | 1178 | 1122 | 1068 | 1013 | 960 | 910 | 858 | 807 | 758 | 711 | 651 | 603 | 554 |
| -13 | 1308 | 1244 | 1180 | 1125 | 1071 | 1017 | 963 | 911 | 861 | 797 | 748 | 700 | 652 | 604 | 555 |
| -12 | 1311 | 1245 | 1182 | 1127 | 1059 | 1005 | 950 | 900 | 849 | 798 | 748 | 701 | 653 | 604 | 556 |
| -11 | 1296 | 1232 | 1171 | 1115 | 1061 | 1006 | 952 | 901 | 850 | 800 | 750 | 702 | 654 | 606 | 558 |
| -10 | 1299 | 1235 | 1175 | 1118 | 1062 | 1008 | 955 | 902 | 851 | 801 | 750 | 701 | 655 | 607 | 559 |
| -9 | 1302 | 1238 | 1177 | 1120 | 1064 | 1010 | 957 | 904 | 852 | 802 | 751 | 702 | 655 | 609 | 561 |
| -8 | 1307 | 1240 | 1181 | 1123 | 1067 | 1013 | 960 | 906 | 854 | 803 | 753 | 703 | 655 | 609 | 563 |
| -7 | 1310 | 1244 | 1185 | 1128 | 1071 | 1016 | 962 | 909 | 855 | 805 | 754 | 705 | 655 | 609 | 563 |
| -6 | 1316 | 1248 | 1189 | 1132 | 1075 | 1019 | 964 | 911 | 858 | 806 | 756 | 706 | 657 | 609 | 563 |
| -5 | 1322 | 1254 | 1194 | 1136 | 1078 | 1021 | 967 | 914 | 861 | 809 | 758 | 708 | 659 | 610 | 564 |
| -4 | 1326 | 1256 | 1197 | 1139 | 1082 | 1025 | 970 | 917 | 864 | 811 | 760 | 710 | 660 | 610 | 563 |
| -3 | 1328 | 1261 | 1201 | 1144 | 1087 | 1026 | 971 | 918 | 865 | 810 | 759 | 710 | 660 | 610 | 563 |
| -2 | 1333 | 1266 | 1206 | 1149 | 1088 | 1028 | 973 | 920 | 865 | 810 | 759 | 710 | 660 | 610 | 562 |
| -1 | 1338 | 1271 | 1211 | 1151 | 1090 | 1030 | 975 | 920 | 865 | 810 | 759 | 710 | 660 | 610 | 563 |
| 0 | 1325 | 1258 | 1197 | 1137 | 1077 | 1016 | 961 | 906 | 851 | 796 | 747 | 697 | 648 | 598 | 551 |
| 1 | 1340 | 1273 | 1212 | 1152 | 1092 | 1032 | 975 | 920 | 865 | 810 | 760 | 710 | 660 | 611 | 563 |
| 2 | 1339 | 1272 | 1211 | 1152 | 1092 | 1032 | 976 | 921 | 866 | 811 | 760 | 710 | 661 | 611 | 564 |
| 3 | 1339 | 1271 | 1210 | 1151 | 1092 | 1032 | 976 | 921 | 866 | 811 | 760 | 711 | 661 | 612 | 564 |
| 4 | 1341 | 1271 | 1209 | 1150 | 1091 | 1032 | 977 | 922 | 867 | 812 | 761 | 712 | 662 | 612 | 565 |
| 5 | 1342 | 1271 | 1208 | 1149 | 1090 | 1031 | 976 | 921 | 866 | 811 | 760 | 710 | 660 | 611 | 563 |
| 6 | 1342 | 1270 | 1207 | 1148 | 1089 | 1031 | 974 | 919 | 864 | 809 | 758 | 709 | 659 | 609 | 562 |
| 7 | 1341 | 1269 | 1206 | 1147 | 1088 | 1029 | 973 | 919 | 863 | 808 | 757 | 707 | 657 | 607 | 560 |
| 8 | 1341 | 1269 | 1205 | 1146 | 1087 | 1028 | 972 | 918 | 863 | 807 | 756 | 706 | 656 | 605 | 558 |
| 9 | 1341 | 1269 | 1204 | 1145 | 1086 | 1027 | 971 | 917 | 862 | 807 | 755 | 705 | 654 | 604 | 556 |
| 10 | 1341 | 1269 | 1205 | 1145 | 1086 | 1027 | 970 | 916 | 861 | 806 | 754 | 703 | 653 | 602 | 555 |
| 11 | 1342 | 1270 | 1205 | 1145 | 1086 | 1026 | 970 | 915 | 860 | 806 | 753 | 703 | 651 | 601 | 553 |
| 12 | 1361 | 1289 | 1219 | 1160 | 1086 | 1027 | 969 | 915 | 860 | 805 | 753 | 702 | 651 | 599 | 552 |
| 13 | 1363 | 1290 | 1220 | 1160 | 1101 | 1041 | 983 | 927 | 872 | 804 | 752 | 701 | 650 | 599 | 550 |
| 14 | 1361 | 1288 | 1219 | 1158 | 1098 | 1039 | 980 | 926 | 870 | 815 | 762 | 712 | 649 | 598 | 549 |
| 15 | 1353 | 1280 | 1217 | 1157 | 1096 | 1036 | 977 | 923 | 870 | 814 | 760 | 710 | 660 | 610 | 560 |
| 16 | 1349 | 1275 | 1215 | 1155 | 1095 | 1033 | 974 | 920 | 867 | 811 | 757 | 708 | 658 | 608 | 559 |
| 17 | 1361 | 1289 | 1228 | 1168 | 1093 | 1033 | 971 | 917 | 864 | 810 | 755 | 706 | 656 | 606 | 557 |
| 18 | 1355 | 1289 | 1228 | 1166 | 1106 | 1046 | 980 | 927 | 860 | 806 | 752 | 703 | 653 | 603 | 554 |
| 19 | 1349 | 1288 | 1227 | 1165 | 1103 | 1039 | 977 | 922 | 869 | 815 | 761 | 700 | 650 | 600 | 551 |
| 20 | 1365 | 1302 | 1226 | 1164 | 1096 | 1031 | 973 | 920 | 865 | 811 | 757 | 708 | 659 | 597 | 548 |
| 21 | 1363 | 1302 | 1237 | 1172 | 1107 | 1025 | 968 | 915 | 861 | 807 | 753 | 704 | 656 | 607 | 557 |
| 22 | 1363 | 1299 | 1234 | 1165 | 1100 | 1034 | 978 | 923 | 858 | 804 | 749 | 699 | 653 | 604 | 554 |
| 23 | 1377 | 1308 | 1227 | 1159 | 1093 | 1033 | 974 | 918 | 866 | 812 | 759 | 695 | 649 | 600 | 552 |
| 24 | 1371 | 1302 | 1237 | 1171 | 1102 | 1029 | 972 | 914 | 862 | 810 | 755 | 703 | 656 | 596 | 549 |
| 25 | 1366 | 1300 | 1231 | 1165 | 1096 | 1039 | 983 | 924 | 858 | 805 | 751 | 699 | 652 | 605 | 544 |
| 26 | 1377 | 1311 | 1244 | 1159 | 1096 | 1036 | 979 | 922 | 866 | 813 | 748 | 693 | 646 | 599 | 549 |
| 27 | 1375 | 1306 | 1239 | 1169 | 1104 | 1045 | 974 | 916 | 861 | 809 | 756 | 700 | 640 | 593 | 544 |
| 28 | 1387 | 1319 | 1233 | 1166 | 1096 | 1042 | 981 | 924 | 857 | 804 | 751 | 695 | 646 | 587 | 541 |
| 29 | 1377 | 1311 | 1242 | 1175 | 1104 | 1035 | 976 | 918 | 861 | 810 | 746 | 693 | 640 | 593 | 547 |
| 30 | 1391 | 1314 | 1240 | 1168 | 1101 | 1039 | 986 | 926 | 855 | 803 | 754 | 701 | 635 | 587 | 543 |
| 31 | 1383 | 1303 | 1244 | 1178 | 1110 | 1031 | 978 | 921 | 862 | 811 | 748 | 697 | 641 | 583 | 538 |
| 32 | 1392 | 1315 | 1233 | 1174 | 1104 | 1039 | 982 | 917 | 857 | 804 | 752 | 692 | 639 | 590 | 545 |
| 33 | 1390 | 1308 | 1236 | 1177 | 1112 | 1031 | 974 | 921 | 864 | 798 | 749 | 696 | 646 | 586 | 541 |
| 34 | 1400 | 1318 | 1244 | 1167 | 1107 | 1041 | 982 | 913 | 860 | 803 | 754 | 690 | 640 | 592 | 537 |
| 35 | 1406 | 1316 | 1234 | 1170 | 1110 | 1045 | 975 | 916 | 862 | 794 | 747 | 694 | 633 | 588 | 543 |
| 36 | 1398 | 1321 | 1242 | 1178 | 1100 | 1040 | 976 | 917 | 851 | 795 | 751 | 691 | 638 | 589 | 535 |
| 37 | 1399 | 1322 | 1252 | 1169 | 1104 | 1044 | 966 | 911 | 852 | 800 | 742 | 696 | 631 | 581 | 537 |
| 38 | 1404 | 1333 | 1241 | 1167 | 1114 | 1035 | 969 | 912 | 840 | 791 | 741 | 690 | 636 | 574 | 532 |
| 39 | 1396 | 1323 | 1247 | 1176 | 1101 | 1040 | 979 | 902 | 847 | 793 | 731 | 691 | 630 | 579 | 533 |
| 40 | 1388 | 1324 | 1258 | 1166 | 1099 | 1046 | 970 | 904 | 849 | 784 | 736 | 680 | 639 | 584 | 526 |
| 41 | 1387 | 1325 | 1260 | 1172 | 1102 | 1044 | 974 | 913 | 839 | 786 | 738 | 679 | 638 | 578 | 528 |
| 42 | 1396 | 1317 | 1242 | 1184 | 1106 | 1031 | 972 | 917 | 842 | 793 | 730 | 678 | 629 | 582 | 521 |
| 43 | 1397 | 1314 | 1242 | 1178 | 1110 | 1029 | 976 | 905 | 845 | 785 | 732 | 673 | 627 | 577 | 526 |
| 44 | 1407 | 1308 | 1235 | 1171 | 1094 | 1033 | 974 | 904 | 850 | 789 | 735 | 675 | 627 | 576 | 521 |
| 45 | 1405 | 1310 | 1228 | 1171 | 1095 | 1032 | 963 | 909 | 839 | 785 | 732 | 678 | 618 | 576 | 521 |
| 46 | 1408 | 1314 | 1234 | 1167 | 1092 | 1029 | 967 | 910 | 839 | 781 | 727 | 678 | 621 | 566 | 519 |
| 47 | 1402 | 1310 | 1233 | 1160 | 1091 | 1034 | 959 | 907 | 839 | 785 | 723 | 664 | 619 | 565 | 513 |
| 48 | 1395 | 1310 | 1241 | 1163 | 1094 | 1032 | 957 | 900 | 846 | 785 | 719 | 665 | 617 | 563 | 511 |
| 49 | 1393 | 1327 | 1241 | 1163 | 1082 | 1028 | 955 | 900 | 839 | 785 | 715 | 661 | 607 | 561 | 511 |
| 50 | 1381 | 1312 | 1238 | 1163 | 1079 | 1025 | 961 | 896 | 832 | 779 | 721 | 657 | 605 | 564 | 511 |
| 51 | 1380 | 1308 | 1234 | 1169 | 1093 | 1013 | 950 | 894 | 825 | 779 | 717 | 656 | 602 | 552 | 507 |
| 52 | 1368 | 1297 | 1227 | 1154 | 1089 | 1002 | 956 | 892 | 825 | 772 | 711 | 651 | 601 | 546 | 503 |
| 53 | 1380 | 1297 | 1224 | 1158 | 1085 | 1011 | 945 | 881 | 826 | 766 | 708 | 653 | 595 | 541 | 503 |
| 54 | 1367 | 1296 | 1214 | 1142 | 1069 | 1005 | 934 | 881 | 817 | 763 | 708 | 650 | 590 | 544 | 498 |
| 55 | 1366 | 1293 | 1222 | 1142 | 1074 | 998 | 936 | 880 | 816 | 753 | 696 | 648 | 592 | 545 | 492 |
| 56 | 1356 | 1281 | 1206 | 1136 | 1056 | 989 | 924 | 868 | 809 | 753 | 692 | 636 | 588 | 537 | 483 |
| 57 | 1343 | 1272 | 1209 | 1131 | 1060 | 971 | 911 | 862 | 807 | 745 | 688 | 632 | 576 | 528 | 475 |
| 58 | 1334 | 1271 | 1187 | 1123 | 1045 | 975 | 915 | 858 | 802 | 746 | 689 | 628 | 565 | 518 | 474 |
| 59 | 1322 | 1250 | 1187 | 1114 | 1050 | 961 | 904 | 854 | 797 | 742 | 683 | 618 | 561 | 521 | 464 |
| 60 | 1315 | 1251 | 1175 | 1102 | 1037 | 965 | 908 | 851 | 801 | 743 | 689 | 626 | 558 | 512 | 461 |
| 61 | 1317 | 1241 | 1168 | 1103 | 1030 | 957 | 903 | 848 | 790 | 732 | 676 | 620 | 556 | 508 | 452 |
| 62 | 1308 | 1243 | 1170 | 1094 | 1023 | 954 | 905 | 850 | 786 | 729 | 670 | 615 | 556 | 499 | 450 |
| 63 | 1311 | 1247 | 1162 | 1099 | 1028 | 957 | 899 | 844 | 789 | 733 | 675 | 617 | 555 | 498 | 455 |
| 64 | 1315 | 1227 | 1164 | 1101 | 1018 | 946 | 894 | 845 | 776 | 726 | 670 | 613 | 549 | 490 | 443 |
| 65 | 1316 | 1230 | 1158 | 1080 | 1010 | 939 | 851 | 794 | 745 | 696 | 648 | 583 | 536 | 471 | 413 |
| 66 | 1301 | 1227 | 1152 | 1074 | 1003 | 922 | 842 | 763 | 690 | 641 | 592 | 544 | 495 | 465 | 391 |
| 67 | 1303 | 1229 | 1141 | 1068 | 989 | 916 | 835 | 738 | 658 | 604 | 555 | 488 | 440 | 412 | 370 |
| 68 | 1306 | 1232 | 1141 | 1064 | 984 | 911 | 821 | 739 | 658 | 560 | 502 | 450 | 401 | 343 | 338 |
| 69 | 1319 | 1229 | 1145 | 1054 | 981 | 898 | 816 | 735 | 636 | 554 | 472 | 397 | 346 | 301 | 283 |
| 70 | 1316 | 1232 | 1158 | 1058 | 985 | 895 | 821 | 722 | 640 | 540 | 457 | 373 | 311 | 246 | 210 |
| 71 | 1301 | 1219 | 1152 | 1068 | 996 | 901 | 827 | 728 | 645 | 545 | 461 | 352 | 267 | 208 | 164 |
| 72 | 1287 | 1200 | 1136 | 1054 | 968 | 906 | 797 | 720 | 614 | 532 | 427 | 341 | 258 | 172 | 110 |
| 73 | 1281 | 1193 | 1123 | 1039 | 956 | 900 | 827 | 718 | 589 | 506 | 400 | 317 | 207 | 141 | 184 |
| 74 | 1255 | 1168 | 1104 | 1018 | 949 | 895 | 826 | 749 | 628 | 483 | 375 | 269 | 181 | 102 | 153 |
| 75 | 1201 | 1124 | 1063 | 984 | 915 | 871 | 811 | 747 | 690 | 554 | 424 | 263 | 160 | 70 | 119 |
| 76 | 1135 | 1079 | 1003 | 940 | 873 | 832 | 766 | 703 | 665 | 610 | 464 | 333 | 182 | 36 | 79 |
| 77 | 1067 | 982 | 909 | 834 | 794 | 760 | 693 | 639 | 607 | 568 | 528 | 398 | 246 | 98 | 34 |
| 78 | 899 | 855 | 807 | 744 | 703 | 663 | 638 | 593 | 564 | 525 | 486 | 459 | 315 | 161 | 27 |
| 79 | 672 | 661 | 647 | 632 | 576 | 536 | 535 | 538 | 530 | 482 | 440 | 401 | 377 | 234 | 97 |
| 80 | 508 | 489 | 459 | 444 | 414 | 404 | 427 | 443 | 448 | 443 | 410 | 371 | 332 | 293 | 170 |
| 81 | 424 | 393 | 373 | 344 | 325 | 313 | 320 | 318 | 337 | 343 | 330 | 339 | 302 | 248 | 208 |
| 82 | 323 | 324 | 293 | 262 | 242 | 218 | 224 | 227 | 211 | 211 | 245 | 235 | 224 | 222 | 183 |
| 83 | 173 | 169 | 168 | 176 | 144 | 120 | 125 | 129 | 134 | 125 | 109 | 120 | 117 | 113 | 116 |
| 84 | 55 | 46 | 44 | 43 | 44 | 44 | 44 | 41 | 42 | 37 | 29 | 22 | 16 | 13 | 21 |
| 85 | 38 | 37 | 37 | 35 | 37 | 39 | 41 | 44 | 43 | 39 | 34 | 27 | 24 | 24 | 25 |
| 86 | 31 | 28 | 27 | 28 | 32 | 36 | 38 | 41 | 38 | 33 | 26 | 25 | 27 | 24 | 26 |
| 87 | 22 | 20 | 20 | 21 | 24 | 29 | 35 | 37 | 32 | 23 | 18 | 22 | 26 | 30 | 21 |
| 88 | 14 | 13 | 13 | 14 | 16 | 20 | 26 | 36 | 23 | 10 | 14 | 19 | 23 | 27 | 32 |
| 89 | 6 | 6 | 6 | 7 | 8 | 12 | 24 | 37 | 21 | 5 | 5 | 5 | 9 | 13 | 19 |
| 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Table 17f.

Relative Spatial Response of ULIS (times 10,000)

| AZ | ZA= 76 | ZA= 77 | ZA= 78 | ZA= 79 | ZA= 80 | ZA= 81 | ZA= 82 | ZA= 83 | ZA= 84 | ZA= 85 | ZA= 86 | ZA= 87 | ZA= 88 | ZA= 89 | ZA= 90 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| -89 | 24 | 25 | 19 | 8 | 7 | 8 | 10 | 13 | 16 | 20 | 26 | 35 | 33 | 22 | 6 |
| -88 | 32 | 22 | 13 | 13 | 16 | 20 | 25 | 32 | 36 | 27 | 13 | 33 | 26 | 17 | 11 |
| -87 | 23 | 27 | 30 | 34 | 35 | 36 | 33 | 24 | 19 | 21 | 33 | 30 | 23 | 17 | 6 |
| -86 | 35 | 36 | 37 | 35 | 31 | 24 | 26 | 26 | 29 | 31 | 33 | 28 | 23 | 17 | 4 |
| -85 | 41 | 34 | 31 | 32 | 34 | 32 | 31 | 30 | 30 | 49 | 27 | 30 | 24 | 17 | 0 |
| -84 | 40 | 34 | 40 | 38 | 33 | 30 | 28 | 51 | 77 | 65 | 25 | 34 | 24 | 22 | 0 |
| -83 | 93 | 86 | 32 | 33 | 30 | 89 | 89 | 110 | 107 | 70 | 56 | 30 | 24 | 21 | 0 |
| -82 | 177 | 118 | 60 | 93 | 151 | 128 | 143 | 148 | 108 | 63 | 62 | 30 | 27 | 22 | 0 |
| -81 | 210 | 152 | 156 | 214 | 161 | 177 | 189 | 151 | 103 | 94 | 58 | 30 | 26 | 18 | 0 |
| -80 | 247 | 222 | 277 | 225 | 210 | 221 | 192 | 151 | 107 | 95 | 54 | 30 | 26 | 14 | 0 |
| -79 | 299 | 329 | 290 | 250 | 251 | 234 | 194 | 146 | 128 | 94 | 56 | 30 | 26 | 13 | 8 |
| -78 | 394 | 354 | 305 | 280 | 275 | 231 | 187 | 144 | 124 | 76 | 58 | 30 | 24 | 15 | 17 |
| -77 | 420 | 371 | 332 | 309 | 272 | 230 | 185 | 157 | 112 | 74 | 62 | 30 | 24 | 20 | 25 |
| -76 | 435 | 398 | 350 | 311 | 273 | 231 | 190 | 140 | 93 | 78 | 66 | 36 | 28 | 25 | 34 |
| -75 | 439 | 401 | 356 | 317 | 272 | 232 | 186 | 124 | 94 | 81 | 70 | 41 | 33 | 34 | 0 |
| -74 | 452 | 407 | 357 | 320 | 282 | 234 | 189 | 112 | 97 | 85 | 73 | 48 | 38 | 28 | 0 |
| -73 | 451 | 409 | 368 | 330 | 286 | 239 | 196 | 129 | 99 | 90 | 77 | 55 | 45 | 20 | 0 |
| -72 | 461 | 420 | 378 | 327 | 289 | 248 | 205 | 173 | 140 | 113 | 85 | 62 | 39 | 20 | 0 |
| -71 | 460 | 419 | 374 | 331 | 288 | 250 | 210 | 181 | 150 | 125 | 89 | 69 | 46 | 21 | 0 |
| -70 | 469 | 426 | 378 | 338 | 298 | 253 | 212 | 187 | 159 | 129 | 95 | 76 | 53 | 28 | 0 |
| -69 | 475 | 428 | 386 | 343 | 307 | 250 | 217 | 193 | 165 | 131 | 101 | 84 | 62 | 32 | 0 |
| -68 | 474 | 429 | 388 | 348 | 302 | 255 | 225 | 202 | 168 | 133 | 107 | 89 | 59 | 31 | 0 |
| -67 | 481 | 436 | 397 | 346 | 293 | 263 | 232 | 202 | 170 | 136 | 114 | 91 | 59 | 30 | 0 |
| -66 | 485 | 441 | 401 | 348 | 295 | 268 | 240 | 203 | 169 | 137 | 117 | 90 | 59 | 28 | 0 |
| -65 | 483 | 441 | 400 | 347 | 301 | 271 | 238 | 205 | 172 | 141 | 117 | 93 | 62 | 32 | 0 |
| -64 | 489 | 446 | 394 | 346 | 304 | 274 | 240 | 208 | 175 | 143 | 121 | 93 | 66 | 37 | 0 |
| -63 | 493 | 445 | 399 | 346 | 311 | 276 | 242 | 211 | 182 | 146 | 123 | 97 | 68 | 42 | 0 |
| -62 | 495 | 450 | 399 | 345 | 317 | 277 | 245 | 214 | 182 | 151 | 125 | 98 | 72 | 45 | 0 |
| -61 | 495 | 449 | 404 | 351 | 322 | 280 | 245 | 217 | 180 | 154 | 129 | 103 | 77 | 44 | 0 |
| -60 | 500 | 450 | 410 | 358 | 320 | 282 | 248 | 217 | 179 | 156 | 130 | 104 | 76 | 43 | 0 |
| -59 | 500 | 457 | 410 | 364 | 322 | 285 | 252 | 216 | 185 | 160 | 135 | 108 | 80 | 58 | 16 |
| -58 | 507 | 461 | 411 | 372 | 325 | 288 | 252 | 219 | 186 | 161 | 137 | 109 | 82 | 69 | 32 |
| -57 | 510 | 456 | 412 | 373 | 328 | 292 | 250 | 218 | 187 | 163 | 135 | 107 | 85 | 74 | 48 |
| -56 | 507 | 458 | 414 | 376 | 337 | 295 | 257 | 220 | 187 | 160 | 136 | 109 | 87 | 59 | 65 |
| -55 | 502 | 460 | 416 | 376 | 340 | 295 | 260 | 228 | 189 | 160 | 133 | 109 | 78 | 42 | 0 |
| -54 | 504 | 459 | 418 | 379 | 344 | 305 | 264 | 232 | 198 | 162 | 133 | 110 | 79 | 43 | 0 |
| -53 | 511 | 460 | 419 | 379 | 345 | 309 | 274 | 240 | 202 | 171 | 141 | 107 | 80 | 43 | 0 |
| -52 | 512 | 467 | 427 | 384 | 342 | 306 | 272 | 238 | 206 | 176 | 142 | 112 | 80 | 43 | 0 |
| -51 | 514 | 468 | 429 | 386 | 347 | 309 | 276 | 242 | 201 | 170 | 141 | 112 | 79 | 42 | 0 |
| -50 | 512 | 470 | 427 | 388 | 349 | 307 | 274 | 240 | 205 | 171 | 140 | 111 | 79 | 42 | 0 |
| -49 | 512 | 468 | 426 | 395 | 355 | 316 | 279 | 243 | 209 | 176 | 142 | 111 | 78 | 42 | 0 |
| -48 | 511 | 474 | 428 | 390 | 351 | 313 | 272 | 237 | 205 | 172 | 139 | 106 | 72 | 36 | 0 |
| -47 | 518 | 472 | 424 | 386 | 347 | 314 | 271 | 233 | 201 | 168 | 136 | 106 | 72 | 36 | 0 |
| -46 | 518 | 471 | 430 | 389 | 348 | 305 | 264 | 232 | 197 | 165 | 134 | 107 | 73 | 36 | 0 |
| -45 | 514 | 474 | 428 | 384 | 346 | 304 | 264 | 226 | 188 | 162 | 135 | 108 | 73 | 40 | 0 |
| -44 | 519 | 470 | 425 | 389 | 344 | 302 | 267 | 227 | 189 | 158 | 135 | 108 | 88 | 78 | 71 |
| -43 | 516 | 475 | 429 | 385 | 350 | 308 | 266 | 231 | 193 | 160 | 131 | 109 | 77 | 60 | 53 |
| -42 | 521 | 472 | 427 | 390 | 349 | 315 | 272 | 228 | 197 | 164 | 132 | 110 | 74 | 42 | 36 |
| -41 | 515 | 477 | 432 | 389 | 355 | 314 | 277 | 235 | 195 | 168 | 135 | 106 | 75 | 37 | 18 |
| -40 | 518 | 473 | 435 | 388 | 348 | 314 | 272 | 238 | 197 | 171 | 139 | 105 | 80 | 42 | 0 |
| -39 | 515 | 477 | 430 | 389 | 346 | 313 | 270 | 229 | 196 | 165 | 139 | 108 | 94 | 61 | 19 |
| -38 | 521 | 472 | 432 | 383 | 345 | 305 | 269 | 227 | 195 | 165 | 139 | 109 | 99 | 80 | 38 |
| -37 | 518 | 478 | 427 | 388 | 340 | 304 | 261 | 227 | 189 | 164 | 139 | 110 | 100 | 88 | 57 |
| -36 | 524 | 471 | 428 | 389 | 341 | 304 | 261 | 226 | 187 | 162 | 132 | 110 | 85 | 69 | 77 |
| -35 | 517 | 473 | 430 | 382 | 343 | 295 | 261 | 227 | 188 | 159 | 128 | 103 | 69 | 38 | 0 |
| -34 | 519 | 474 | 422 | 383 | 345 | 296 | 262 | 220 | 189 | 159 | 124 | 99 | 68 | 38 | 0 |
| -33 | 521 | 466 | 423 | 385 | 338 | 297 | 262 | 221 | 190 | 154 | 124 | 96 | 65 | 38 | 0 |
| -32 | 514 | 469 | 426 | 378 | 340 | 298 | 255 | 221 | 191 | 155 | 126 | 97 | 62 | 36 | 0 |
| -31 | 517 | 472 | 418 | 380 | 342 | 291 | 256 | 222 | 185 | 156 | 127 | 92 | 63 | 35 | 0 |
| -30 | 518 | 465 | 421 | 382 | 344 | 293 | 256 | 223 | 187 | 158 | 129 | 94 | 65 | 36 | 0 |
| -29 | 511 | 468 | 424 | 384 | 335 | 296 | 257 | 216 | 188 | 159 | 131 | 95 | 67 | 38 | 0 |
| -28 | 514 | 471 | 428 | 377 | 338 | 298 | 259 | 218 | 189 | 161 | 132 | 97 | 68 | 39 | 0 |
| -27 | 518 | 474 | 421 | 380 | 340 | 301 | 252 | 219 | 191 | 162 | 127 | 98 | 70 | 39 | 0 |
| -26 | 521 | 468 | 425 | 383 | 343 | 294 | 255 | 220 | 192 | 164 | 128 | 100 | 70 | 39 | 0 |
| -25 | 514 | 472 | 429 | 385 | 345 | 297 | 258 | 222 | 193 | 165 | 129 | 101 | 70 | 43 | 0 |
| -24 | 515 | 473 | 430 | 387 | 337 | 298 | 260 | 223 | 194 | 159 | 131 | 107 | 98 | 88 | 82 |
| -23 | 517 | 474 | 431 | 378 | 337 | 299 | 260 | 223 | 195 | 160 | 131 | 100 | 77 | 68 | 61 |
| -22 | 519 | 464 | 422 | 379 | 338 | 299 | 260 | 224 | 189 | 161 | 131 | 100 | 69 | 47 | 41 |
| -21 | 510 | 465 | 423 | 381 | 339 | 299 | 261 | 218 | 190 | 161 | 131 | 100 | 69 | 38 | 20 |
| -20 | 512 | 468 | 424 | 382 | 341 | 300 | 261 | 219 | 191 | 161 | 131 | 100 | 69 | 38 | 0 |
| -19 | 515 | 470 | 425 | 384 | 342 | 301 | 252 | 219 | 191 | 161 | 131 | 100 | 70 | 38 | 20 |
| -18 | 517 | 471 | 427 | 385 | 344 | 293 | 253 | 220 | 191 | 161 | 131 | 100 | 69 | 58 | 40 |
| -17 | 520 | 474 | 430 | 377 | 336 | 295 | 254 | 221 | 191 | 161 | 131 | 100 | 89 | 78 | 60 |
| -16 | 520 | 476 | 421 | 378 | 337 | 296 | 255 | 220 | 190 | 161 | 131 | 109 | 89 | 73 | 80 |
| -15 | 509 | 465 | 421 | 378 | 337 | 296 | 255 | 220 | 190 | 159 | 130 | 100 | 69 | 38 | 0 |
| -14 | 509 | 465 | 421 | 378 | 337 | 296 | 255 | 221 | 190 | 159 | 129 | 99 | 69 | 38 | 0 |
| -13 | 509 | 465 | 421 | 378 | 337 | 296 | 255 | 221 | 190 | 159 | 129 | 98 | 69 | 38 | 0 |
| -12 | 509 | 465 | 421 | 378 | 337 | 296 | 255 | 221 | 191 | 160 | 129 | 98 | 68 | 38 | 0 |
| -11 | 510 | 465 | 422 | 379 | 337 | 296 | 256 | 222 | 191 | 160 | 130 | 99 | 68 | 38 | 0 |
| -10 | 512 | 466 | 422 | 380 | 337 | 296 | 256 | 222 | 192 | 161 | 130 | 100 | 69 | 38 | 0 |
| -9 | 513 | 466 | 423 | 380 | 338 | 297 | 256 | 223 | 192 | 161 | 131 | 100 | 70 | 39 | 0 |
| -8 | 515 | 468 | 423 | 381 | 339 | 297 | 257 | 223 | 193 | 162 | 132 | 101 | 69 | 38 | 0 |
| -7 | 517 | 469 | 424 | 382 | 340 | 298 | 257 | 224 | 193 | 163 | 132 | 100 | 68 | 38 | 0 |
| -6 | 518 | 471 | 425 | 383 | 341 | 299 | 257 | 224 | 194 | 163 | 131 | 99 | 68 | 38 | 0 |
| -5 | 518 | 472 | 427 | 384 | 342 | 300 | 258 | 225 | 193 | 161 | 129 | 98 | 68 | 42 | 0 |
| -4 | 518 | 472 | 427 | 384 | 342 | 300 | 259 | 224 | 192 | 160 | 128 | 107 | 97 | 87 | 79 |
| -3 | 517 | 472 | 426 | 383 | 341 | 300 | 258 | 224 | 192 | 160 | 128 | 97 | 77 | 67 | 59 |
| -2 | 516 | 471 | 425 | 383 | 341 | 299 | 257 | 223 | 192 | 160 | 128 | 97 | 67 | 47 | 39 |
| -1 | 516 | 470 | 425 | 382 | 340 | 298 | 256 | 223 | 191 | 159 | 127 | 97 | 67 | 37 | 20 |
| 0 | 504 | 459 | 413 | 371 | 329 | 287 | 245 | 214 | 183 | 151 | 119 | 89 | 60 | 30 | 0 |
| 1 | 516 | 470 | 424 | 381 | 339 | 297 | 255 | 221 | 190 | 159 | 127 | 97 | 67 | 37 | 21 |
| 2 | 517 | 470 | 424 | 381 | 339 | 297 | 255 | 221 | 190 | 159 | 127 | 97 | 67 | 53 | 42 |
| 3 | 517 | 470 | 424 | 381 | 338 | 296 | 254 | 220 | 189 | 158 | 127 | 96 | 82 | 74 | 63 |
| 4 | 518 | 471 | 424 | 380 | 338 | 296 | 254 | 220 | 189 | 158 | 126 | 112 | 103 | 95 | 84 |
| 5 | 516 | 469 | 422 | 379 | 337 | 295 | 253 | 220 | 190 | 159 | 127 | 97 | 67 | 40 | 0 |
| 6 | 515 | 468 | 421 | 378 | 336 | 294 | 251 | 219 | 189 | 159 | 129 | 97 | 67 | 37 | 0 |
| 7 | 513 | 466 | 419 | 377 | 335 | 293 | 250 | 217 | 188 | 159 | 129 | 98 | 67 | 37 | 0 |
| 8 | 512 | 465 | 418 | 375 | 334 | 292 | 249 | 216 | 187 | 157 | 128 | 98 | 68 | 37 | 0 |
| 9 | 511 | 464 | 416 | 374 | 333 | 291 | 248 | 215 | 186 | 156 | 127 | 97 | 68 | 38 | 0 |
| 10 | 509 | 462 | 415 | 372 | 332 | 290 | 247 | 214 | 184 | 155 | 125 | 96 | 66 | 37 | 0 |
| 11 | 508 | 461 | 414 | 371 | 330 | 289 | 246 | 213 | 183 | 154 | 124 | 95 | 65 | 36 | 0 |
| 12 | 506 | 460 | 413 | 370 | 329 | 288 | 246 | 212 | 182 | 152 | 123 | 93 | 64 | 37 | 0 |
| 13 | 505 | 459 | 412 | 368 | 328 | 287 | 245 | 211 | 181 | 151 | 122 | 93 | 66 | 37 | 0 |
| 14 | 503 | 457 | 411 | 367 | 326 | 286 | 244 | 210 | 180 | 150 | 122 | 95 | 66 | 37 | 0 |
| 15 | 502 | 456 | 410 | 366 | 325 | 284 | 243 | 209 | 179 | 152 | 124 | 96 | 67 | 37 | 0 |
| 16 | 512 | 466 | 409 | 365 | 324 | 283 | 242 | 209 | 181 | 153 | 125 | 108 | 98 | 91 | 79 |
| 17 | 511 | 464 | 418 | 364 | 323 | 282 | 241 | 208 | 181 | 153 | 125 | 98 | 81 | 71 | 59 |
| 18 | 509 | 462 | 416 | 371 | 332 | 281 | 240 | 207 | 180 | 152 | 124 | 95 | 69 | 51 | 39 |
| 19 | 507 | 461 | 413 | 369 | 330 | 290 | 239 | 205 | 179 | 151 | 123 | 95 | 66 | 39 | 20 |
| 20 | 504 | 459 | 411 | 366 | 328 | 289 | 248 | 204 | 178 | 151 | 122 | 94 | 65 | 36 | 0 |
| 21 | 501 | 456 | 409 | 364 | 325 | 287 | 247 | 202 | 176 | 150 | 122 | 93 | 64 | 35 | 20 |
| 22 | 508 | 453 | 408 | 361 | 323 | 285 | 246 | 207 | 175 | 148 | 121 | 92 | 63 | 49 | 40 |
| 23 | 505 | 461 | 417 | 359 | 320 | 282 | 243 | 206 | 179 | 147 | 120 | 92 | 76 | 69 | 60 |
| 24 | 502 | 458 | 413 | 367 | 318 | 280 | 241 | 204 | 178 | 145 | 119 | 103 | 96 | 89 | 79 |
| 25 | 500 | 455 | 410 | 364 | 325 | 277 | 239 | 202 | 177 | 151 | 118 | 92 | 63 | 37 | 0 |
| 26 | 506 | 452 | 407 | 362 | 323 | 275 | 236 | 199 | 174 | 149 | 117 | 91 | 64 | 35 | 0 |
| 27 | 501 | 459 | 404 | 359 | 319 | 282 | 234 | 196 | 171 | 147 | 116 | 90 | 63 | 35 | 0 |
| 28 | 496 | 456 | 411 | 355 | 316 | 281 | 241 | 194 | 169 | 144 | 119 | 88 | 62 | 35 | 0 |
| 29 | 490 | 451 | 408 | 363 | 312 | 278 | 239 | 191 | 166 | 141 | 116 | 85 | 60 | 34 | 0 |
| 30 | 496 | 447 | 406 | 360 | 317 | 274 | 237 | 198 | 163 | 138 | 112 | 81 | 56 | 31 | 0 |
| 31 | 494 | 451 | 402 | 357 | 313 | 271 | 235 | 196 | 160 | 135 | 109 | 78 | 53 | 28 | 0 |
| 32 | 489 | 446 | 406 | 354 | 310 | 275 | 233 | 194 | 163 | 132 | 106 | 81 | 51 | 33 | 0 |
| 33 | 496 | 441 | 401 | 361 | 306 | 271 | 237 | 192 | 161 | 129 | 103 | 79 | 57 | 33 | 0 |
| 34 | 492 | 445 | 397 | 357 | 313 | 267 | 233 | 190 | 158 | 132 | 103 | 85 | 61 | 34 | 0 |
| 35 | 488 | 443 | 401 | 352 | 310 | 263 | 229 | 195 | 155 | 132 | 109 | 91 | 61 | 34 | 42 |
| 36 | 493 | 439 | 396 | 356 | 306 | 269 | 226 | 191 | 156 | 137 | 115 | 97 | 88 | 82 | 74 |
| 37 | 485 | 443 | 392 | 352 | 302 | 267 | 224 | 189 | 156 | 137 | 119 | 95 | 73 | 63 | 55 |
| 38 | 487 | 436 | 397 | 348 | 308 | 264 | 230 | 188 | 156 | 137 | 118 | 94 | 68 | 45 | 37 |
| 39 | 479 | 437 | 394 | 354 | 304 | 266 | 229 | 186 | 154 | 136 | 117 | 93 | 67 | 35 | 18 |
| 40 | 480 | 431 | 394 | 350 | 301 | 261 | 225 | 192 | 151 | 137 | 116 | 93 | 66 | 33 | 0 |
| 41 | 473 | 433 | 387 | 344 | 306 | 256 | 227 | 191 | 148 | 134 | 116 | 92 | 65 | 32 | 17 |
| 42 | 478 | 427 | 380 | 344 | 302 | 261 | 223 | 187 | 154 | 131 | 113 | 96 | 64 | 41 | 34 |
| 43 | 471 | 429 | 382 | 337 | 302 | 258 | 219 | 189 | 152 | 128 | 110 | 95 | 64 | 58 | 51 |
| 44 | 473 | 423 | 377 | 340 | 295 | 253 | 220 | 184 | 148 | 126 | 111 | 93 | 82 | 75 | 67 |
| 45 | 468 | 426 | 379 | 334 | 297 | 255 | 216 | 179 | 147 | 127 | 107 | 89 | 61 | 31 | 0 |
| 46 | 470 | 422 | 378 | 336 | 294 | 252 | 212 | 181 | 153 | 128 | 104 | 85 | 58 | 27 | 0 |
| 47 | 469 | 421 | 371 | 332 | 290 | 256 | 214 | 180 | 154 | 129 | 104 | 82 | 54 | 27 | 0 |
| 48 | 460 | 420 | 374 | 334 | 291 | 251 | 210 | 182 | 156 | 131 | 105 | 80 | 54 | 27 | 0 |
| 49 | 460 | 412 | 370 | 336 | 292 | 251 | 205 | 186 | 161 | 135 | 109 | 83 | 58 | 32 | 0 |
| 50 | 459 | 413 | 367 | 326 | 284 | 230 | 180 | 179 | 160 | 134 | 108 | 82 | 56 | 30 | 0 |
| 51 | 458 | 409 | 368 | 322 | 279 | 207 | 160 | 159 | 153 | 134 | 108 | 81 | 55 | 29 | 0 |
| 52 | 454 | 407 | 365 | 318 | 247 | 174 | 140 | 139 | 134 | 132 | 107 | 80 | 54 | 27 | 0 |
| 53 | 450 | 397 | 355 | 289 | 225 | 151 | 121 | 118 | 114 | 115 | 111 | 80 | 53 | 26 | 0 |
| 54 | 441 | 395 | 352 | 268 | 203 | 129 | 101 | 98 | 94 | 95 | 96 | 83 | 52 | 26 | 0 |
| 55 | 440 | 397 | 351 | 262 | 182 | 95 | 81 | 78 | 73 | 74 | 76 | 75 | 50 | 27 | 0 |
| 56 | 429 | 396 | 349 | 281 | 176 | 86 | 62 | 58 | 53 | 54 | 55 | 57 | 55 | 56 | 56 |
| 57 | 427 | 395 | 348 | 297 | 175 | 75 | 50 | 50 | 50 | 53 | 55 | 58 | 52 | 43 | 42 |
| 58 | 419 | 371 | 347 | 302 | 191 | 87 | 57 | 56 | 55 | 56 | 58 | 61 | 50 | 29 | 21 |
| 59 | 412 | 356 | 346 | 301 | 207 | 102 | 68 | 62 | 61 | 58 | 61 | 64 | 47 | 25 | 14 |
| 60 | 410 | 348 | 321 | 301 | 225 | 118 | 84 | 68 | 67 | 61 | 64 | 67 | 44 | 22 | 0 |
| 61 | 400 | 348 | 295 | 289 | 247 | 135 | 100 | 80 | 74 | 68 | 66 | 68 | 45 | 22 | 0 |
| 62 | 398 | 346 | 283 | 265 | 254 | 153 | 113 | 96 | 80 | 76 | 69 | 65 | 42 | 22 | 0 |
| 63 | 396 | 337 | 283 | 239 | 240 | 173 | 129 | 112 | 91 | 83 | 74 | 66 | 42 | 22 | 0 |
| 64 | 395 | 335 | 275 | 219 | 214 | 193 | 146 | 128 | 104 | 91 | 81 | 66 | 45 | 23 | 0 |
| 65 | 367 | 318 | 279 | 218 | 188 | 183 | 162 | 138 | 112 | 89 | 81 | 67 | 45 | 22 | 0 |
| 66 | 347 | 297 | 248 | 199 | 168 | 172 | 163 | 136 | 110 | 85 | 81 | 67 | 44 | 21 | 0 |
| 67 | 315 | 265 | 215 | 161 | 158 | 163 | 155 | 135 | 110 | 83 | 75 | 68 | 45 | 23 | 0 |
| 68 | 284 | 233 | 182 | 141 | 148 | 154 | 147 | 132 | 108 | 81 | 68 | 68 | 46 | 24 | 0 |
| 69 | 277 | 214 | 163 | 131 | 137 | 144 | 138 | 122 | 103 | 79 | 61 | 64 | 49 | 24 | 0 |
| 70 | 209 | 189 | 132 | 119 | 126 | 133 | 131 | 115 | 95 | 77 | 54 | 59 | 49 | 25 | 0 |
| 71 | 139 | 139 | 105 | 107 | 114 | 121 | 127 | 107 | 86 | 70 | 47 | 54 | 44 | 28 | 0 |
| 72 | 80 | 79 | 96 | 97 | 105 | 112 | 120 | 102 | 81 | 62 | 42 | 50 | 39 | 28 | 0 |
| 73 | 75 | 95 | 136 | 104 | 100 | 108 | 115 | 100 | 80 | 60 | 40 | 48 | 34 | 23 | 0 |
| 74 | 203 | 125 | 174 | 142 | 108 | 103 | 109 | 97 | 80 | 59 | 38 | 48 | 33 | 18 | 0 |
| 75 | 173 | 221 | 210 | 177 | 146 | 108 | 104 | 94 | 78 | 57 | 36 | 48 | 33 | 16 | 18 |
| 76 | 131 | 185 | 244 | 212 | 177 | 140 | 103 | 89 | 75 | 53 | 34 | 49 | 43 | 40 | 37 |
| 77 | 86 | 142 | 204 | 221 | 187 | 148 | 106 | 84 | 72 | 51 | 31 | 44 | 37 | 32 | 27 |
| 78 | 35 | 95 | 153 | 209 | 198 | 157 | 117 | 77 | 67 | 50 | 30 | 41 | 37 | 23 | 14 |
| 79 | 27 | 42 | 98 | 162 | 194 | 171 | 130 | 86 | 63 | 51 | 28 | 37 | 38 | 20 | 9 |
| 80 | 27 | 26 | 40 | 106 | 167 | 178 | 141 | 99 | 56 | 45 | 28 | 32 | 35 | 20 | 0 |
| 81 | 116 | 28 | 24 | 39 | 111 | 154 | 157 | 113 | 64 | 38 | 29 | 28 | 31 | 20 | 0 |
| 82 | 139 | 60 | 27 | 23 | 36 | 102 | 136 | 131 | 82 | 33 | 24 | 25 | 27 | 22 | 0 |
| 83 | 125 | 75 | 30 | 27 | 22 | 31 | 92 | 118 | 104 | 52 | 19 | 23 | 23 | 23 | 0 |
| 84 | 23 | 26 | 23 | 27 | 30 | 24 | 20 | 80 | 100 | 74 | 14 | 19 | 18 | 21 | 0 |
| 85 | 29 | 36 | 35 | 28 | 20 | 27 | 31 | 28 | 64 | 86 | 42 | 12 | 16 | 17 | 2 |
| 86 | 27 | 29 | 33 | 35 | 34 | 27 | 17 | 25 | 28 | 40 | 68 | 9 | 13 | 14 | 5 |
| 87 | 17 | 19 | 21 | 25 | 29 | 33 | 35 | 25 | 15 | 18 | 26 | 50 | 7 | 9 | 7 |
| 88 | 28 | 18 | 10 | 10 | 14 | 17 | 22 | 27 | 33 | 27 | 11 | 12 | 32 | 6 | 8 |
| 89 | 23 | 24 | 15 | 5 | 6 | 6 | 6 | 6 | 5 | 4 | 10 | 33 | 5 | 14 | 3 |
| 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

The average intensity over the field of view is evaluated by integrating the diffuse intensity multiplied by this response function and dividing by the integral of the response function.  The direct solar beam must also be included. This is done by multiplying the direct solar beam by the relative response at the location of the sun and dividing by the integral of the response function over the field of view.

5.2 Field of view of the DLIS

The field of view of the DLIS is centered at an azimuth angle of 1.151 degrees and at a zenith angle of  158.6 degrees.  Constants for parabolic fits for the edges of the field of view are contained in the following table.

| Boundary | Constant | Linear Term | Quadratic Term |
| --- | --- | --- | --- |
| Left | 134.22 | 6.3309 | -0.32633 |
| Top | 153.73 | -0.018787 | -0.0037792 |
| Right | -2609.6 | -1777.1 | -284.53 |
| Bottom | 163.51 | -0.011932 | -0.0024252 |

With these constants the zenith angle of each of the four boundaries can be computed at the azimuth angle of interest according to

Nadir angle = a +b*azimuth angle + c*(azimuth angle)2.

Note that the limits of the azimuth angle range to be used for the left boundary are 3.5 to 8 degrees, for the top boundary -2.5 to +4 degrees, for the right boundary are -2.9 to -3.1 degrees, and for the bottom boundary are -4 to +8 degrees.  The response over this field is assumed to be constant.

- Azimuth of the instrument during each IR measurement cycle

For the DLIS, the instrument collected data in 8 regions.  The locations of each region for each cycle are given below.

| DLIS azimuth for each region |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| cycle | region | samp.t. | # reads | altitude | mt start | mt center | mt end | az start | az center | az end | weights |
| 1 | 1 | 1 | 10 | 142.08 | 148.90 | 148.94 | 148.99 | 45.63 | 44.45 | 43.27 | 0.03030 |
| 1 | 1 | 1 | 10 | 142.07 | 149.07 | 149.11 | 149.15 | 40.91 | 39.73 | 38.55 | 0.03030 |
| 1 | 1 | 1 | 10 | 142.06 | 149.23 | 149.27 | 149.31 | 36.19 | 35.01 | 33.84 | 0.03030 |
| 1 | 1 | 1 | 10 | 142.05 | 149.39 | 149.43 | 149.47 | 31.48 | 30.30 | 29.12 | 0.03030 |
| 1 | 1 | 1 | 12 | 141.60 | 158.16 | 158.21 | 158.26 | 137.85 | 136.47 | 135.08 | 0.03030 |
| 1 | 1 | 1 | 12 | 141.59 | 158.36 | 158.40 | 158.45 | 132.32 | 130.94 | 129.55 | 0.03030 |
| 1 | 1 | 1 | 12 | 141.58 | 158.55 | 158.60 | 158.65 | 126.79 | 125.41 | 124.03 | 0.03030 |
| 1 | 1 | 1 | 12 | 141.57 | 158.74 | 158.79 | 158.84 | 121.26 | 119.88 | 118.50 | 0.03030 |
| 1 | 1 | 1 | 10 | 141.10 | 167.78 | 167.82 | 167.86 | 226.03 | 224.90 | 223.78 | 0.03030 |
| 1 | 1 | 1 | 10 | 141.10 | 167.94 | 167.98 | 168.02 | 221.53 | 220.40 | 219.28 | 0.03030 |
| 1 | 1 | 1 | 10 | 141.09 | 168.11 | 168.15 | 168.19 | 217.03 | 215.90 | 214.78 | 0.03030 |
| 1 | 1 | 1 | 10 | 141.08 | 168.27 | 168.31 | 168.35 | 212.53 | 211.41 | 210.28 | 0.03030 |
| 1 | 1 | 1 | 10 | 141.07 | 168.43 | 168.47 | 168.51 | 208.03 | 206.91 | 205.78 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.59 | 177.76 | 177.80 | 177.84 | 310.84 | 309.74 | 308.65 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.59 | 177.92 | 177.96 | 178.00 | 306.45 | 305.35 | 304.25 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.58 | 178.08 | 178.12 | 178.16 | 302.05 | 300.96 | 299.86 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.57 | 178.24 | 178.28 | 178.32 | 297.66 | 296.56 | 295.47 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.56 | 178.40 | 178.44 | 178.48 | 293.27 | 292.17 | 291.08 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.07 | 188.04 | 188.08 | 188.12 | 34.20 | 33.13 | 32.06 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.06 | 188.20 | 188.24 | 188.28 | 29.92 | 28.85 | 27.78 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.06 | 188.36 | 188.40 | 188.44 | 25.64 | 24.57 | 23.50 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.05 | 188.52 | 188.56 | 188.60 | 21.37 | 20.30 | 19.23 | 0.03030 |
| 1 | 1 | 1 | 10 | 140.04 | 188.68 | 188.72 | 188.77 | 17.09 | 16.02 | 14.95 | 0.03030 |
| 1 | 1 | 1 | 10 | 139.55 | 198.46 | 198.50 | 198.54 | 121.57 | 120.53 | 119.49 | 0.03030 |
| 1 | 1 | 1 | 10 | 139.54 | 198.62 | 198.66 | 198.70 | 117.41 | 116.37 | 115.33 | 0.03030 |
| 1 | 1 | 1 | 10 | 139.53 | 198.78 | 198.82 | 198.86 | 113.26 | 112.22 | 111.18 | 0.03030 |
| 1 | 1 | 1 | 10 | 139.52 | 198.94 | 198.98 | 199.02 | 109.10 | 108.07 | 107.03 | 0.03030 |
| 1 | 1 | 1 | 10 | 139.51 | 199.10 | 199.14 | 199.18 | 104.95 | 103.92 | 102.88 | 0.03030 |
| 1 | 1 | 1 | 10 | 139.00 | 209.39 | 209.43 | 209.47 | 204.46 | 203.46 | 202.46 | 0.03030 |
| 1 | 1 | 1 | 10 | 138.99 | 209.55 | 209.59 | 209.63 | 200.45 | 199.45 | 198.44 | 0.03030 |
| 1 | 1 | 1 | 10 | 138.98 | 209.72 | 209.76 | 209.80 | 196.44 | 195.43 | 194.43 | 0.03030 |
| 1 | 1 | 1 | 10 | 138.98 | 209.88 | 209.92 | 209.96 | 192.43 | 191.42 | 190.42 | 0.03030 |
| 1 | 1 | 1 | 10 | 138.97 | 210.04 | 210.08 | 210.12 | 188.42 | 187.42 | 186.42 | 0.03030 |
| 2 | 1 | 16 | 2 | 134.98 | 292.26 | 292.39 | 292.52 | 178.07 | 175.59 | 173.12 | 0.09249 |
| 2 | 1 | 16 | 2 | 134.95 | 292.78 | 292.91 | 293.04 | 168.17 | 165.70 | 163.23 | 0.09249 |
| 2 | 1 | 21 | 2 | 134.57 | 300.88 | 301.05 | 301.22 | 15.94 | 12.83 | 9.72 | 0.12139 |
| 2 | 1 | 21 | 2 | 134.54 | 301.56 | 301.73 | 301.90 | 3.50 | 0.39 | 357.29 | 0.12139 |
| 2 | 1 | 21 | 2 | 133.52 | 323.54 | 323.71 | 323.88 | 344.13 | 341.40 | 338.67 | 0.12139 |
| 2 | 1 | 21 | 2 | 133.48 | 324.22 | 324.39 | 324.56 | 333.22 | 330.49 | 327.78 | 0.12139 |
| 2 | 1 | 21 | 2 | 133.45 | 324.89 | 325.06 | 325.23 | 322.35 | 319.64 | 316.93 | 0.12139 |
| 2 | 1 | 18 | 2 | 132.82 | 338.76 | 338.91 | 339.05 | 110.00 | 107.88 | 105.76 | 0.10405 |
| 2 | 1 | 18 | 2 | 132.79 | 339.35 | 339.49 | 339.64 | 101.53 | 99.42 | 97.31 | 0.10405 |
| 3 | 1 | 23 | 2 | 128.47 | 436.63 | 436.81 | 437.00 | 143.69 | 142.34 | 141.00 | 0.10360 |
| 3 | 1 | 23 | 2 | 128.43 | 437.37 | 437.55 | 437.74 | 138.32 | 136.99 | 135.65 | 0.10360 |
| 3 | 1 | 23 | 2 | 128.40 | 438.11 | 438.29 | 438.48 | 132.99 | 131.67 | 130.35 | 0.10360 |
| 3 | 1 | 23 | 2 | 127.40 | 461.62 | 461.81 | 461.99 | 346.99 | 346.02 | 345.05 | 0.10360 |
| 3 | 1 | 23 | 2 | 127.37 | 462.37 | 462.55 | 462.74 | 343.11 | 342.15 | 341.18 | 0.10360 |
| 3 | 1 | 23 | 2 | 127.34 | 463.11 | 463.29 | 463.48 | 339.27 | 338.31 | 337.36 | 0.10360 |
| 3 | 1 | 21 | 2 | 126.33 | 487.15 | 487.32 | 487.49 | 237.25 | 236.68 | 236.11 | 0.09459 |
| 3 | 1 | 21 | 2 | 126.30 | 487.82 | 487.99 | 488.16 | 234.98 | 234.42 | 233.86 | 0.09459 |
| 3 | 1 | 21 | 2 | 126.27 | 488.50 | 488.67 | 488.84 | 232.75 | 232.19 | 231.64 | 0.09459 |
| 3 | 1 | 21 | 2 | 126.25 | 489.18 | 489.35 | 489.52 | 230.54 | 230.00 | 229.45 | 0.09459 |
| 4 | 1 | 23 | 2 | 123.31 | 561.10 | 561.28 | 561.47 | 160.03 | 160.24 | 160.45 | 0.09312 |
| 4 | 1 | 23 | 2 | 123.28 | 561.84 | 562.02 | 562.21 | 160.88 | 161.09 | 161.31 | 0.09312 |
| 4 | 1 | 23 | 2 | 123.26 | 562.58 | 562.77 | 562.95 | 161.76 | 161.99 | 162.21 | 0.09312 |
| 4 | 1 | 23 | 2 | 123.23 | 563.32 | 563.51 | 563.69 | 162.68 | 162.91 | 163.15 | 0.09312 |
| 4 | 1 | 23 | 2 | 123.20 | 564.06 | 564.25 | 564.43 | 163.63 | 163.87 | 164.12 | 0.09312 |
| 4 | 1 | 22 | 2 | 121.52 | 606.56 | 606.74 | 606.91 | 272.08 | 272.75 | 273.42 | 0.08907 |
| 4 | 1 | 22 | 2 | 121.49 | 607.27 | 607.45 | 607.62 | 274.78 | 275.46 | 276.14 | 0.08907 |
| 4 | 1 | 22 | 2 | 121.46 | 607.98 | 608.16 | 608.33 | 277.51 | 278.20 | 278.89 | 0.08907 |
| 4 | 1 | 22 | 2 | 121.44 | 608.69 | 608.87 | 609.04 | 280.27 | 280.96 | 281.66 | 0.08907 |
| 4 | 1 | 22 | 2 | 121.41 | 609.40 | 609.57 | 609.75 | 283.06 | 283.76 | 284.46 | 0.08907 |
| 4 | 1 | 22 | 2 | 121.38 | 610.11 | 610.28 | 610.46 | 285.87 | 286.58 | 287.29 | 0.08907 |
| 8 | 1 | 20 | 2 | 113.05 | 837.13 | 837.29 | 837.45 | 344.72 | 347.20 | 349.69 | 0.10582 |
| 8 | 1 | 20 | 2 | 113.03 | 837.77 | 837.93 | 838.09 | 354.66 | 357.14 | 359.63 | 0.10582 |
| 8 | 1 | 20 | 2 | 113.01 | 838.42 | 838.58 | 838.74 | 4.61 | 7.10 | 9.59 | 0.10582 |
| 8 | 1 | 22 | 2 | 112.17 | 862.94 | 863.12 | 863.29 | 37.11 | 40.05 | 42.99 | 0.11640 |
| 8 | 1 | 22 | 2 | 112.14 | 863.65 | 863.83 | 864.00 | 48.87 | 51.82 | 54.76 | 0.11640 |
| 8 | 1 | 22 | 2 | 112.12 | 864.36 | 864.54 | 864.71 | 60.66 | 63.61 | 66.56 | 0.11640 |
| 8 | 1 | 21 | 2 | 111.41 | 885.32 | 885.49 | 885.66 | 59.17 | 62.14 | 65.12 | 0.11111 |
| 8 | 1 | 21 | 2 | 111.39 | 886.00 | 886.16 | 886.33 | 71.09 | 74.07 | 77.06 | 0.11111 |
| 8 | 1 | 21 | 2 | 111.37 | 886.67 | 886.84 | 887.01 | 83.03 | 86.03 | 89.02 | 0.11111 |
| 9 | 1 | 20 | 2 | 107.52 | 961.12 | 961.28 | 961.44 | 170.43 | 174.38 | 178.34 | 0.10811 |
| 9 | 1 | 20 | 2 | 107.48 | 961.76 | 961.92 | 962.09 | 186.28 | 190.26 | 194.24 | 0.10811 |
| 9 | 1 | 19 | 2 | 106.36 | 977.37 | 977.53 | 977.68 | 237.97 | 242.22 | 246.48 | 0.10270 |
| 9 | 1 | 19 | 2 | 106.32 | 977.99 | 978.14 | 978.29 | 255.01 | 259.28 | 263.56 | 0.10270 |
| 9 | 1 | 18 | 2 | 105.86 | 984.30 | 984.45 | 984.59 | 74.85 | 79.08 | 83.32 | 0.09730 |
| 9 | 1 | 18 | 2 | 105.81 | 984.88 | 985.03 | 985.17 | 91.82 | 96.07 | 100.33 | 0.09730 |
| 9 | 1 | 20 | 2 | 105.18 | 993.53 | 993.69 | 993.85 | 353.95 | 359.00 | 4.06 | 0.10811 |
| 9 | 1 | 21 | 2 | 104.26 | 1006.00 | 1006.17 | 1006.34 | 36.32 | 41.92 | 47.53 | 0.11351 |
| 9 | 1 | 15 | 2 | 103.39 | 1017.87 | 1017.99 | 1018.11 | 76.72 | 80.87 | 85.03 | 0.08108 |
| 9 | 1 | 15 | 2 | 103.35 | 1018.36 | 1018.48 | 1018.60 | 93.34 | 97.49 | 101.65 | 0.08108 |
| 10 | 1 | 19 | 2 | 100.93 | 1051.77 | 1051.92 | 1052.07 | 209.23 | 214.94 | 220.65 | 0.13768 |
| 10 | 1 | 19 | 2 | 100.19 | 1062.25 | 1062.40 | 1062.56 | 245.09 | 250.97 | 256.84 | 0.13768 |
| 10 | 1 | 18 | 2 | 99.49 | 1072.25 | 1072.40 | 1072.54 | 273.65 | 279.38 | 285.11 | 0.13043 |
| 10 | 1 | 17 | 2 | 98.81 | 1081.93 | 1082.06 | 1082.20 | 300.99 | 306.56 | 312.14 | 0.12319 |
| 10 | 1 | 17 | 2 | 98.65 | 1084.36 | 1084.50 | 1084.64 | 40.42 | 46.04 | 51.66 | 0.12319 |
| 10 | 1 | 17 | 2 | 98.08 | 1092.53 | 1092.67 | 1092.81 | 19.94 | 25.72 | 31.50 | 0.12319 |
| 10 | 1 | 16 | 2 | 97.47 | 1101.56 | 1101.68 | 1101.81 | 46.11 | 51.71 | 57.32 | 0.11594 |
| 10 | 1 | 15 | 2 | 96.88 | 1110.34 | 1110.47 | 1110.59 | 73.20 | 78.60 | 84.01 | 0.10870 |
| 11 | 1 | 14 | 2 | 94.70 | 1143.51 | 1143.62 | 1143.73 | 184.84 | 190.36 | 195.89 | 0.12069 |
| 11 | 1 | 14 | 2 | 94.19 | 1151.40 | 1151.51 | 1151.63 | 214.91 | 220.55 | 226.18 | 0.12069 |
| 11 | 1 | 13 | 2 | 93.70 | 1159.17 | 1159.27 | 1159.38 | 246.33 | 251.66 | 257.00 | 0.11207 |
| 11 | 1 | 13 | 2 | 93.52 | 1161.97 | 1162.08 | 1162.18 | 29.59 | 34.96 | 40.33 | 0.11207 |
| 11 | 1 | 13 | 2 | 93.06 | 1169.26 | 1169.37 | 1169.47 | 46.11 | 51.58 | 57.05 | 0.11207 |
| 11 | 1 | 13 | 2 | 92.60 | 1176.58 | 1176.68 | 1176.79 | 70.35 | 75.90 | 81.46 | 0.11207 |
| 11 | 1 | 12 | 2 | 92.15 | 1183.88 | 1183.98 | 1184.08 | 100.91 | 106.13 | 111.35 | 0.10345 |
| 11 | 1 | 12 | 2 | 91.72 | 1190.92 | 1191.02 | 1191.12 | 124.08 | 129.40 | 134.71 | 0.10345 |
| 11 | 1 | 12 | 2 | 91.30 | 1197.83 | 1197.93 | 1198.03 | 146.86 | 152.26 | 157.66 | 0.10345 |
| 12 | 1 | 10 | 2 | 89.20 | 1233.06 | 1233.14 | 1233.22 | 15.06 | 19.78 | 24.51 | 0.05051 |
| 12 | 1 | 21 | 2 | 88.85 | 1239.08 | 1239.25 | 1239.42 | 7.84 | 17.76 | 27.67 | 0.10606 |
| 12 | 1 | 21 | 2 | 88.50 | 1245.05 | 1245.22 | 1245.39 | 356.75 | 6.64 | 16.53 | 0.10606 |
| 12 | 1 | 10 | 2 | 88.14 | 1251.34 | 1251.42 | 1251.50 | 3.14 | 7.83 | 12.51 | 0.05051 |
| 12 | 1 | 10 | 2 | 87.80 | 1257.41 | 1257.49 | 1257.57 | 354.61 | 359.27 | 3.93 | 0.05051 |
| 12 | 1 | 21 | 2 | 87.45 | 1263.44 | 1263.61 | 1263.78 | 342.43 | 352.17 | 1.92 | 0.10606 |
| 12 | 1 | 21 | 2 | 87.11 | 1269.44 | 1269.61 | 1269.78 | 326.84 | 336.54 | 346.25 | 0.10606 |
| 12 | 1 | 21 | 2 | 86.78 | 1275.32 | 1275.49 | 1275.66 | 303.59 | 313.26 | 322.93 | 0.10606 |
| 12 | 1 | 21 | 2 | 86.46 | 1281.15 | 1281.32 | 1281.49 | 276.14 | 285.78 | 295.43 | 0.10606 |
| 12 | 1 | 21 | 2 | 86.14 | 1286.97 | 1287.14 | 1287.31 | 246.84 | 256.46 | 266.09 | 0.10606 |
| 12 | 1 | 21 | 2 | 86.04 | 1288.65 | 1288.81 | 1288.98 | 342.10 | 351.71 | 1.33 | 0.10606 |
| 13 | 1 | 20 | 2 | 84.16 | 1323.64 | 1323.81 | 1323.97 | 158.30 | 167.38 | 176.45 | 0.08929 |
| 13 | 1 | 20 | 2 | 83.99 | 1326.82 | 1326.98 | 1327.14 | 336.95 | 346.01 | 355.06 | 0.08929 |
| 13 | 1 | 22 | 2 | 83.63 | 1333.72 | 1333.90 | 1334.08 | 3.86 | 13.77 | 23.68 | 0.09821 |
| 13 | 1 | 20 | 2 | 83.32 | 1339.64 | 1339.80 | 1339.97 | 333.65 | 342.60 | 351.55 | 0.08929 |
| 13 | 1 | 20 | 2 | 83.02 | 1345.36 | 1345.52 | 1345.68 | 289.87 | 298.75 | 307.62 | 0.08929 |
| 13 | 1 | 20 | 2 | 82.73 | 1351.07 | 1351.23 | 1351.39 | 242.62 | 251.41 | 260.19 | 0.08929 |
| 13 | 1 | 20 | 2 | 82.44 | 1356.79 | 1356.95 | 1357.11 | 192.19 | 200.86 | 209.53 | 0.08929 |
| 13 | 1 | 20 | 2 | 82.14 | 1362.58 | 1362.75 | 1362.91 | 141.70 | 150.23 | 158.76 | 0.08929 |
| 13 | 1 | 20 | 2 | 81.85 | 1368.35 | 1368.51 | 1368.67 | 84.56 | 92.96 | 101.35 | 0.08929 |
| 13 | 1 | 21 | 2 | 81.56 | 1374.17 | 1374.34 | 1374.51 | 25.24 | 33.91 | 42.58 | 0.09375 |
| 13 | 1 | 21 | 2 | 81.27 | 1380.00 | 1380.17 | 1380.34 | 321.31 | 329.84 | 338.37 | 0.09375 |
| 14 | 1 | 13 | 2 | 79.77 | 1410.54 | 1410.65 | 1410.75 | 12.52 | 17.54 | 22.55 | 0.06311 |
| 14 | 1 | 13 | 2 | 79.44 | 1417.61 | 1417.72 | 1417.82 | 350.21 | 355.21 | 0.20 | 0.06311 |
| 14 | 1 | 22 | 2 | 79.18 | 1422.90 | 1423.07 | 1423.25 | 241.75 | 250.19 | 258.63 | 0.10680 |
| 14 | 1 | 22 | 2 | 78.89 | 1429.02 | 1429.19 | 1429.37 | 172.61 | 181.02 | 189.44 | 0.10680 |
| 14 | 1 | 22 | 2 | 78.60 | 1435.27 | 1435.44 | 1435.62 | 108.70 | 117.08 | 125.47 | 0.10680 |
| 14 | 1 | 22 | 2 | 78.30 | 1441.55 | 1441.73 | 1441.90 | 44.96 | 53.30 | 61.63 | 0.10680 |
| 14 | 1 | 23 | 2 | 78.01 | 1447.85 | 1448.04 | 1448.22 | 340.33 | 348.99 | 357.64 | 0.11165 |
| 14 | 1 | 23 | 2 | 77.72 | 1454.17 | 1454.35 | 1454.54 | 273.48 | 282.05 | 290.62 | 0.11165 |
| 14 | 1 | 23 | 2 | 77.42 | 1460.59 | 1460.78 | 1460.97 | 208.68 | 217.15 | 225.60 | 0.11165 |
| 14 | 1 | 23 | 2 | 77.13 | 1467.10 | 1467.29 | 1467.47 | 143.88 | 152.23 | 160.58 | 0.11165 |
| 15 | 1 | 16 | 2 | 73.94 | 1540.75 | 1540.88 | 1541.01 | 14.59 | 19.80 | 25.03 | 0.15094 |
| 15 | 1 | 12 | 2 | 73.69 | 1546.84 | 1546.94 | 1547.04 | 260.39 | 264.26 | 268.13 | 0.11321 |
| 15 | 1 | 13 | 2 | 73.38 | 1554.26 | 1554.37 | 1554.47 | 193.84 | 197.94 | 202.04 | 0.12264 |
| 15 | 1 | 13 | 2 | 73.08 | 1561.67 | 1561.78 | 1561.88 | 118.89 | 122.86 | 126.82 | 0.12264 |
| 15 | 1 | 13 | 2 | 72.77 | 1569.23 | 1569.34 | 1569.44 | 39.45 | 43.29 | 47.12 | 0.12264 |
| 15 | 1 | 13 | 2 | 72.46 | 1576.87 | 1576.97 | 1577.08 | 317.86 | 321.63 | 325.40 | 0.12264 |
| 15 | 1 | 13 | 2 | 72.15 | 1584.62 | 1584.73 | 1584.83 | 233.52 | 237.23 | 240.93 | 0.12264 |
| 15 | 1 | 13 | 2 | 71.84 | 1592.48 | 1592.58 | 1592.69 | 148.34 | 152.00 | 155.65 | 0.12264 |
| 16 | 1 | 14 | 2 | 70.49 | 1627.50 | 1627.61 | 1627.72 | 301.05 | 305.01 | 308.98 | 0.13462 |
| 16 | 1 | 15 | 2 | 70.18 | 1635.95 | 1636.07 | 1636.19 | 237.02 | 241.25 | 245.47 | 0.14423 |
| 16 | 1 | 15 | 2 | 69.86 | 1644.53 | 1644.65 | 1644.77 | 172.77 | 176.93 | 181.12 | 0.14423 |
| 16 | 1 | 15 | 2 | 69.55 | 1653.10 | 1653.22 | 1653.34 | 110.18 | 114.37 | 118.56 | 0.14423 |
| 16 | 1 | 15 | 2 | 69.23 | 1661.87 | 1661.99 | 1662.11 | 45.46 | 49.33 | 53.22 | 0.14423 |
| 16 | 1 | 15 | 2 | 68.92 | 1670.57 | 1670.69 | 1670.81 | 326.66 | 330.58 | 334.50 | 0.14423 |
| 16 | 1 | 15 | 2 | 68.61 | 1679.46 | 1679.58 | 1679.70 | 255.56 | 259.51 | 263.45 | 0.14423 |
| 17 | 1 | 17 | 2 | 67.32 | 1716.60 | 1716.73 | 1716.87 | 29.67 | 34.10 | 38.51 | 0.14407 |
| 17 | 1 | 16 | 2 | 67.01 | 1725.99 | 1726.12 | 1726.25 | 326.10 | 330.04 | 333.97 | 0.13559 |
| 17 | 1 | 17 | 2 | 66.70 | 1735.35 | 1735.48 | 1735.62 | 248.08 | 252.23 | 256.38 | 0.14407 |
| 17 | 1 | 17 | 2 | 66.39 | 1744.96 | 1745.09 | 1745.23 | 185.04 | 189.29 | 193.54 | 0.14407 |
| 17 | 1 | 17 | 2 | 66.09 | 1754.48 | 1754.62 | 1754.76 | 110.60 | 114.62 | 118.65 | 0.14407 |
| 17 | 1 | 17 | 2 | 65.78 | 1764.30 | 1764.44 | 1764.58 | 36.39 | 40.35 | 44.31 | 0.14407 |
| 17 | 1 | 17 | 2 | 65.47 | 1774.25 | 1774.38 | 1774.52 | 319.46 | 323.35 | 327.24 | 0.14407 |
| 18 | 1 | 18 | 2 | 64.49 | 1806.56 | 1806.71 | 1806.85 | 171.69 | 175.85 | 180.01 | 0.16216 |
| 18 | 1 | 18 | 2 | 64.19 | 1816.91 | 1817.05 | 1817.20 | 103.90 | 107.98 | 112.06 | 0.16216 |
| 18 | 1 | 18 | 2 | 63.89 | 1827.25 | 1827.40 | 1827.54 | 37.14 | 41.29 | 45.44 | 0.16216 |
| 18 | 1 | 19 | 2 | 63.58 | 1837.92 | 1838.08 | 1838.23 | 339.58 | 343.91 | 348.23 | 0.17117 |
| 18 | 1 | 19 | 2 | 63.28 | 1848.75 | 1848.91 | 1849.06 | 281.25 | 285.44 | 289.63 | 0.17117 |
| 18 | 1 | 19 | 2 | 62.98 | 1859.79 | 1859.94 | 1860.10 | 222.02 | 226.26 | 230.50 | 0.17117 |
| 22 | 1 | 26 | 2 | 59.08 | 2013.72 | 2013.93 | 2014.14 | 14.32 | 19.51 | 24.71 | 0.22034 |
| 22 | 1 | 23 | 2 | 58.79 | 2026.33 | 2026.51 | 2026.70 | 326.83 | 331.45 | 336.07 | 0.19492 |
| 22 | 1 | 23 | 2 | 58.50 | 2038.83 | 2039.01 | 2039.20 | 281.63 | 286.37 | 291.12 | 0.19492 |
| 22 | 1 | 23 | 2 | 58.21 | 2051.33 | 2051.51 | 2051.70 | 237.64 | 242.25 | 246.85 | 0.19492 |
| 22 | 1 | 23 | 2 | 57.93 | 2063.83 | 2064.01 | 2064.20 | 195.80 | 200.51 | 205.22 | 0.19492 |
| 23 | 1 | 23 | 2 | 57.08 | 2102.09 | 2102.27 | 2102.46 | 47.59 | 51.97 | 56.34 | 0.16547 |
| 23 | 1 | 23 | 2 | 56.80 | 2114.86 | 2115.05 | 2115.23 | 344.90 | 349.20 | 353.50 | 0.16547 |
| 23 | 1 | 23 | 2 | 56.51 | 2128.31 | 2128.50 | 2128.68 | 300.42 | 304.76 | 309.09 | 0.16547 |
| 23 | 1 | 23 | 2 | 56.24 | 2141.08 | 2141.26 | 2141.45 | 237.67 | 241.97 | 246.28 | 0.16547 |
| 23 | 1 | 23 | 2 | 56.17 | 2144.20 | 2144.38 | 2144.57 | 309.90 | 314.20 | 318.49 | 0.16547 |
| 23 | 1 | 24 | 2 | 55.90 | 2157.23 | 2157.42 | 2157.62 | 240.88 | 245.06 | 249.24 | 0.17266 |
| 24 | 1 | 25 | 2 | 54.25 | 2237.99 | 2238.19 | 2238.39 | 93.90 | 97.71 | 101.54 | 0.15723 |
| 24 | 1 | 26 | 2 | 53.97 | 2252.31 | 2252.52 | 2252.73 | 3.21 | 7.08 | 10.95 | 0.16352 |
| 24 | 1 | 26 | 2 | 53.69 | 2266.60 | 2266.81 | 2267.02 | 270.12 | 274.05 | 277.98 | 0.16352 |
| 24 | 1 | 27 | 2 | 53.40 | 2281.49 | 2281.70 | 2281.92 | 191.20 | 195.39 | 199.58 | 0.16981 |
| 24 | 1 | 27 | 2 | 53.11 | 2296.74 | 2296.95 | 2297.17 | 114.17 | 118.13 | 122.09 | 0.16981 |
| 24 | 1 | 28 | 2 | 52.92 | 2306.23 | 2306.45 | 2306.68 | 284.48 | 288.41 | 292.35 | 0.17610 |
| 25 | 1 | 29 | 2 | 52.54 | 2326.82 | 2327.06 | 2327.29 | 276.06 | 280.13 | 284.21 | 0.24370 |
| 25 | 1 | 29 | 2 | 52.24 | 2342.61 | 2342.85 | 2343.08 | 193.43 | 197.45 | 201.47 | 0.24370 |
| 25 | 1 | 30 | 2 | 51.94 | 2359.06 | 2359.30 | 2359.54 | 109.06 | 113.03 | 116.98 | 0.25210 |
| 25 | 1 | 31 | 2 | 51.63 | 2375.86 | 2376.11 | 2376.36 | 17.22 | 21.12 | 25.02 | 0.26050 |
| 26 | 1 | 31 | 2 | 51.01 | 2410.28 | 2410.53 | 2410.78 | 194.68 | 198.67 | 202.66 | 0.24219 |
| 26 | 1 | 32 | 2 | 50.71 | 2427.40 | 2427.66 | 2427.92 | 108.45 | 112.59 | 116.75 | 0.25000 |
| 26 | 1 | 32 | 2 | 50.39 | 2445.40 | 2445.66 | 2445.92 | 36.90 | 40.99 | 45.08 | 0.25000 |
| 26 | 1 | 33 | 2 | 50.08 | 2463.03 | 2463.29 | 2463.56 | 307.13 | 311.03 | 314.93 | 0.25781 |
| 30 | 1 | 64 | 2 | 47.43 | 2622.40 | 2622.92 | 2623.43 | 280.85 | 289.00 | 297.17 | 0.39752 |
| 30 | 1 | 32 | 2 | 47.17 | 2639.27 | 2639.53 | 2639.79 | 187.95 | 191.98 | 196.01 | 0.19876 |
| 30 | 1 | 32 | 2 | 46.90 | 2656.41 | 2656.67 | 2656.93 | 91.21 | 95.12 | 99.02 | 0.19876 |
| 30 | 1 | 33 | 2 | 46.63 | 2674.25 | 2674.52 | 2674.78 | 357.09 | 0.97 | 4.86 | 0.20497 |
| 31 | 1 | 69 | 2 | 45.96 | 2717.64 | 2718.20 | 2718.75 | 311.79 | 320.93 | 330.05 | 0.23875 |
| 31 | 1 | 66 | 2 | 45.69 | 2735.24 | 2735.78 | 2736.31 | 234.16 | 242.54 | 250.93 | 0.22837 |
| 31 | 1 | 73 | 2 | 45.60 | 2741.24 | 2741.83 | 2742.42 | 328.35 | 337.54 | 346.73 | 0.25260 |
| 31 | 1 | 81 | 2 | 45.26 | 2763.95 | 2764.60 | 2765.26 | 331.23 | 342.02 | 352.80 | 0.28028 |
| 32 | 1 | 64 | 2 | 44.81 | 2794.90 | 2795.41 | 2795.93 | 133.95 | 142.95 | 151.97 | 0.25498 |
| 32 | 1 | 63 | 2 | 44.56 | 2811.59 | 2812.10 | 2812.60 | 62.50 | 71.08 | 79.63 | 0.25100 |
| 32 | 1 | 63 | 2 | 44.32 | 2828.25 | 2828.76 | 2829.26 | 337.66 | 346.14 | 354.63 | 0.25100 |
| 32 | 1 | 61 | 2 | 44.09 | 2844.19 | 2844.68 | 2845.18 | 242.31 | 250.32 | 258.31 | 0.24303 |
| 33 | 1 | 85 | 2 | 43.34 | 2897.23 | 2897.92 | 2898.60 | 352.03 | 2.69 | 13.36 | 0.42289 |
| 33 | 1 | 49 | 2 | 42.98 | 2923.42 | 2923.81 | 2924.21 | 57.13 | 63.79 | 70.43 | 0.24378 |
| 33 | 1 | 67 | 2 | 42.77 | 2938.76 | 2939.30 | 2939.84 | 310.94 | 319.39 | 327.83 | 0.33333 |
| 34 | 1 | 67 | 2 | 42.05 | 2991.53 | 2992.07 | 2992.61 | 25.97 | 34.77 | 43.55 | 0.21136 |
| 34 | 1 | 67 | 2 | 41.81 | 3009.47 | 3010.01 | 3010.55 | 322.11 | 330.98 | 339.85 | 0.21136 |
| 34 | 1 | 64 | 2 | 41.58 | 3026.57 | 3027.09 | 3027.60 | 239.56 | 247.82 | 256.07 | 0.20189 |
| 34 | 1 | 60 | 2 | 41.55 | 3029.03 | 3029.52 | 3030.00 | 278.87 | 286.58 | 294.29 | 0.18927 |
| 34 | 1 | 59 | 2 | 41.34 | 3044.88 | 3045.35 | 3045.83 | 169.69 | 177.23 | 184.78 | 0.18612 |
| 35 | 1 | 61 | 2 | 40.42 | 3115.97 | 3116.47 | 3116.96 | 179.57 | 186.72 | 193.89 | 0.24797 |
| 35 | 1 | 61 | 2 | 40.22 | 3132.33 | 3132.82 | 3133.31 | 59.02 | 66.23 | 73.43 | 0.24797 |
| 35 | 1 | 62 | 2 | 40.01 | 3148.87 | 3149.37 | 3149.87 | 296.33 | 302.78 | 309.18 | 0.25203 |
| 35 | 1 | 62 | 2 | 39.80 | 3165.51 | 3166.01 | 3166.51 | 153.89 | 160.94 | 168.00 | 0.25203 |
| 36 | 1 | 64 | 2 | 39.17 | 3216.06 | 3216.57 | 3217.09 | 131.25 | 138.43 | 145.62 | 0.24903 |
| 36 | 1 | 64 | 2 | 38.96 | 3233.09 | 3233.60 | 3234.12 | 9.09 | 16.31 | 23.53 | 0.24903 |
| 36 | 1 | 64 | 2 | 38.75 | 3250.27 | 3250.79 | 3251.30 | 244.92 | 251.68 | 258.40 | 0.24903 |
| 36 | 1 | 65 | 2 | 38.55 | 3267.53 | 3268.05 | 3268.58 | 100.10 | 106.69 | 113.25 | 0.25292 |
| 37 | 1 | 66 | 2 | 38.13 | 3302.42 | 3302.96 | 3303.49 | 235.64 | 243.39 | 251.07 | 0.24906 |
| 37 | 1 | 66 | 2 | 37.92 | 3319.98 | 3320.51 | 3321.04 | 115.10 | 122.16 | 129.19 | 0.24906 |
| 37 | 1 | 66 | 2 | 37.71 | 3337.66 | 3338.20 | 3338.73 | 350.15 | 357.32 | 4.52 | 0.24906 |
| 37 | 1 | 67 | 2 | 37.51 | 3355.46 | 3356.00 | 3356.54 | 237.67 | 245.50 | 253.39 | 0.25283 |
| 38 | 1 | 68 | 2 | 36.99 | 3400.15 | 3400.69 | 3401.24 | 205.47 | 213.81 | 222.11 | 0.24818 |
| 38 | 1 | 68 | 2 | 36.78 | 3418.35 | 3418.89 | 3419.44 | 105.95 | 113.35 | 120.71 | 0.24818 |
| 38 | 1 | 69 | 2 | 36.57 | 3436.66 | 3437.22 | 3437.77 | 348.76 | 356.15 | 3.53 | 0.25182 |
| 38 | 1 | 69 | 2 | 36.37 | 3455.09 | 3455.64 | 3456.20 | 229.98 | 237.30 | 244.65 | 0.25182 |
| 39 | 1 | 70 | 2 | 35.95 | 3492.24 | 3492.80 | 3493.37 | 36.20 | 44.70 | 53.21 | 0.19718 |
| 39 | 1 | 71 | 2 | 35.76 | 3509.34 | 3509.91 | 3510.49 | 294.26 | 302.69 | 311.11 | 0.20000 |
| 39 | 1 | 71 | 2 | 35.55 | 3528.22 | 3528.79 | 3529.36 | 195.07 | 202.45 | 209.80 | 0.20000 |
| 39 | 1 | 71 | 2 | 35.51 | 3532.41 | 3532.99 | 3533.56 | 248.71 | 255.98 | 263.23 | 0.20000 |
| 39 | 1 | 72 | 2 | 35.30 | 3551.36 | 3551.94 | 3552.52 | 125.89 | 133.67 | 141.47 | 0.20282 |
| 40 | 1 | 72 | 2 | 35.09 | 3570.56 | 3571.14 | 3571.72 | 19.15 | 26.60 | 34.06 | 0.24828 |
| 40 | 1 | 72 | 2 | 34.88 | 3589.75 | 3590.33 | 3590.91 | 269.29 | 277.35 | 285.46 | 0.24828 |
| 40 | 1 | 73 | 2 | 34.68 | 3609.12 | 3609.71 | 3610.30 | 195.03 | 203.98 | 212.90 | 0.25172 |
| 40 | 1 | 73 | 2 | 34.47 | 3628.63 | 3629.22 | 3629.80 | 117.54 | 126.33 | 135.16 | 0.25172 |
| 41 | 1 | 74 | 2 | 34.09 | 3663.96 | 3664.56 | 3665.16 | 268.63 | 277.19 | 285.80 | 0.24667 |
| 41 | 1 | 75 | 2 | 33.89 | 3683.75 | 3684.36 | 3684.96 | 187.75 | 195.90 | 204.05 | 0.25000 |
| 41 | 1 | 75 | 2 | 33.68 | 3703.73 | 3704.34 | 3704.94 | 107.68 | 116.26 | 124.83 | 0.25000 |
| 41 | 1 | 76 | 2 | 33.47 | 3723.81 | 3724.43 | 3725.04 | 28.88 | 37.39 | 45.91 | 0.25333 |
| 42 | 1 | 78 | 2 | 32.65 | 3804.69 | 3805.31 | 3805.94 | 54.90 | 62.44 | 69.93 | 0.33191 |
| 42 | 1 | 78 | 2 | 32.44 | 3825.35 | 3825.98 | 3826.61 | 291.17 | 299.26 | 307.42 | 0.33191 |
| 42 | 1 | 79 | 2 | 32.23 | 3846.22 | 3846.85 | 3847.49 | 203.12 | 211.21 | 219.29 | 0.33617 |
| 43 | 1 | 80 | 2 | 31.81 | 3888.40 | 3889.05 | 3889.69 | 331.84 | 337.60 | 343.35 | 0.33333 |
| 43 | 1 | 80 | 2 | 31.61 | 3909.56 | 3910.21 | 3910.85 | 189.59 | 197.57 | 205.62 | 0.33333 |
| 43 | 1 | 80 | 2 | 31.40 | 3930.93 | 3931.58 | 3932.22 | 143.81 | 153.20 | 162.57 | 0.33333 |
| 44 | 1 | 82 | 2 | 30.77 | 3995.86 | 3996.52 | 3997.18 | 35.72 | 47.25 | 58.75 | 0.33065 |
| 44 | 1 | 83 | 2 | 30.58 | 4016.26 | 4016.93 | 4017.60 | 32.04 | 43.34 | 54.59 | 0.33468 |
| 44 | 1 | 83 | 2 | 30.37 | 4038.31 | 4038.98 | 4039.64 | 23.19 | 33.27 | 43.32 | 0.33468 |
| 45 | 1 | 84 | 2 | 29.95 | 4082.81 | 4083.49 | 4084.17 | 285.04 | 294.50 | 303.96 | 0.33071 |
| 45 | 1 | 85 | 2 | 29.75 | 4105.25 | 4105.93 | 4106.62 | 232.35 | 241.46 | 250.59 | 0.33465 |
| 45 | 1 | 85 | 2 | 29.54 | 4127.84 | 4128.53 | 4129.21 | 165.53 | 174.03 | 182.50 | 0.33465 |
| 46 | 1 | 86 | 2 | 29.12 | 4173.48 | 4174.18 | 4174.87 | 290.53 | 298.17 | 305.85 | 0.33077 |
| 46 | 1 | 87 | 2 | 28.91 | 4196.49 | 4197.19 | 4197.89 | 205.17 | 213.99 | 222.84 | 0.33462 |
| 46 | 1 | 87 | 2 | 28.70 | 4219.67 | 4220.37 | 4221.07 | 139.86 | 149.25 | 158.63 | 0.33462 |
| 47 | 1 | 89 | 2 | 28.28 | 4266.44 | 4267.16 | 4267.88 | 354.70 | 2.86 | 11.00 | 0.33209 |
| 47 | 1 | 89 | 2 | 28.08 | 4290.07 | 4290.78 | 4291.50 | 270.74 | 279.40 | 288.15 | 0.33209 |
| 47 | 1 | 90 | 2 | 27.87 | 4313.81 | 4314.53 | 4315.26 | 167.14 | 173.71 | 180.28 | 0.33582 |
| 48 | 1 | 90 | 2 | 27.66 | 4337.84 | 4338.57 | 4339.29 | 39.87 | 47.37 | 54.89 | 0.33088 |
| 48 | 1 | 91 | 2 | 27.45 | 4361.77 | 4362.50 | 4363.24 | 299.08 | 306.84 | 314.60 | 0.33456 |
| 48 | 1 | 91 | 2 | 27.24 | 4385.97 | 4386.70 | 4387.44 | 186.59 | 193.92 | 201.28 | 0.33456 |
| 50 | 1 | 129 | 2 | 23.15 | 4887.35 | 4888.39 | 4889.43 | 123.64 | 133.86 | 144.05 | 0.50000 |
| 50 | 1 | 129 | 2 | 22.89 | 4921.04 | 4922.08 | 4923.13 | 101.29 | 111.97 | 122.61 | 0.50000 |
| 51 | 1 | 132 | 2 | 20.58 | 5230.56 | 5231.62 | 5232.68 | 31.82 | 41.83 | 51.78 | 0.49811 |
| 51 | 1 | 133 | 2 | 20.33 | 5265.50 | 5266.58 | 5267.65 | 342.61 | 352.06 | 1.54 | 0.50189 |
| 54 | 1 | 131 | 2 | 16.93 | 5758.28 | 5759.33 | 5760.39 | 144.74 | 154.54 | 164.34 | 0.51172 |
| 54 | 1 | 125 | 2 | 16.72 | 5790.14 | 5791.15 | 5792.15 | 82.83 | 92.54 | 102.27 | 0.48828 |
| 55 | 1 | 117 | 2 | 14.75 | 6097.44 | 6098.38 | 6099.33 | 114.43 | 123.21 | 132.02 | 0.49787 |
| 55 | 1 | 118 | 2 | 14.56 | 6128.50 | 6129.46 | 6130.41 | 55.66 | 65.07 | 74.51 | 0.50213 |
| 56 | 1 | 124 | 2 | 12.80 | 6416.28 | 6417.28 | 6418.28 | 158.75 | 168.21 | 177.67 | 0.50000 |
| 56 | 1 | 124 | 2 | 12.60 | 6448.93 | 6449.93 | 6450.93 | 94.48 | 103.21 | 111.91 | 0.50000 |
| 57 | 1 | 130 | 2 | 10.83 | 6751.47 | 6752.52 | 6753.57 | 53.06 | 61.70 | 70.37 | 0.49808 |
| 57 | 1 | 131 | 2 | 10.63 | 6785.79 | 6786.85 | 6787.90 | 335.93 | 344.22 | 352.49 | 0.50192 |
| 58 | 1 | 135 | 2 | 9.06 | 7066.43 | 7067.51 | 7068.60 | 99.79 | 107.89 | 115.96 | 0.49815 |
| 58 | 1 | 136 | 2 | 8.86 | 7101.97 | 7103.07 | 7104.16 | 357.64 | 5.52 | 13.45 | 0.50185 |
| 59 | 1 | 141 | 2 | 7.28 | 7393.94 | 7395.07 | 7396.21 | 17.44 | 25.71 | 33.96 | 0.50000 |
| 59 | 1 | 141 | 2 | 7.09 | 7430.93 | 7432.07 | 7433.21 | 283.94 | 292.38 | 300.82 | 0.50000 |
| 60 | 1 | 146 | 2 | 5.51 | 7733.25 | 7734.43 | 7735.61 | 8.84 | 17.05 | 25.20 | 0.50000 |
| 60 | 1 | 146 | 2 | 5.32 | 7771.50 | 7772.68 | 7773.86 | 266.06 | 274.03 | 282.04 | 0.50000 |
| 1 | 2 | 1 | 10 | 142.02 | 150.10 | 150.14 | 150.18 | 10.74 | 9.57 | 8.39 | 0.07143 |
| 1 | 2 | 1 | 10 | 142.01 | 150.26 | 150.30 | 150.34 | 6.04 | 4.86 | 3.69 | 0.07143 |
| 1 | 2 | 1 | 10 | 141.53 | 159.48 | 159.52 | 159.56 | 100.34 | 99.19 | 98.04 | 0.07143 |
| 1 | 2 | 1 | 10 | 141.52 | 159.64 | 159.68 | 159.72 | 95.74 | 94.59 | 93.44 | 0.07143 |
| 1 | 2 | 1 | 10 | 141.03 | 169.17 | 169.21 | 169.25 | 187.37 | 186.25 | 185.12 | 0.07143 |
| 1 | 2 | 1 | 10 | 141.03 | 169.33 | 169.37 | 169.41 | 182.88 | 181.76 | 180.64 | 0.07143 |
| 1 | 2 | 1 | 12 | 140.52 | 179.11 | 179.16 | 179.21 | 273.97 | 272.65 | 271.34 | 0.07143 |
| 1 | 2 | 1 | 12 | 140.51 | 179.31 | 179.35 | 179.40 | 268.71 | 267.40 | 266.08 | 0.07143 |
| 1 | 2 | 1 | 12 | 140.00 | 189.38 | 189.43 | 189.47 | 358.73 | 357.45 | 356.17 | 0.07143 |
| 1 | 2 | 1 | 12 | 139.99 | 189.57 | 189.62 | 189.67 | 353.61 | 352.33 | 351.05 | 0.07143 |
| 1 | 2 | 1 | 12 | 139.47 | 199.97 | 200.01 | 200.06 | 82.77 | 81.53 | 80.29 | 0.07143 |
| 1 | 2 | 1 | 12 | 139.46 | 200.16 | 200.21 | 200.26 | 77.81 | 76.57 | 75.33 | 0.07143 |
| 1 | 2 | 1 | 12 | 138.93 | 210.87 | 210.92 | 210.97 | 167.81 | 166.61 | 165.42 | 0.07143 |
| 1 | 2 | 1 | 12 | 138.92 | 211.06 | 211.11 | 211.16 | 163.02 | 161.82 | 160.62 | 0.07143 |
| 2 | 2 | 20 | 2 | 134.39 | 304.73 | 304.89 | 305.05 | 305.83 | 302.93 | 300.02 | 0.55556 |
| 2 | 2 | 16 | 2 | 132.71 | 341.00 | 341.13 | 341.26 | 77.60 | 75.75 | 73.90 | 0.44444 |
| 3 | 2 | 16 | 2 | 128.34 | 439.74 | 439.87 | 440.00 | 121.46 | 120.56 | 119.65 | 0.21053 |
| 3 | 2 | 16 | 2 | 128.31 | 440.25 | 440.38 | 440.51 | 117.85 | 116.95 | 116.06 | 0.21053 |
| 3 | 2 | 22 | 2 | 126.08 | 493.07 | 493.25 | 493.43 | 218.50 | 217.98 | 217.46 | 0.28947 |
| 3 | 2 | 22 | 2 | 126.05 | 493.78 | 493.96 | 494.14 | 216.42 | 215.90 | 215.39 | 0.28947 |
| 4 | 2 | 18 | 2 | 124.64 | 528.30 | 528.45 | 528.59 | 154.21 | 154.10 | 153.99 | 0.15385 |
| 4 | 2 | 18 | 2 | 124.62 | 528.88 | 529.03 | 529.17 | 153.78 | 153.67 | 153.57 | 0.15385 |
| 4 | 2 | 18 | 2 | 124.59 | 529.46 | 529.61 | 529.75 | 153.37 | 153.27 | 153.17 | 0.15385 |
| 4 | 2 | 21 | 2 | 123.00 | 569.05 | 569.21 | 569.38 | 170.79 | 171.06 | 171.33 | 0.17949 |
| 4 | 2 | 21 | 2 | 122.97 | 569.72 | 569.89 | 570.06 | 171.88 | 172.15 | 172.43 | 0.17949 |
| 4 | 2 | 21 | 2 | 122.94 | 570.40 | 570.57 | 570.74 | 173.00 | 173.28 | 173.57 | 0.17949 |
| 8 | 2 | 16 | 2 | 112.90 | 841.39 | 841.52 | 841.65 | 50.74 | 52.75 | 54.75 | 0.21918 |
| 8 | 2 | 16 | 2 | 112.89 | 841.91 | 842.04 | 842.17 | 58.78 | 60.79 | 62.80 | 0.21918 |
| 8 | 2 | 22 | 2 | 112.06 | 865.96 | 866.14 | 866.32 | 87.42 | 90.39 | 93.35 | 0.30137 |
| 8 | 2 | 19 | 2 | 111.32 | 888.20 | 888.36 | 888.51 | 110.15 | 112.86 | 115.58 | 0.26027 |
| 9 | 2 | 15 | 2 | 107.38 | 963.17 | 963.30 | 963.42 | 221.33 | 224.35 | 227.38 | 0.37500 |
| 9 | 2 | 11 | 2 | 105.04 | 995.49 | 995.58 | 995.67 | 55.85 | 58.66 | 61.47 | 0.27500 |
| 9 | 2 | 14 | 2 | 103.27 | 1019.48 | 1019.59 | 1019.70 | 131.91 | 135.80 | 139.69 | 0.35000 |
| 10 | 2 | 18 | 2 | 100.84 | 1053.11 | 1053.26 | 1053.40 | 259.47 | 264.90 | 270.33 | 0.21687 |
| 10 | 2 | 12 | 2 | 100.10 | 1063.55 | 1063.65 | 1063.74 | 294.91 | 298.62 | 302.34 | 0.14458 |
| 10 | 2 | 12 | 2 | 99.40 | 1073.49 | 1073.59 | 1073.69 | 322.72 | 326.54 | 330.38 | 0.14458 |
| 10 | 2 | 13 | 2 | 98.00 | 1093.72 | 1093.83 | 1093.93 | 70.35 | 74.79 | 79.23 | 0.15663 |
| 10 | 2 | 14 | 2 | 97.39 | 1102.70 | 1102.81 | 1102.93 | 95.94 | 100.86 | 105.79 | 0.16867 |
| 10 | 2 | 14 | 2 | 96.81 | 1111.43 | 1111.55 | 1111.66 | 121.88 | 126.94 | 132.00 | 0.16867 |
| 11 | 2 | 14 | 2 | 94.63 | 1144.53 | 1144.64 | 1144.76 | 235.01 | 240.55 | 246.09 | 0.17073 |
| 11 | 2 | 12 | 2 | 94.13 | 1152.41 | 1152.51 | 1152.60 | 265.29 | 270.12 | 274.97 | 0.14634 |
| 11 | 2 | 4 | 2 | 93.47 | 1162.87 | 1162.90 | 1162.93 | 75.49 | 77.14 | 78.80 | 0.04878 |
| 11 | 2 | 10 | 2 | 93.00 | 1170.21 | 1170.29 | 1170.37 | 95.33 | 99.53 | 103.74 | 0.12195 |
| 11 | 2 | 12 | 2 | 92.54 | 1177.53 | 1177.63 | 1177.72 | 120.81 | 125.95 | 131.09 | 0.14634 |
| 11 | 2 | 10 | 2 | 92.10 | 1184.76 | 1184.84 | 1184.92 | 148.38 | 152.74 | 157.09 | 0.12195 |
| 11 | 2 | 10 | 2 | 91.67 | 1191.81 | 1191.89 | 1191.97 | 172.85 | 177.30 | 181.73 | 0.12195 |
| 11 | 2 | 10 | 2 | 91.25 | 1198.71 | 1198.79 | 1198.87 | 195.99 | 200.50 | 205.02 | 0.12195 |
| 12 | 2 | 7 | 2 | 89.16 | 1233.82 | 1233.87 | 1233.93 | 59.44 | 62.75 | 66.06 | 0.07609 |
| 12 | 2 | 10 | 2 | 88.81 | 1239.76 | 1239.84 | 1239.92 | 47.49 | 52.21 | 56.92 | 0.10870 |
| 12 | 2 | 11 | 2 | 88.46 | 1245.93 | 1246.02 | 1246.11 | 48.05 | 53.23 | 58.41 | 0.11957 |
| 12 | 2 | 9 | 2 | 88.10 | 1252.10 | 1252.17 | 1252.24 | 47.20 | 51.42 | 55.64 | 0.09783 |
| 12 | 2 | 9 | 2 | 87.75 | 1258.16 | 1258.24 | 1258.31 | 38.42 | 42.61 | 46.81 | 0.09783 |
| 12 | 2 | 10 | 2 | 87.42 | 1264.12 | 1264.20 | 1264.28 | 21.85 | 26.48 | 31.12 | 0.10870 |
| 12 | 2 | 8 | 2 | 87.08 | 1270.11 | 1270.17 | 1270.24 | 5.18 | 8.87 | 12.57 | 0.08696 |
| 12 | 2 | 8 | 2 | 86.75 | 1275.99 | 1276.06 | 1276.12 | 341.81 | 345.49 | 349.17 | 0.08696 |
| 12 | 2 | 9 | 2 | 86.43 | 1281.83 | 1281.90 | 1281.98 | 314.71 | 318.85 | 322.98 | 0.09783 |
| 12 | 2 | 11 | 2 | 85.99 | 1289.77 | 1289.85 | 1289.94 | 45.69 | 50.73 | 55.76 | 0.11957 |
| 13 | 2 | 10 | 2 | 84.43 | 1318.58 | 1318.66 | 1318.74 | 233.14 | 237.68 | 242.22 | 0.09901 |
| 13 | 2 | 11 | 2 | 83.92 | 1328.18 | 1328.26 | 1328.35 | 53.03 | 58.01 | 62.99 | 0.10891 |
| 13 | 2 | 10 | 2 | 83.59 | 1334.57 | 1334.65 | 1334.73 | 51.16 | 55.66 | 60.16 | 0.09901 |
| 13 | 2 | 10 | 2 | 83.29 | 1340.30 | 1340.38 | 1340.46 | 9.91 | 14.38 | 18.85 | 0.09901 |
| 13 | 2 | 10 | 2 | 82.99 | 1346.01 | 1346.09 | 1346.17 | 325.82 | 330.25 | 334.69 | 0.09901 |
| 13 | 2 | 10 | 2 | 82.70 | 1351.72 | 1351.80 | 1351.88 | 278.20 | 282.58 | 286.97 | 0.09901 |
| 13 | 2 | 10 | 2 | 82.41 | 1357.46 | 1357.54 | 1357.62 | 228.15 | 232.48 | 236.81 | 0.09901 |
| 13 | 2 | 10 | 2 | 82.11 | 1363.24 | 1363.32 | 1363.40 | 176.25 | 180.51 | 184.78 | 0.09901 |
| 13 | 2 | 10 | 2 | 81.82 | 1369.01 | 1369.09 | 1369.17 | 118.97 | 123.16 | 127.35 | 0.09901 |
| 13 | 2 | 10 | 2 | 81.53 | 1374.85 | 1374.93 | 1375.01 | 59.89 | 64.02 | 68.13 | 0.09901 |
| 14 | 2 | 5 | 2 | 79.73 | 1411.44 | 1411.48 | 1411.52 | 55.33 | 57.26 | 59.18 | 0.05319 |
| 14 | 2 | 10 | 2 | 79.15 | 1423.61 | 1423.69 | 1423.77 | 275.51 | 279.34 | 283.17 | 0.10638 |
| 14 | 2 | 11 | 2 | 78.86 | 1429.77 | 1429.85 | 1429.94 | 208.18 | 212.40 | 216.61 | 0.11702 |
| 14 | 2 | 11 | 2 | 78.57 | 1436.00 | 1436.09 | 1436.18 | 143.38 | 147.57 | 151.76 | 0.11702 |
| 14 | 2 | 11 | 2 | 78.27 | 1442.26 | 1442.35 | 1442.44 | 78.31 | 82.48 | 86.65 | 0.11702 |
| 14 | 2 | 10 | 2 | 77.98 | 1448.59 | 1448.67 | 1448.75 | 14.55 | 18.31 | 22.07 | 0.10638 |
| 14 | 2 | 12 | 2 | 77.69 | 1454.93 | 1455.02 | 1455.12 | 308.48 | 312.94 | 317.41 | 0.12766 |
| 14 | 2 | 12 | 2 | 77.39 | 1461.37 | 1461.47 | 1461.56 | 243.98 | 248.39 | 252.80 | 0.12766 |
| 14 | 2 | 12 | 2 | 77.10 | 1467.86 | 1467.96 | 1468.05 | 177.99 | 182.34 | 186.69 | 0.12766 |
| 15 | 2 | 17 | 2 | 74.26 | 1532.94 | 1533.08 | 1533.22 | 58.74 | 64.30 | 69.86 | 0.16832 |
| 15 | 2 | 14 | 2 | 73.65 | 1547.80 | 1547.92 | 1548.03 | 298.76 | 303.27 | 307.78 | 0.13861 |
| 15 | 2 | 13 | 2 | 73.34 | 1555.23 | 1555.34 | 1555.44 | 231.63 | 235.73 | 239.81 | 0.12871 |
| 15 | 2 | 14 | 2 | 73.04 | 1562.67 | 1562.79 | 1562.90 | 156.57 | 160.81 | 165.05 | 0.13861 |
| 15 | 2 | 14 | 2 | 72.73 | 1570.23 | 1570.34 | 1570.46 | 75.99 | 80.12 | 84.25 | 0.13861 |
| 15 | 2 | 14 | 2 | 72.42 | 1577.90 | 1578.01 | 1578.13 | 354.83 | 358.85 | 2.88 | 0.13861 |
| 15 | 2 | 15 | 2 | 72.11 | 1585.66 | 1585.78 | 1585.90 | 269.96 | 274.22 | 278.49 | 0.14851 |
| 16 | 2 | 16 | 2 | 70.45 | 1628.58 | 1628.70 | 1628.83 | 338.94 | 343.46 | 347.99 | 0.13913 |
| 16 | 2 | 16 | 2 | 70.14 | 1637.04 | 1637.17 | 1637.30 | 275.22 | 279.70 | 284.18 | 0.13913 |
| 16 | 2 | 16 | 2 | 69.82 | 1645.62 | 1645.75 | 1645.88 | 210.60 | 215.05 | 219.51 | 0.13913 |
| 16 | 2 | 17 | 2 | 69.50 | 1654.27 | 1654.41 | 1654.54 | 150.66 | 155.41 | 160.16 | 0.14783 |
| 16 | 2 | 16 | 2 | 69.19 | 1662.98 | 1663.11 | 1663.24 | 81.02 | 85.17 | 89.32 | 0.13913 |
| 16 | 2 | 17 | 2 | 68.88 | 1671.74 | 1671.88 | 1672.02 | 4.80 | 9.24 | 13.69 | 0.14783 |
| 16 | 2 | 17 | 2 | 68.56 | 1680.71 | 1680.85 | 1680.98 | 296.30 | 300.78 | 305.25 | 0.14783 |
| 17 | 2 | 17 | 2 | 67.28 | 1717.81 | 1717.95 | 1718.09 | 68.83 | 73.24 | 77.63 | 0.13281 |
| 17 | 2 | 18 | 2 | 66.97 | 1727.17 | 1727.31 | 1727.46 | 1.84 | 6.22 | 10.59 | 0.14063 |
| 17 | 2 | 18 | 2 | 66.66 | 1736.64 | 1736.79 | 1736.93 | 287.55 | 291.99 | 296.44 | 0.14063 |
| 17 | 2 | 18 | 2 | 66.35 | 1746.18 | 1746.33 | 1746.47 | 222.98 | 227.44 | 231.90 | 0.14063 |
| 17 | 2 | 19 | 2 | 66.04 | 1755.84 | 1756.00 | 1756.15 | 150.56 | 155.04 | 159.53 | 0.14844 |
| 17 | 2 | 19 | 2 | 65.73 | 1765.65 | 1765.80 | 1765.96 | 75.19 | 79.59 | 84.00 | 0.14844 |
| 17 | 2 | 19 | 2 | 65.43 | 1775.58 | 1775.73 | 1775.88 | 357.20 | 1.55 | 5.90 | 0.14844 |
| 18 | 2 | 20 | 2 | 64.45 | 1807.87 | 1808.03 | 1808.19 | 209.17 | 213.80 | 218.42 | 0.16129 |
| 18 | 2 | 20 | 2 | 64.15 | 1818.21 | 1818.37 | 1818.53 | 140.42 | 144.96 | 149.50 | 0.16129 |
| 18 | 2 | 20 | 2 | 63.84 | 1828.74 | 1828.90 | 1829.06 | 79.53 | 84.13 | 88.72 | 0.16129 |
| 18 | 2 | 21 | 2 | 63.54 | 1839.39 | 1839.56 | 1839.73 | 20.98 | 25.75 | 30.52 | 0.16935 |
| 18 | 2 | 21 | 2 | 63.24 | 1850.25 | 1850.42 | 1850.59 | 322.08 | 326.67 | 331.26 | 0.16935 |
| 18 | 2 | 22 | 2 | 62.93 | 1861.36 | 1861.54 | 1861.72 | 265.53 | 270.45 | 275.36 | 0.17742 |
| 22 | 2 | 22 | 2 | 59.04 | 2015.49 | 2015.67 | 2015.84 | 58.30 | 62.69 | 67.08 | 0.18487 |
| 22 | 2 | 24 | 2 | 58.75 | 2027.94 | 2028.13 | 2028.33 | 7.03 | 11.85 | 16.69 | 0.20168 |
| 22 | 2 | 24 | 2 | 58.46 | 2040.44 | 2040.63 | 2040.83 | 322.97 | 327.95 | 332.93 | 0.20168 |
| 22 | 2 | 24 | 2 | 58.17 | 2052.94 | 2053.13 | 2053.33 | 277.95 | 282.83 | 287.71 | 0.20168 |
| 22 | 2 | 25 | 2 | 57.89 | 2065.45 | 2065.66 | 2065.86 | 237.12 | 242.22 | 247.32 | 0.21008 |
| 23 | 2 | 25 | 2 | 57.04 | 2103.71 | 2103.91 | 2104.11 | 85.76 | 90.49 | 95.23 | 0.19841 |
| 23 | 2 | 25 | 2 | 56.75 | 2117.16 | 2117.36 | 2117.56 | 38.18 | 42.87 | 47.55 | 0.19841 |
| 23 | 2 | 25 | 2 | 56.48 | 2129.93 | 2130.14 | 2130.34 | 338.22 | 342.91 | 347.60 | 0.19841 |
| 23 | 2 | 26 | 2 | 56.14 | 2145.84 | 2146.05 | 2146.26 | 347.95 | 352.79 | 357.63 | 0.20635 |
| 23 | 2 | 25 | 2 | 55.86 | 2158.91 | 2159.11 | 2159.31 | 277.10 | 281.45 | 285.81 | 0.19841 |
| 24 | 2 | 29 | 2 | 54.21 | 2239.84 | 2240.07 | 2240.30 | 128.94 | 133.37 | 137.81 | 0.19728 |
| 24 | 2 | 28 | 2 | 53.93 | 2254.13 | 2254.36 | 2254.58 | 36.84 | 41.01 | 45.18 | 0.19048 |
| 24 | 2 | 29 | 2 | 53.65 | 2268.71 | 2268.95 | 2269.18 | 309.63 | 314.00 | 318.36 | 0.19728 |
| 24 | 2 | 30 | 2 | 53.36 | 2283.63 | 2283.87 | 2284.11 | 232.34 | 236.97 | 241.57 | 0.20408 |
| 24 | 2 | 31 | 2 | 52.89 | 2308.17 | 2308.42 | 2308.67 | 318.19 | 322.50 | 326.81 | 0.21088 |
| 25 | 2 | 31 | 2 | 52.50 | 2328.84 | 2329.09 | 2329.34 | 311.41 | 315.82 | 320.25 | 0.23308 |
| 25 | 2 | 33 | 2 | 52.20 | 2344.90 | 2345.16 | 2345.43 | 232.56 | 237.11 | 241.64 | 0.24812 |
| 25 | 2 | 34 | 2 | 51.90 | 2361.31 | 2361.59 | 2361.86 | 145.86 | 150.30 | 154.75 | 0.25564 |
| 25 | 2 | 35 | 2 | 51.59 | 2378.14 | 2378.43 | 2378.71 | 52.75 | 57.13 | 61.51 | 0.26316 |
| 26 | 2 | 35 | 2 | 50.97 | 2412.43 | 2412.72 | 2413.00 | 229.05 | 233.57 | 238.07 | 0.24648 |
| 26 | 2 | 36 | 2 | 50.66 | 2429.97 | 2430.26 | 2430.55 | 149.79 | 154.46 | 159.13 | 0.25352 |
| 26 | 2 | 36 | 2 | 50.35 | 2447.62 | 2447.91 | 2448.20 | 71.89 | 76.45 | 81.00 | 0.25352 |
| 26 | 2 | 35 | 2 | 50.04 | 2465.29 | 2465.58 | 2465.86 | 340.26 | 344.37 | 348.47 | 0.24648 |
| 30 | 2 | 34 | 2 | 47.41 | 2624.32 | 2624.59 | 2624.87 | 311.25 | 315.62 | 319.98 | 0.25185 |
| 30 | 2 | 34 | 2 | 47.14 | 2641.47 | 2641.75 | 2642.02 | 222.29 | 226.55 | 230.82 | 0.25185 |
| 30 | 2 | 36 | 2 | 46.86 | 2658.84 | 2659.13 | 2659.42 | 127.92 | 132.31 | 136.69 | 0.26667 |
| 30 | 2 | 31 | 2 | 46.59 | 2676.49 | 2676.74 | 2676.99 | 29.77 | 33.41 | 37.06 | 0.22963 |
| 31 | 2 | 32 | 2 | 45.93 | 2719.66 | 2719.91 | 2720.17 | 344.83 | 349.04 | 353.24 | 0.42105 |
| 31 | 2 | 44 | 2 | 45.55 | 2745.05 | 2745.41 | 2745.76 | 27.85 | 33.41 | 38.96 | 0.57895 |
| 32 | 2 | 33 | 2 | 44.78 | 2796.80 | 2797.07 | 2797.33 | 167.19 | 171.84 | 176.49 | 0.27049 |
| 32 | 2 | 33 | 2 | 44.54 | 2813.47 | 2813.73 | 2814.00 | 94.09 | 98.52 | 102.96 | 0.27049 |
| 32 | 2 | 22 | 2 | 44.30 | 2830.05 | 2830.22 | 2830.40 | 7.70 | 10.66 | 13.62 | 0.18033 |
| 32 | 2 | 34 | 2 | 44.06 | 2846.44 | 2846.72 | 2846.99 | 278.78 | 283.21 | 287.63 | 0.27869 |
| 33 | 2 | 51 | 2 | 43.29 | 2901.35 | 2901.76 | 2902.17 | 56.21 | 62.66 | 69.14 | 0.60714 |
| 33 | 2 | 33 | 2 | 42.74 | 2940.75 | 2941.01 | 2941.28 | 341.87 | 345.98 | 350.11 | 0.39286 |
| 34 | 2 | 36 | 2 | 42.27 | 2975.84 | 2976.13 | 2976.42 | 144.78 | 149.02 | 153.27 | 0.26866 |
| 34 | 2 | 36 | 2 | 42.02 | 2993.78 | 2994.07 | 2994.36 | 62.75 | 67.48 | 72.21 | 0.26866 |
| 34 | 2 | 31 | 2 | 41.79 | 3011.43 | 3011.68 | 3011.93 | 354.26 | 358.36 | 2.46 | 0.23134 |
| 34 | 2 | 31 | 2 | 41.53 | 3030.83 | 3031.08 | 3031.33 | 307.49 | 311.44 | 315.41 | 0.23134 |
| 35 | 2 | 33 | 2 | 40.40 | 3117.81 | 3118.07 | 3118.34 | 206.22 | 210.10 | 213.98 | 0.24627 |
| 35 | 2 | 33 | 2 | 40.20 | 3134.25 | 3134.51 | 3134.78 | 87.13 | 91.04 | 94.95 | 0.24627 |
| 35 | 2 | 34 | 2 | 39.99 | 3150.79 | 3151.07 | 3151.34 | 320.98 | 324.45 | 327.92 | 0.25373 |
| 35 | 2 | 34 | 2 | 39.78 | 3167.44 | 3167.71 | 3167.99 | 181.12 | 185.00 | 188.90 | 0.25373 |
| 36 | 2 | 35 | 2 | 39.15 | 3218.01 | 3218.29 | 3218.57 | 158.41 | 162.33 | 166.26 | 0.25000 |
| 36 | 2 | 35 | 2 | 38.94 | 3235.09 | 3235.37 | 3235.65 | 37.02 | 40.96 | 44.88 | 0.25000 |
| 36 | 2 | 35 | 2 | 38.73 | 3252.26 | 3252.54 | 3252.82 | 270.70 | 274.31 | 277.91 | 0.25000 |
| 36 | 2 | 35 | 2 | 38.53 | 3269.53 | 3269.81 | 3270.09 | 125.16 | 128.68 | 132.21 | 0.25000 |
| 37 | 2 | 36 | 2 | 38.11 | 3304.41 | 3304.70 | 3304.99 | 264.23 | 268.35 | 272.45 | 0.25000 |
| 37 | 2 | 36 | 2 | 37.90 | 3322.01 | 3322.30 | 3322.59 | 141.96 | 145.78 | 149.61 | 0.25000 |
| 37 | 2 | 36 | 2 | 37.69 | 3339.72 | 3340.01 | 3340.30 | 17.97 | 21.92 | 25.86 | 0.25000 |
| 37 | 2 | 36 | 2 | 37.48 | 3357.53 | 3357.82 | 3358.11 | 267.98 | 272.28 | 276.59 | 0.25000 |
| 38 | 2 | 37 | 2 | 37.18 | 3384.16 | 3384.46 | 3384.76 | 317.89 | 322.53 | 327.18 | 0.25000 |
| 38 | 2 | 37 | 2 | 36.97 | 3402.25 | 3402.55 | 3402.85 | 237.28 | 241.74 | 246.19 | 0.25000 |
| 38 | 2 | 37 | 2 | 36.76 | 3420.47 | 3420.76 | 3421.06 | 134.40 | 138.37 | 142.34 | 0.25000 |
| 38 | 2 | 37 | 2 | 36.55 | 3438.79 | 3439.09 | 3439.38 | 17.00 | 20.95 | 24.91 | 0.25000 |
| 39 | 2 | 38 | 2 | 36.14 | 3475.76 | 3476.07 | 3476.38 | 153.48 | 157.90 | 162.30 | 0.33333 |
| 39 | 2 | 38 | 2 | 35.74 | 3511.53 | 3511.83 | 3512.14 | 326.36 | 330.85 | 335.32 | 0.33333 |
| 39 | 2 | 38 | 2 | 35.49 | 3534.53 | 3534.84 | 3535.15 | 275.50 | 279.32 | 283.13 | 0.33333 |
| 40 | 2 | 39 | 2 | 35.07 | 3572.72 | 3573.03 | 3573.35 | 46.92 | 50.97 | 55.02 | 0.24684 |
| 40 | 2 | 39 | 2 | 34.86 | 3591.99 | 3592.31 | 3592.62 | 300.71 | 305.17 | 309.64 | 0.24684 |
| 40 | 2 | 40 | 2 | 34.65 | 3611.39 | 3611.71 | 3612.03 | 229.29 | 234.12 | 238.93 | 0.25316 |
| 40 | 2 | 40 | 2 | 34.45 | 3630.90 | 3631.22 | 3631.55 | 151.63 | 156.48 | 161.32 | 0.25316 |
| 41 | 2 | 40 | 2 | 34.07 | 3666.22 | 3666.54 | 3666.87 | 301.25 | 305.94 | 310.64 | 0.28369 |
| 41 | 2 | 41 | 2 | 33.87 | 3686.07 | 3686.40 | 3686.74 | 219.09 | 223.59 | 228.09 | 0.29078 |
| 41 | 2 | 41 | 2 | 33.66 | 3706.05 | 3706.38 | 3706.71 | 140.52 | 145.20 | 149.88 | 0.29078 |
| 41 | 2 | 19 | 2 | 33.45 | 3725.89 | 3726.05 | 3726.20 | 57.84 | 59.99 | 62.13 | 0.13475 |
| 42 | 2 | 42 | 2 | 32.63 | 3807.03 | 3807.37 | 3807.71 | 82.69 | 86.62 | 90.53 | 0.32813 |
| 42 | 2 | 43 | 2 | 32.42 | 3827.77 | 3828.12 | 3828.47 | 322.68 | 327.28 | 331.89 | 0.33594 |
| 42 | 2 | 43 | 2 | 32.21 | 3848.64 | 3848.99 | 3849.34 | 233.92 | 238.32 | 242.72 | 0.33594 |
| 43 | 2 | 43 | 2 | 31.79 | 3890.78 | 3891.13 | 3891.47 | 353.06 | 356.15 | 359.25 | 0.32824 |
| 43 | 2 | 44 | 2 | 31.59 | 3912.03 | 3912.38 | 3912.74 | 220.57 | 225.14 | 229.73 | 0.33588 |
| 43 | 2 | 44 | 2 | 31.38 | 3933.43 | 3933.78 | 3934.14 | 179.99 | 185.09 | 190.17 | 0.33588 |
| 44 | 2 | 45 | 2 | 30.96 | 3976.61 | 3976.97 | 3977.33 | 88.85 | 94.33 | 99.78 | 0.49451 |
| 44 | 2 | 1 | 2 | 30.76 | 3997.95 | 3997.95 | 3997.96 | 72.06 | 72.20 | 72.33 | 0.01099 |
| 44 | 2 | 45 | 2 | 30.56 | 4018.81 | 4019.18 | 4019.54 | 74.93 | 80.96 | 86.98 | 0.49451 |
| 45 | 2 | 46 | 2 | 29.93 | 4085.40 | 4085.77 | 4086.14 | 321.26 | 326.46 | 331.68 | 0.33094 |
| 45 | 2 | 46 | 2 | 29.72 | 4107.87 | 4108.25 | 4108.62 | 267.31 | 272.24 | 277.19 | 0.33094 |
| 45 | 2 | 47 | 2 | 29.52 | 4130.49 | 4130.86 | 4131.24 | 198.14 | 202.76 | 207.37 | 0.33813 |
| 46 | 2 | 47 | 2 | 29.10 | 4176.15 | 4176.52 | 4176.90 | 320.13 | 324.42 | 328.73 | 0.32867 |
| 46 | 2 | 48 | 2 | 28.89 | 4199.18 | 4199.57 | 4199.96 | 239.05 | 243.90 | 248.74 | 0.33566 |
| 46 | 2 | 48 | 2 | 28.68 | 4222.38 | 4222.76 | 4223.15 | 176.13 | 181.30 | 186.46 | 0.33566 |
| 47 | 2 | 48 | 2 | 28.47 | 4245.71 | 4246.09 | 4246.48 | 108.69 | 113.44 | 118.21 | 0.33103 |
| 47 | 2 | 48 | 2 | 28.26 | 4269.20 | 4269.58 | 4269.97 | 26.08 | 30.50 | 34.92 | 0.33103 |
| 47 | 2 | 49 | 2 | 28.05 | 4292.82 | 4293.21 | 4293.61 | 304.27 | 309.11 | 313.98 | 0.33793 |
| 48 | 2 | 49 | 2 | 27.64 | 4340.51 | 4340.91 | 4341.30 | 67.59 | 71.74 | 75.89 | 0.32886 |
| 48 | 2 | 50 | 2 | 27.43 | 4364.58 | 4364.99 | 4365.39 | 328.80 | 333.05 | 337.28 | 0.33557 |
| 48 | 2 | 50 | 2 | 27.22 | 4388.80 | 4389.20 | 4389.60 | 214.96 | 219.05 | 223.12 | 0.33557 |
| 50 | 2 | 69 | 2 | 23.12 | 4891.12 | 4891.68 | 4892.23 | 160.54 | 165.95 | 171.35 | 0.49286 |
| 50 | 2 | 71 | 2 | 22.86 | 4925.05 | 4925.62 | 4926.20 | 142.29 | 148.11 | 153.91 | 0.50714 |
| 51 | 2 | 72 | 2 | 20.81 | 5199.82 | 5200.40 | 5200.98 | 83.84 | 89.91 | 95.99 | 0.49655 |
| 51 | 2 | 73 | 2 | 20.55 | 5234.65 | 5235.24 | 5235.83 | 70.02 | 75.44 | 80.84 | 0.50345 |
| 54 | 2 | 72 | 2 | 17.14 | 5727.77 | 5728.35 | 5728.93 | 218.43 | 223.84 | 229.24 | 0.65455 |
| 54 | 2 | 38 | 2 | 16.91 | 5761.84 | 5762.15 | 5762.45 | 177.76 | 180.59 | 183.44 | 0.34545 |
| 55 | 2 | 64 | 2 | 14.93 | 6070.07 | 6070.59 | 6071.10 | 221.08 | 225.70 | 230.33 | 0.50000 |
| 55 | 2 | 64 | 2 | 14.73 | 6101.07 | 6101.59 | 6102.10 | 148.29 | 153.13 | 157.95 | 0.50000 |
| 56 | 2 | 68 | 2 | 12.78 | 6419.98 | 6420.53 | 6421.08 | 193.71 | 198.85 | 203.98 | 0.50000 |
| 56 | 2 | 68 | 2 | 12.58 | 6452.73 | 6453.28 | 6453.83 | 127.49 | 132.23 | 136.98 | 0.50000 |
| 57 | 2 | 72 | 2 | 10.81 | 6755.35 | 6755.93 | 6756.51 | 85.23 | 90.10 | 94.98 | 0.50000 |
| 57 | 2 | 72 | 2 | 10.61 | 6789.69 | 6790.27 | 6790.85 | 6.38 | 10.89 | 15.40 | 0.50000 |
| 58 | 2 | 74 | 2 | 9.04 | 7070.39 | 7070.99 | 7071.59 | 129.20 | 133.60 | 138.03 | 0.49664 |
| 58 | 2 | 75 | 2 | 8.84 | 7106.23 | 7106.83 | 7107.44 | 28.50 | 32.95 | 37.42 | 0.50336 |
| 59 | 2 | 77 | 2 | 7.27 | 7398.12 | 7398.74 | 7399.36 | 47.86 | 52.36 | 56.84 | 0.49677 |
| 59 | 2 | 78 | 2 | 7.07 | 7435.23 | 7435.86 | 7436.49 | 315.85 | 320.55 | 325.25 | 0.50323 |
| 60 | 2 | 80 | 2 | 5.49 | 7737.61 | 7738.25 | 7738.90 | 39.02 | 43.48 | 47.92 | 0.50000 |
| 60 | 2 | 80 | 2 | 5.30 | 7775.91 | 7776.56 | 7777.20 | 296.02 | 300.41 | 304.80 | 0.50000 |
| 1 | 3 | 1 | 12 | 141.96 | 151.24 | 151.29 | 151.34 | 337.38 | 335.98 | 334.57 | 0.02857 |
| 1 | 3 | 1 | 12 | 141.95 | 151.44 | 151.49 | 151.53 | 331.76 | 330.35 | 328.94 | 0.02857 |
| 1 | 3 | 1 | 12 | 141.94 | 151.63 | 151.68 | 151.73 | 326.13 | 324.73 | 323.32 | 0.02857 |
| 1 | 3 | 1 | 12 | 141.93 | 151.82 | 151.87 | 151.92 | 320.51 | 319.10 | 317.70 | 0.02857 |
| 1 | 3 | 1 | 10 | 141.47 | 160.64 | 160.68 | 160.72 | 67.29 | 66.15 | 65.00 | 0.02857 |
| 1 | 3 | 1 | 10 | 141.46 | 160.80 | 160.84 | 160.88 | 62.71 | 61.57 | 60.42 | 0.02857 |
| 1 | 3 | 1 | 10 | 141.46 | 160.96 | 161.00 | 161.04 | 58.13 | 56.99 | 55.84 | 0.02857 |
| 1 | 3 | 1 | 10 | 141.45 | 161.12 | 161.16 | 161.20 | 53.55 | 52.41 | 51.27 | 0.02857 |
| 1 | 3 | 1 | 10 | 141.44 | 161.28 | 161.32 | 161.36 | 48.98 | 47.84 | 46.69 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.97 | 170.38 | 170.42 | 170.46 | 153.75 | 152.63 | 151.51 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.96 | 170.54 | 170.58 | 170.62 | 149.27 | 148.16 | 147.04 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.95 | 170.70 | 170.74 | 170.78 | 144.80 | 143.68 | 142.57 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.95 | 170.86 | 170.90 | 170.94 | 140.33 | 139.21 | 138.10 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.94 | 171.02 | 171.06 | 171.10 | 135.86 | 134.75 | 133.63 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.46 | 180.41 | 180.45 | 180.49 | 238.75 | 237.66 | 236.57 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.45 | 180.57 | 180.61 | 180.65 | 234.39 | 233.30 | 232.21 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.44 | 180.73 | 180.77 | 180.81 | 230.03 | 228.94 | 227.85 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.43 | 180.89 | 180.93 | 180.98 | 225.67 | 224.58 | 223.49 | 0.02857 |
| 1 | 3 | 1 | 10 | 140.43 | 181.06 | 181.10 | 181.14 | 221.31 | 220.22 | 219.13 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.94 | 190.70 | 190.74 | 190.78 | 323.81 | 322.75 | 321.69 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.93 | 190.86 | 190.90 | 190.94 | 319.56 | 318.50 | 317.44 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.92 | 191.02 | 191.06 | 191.10 | 315.31 | 314.25 | 313.19 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.91 | 191.18 | 191.22 | 191.26 | 311.07 | 310.01 | 308.95 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.91 | 191.35 | 191.39 | 191.43 | 306.82 | 305.76 | 304.70 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.40 | 201.35 | 201.39 | 201.43 | 47.26 | 46.23 | 45.20 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.39 | 201.51 | 201.55 | 201.59 | 43.14 | 42.11 | 41.08 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.39 | 201.68 | 201.72 | 201.76 | 39.02 | 37.99 | 36.96 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.38 | 201.84 | 201.88 | 201.92 | 34.91 | 33.88 | 32.85 | 0.02857 |
| 1 | 3 | 1 | 10 | 139.37 | 202.00 | 202.04 | 202.08 | 30.79 | 29.76 | 28.74 | 0.02857 |
| 1 | 3 | 1 | 10 | 138.86 | 212.22 | 212.26 | 212.30 | 134.51 | 133.52 | 132.52 | 0.02857 |
| 1 | 3 | 1 | 10 | 138.85 | 212.38 | 212.42 | 212.46 | 130.53 | 129.54 | 128.55 | 0.02857 |
| 1 | 3 | 1 | 10 | 138.84 | 212.54 | 212.58 | 212.62 | 126.56 | 125.56 | 124.57 | 0.02857 |
| 1 | 3 | 1 | 10 | 138.84 | 212.70 | 212.74 | 212.78 | 122.58 | 121.59 | 120.59 | 0.02857 |
| 1 | 3 | 1 | 10 | 138.83 | 212.86 | 212.90 | 212.94 | 118.61 | 117.61 | 116.62 | 0.02857 |
| 1 | 3 | 1 | 10 | 138.82 | 213.02 | 213.06 | 213.10 | 114.64 | 113.64 | 112.65 | 0.02857 |
| 2 | 3 | 19 | 2 | 134.27 | 307.25 | 307.40 | 307.56 | 260.63 | 257.91 | 255.18 | 0.14504 |
| 2 | 3 | 19 | 2 | 134.25 | 307.86 | 308.02 | 308.17 | 249.74 | 247.03 | 244.31 | 0.14504 |
| 2 | 3 | 19 | 2 | 134.22 | 308.48 | 308.63 | 308.78 | 238.89 | 236.19 | 233.48 | 0.14504 |
| 2 | 3 | 18 | 2 | 133.35 | 327.23 | 327.37 | 327.52 | 285.30 | 283.01 | 280.72 | 0.13740 |
| 2 | 3 | 18 | 2 | 133.32 | 327.81 | 327.95 | 328.10 | 276.15 | 273.87 | 271.59 | 0.13740 |
| 2 | 3 | 19 | 2 | 132.62 | 343.07 | 343.22 | 343.38 | 48.00 | 45.83 | 43.65 | 0.14504 |
| 2 | 3 | 19 | 2 | 132.59 | 343.68 | 343.84 | 343.99 | 39.32 | 37.15 | 34.99 | 0.14504 |
| 3 | 3 | 23 | 2 | 128.20 | 442.88 | 443.06 | 443.25 | 99.85 | 98.59 | 97.34 | 0.09237 |
| 3 | 3 | 23 | 2 | 128.17 | 443.62 | 443.80 | 443.99 | 94.84 | 93.59 | 92.35 | 0.09237 |
| 3 | 3 | 23 | 2 | 128.14 | 444.36 | 444.54 | 444.73 | 89.88 | 88.65 | 87.42 | 0.09237 |
| 3 | 3 | 21 | 2 | 127.23 | 465.73 | 465.90 | 466.07 | 326.02 | 325.19 | 324.35 | 0.08434 |
| 3 | 3 | 21 | 2 | 127.20 | 466.41 | 466.58 | 466.74 | 322.69 | 321.87 | 321.04 | 0.08434 |
| 3 | 3 | 21 | 2 | 127.17 | 467.08 | 467.25 | 467.42 | 319.40 | 318.58 | 317.77 | 0.08434 |
| 3 | 3 | 21 | 2 | 127.14 | 467.76 | 467.93 | 468.10 | 316.14 | 315.33 | 314.53 | 0.08434 |
| 3 | 3 | 24 | 2 | 125.91 | 497.27 | 497.47 | 497.66 | 206.66 | 206.14 | 205.63 | 0.09639 |
| 3 | 3 | 24 | 2 | 125.88 | 498.05 | 498.24 | 498.44 | 204.60 | 204.10 | 203.59 | 0.09639 |
| 3 | 3 | 24 | 2 | 125.85 | 498.82 | 499.02 | 499.21 | 202.59 | 202.10 | 201.60 | 0.09639 |
| 3 | 3 | 24 | 2 | 125.81 | 499.60 | 499.79 | 499.98 | 200.63 | 200.14 | 199.66 | 0.09639 |
| 4 | 3 | 23 | 2 | 124.45 | 532.97 | 533.15 | 533.34 | 151.27 | 151.18 | 151.10 | 0.09312 |
| 4 | 3 | 23 | 2 | 124.42 | 533.71 | 533.90 | 534.08 | 150.93 | 150.85 | 150.77 | 0.09312 |
| 4 | 3 | 23 | 2 | 124.39 | 534.45 | 534.64 | 534.82 | 150.61 | 150.54 | 150.47 | 0.09312 |
| 4 | 3 | 23 | 2 | 124.36 | 535.19 | 535.38 | 535.57 | 150.33 | 150.26 | 150.19 | 0.09312 |
| 4 | 3 | 23 | 2 | 124.33 | 535.94 | 536.12 | 536.31 | 150.07 | 150.01 | 149.95 | 0.09312 |
| 4 | 3 | 22 | 2 | 122.78 | 574.42 | 574.59 | 574.77 | 180.18 | 180.52 | 180.86 | 0.08907 |
| 4 | 3 | 22 | 2 | 122.76 | 575.13 | 575.30 | 575.48 | 181.55 | 181.89 | 182.24 | 0.08907 |
| 4 | 3 | 22 | 2 | 122.73 | 575.84 | 576.01 | 576.19 | 182.95 | 183.30 | 183.66 | 0.08907 |
| 4 | 3 | 22 | 2 | 122.70 | 576.54 | 576.72 | 576.90 | 184.38 | 184.74 | 185.11 | 0.08907 |
| 4 | 3 | 22 | 2 | 122.67 | 577.25 | 577.43 | 577.61 | 185.84 | 186.21 | 186.59 | 0.08907 |
| 4 | 3 | 22 | 2 | 122.64 | 577.96 | 578.14 | 578.32 | 187.34 | 187.72 | 188.10 | 0.08907 |
| 8 | 3 | 22 | 2 | 112.80 | 844.52 | 844.70 | 844.88 | 99.71 | 102.50 | 105.29 | 0.11828 |
| 8 | 3 | 22 | 2 | 112.77 | 845.23 | 845.41 | 845.59 | 110.87 | 113.66 | 116.46 | 0.11828 |
| 8 | 3 | 22 | 2 | 112.75 | 845.94 | 846.12 | 846.30 | 122.05 | 124.85 | 127.65 | 0.11828 |
| 8 | 3 | 21 | 2 | 112.00 | 868.01 | 868.18 | 868.35 | 121.71 | 124.55 | 127.40 | 0.11290 |
| 8 | 3 | 21 | 2 | 111.97 | 868.69 | 868.86 | 869.03 | 133.10 | 135.95 | 138.80 | 0.11290 |
| 8 | 3 | 21 | 2 | 111.95 | 869.37 | 869.54 | 869.71 | 144.51 | 147.37 | 150.23 | 0.11290 |
| 8 | 3 | 19 | 2 | 111.26 | 890.11 | 890.26 | 890.41 | 143.94 | 146.67 | 149.39 | 0.10215 |
| 8 | 3 | 19 | 2 | 111.24 | 890.72 | 890.87 | 891.03 | 154.86 | 157.59 | 160.32 | 0.10215 |
| 8 | 3 | 19 | 2 | 111.21 | 891.33 | 891.49 | 891.64 | 165.79 | 168.52 | 171.26 | 0.10215 |
| 9 | 3 | 19 | 2 | 107.23 | 965.21 | 965.37 | 965.52 | 272.84 | 276.74 | 280.66 | 0.16964 |
| 9 | 3 | 19 | 2 | 107.19 | 965.83 | 965.98 | 966.13 | 288.50 | 292.43 | 296.37 | 0.16964 |
| 9 | 3 | 22 | 2 | 104.92 | 997.07 | 997.24 | 997.42 | 105.79 | 111.45 | 117.11 | 0.19643 |
| 9 | 3 | 15 | 2 | 104.11 | 1008.09 | 1008.21 | 1008.33 | 105.65 | 109.69 | 113.72 | 0.13393 |
| 9 | 3 | 15 | 2 | 104.07 | 1008.57 | 1008.70 | 1008.82 | 121.79 | 125.83 | 129.87 | 0.13393 |
| 9 | 3 | 22 | 2 | 103.17 | 1020.75 | 1020.93 | 1021.11 | 175.89 | 182.03 | 188.17 | 0.19643 |
| 10 | 3 | 19 | 2 | 100.75 | 1054.36 | 1054.51 | 1054.66 | 305.97 | 311.72 | 317.47 | 0.16522 |
| 10 | 3 | 18 | 2 | 100.02 | 1064.62 | 1064.77 | 1064.91 | 336.19 | 341.79 | 347.39 | 0.15652 |
| 10 | 3 | 17 | 2 | 99.33 | 1074.55 | 1074.69 | 1074.82 | 4.59 | 10.04 | 15.49 | 0.14783 |
| 10 | 3 | 15 | 2 | 98.56 | 1085.59 | 1085.71 | 1085.83 | 90.76 | 95.74 | 100.72 | 0.13043 |
| 10 | 3 | 16 | 2 | 97.93 | 1094.73 | 1094.86 | 1094.99 | 113.10 | 118.58 | 124.05 | 0.13913 |
| 10 | 3 | 15 | 2 | 97.32 | 1103.72 | 1103.85 | 1103.97 | 140.65 | 145.94 | 151.24 | 0.13043 |
| 10 | 3 | 15 | 2 | 96.74 | 1112.45 | 1112.57 | 1112.69 | 167.46 | 172.90 | 178.33 | 0.13043 |
| 11 | 3 | 14 | 2 | 94.57 | 1145.48 | 1145.59 | 1145.71 | 281.74 | 287.28 | 292.84 | 0.13861 |
| 11 | 3 | 13 | 2 | 94.07 | 1153.32 | 1153.42 | 1153.53 | 310.92 | 316.19 | 321.45 | 0.12871 |
| 11 | 3 | 13 | 2 | 93.42 | 1163.56 | 1163.67 | 1163.77 | 111.11 | 116.50 | 121.89 | 0.12871 |
| 11 | 3 | 13 | 2 | 92.95 | 1171.02 | 1171.13 | 1171.23 | 137.87 | 143.35 | 148.84 | 0.12871 |
| 11 | 3 | 12 | 2 | 92.49 | 1178.39 | 1178.49 | 1178.58 | 166.67 | 171.82 | 176.97 | 0.11881 |
| 11 | 3 | 12 | 2 | 92.05 | 1185.57 | 1185.66 | 1185.76 | 192.01 | 197.26 | 202.49 | 0.11881 |
| 11 | 3 | 12 | 2 | 91.62 | 1192.60 | 1192.70 | 1192.79 | 216.40 | 221.74 | 227.07 | 0.11881 |
| 11 | 3 | 12 | 2 | 91.20 | 1199.49 | 1199.58 | 1199.68 | 239.35 | 244.78 | 250.20 | 0.11881 |
| 12 | 3 | 10 | 2 | 89.12 | 1234.48 | 1234.56 | 1234.64 | 98.17 | 102.90 | 107.62 | 0.06024 |
| 12 | 3 | 21 | 2 | 88.76 | 1240.57 | 1240.74 | 1240.91 | 94.67 | 104.59 | 114.49 | 0.12651 |
| 12 | 3 | 10 | 2 | 88.41 | 1246.70 | 1246.79 | 1246.87 | 93.22 | 97.93 | 102.63 | 0.06024 |
| 12 | 3 | 10 | 2 | 88.06 | 1252.81 | 1252.89 | 1252.97 | 88.40 | 93.08 | 97.76 | 0.06024 |
| 12 | 3 | 21 | 2 | 87.71 | 1258.94 | 1259.11 | 1259.28 | 83.13 | 92.91 | 102.68 | 0.12651 |
| 12 | 3 | 21 | 2 | 87.36 | 1264.94 | 1265.11 | 1265.28 | 68.67 | 78.40 | 88.13 | 0.12651 |
| 12 | 3 | 21 | 2 | 87.03 | 1270.87 | 1271.04 | 1271.21 | 49.05 | 58.75 | 68.44 | 0.12651 |
| 12 | 3 | 21 | 2 | 86.70 | 1276.74 | 1276.91 | 1277.08 | 24.62 | 34.29 | 43.95 | 0.12651 |
| 12 | 3 | 21 | 2 | 86.38 | 1282.60 | 1282.77 | 1282.94 | 358.34 | 7.98 | 17.62 | 0.12651 |
| 12 | 3 | 10 | 2 | 85.95 | 1290.53 | 1290.61 | 1290.69 | 89.15 | 93.73 | 98.29 | 0.06024 |
| 13 | 3 | 20 | 2 | 84.38 | 1319.36 | 1319.52 | 1319.68 | 277.21 | 286.29 | 295.38 | 0.09756 |
| 13 | 3 | 22 | 2 | 83.87 | 1329.05 | 1329.22 | 1329.40 | 101.91 | 111.85 | 121.81 | 0.10732 |
| 13 | 3 | 20 | 2 | 83.54 | 1335.36 | 1335.52 | 1335.68 | 95.27 | 104.27 | 113.26 | 0.09756 |
| 13 | 3 | 20 | 2 | 83.24 | 1341.07 | 1341.23 | 1341.39 | 52.83 | 61.77 | 70.70 | 0.09756 |
| 13 | 3 | 20 | 2 | 82.95 | 1346.79 | 1346.95 | 1347.11 | 8.37 | 17.22 | 26.07 | 0.09756 |
| 13 | 3 | 20 | 2 | 82.66 | 1352.50 | 1352.66 | 1352.82 | 320.29 | 329.05 | 337.80 | 0.09756 |
| 13 | 3 | 20 | 2 | 82.36 | 1358.25 | 1358.42 | 1358.58 | 270.97 | 279.60 | 288.24 | 0.09756 |
| 13 | 3 | 21 | 2 | 82.07 | 1364.02 | 1364.19 | 1364.36 | 217.54 | 226.47 | 235.40 | 0.10244 |
| 13 | 3 | 21 | 2 | 81.78 | 1369.81 | 1369.98 | 1370.15 | 160.44 | 169.22 | 178.00 | 0.10244 |
| 13 | 3 | 21 | 2 | 81.49 | 1375.63 | 1375.80 | 1375.97 | 99.84 | 108.48 | 117.11 | 0.10244 |
| 14 | 3 | 12 | 2 | 79.70 | 1412.14 | 1412.24 | 1412.33 | 88.85 | 93.47 | 98.10 | 0.05911 |
| 14 | 3 | 11 | 2 | 79.39 | 1418.57 | 1418.65 | 1418.74 | 35.56 | 39.79 | 44.02 | 0.05419 |
| 14 | 3 | 22 | 2 | 79.11 | 1424.43 | 1424.61 | 1424.78 | 314.62 | 323.06 | 331.49 | 0.10837 |
| 14 | 3 | 22 | 2 | 78.81 | 1430.62 | 1430.80 | 1430.98 | 248.75 | 257.16 | 265.57 | 0.10837 |
| 14 | 3 | 22 | 2 | 78.52 | 1436.85 | 1437.03 | 1437.21 | 183.77 | 192.15 | 200.53 | 0.10837 |
| 14 | 3 | 22 | 2 | 78.23 | 1443.12 | 1443.30 | 1443.48 | 118.83 | 127.15 | 135.48 | 0.10837 |
| 14 | 3 | 23 | 2 | 77.94 | 1449.43 | 1449.62 | 1449.81 | 54.00 | 62.63 | 71.27 | 0.11330 |
| 14 | 3 | 23 | 2 | 77.64 | 1455.81 | 1455.99 | 1456.18 | 349.00 | 357.55 | 6.09 | 0.11330 |
| 14 | 3 | 23 | 2 | 77.35 | 1462.26 | 1462.45 | 1462.64 | 284.74 | 293.17 | 301.61 | 0.11330 |
| 14 | 3 | 23 | 2 | 77.06 | 1468.76 | 1468.95 | 1469.14 | 218.58 | 226.90 | 235.23 | 0.11330 |
| 15 | 3 | 16 | 2 | 74.22 | 1534.04 | 1534.17 | 1534.30 | 103.17 | 108.39 | 113.62 | 0.15385 |
| 15 | 3 | 10 | 2 | 73.89 | 1541.85 | 1541.93 | 1542.01 | 59.32 | 62.59 | 65.85 | 0.09615 |
| 15 | 3 | 13 | 2 | 73.61 | 1548.71 | 1548.81 | 1548.92 | 334.75 | 338.92 | 343.09 | 0.12500 |
| 15 | 3 | 13 | 2 | 73.30 | 1556.12 | 1556.22 | 1556.33 | 266.17 | 270.24 | 274.31 | 0.12500 |
| 15 | 3 | 13 | 2 | 73.00 | 1563.60 | 1563.71 | 1563.81 | 191.33 | 195.25 | 199.17 | 0.12500 |
| 15 | 3 | 13 | 2 | 72.69 | 1571.17 | 1571.28 | 1571.38 | 110.49 | 114.32 | 118.16 | 0.12500 |
| 15 | 3 | 13 | 2 | 72.38 | 1578.85 | 1578.96 | 1579.06 | 28.75 | 32.48 | 36.22 | 0.12500 |
| 15 | 3 | 13 | 2 | 72.07 | 1586.63 | 1586.74 | 1586.84 | 304.29 | 307.98 | 311.66 | 0.12500 |
| 16 | 3 | 15 | 2 | 70.41 | 1629.61 | 1629.73 | 1629.85 | 15.12 | 19.35 | 23.59 | 0.14151 |
| 16 | 3 | 15 | 2 | 70.10 | 1638.10 | 1638.22 | 1638.34 | 311.86 | 316.03 | 320.21 | 0.14151 |
| 16 | 3 | 15 | 2 | 69.78 | 1646.66 | 1646.78 | 1646.91 | 246.52 | 250.71 | 254.89 | 0.14151 |
| 16 | 3 | 15 | 2 | 69.47 | 1655.35 | 1655.47 | 1655.59 | 188.11 | 192.30 | 196.49 | 0.14151 |
| 16 | 3 | 15 | 2 | 69.15 | 1664.04 | 1664.16 | 1664.28 | 115.28 | 119.19 | 123.10 | 0.14151 |
| 16 | 3 | 15 | 2 | 68.84 | 1672.85 | 1672.97 | 1673.09 | 40.61 | 44.53 | 48.45 | 0.14151 |
| 16 | 3 | 16 | 2 | 68.52 | 1681.83 | 1681.96 | 1682.09 | 332.88 | 337.09 | 341.30 | 0.15094 |
| 17 | 3 | 16 | 2 | 67.25 | 1718.96 | 1719.09 | 1719.22 | 105.49 | 109.61 | 113.72 | 0.13445 |
| 17 | 3 | 17 | 2 | 66.93 | 1728.31 | 1728.45 | 1728.59 | 36.30 | 40.44 | 44.57 | 0.14286 |
| 17 | 3 | 17 | 2 | 66.62 | 1737.81 | 1737.95 | 1738.09 | 323.45 | 327.67 | 331.91 | 0.14286 |
| 17 | 3 | 17 | 2 | 66.31 | 1747.34 | 1747.48 | 1747.62 | 258.60 | 262.79 | 266.98 | 0.14286 |
| 17 | 3 | 17 | 2 | 66.01 | 1757.05 | 1757.18 | 1757.32 | 185.68 | 189.68 | 193.68 | 0.14286 |
| 17 | 3 | 17 | 2 | 65.70 | 1766.88 | 1767.01 | 1767.15 | 110.34 | 114.26 | 118.19 | 0.14286 |
| 17 | 3 | 18 | 2 | 65.39 | 1776.80 | 1776.95 | 1777.09 | 32.02 | 36.15 | 40.28 | 0.15126 |
| 18 | 3 | 18 | 2 | 64.41 | 1809.15 | 1809.29 | 1809.44 | 245.83 | 249.97 | 254.11 | 0.15929 |
| 18 | 3 | 18 | 2 | 64.11 | 1819.50 | 1819.64 | 1819.79 | 176.76 | 180.85 | 184.94 | 0.15929 |
| 18 | 3 | 19 | 2 | 63.81 | 1830.03 | 1830.18 | 1830.33 | 116.27 | 120.63 | 124.98 | 0.16814 |
| 18 | 3 | 19 | 2 | 63.50 | 1840.72 | 1840.87 | 1841.03 | 58.39 | 62.69 | 66.99 | 0.16814 |
| 18 | 3 | 19 | 2 | 63.20 | 1851.61 | 1851.77 | 1851.92 | 358.94 | 3.07 | 7.19 | 0.16814 |
| 18 | 3 | 20 | 2 | 62.89 | 1862.74 | 1862.90 | 1863.06 | 303.74 | 308.21 | 312.68 | 0.17699 |
| 22 | 3 | 22 | 2 | 59.01 | 2016.97 | 2017.15 | 2017.33 | 95.08 | 99.48 | 103.89 | 0.19298 |
| 22 | 3 | 23 | 2 | 58.71 | 2029.46 | 2029.64 | 2029.83 | 44.90 | 49.54 | 54.19 | 0.20175 |
| 22 | 3 | 23 | 2 | 58.43 | 2041.96 | 2042.14 | 2042.33 | 2.10 | 6.89 | 11.67 | 0.20175 |
| 22 | 3 | 23 | 2 | 58.14 | 2054.46 | 2054.64 | 2054.83 | 316.35 | 321.08 | 325.81 | 0.20175 |
| 22 | 3 | 23 | 2 | 57.86 | 2066.99 | 2067.18 | 2067.37 | 276.00 | 280.67 | 285.34 | 0.20175 |
| 23 | 3 | 23 | 2 | 57.01 | 2105.28 | 2105.47 | 2105.65 | 122.62 | 126.96 | 131.29 | 0.19658 |
| 23 | 3 | 23 | 2 | 56.72 | 2118.73 | 2118.92 | 2119.10 | 74.78 | 79.11 | 83.45 | 0.19658 |
| 23 | 3 | 23 | 2 | 56.44 | 2131.51 | 2131.69 | 2131.88 | 14.80 | 19.11 | 23.42 | 0.19658 |
| 23 | 3 | 24 | 2 | 56.10 | 2147.45 | 2147.64 | 2147.84 | 24.89 | 29.33 | 33.77 | 0.20513 |
| 23 | 3 | 24 | 2 | 55.83 | 2160.49 | 2160.68 | 2160.87 | 311.22 | 315.38 | 319.55 | 0.20513 |
| 24 | 3 | 26 | 2 | 54.18 | 2241.60 | 2241.81 | 2242.02 | 162.48 | 166.46 | 170.45 | 0.19403 |
| 24 | 3 | 26 | 2 | 53.90 | 2255.88 | 2256.09 | 2256.30 | 69.21 | 73.11 | 77.00 | 0.19403 |
| 24 | 3 | 27 | 2 | 53.61 | 2270.51 | 2270.73 | 2270.95 | 343.19 | 347.24 | 351.31 | 0.20149 |
| 24 | 3 | 27 | 2 | 53.32 | 2285.49 | 2285.70 | 2285.92 | 267.51 | 271.59 | 275.67 | 0.20149 |
| 24 | 3 | 28 | 2 | 52.85 | 2310.07 | 2310.30 | 2310.52 | 350.92 | 354.79 | 358.64 | 0.20896 |
| 25 | 3 | 29 | 2 | 52.46 | 2330.77 | 2331.00 | 2331.24 | 345.61 | 349.76 | 353.92 | 0.23770 |
| 25 | 3 | 30 | 2 | 52.16 | 2346.90 | 2347.14 | 2347.38 | 266.61 | 270.71 | 274.81 | 0.24590 |
| 25 | 3 | 31 | 2 | 51.86 | 2363.36 | 2363.61 | 2363.86 | 179.00 | 183.04 | 187.08 | 0.25410 |
| 25 | 3 | 32 | 2 | 51.55 | 2380.26 | 2380.51 | 2380.77 | 85.47 | 89.45 | 93.44 | 0.26230 |
| 26 | 3 | 31 | 2 | 50.93 | 2414.56 | 2414.81 | 2415.06 | 263.08 | 267.07 | 271.07 | 0.24031 |
| 26 | 3 | 32 | 2 | 50.62 | 2432.17 | 2432.42 | 2432.68 | 185.06 | 189.21 | 193.35 | 0.24806 |
| 26 | 3 | 33 | 2 | 50.31 | 2449.79 | 2450.06 | 2450.33 | 105.93 | 110.06 | 114.20 | 0.25581 |
| 26 | 3 | 33 | 2 | 50.01 | 2467.44 | 2467.71 | 2467.97 | 11.47 | 15.34 | 19.20 | 0.25581 |
| 30 | 3 | 65 | 2 | 47.37 | 2626.68 | 2627.21 | 2627.73 | 348.93 | 357.30 | 5.68 | 0.33505 |
| 30 | 3 | 65 | 2 | 47.10 | 2643.83 | 2644.35 | 2644.88 | 258.83 | 266.92 | 275.02 | 0.33505 |
| 30 | 3 | 32 | 2 | 46.83 | 2661.03 | 2661.29 | 2661.55 | 161.00 | 164.89 | 168.78 | 0.16495 |
| 30 | 3 | 32 | 2 | 46.56 | 2678.57 | 2678.82 | 2679.08 | 59.99 | 63.75 | 67.51 | 0.16495 |
| 31 | 3 | 67 | 2 | 45.89 | 2722.00 | 2722.54 | 2723.08 | 23.02 | 31.79 | 40.54 | 0.42675 |
| 31 | 3 | 81 | 2 | 45.50 | 2747.99 | 2748.65 | 2749.30 | 74.12 | 84.46 | 94.82 | 0.51592 |
| 31 | 3 | 9 | 2 | 45.24 | 2766.07 | 2766.15 | 2766.22 | 6.29 | 7.48 | 8.69 | 0.05732 |
| 32 | 3 | 63 | 2 | 44.74 | 2799.09 | 2799.60 | 2800.11 | 207.17 | 216.02 | 224.86 | 0.25200 |
| 32 | 3 | 63 | 2 | 44.50 | 2815.75 | 2816.26 | 2816.77 | 131.87 | 140.21 | 148.55 | 0.25200 |
| 32 | 3 | 61 | 2 | 44.27 | 2832.03 | 2832.52 | 2833.01 | 40.80 | 48.99 | 57.19 | 0.24400 |
| 32 | 3 | 63 | 2 | 44.03 | 2848.76 | 2849.26 | 2849.77 | 315.99 | 324.11 | 332.22 | 0.25200 |
| 33 | 3 | 94 | 2 | 43.23 | 2904.77 | 2905.53 | 2906.29 | 110.45 | 122.51 | 134.60 | 0.50811 |
| 33 | 3 | 26 | 2 | 42.94 | 2926.51 | 2926.72 | 2926.93 | 109.14 | 112.68 | 116.22 | 0.14054 |
| 33 | 3 | 65 | 2 | 42.71 | 2943.08 | 2943.61 | 2944.13 | 17.96 | 26.00 | 34.04 | 0.35135 |
| 34 | 3 | 66 | 2 | 42.23 | 2978.29 | 2978.83 | 2979.36 | 180.71 | 188.56 | 196.41 | 0.25781 |
| 34 | 3 | 66 | 2 | 41.99 | 2996.24 | 2996.77 | 2997.30 | 102.77 | 111.52 | 120.32 | 0.25781 |
| 34 | 3 | 65 | 2 | 41.75 | 3013.71 | 3014.23 | 3014.76 | 31.61 | 40.17 | 48.74 | 0.25391 |
| 34 | 3 | 59 | 2 | 41.50 | 3033.00 | 3033.47 | 3033.95 | 341.81 | 349.33 | 356.84 | 0.23047 |
| 35 | 3 | 62 | 2 | 40.37 | 3120.06 | 3120.56 | 3121.06 | 239.02 | 246.32 | 253.64 | 0.24900 |
| 35 | 3 | 62 | 2 | 40.16 | 3136.52 | 3137.02 | 3137.52 | 120.70 | 128.12 | 135.57 | 0.24900 |
| 35 | 3 | 62 | 2 | 39.95 | 3153.09 | 3153.59 | 3154.09 | 349.91 | 356.18 | 2.44 | 0.24900 |
| 35 | 3 | 63 | 2 | 39.74 | 3169.75 | 3170.25 | 3170.76 | 213.74 | 220.90 | 228.03 | 0.25301 |
| 36 | 3 | 64 | 2 | 39.12 | 3220.36 | 3220.88 | 3221.40 | 191.17 | 198.35 | 205.53 | 0.24806 |
| 36 | 3 | 64 | 2 | 38.91 | 3237.45 | 3237.97 | 3238.48 | 69.81 | 76.94 | 84.07 | 0.24806 |
| 36 | 3 | 65 | 2 | 38.70 | 3254.63 | 3255.16 | 3255.68 | 300.78 | 307.31 | 313.80 | 0.25194 |
| 36 | 3 | 65 | 2 | 38.49 | 3271.93 | 3272.46 | 3272.98 | 155.31 | 161.93 | 168.56 | 0.25194 |
| 37 | 3 | 66 | 2 | 38.08 | 3306.83 | 3307.37 | 3307.90 | 298.25 | 305.60 | 312.90 | 0.24719 |
| 37 | 3 | 67 | 2 | 37.87 | 3324.45 | 3324.99 | 3325.53 | 173.99 | 181.10 | 188.21 | 0.25094 |
| 37 | 3 | 67 | 2 | 37.66 | 3342.17 | 3342.71 | 3343.25 | 51.38 | 58.79 | 66.21 | 0.25094 |
| 37 | 3 | 67 | 2 | 37.45 | 3360.00 | 3360.54 | 3361.08 | 304.93 | 313.13 | 321.35 | 0.25094 |
| 38 | 3 | 68 | 2 | 37.14 | 3386.65 | 3387.19 | 3387.74 | 356.52 | 5.04 | 13.57 | 0.24818 |
| 38 | 3 | 68 | 2 | 36.94 | 3404.76 | 3405.31 | 3405.85 | 274.42 | 282.44 | 290.43 | 0.24818 |
| 38 | 3 | 69 | 2 | 36.73 | 3422.99 | 3423.55 | 3424.10 | 167.87 | 175.22 | 182.55 | 0.25182 |
| 38 | 3 | 69 | 2 | 36.52 | 3441.33 | 3441.88 | 3442.44 | 50.58 | 57.92 | 65.24 | 0.25182 |
| 39 | 3 | 70 | 2 | 36.10 | 3478.34 | 3478.90 | 3479.47 | 190.54 | 198.72 | 206.92 | 0.24823 |
| 39 | 3 | 70 | 2 | 35.91 | 3495.35 | 3495.92 | 3496.48 | 83.10 | 91.60 | 100.15 | 0.24823 |
| 39 | 3 | 71 | 2 | 35.71 | 3514.13 | 3514.70 | 3515.28 | 4.20 | 12.41 | 20.60 | 0.25177 |
| 39 | 3 | 71 | 2 | 35.45 | 3537.15 | 3537.73 | 3538.30 | 307.88 | 314.87 | 321.86 | 0.25177 |
| 40 | 3 | 72 | 2 | 35.04 | 3575.37 | 3575.95 | 3576.53 | 81.19 | 88.71 | 96.20 | 0.24658 |
| 40 | 3 | 73 | 2 | 34.83 | 3594.66 | 3595.25 | 3595.84 | 338.95 | 347.49 | 356.07 | 0.25000 |
| 40 | 3 | 73 | 2 | 34.62 | 3614.07 | 3614.66 | 3615.25 | 268.91 | 277.44 | 285.92 | 0.25000 |
| 40 | 3 | 74 | 2 | 34.41 | 3633.61 | 3634.21 | 3634.80 | 192.41 | 201.40 | 210.36 | 0.25342 |
| 41 | 3 | 75 | 2 | 34.04 | 3668.95 | 3669.56 | 3670.16 | 341.17 | 350.02 | 358.86 | 0.24834 |
| 41 | 3 | 75 | 2 | 33.83 | 3688.83 | 3689.44 | 3690.04 | 256.86 | 265.26 | 273.71 | 0.24834 |
| 41 | 3 | 76 | 2 | 33.63 | 3708.82 | 3709.44 | 3710.05 | 179.81 | 188.49 | 197.17 | 0.25166 |
| 41 | 3 | 76 | 2 | 33.43 | 3728.14 | 3728.76 | 3729.37 | 89.44 | 98.11 | 106.80 | 0.25166 |
| 42 | 3 | 78 | 2 | 32.59 | 3809.89 | 3810.52 | 3811.15 | 115.07 | 122.05 | 128.98 | 0.33051 |
| 42 | 3 | 79 | 2 | 32.39 | 3830.65 | 3831.29 | 3831.93 | 1.21 | 9.78 | 18.37 | 0.33475 |
| 42 | 3 | 79 | 2 | 32.18 | 3851.55 | 3852.18 | 3852.82 | 270.80 | 278.93 | 287.05 | 0.33475 |
| 43 | 3 | 80 | 2 | 31.76 | 3893.71 | 3894.35 | 3895.00 | 19.42 | 25.35 | 31.32 | 0.33058 |
| 43 | 3 | 81 | 2 | 31.55 | 3914.98 | 3915.63 | 3916.29 | 259.56 | 268.57 | 277.80 | 0.33471 |
| 43 | 3 | 81 | 2 | 31.35 | 3936.39 | 3937.05 | 3937.70 | 222.31 | 231.51 | 240.69 | 0.33471 |
| 44 | 3 | 82 | 2 | 30.93 | 3979.62 | 3980.28 | 3980.94 | 133.90 | 143.75 | 153.59 | 0.33065 |
| 44 | 3 | 83 | 2 | 30.73 | 3999.91 | 4000.58 | 4001.25 | 106.44 | 118.22 | 130.06 | 0.33468 |
| 44 | 3 | 83 | 2 | 30.53 | 4021.86 | 4022.52 | 4023.19 | 125.01 | 135.90 | 146.78 | 0.33468 |
| 45 | 3 | 84 | 2 | 30.11 | 4066.15 | 4066.83 | 4067.51 | 56.69 | 65.92 | 75.17 | 0.33071 |
| 45 | 3 | 85 | 2 | 29.90 | 4088.49 | 4089.17 | 4089.86 | 4.56 | 14.14 | 23.70 | 0.33465 |
| 45 | 3 | 85 | 2 | 29.69 | 4110.99 | 4111.67 | 4112.36 | 308.65 | 317.74 | 326.82 | 0.33465 |
| 46 | 3 | 86 | 2 | 29.28 | 4156.41 | 4157.11 | 4157.80 | 124.00 | 130.51 | 136.96 | 0.32950 |
| 46 | 3 | 87 | 2 | 29.07 | 4179.31 | 4180.02 | 4180.72 | 356.50 | 4.69 | 12.89 | 0.33333 |
| 46 | 3 | 88 | 2 | 28.86 | 4202.38 | 4203.09 | 4203.80 | 278.93 | 287.79 | 296.65 | 0.33716 |
| 47 | 3 | 89 | 2 | 28.44 | 4248.94 | 4249.66 | 4250.38 | 148.43 | 157.24 | 166.02 | 0.33209 |
| 47 | 3 | 89 | 2 | 28.23 | 4272.46 | 4273.17 | 4273.89 | 63.47 | 71.80 | 80.16 | 0.33209 |
| 47 | 3 | 90 | 2 | 28.02 | 4296.09 | 4296.82 | 4297.54 | 344.39 | 353.14 | 1.83 | 0.33582 |
| 48 | 3 | 91 | 2 | 27.60 | 4343.83 | 4344.56 | 4345.29 | 102.87 | 110.85 | 118.87 | 0.33212 |
| 48 | 3 | 91 | 2 | 27.39 | 4367.92 | 4368.65 | 4369.39 | 3.78 | 11.43 | 19.05 | 0.33212 |
| 48 | 3 | 92 | 2 | 27.18 | 4392.16 | 4392.90 | 4393.65 | 249.31 | 257.00 | 264.74 | 0.33577 |
| 50 | 3 | 128 | 2 | 23.34 | 4862.15 | 4863.18 | 4864.21 | 212.80 | 223.86 | 234.95 | 0.49805 |
| 50 | 3 | 129 | 2 | 23.08 | 4895.76 | 4896.80 | 4897.84 | 205.74 | 215.94 | 226.20 | 0.50195 |
| 51 | 3 | 132 | 2 | 20.77 | 5204.62 | 5205.69 | 5206.75 | 133.95 | 145.02 | 156.06 | 0.49811 |
| 51 | 3 | 133 | 2 | 20.52 | 5239.47 | 5240.54 | 5241.61 | 113.63 | 123.20 | 132.69 | 0.50189 |
| 54 | 3 | 131 | 2 | 17.10 | 5732.54 | 5733.59 | 5734.65 | 263.02 | 272.99 | 282.96 | 0.51172 |
| 54 | 3 | 125 | 2 | 16.88 | 5765.64 | 5766.65 | 5767.65 | 213.00 | 222.36 | 231.72 | 0.48828 |
| 55 | 3 | 117 | 2 | 14.90 | 6074.35 | 6075.29 | 6076.23 | 259.62 | 268.24 | 276.89 | 0.49787 |
| 55 | 3 | 118 | 2 | 14.70 | 6105.34 | 6106.30 | 6107.25 | 188.44 | 197.40 | 206.44 | 0.50213 |
| 56 | 3 | 124 | 2 | 12.75 | 6424.50 | 6425.50 | 6426.50 | 235.72 | 244.92 | 254.06 | 0.49799 |
| 56 | 3 | 125 | 2 | 12.55 | 6457.27 | 6458.28 | 6459.28 | 166.42 | 174.95 | 183.42 | 0.50201 |
| 57 | 3 | 131 | 2 | 10.78 | 6760.10 | 6761.16 | 6762.21 | 125.26 | 134.23 | 143.20 | 0.50000 |
| 57 | 3 | 131 | 2 | 10.59 | 6794.45 | 6795.51 | 6796.56 | 43.18 | 51.25 | 59.30 | 0.50000 |
| 58 | 3 | 136 | 2 | 9.01 | 7075.32 | 7076.42 | 7077.51 | 165.63 | 173.70 | 181.78 | 0.50000 |
| 58 | 3 | 136 | 2 | 8.81 | 7111.18 | 7112.28 | 7113.37 | 65.64 | 74.10 | 82.66 | 0.50000 |
| 59 | 3 | 141 | 2 | 7.24 | 7403.24 | 7404.38 | 7405.52 | 84.79 | 92.93 | 101.05 | 0.49823 |
| 59 | 3 | 142 | 2 | 7.04 | 7440.38 | 7441.52 | 7442.67 | 354.32 | 2.88 | 11.43 | 0.50177 |
| 60 | 3 | 146 | 2 | 5.46 | 7742.90 | 7744.08 | 7745.25 | 75.22 | 83.19 | 91.10 | 0.50000 |
| 60 | 3 | 146 | 2 | 5.27 | 7781.21 | 7782.39 | 7783.57 | 332.15 | 340.18 | 348.21 | 0.50000 |
| 1 | 4 | 1 | 10 | 142.37 | 143.32 | 143.36 | 143.40 | 210.36 | 209.16 | 207.97 | 0.03030 |
| 1 | 4 | 1 | 10 | 142.36 | 143.48 | 143.52 | 143.56 | 205.57 | 204.38 | 203.18 | 0.03030 |
| 1 | 4 | 1 | 10 | 142.35 | 143.64 | 143.68 | 143.72 | 200.79 | 199.59 | 198.40 | 0.03030 |
| 1 | 4 | 1 | 10 | 142.35 | 143.80 | 143.84 | 143.88 | 196.01 | 194.81 | 193.61 | 0.03030 |
| 1 | 4 | 1 | 12 | 141.90 | 152.40 | 152.44 | 152.49 | 303.89 | 302.49 | 301.08 | 0.03030 |
| 1 | 4 | 1 | 12 | 141.89 | 152.59 | 152.64 | 152.69 | 298.28 | 296.88 | 295.47 | 0.03030 |
| 1 | 4 | 1 | 12 | 141.88 | 152.78 | 152.83 | 152.88 | 292.67 | 291.27 | 289.87 | 0.03030 |
| 1 | 4 | 1 | 12 | 141.87 | 152.98 | 153.03 | 153.07 | 287.06 | 285.66 | 284.26 | 0.03030 |
| 1 | 4 | 1 | 10 | 141.41 | 161.83 | 161.87 | 161.91 | 33.44 | 32.29 | 31.15 | 0.03030 |
| 1 | 4 | 1 | 10 | 141.40 | 161.99 | 162.03 | 162.07 | 28.87 | 27.73 | 26.59 | 0.03030 |
| 1 | 4 | 1 | 10 | 141.39 | 162.15 | 162.19 | 162.23 | 24.30 | 23.16 | 22.02 | 0.03030 |
| 1 | 4 | 1 | 10 | 141.39 | 162.32 | 162.36 | 162.40 | 19.74 | 18.60 | 17.45 | 0.03030 |
| 1 | 4 | 1 | 10 | 141.38 | 162.48 | 162.52 | 162.56 | 15.17 | 14.03 | 12.89 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.91 | 171.61 | 171.65 | 171.69 | 119.57 | 118.45 | 117.34 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.90 | 171.77 | 171.81 | 171.85 | 115.11 | 113.99 | 112.88 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.89 | 171.94 | 171.98 | 172.02 | 110.65 | 109.54 | 108.42 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.88 | 172.10 | 172.14 | 172.18 | 106.19 | 105.08 | 103.97 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.88 | 172.26 | 172.30 | 172.34 | 101.74 | 100.62 | 99.51 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.39 | 181.68 | 181.73 | 181.77 | 204.33 | 203.25 | 202.16 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.39 | 181.85 | 181.89 | 181.93 | 199.99 | 198.90 | 197.82 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.38 | 182.01 | 182.05 | 182.09 | 195.64 | 194.55 | 193.47 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.37 | 182.17 | 182.21 | 182.25 | 191.29 | 190.21 | 189.12 | 0.03030 |
| 1 | 4 | 1 | 10 | 140.36 | 182.33 | 182.37 | 182.41 | 186.95 | 185.86 | 184.77 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.87 | 192.00 | 192.04 | 192.08 | 289.65 | 288.60 | 287.54 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.86 | 192.16 | 192.20 | 192.24 | 285.42 | 284.36 | 283.30 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.86 | 192.32 | 192.36 | 192.40 | 281.19 | 280.13 | 279.07 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.85 | 192.48 | 192.52 | 192.56 | 276.96 | 275.90 | 274.84 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.84 | 192.64 | 192.68 | 192.72 | 272.73 | 271.67 | 270.61 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.33 | 202.69 | 202.73 | 202.77 | 13.13 | 12.10 | 11.08 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.33 | 202.85 | 202.89 | 202.93 | 9.03 | 8.00 | 6.98 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.32 | 203.01 | 203.05 | 203.09 | 4.93 | 3.90 | 2.88 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.31 | 203.18 | 203.22 | 203.26 | 0.83 | 359.81 | 358.78 | 0.03030 |
| 1 | 4 | 1 | 10 | 139.30 | 203.34 | 203.38 | 203.42 | 356.73 | 355.71 | 354.69 | 0.03030 |
| 2 | 4 | 16 | 2 | 135.37 | 283.94 | 284.07 | 284.20 | 340.86 | 338.29 | 335.73 | 0.12800 |
| 2 | 4 | 16 | 2 | 135.35 | 284.46 | 284.58 | 284.71 | 330.60 | 328.04 | 325.48 | 0.12800 |
| 2 | 4 | 19 | 2 | 134.15 | 309.87 | 310.02 | 310.18 | 214.33 | 211.65 | 208.96 | 0.15200 |
| 2 | 4 | 19 | 2 | 134.12 | 310.48 | 310.64 | 310.79 | 203.60 | 200.93 | 198.25 | 0.15200 |
| 2 | 4 | 19 | 2 | 134.09 | 311.10 | 311.25 | 311.40 | 192.91 | 190.25 | 187.58 | 0.15200 |
| 2 | 4 | 18 | 2 | 133.26 | 329.14 | 329.29 | 329.43 | 255.21 | 252.95 | 250.69 | 0.14400 |
| 2 | 4 | 18 | 2 | 133.23 | 329.72 | 329.87 | 330.02 | 246.18 | 243.93 | 241.68 | 0.14400 |
| 3 | 4 | 23 | 2 | 128.07 | 446.00 | 446.18 | 446.37 | 79.12 | 77.91 | 76.71 | 0.09237 |
| 3 | 4 | 23 | 2 | 128.03 | 446.74 | 446.92 | 447.11 | 74.30 | 73.11 | 71.91 | 0.09237 |
| 3 | 4 | 23 | 2 | 128.00 | 447.48 | 447.67 | 447.85 | 69.53 | 68.34 | 67.16 | 0.09237 |
| 3 | 4 | 21 | 2 | 127.08 | 469.29 | 469.46 | 469.63 | 308.89 | 308.10 | 307.31 | 0.08434 |
| 3 | 4 | 21 | 2 | 127.05 | 469.97 | 470.14 | 470.31 | 305.74 | 304.96 | 304.18 | 0.08434 |
| 3 | 4 | 21 | 2 | 127.02 | 470.65 | 470.82 | 470.99 | 302.63 | 301.86 | 301.09 | 0.08434 |
| 3 | 4 | 21 | 2 | 126.99 | 471.32 | 471.49 | 471.66 | 299.56 | 298.80 | 298.04 | 0.08434 |
| 3 | 4 | 24 | 2 | 125.74 | 501.44 | 501.64 | 501.83 | 196.09 | 195.62 | 195.16 | 0.09639 |
| 3 | 4 | 24 | 2 | 125.70 | 502.22 | 502.41 | 502.61 | 194.25 | 193.79 | 193.34 | 0.09639 |
| 3 | 4 | 24 | 2 | 125.67 | 502.99 | 503.19 | 503.38 | 192.44 | 192.00 | 191.56 | 0.09639 |
| 3 | 4 | 24 | 2 | 125.64 | 503.77 | 503.96 | 504.15 | 190.68 | 190.24 | 189.81 | 0.09639 |
| 4 | 4 | 23 | 2 | 124.26 | 537.65 | 537.84 | 538.03 | 149.61 | 149.58 | 149.54 | 0.09312 |
| 4 | 4 | 23 | 2 | 124.23 | 538.40 | 538.58 | 538.77 | 149.47 | 149.43 | 149.40 | 0.09312 |
| 4 | 4 | 23 | 2 | 124.20 | 539.14 | 539.32 | 539.51 | 149.34 | 149.31 | 149.28 | 0.09312 |
| 4 | 4 | 23 | 2 | 124.17 | 539.88 | 540.07 | 540.25 | 149.23 | 149.21 | 149.20 | 0.09312 |
| 4 | 4 | 23 | 2 | 124.14 | 540.62 | 540.81 | 540.99 | 149.17 | 149.16 | 149.15 | 0.09312 |
| 4 | 4 | 22 | 2 | 122.57 | 579.77 | 579.95 | 580.13 | 191.26 | 191.66 | 192.05 | 0.08907 |
| 4 | 4 | 22 | 2 | 122.54 | 580.48 | 580.66 | 580.83 | 192.86 | 193.26 | 193.67 | 0.08907 |
| 4 | 4 | 22 | 2 | 122.52 | 581.19 | 581.37 | 581.54 | 194.48 | 194.89 | 195.31 | 0.08907 |
| 4 | 4 | 22 | 2 | 122.49 | 581.90 | 582.08 | 582.25 | 196.14 | 196.56 | 196.98 | 0.08907 |
| 4 | 4 | 22 | 2 | 122.46 | 582.61 | 582.79 | 582.96 | 197.83 | 198.25 | 198.68 | 0.08907 |
| 4 | 4 | 22 | 2 | 122.43 | 583.32 | 583.50 | 583.67 | 199.54 | 199.97 | 200.40 | 0.08907 |
| 8 | 4 | 22 | 2 | 112.69 | 847.59 | 847.77 | 847.95 | 148.21 | 151.03 | 153.85 | 0.11828 |
| 8 | 4 | 22 | 2 | 112.67 | 848.30 | 848.48 | 848.66 | 159.49 | 162.31 | 165.14 | 0.11828 |
| 8 | 4 | 22 | 2 | 112.64 | 849.01 | 849.19 | 849.37 | 170.79 | 173.62 | 176.46 | 0.11828 |
| 8 | 4 | 21 | 2 | 111.90 | 870.89 | 871.06 | 871.23 | 170.26 | 173.12 | 175.99 | 0.11290 |
| 8 | 4 | 21 | 2 | 111.88 | 871.57 | 871.74 | 871.91 | 181.72 | 184.59 | 187.46 | 0.11290 |
| 8 | 4 | 21 | 2 | 111.85 | 872.25 | 872.42 | 872.58 | 193.21 | 196.09 | 198.97 | 0.11290 |
| 8 | 4 | 19 | 2 | 111.17 | 892.79 | 892.94 | 893.09 | 191.75 | 194.50 | 197.24 | 0.10215 |
| 8 | 4 | 19 | 2 | 111.15 | 893.40 | 893.55 | 893.70 | 202.74 | 205.49 | 208.24 | 0.10215 |
| 8 | 4 | 19 | 2 | 111.13 | 894.01 | 894.16 | 894.32 | 213.75 | 216.50 | 219.26 | 0.10215 |
| 9 | 4 | 19 | 2 | 107.09 | 967.24 | 967.39 | 967.54 | 324.93 | 328.92 | 332.91 | 0.14286 |
| 9 | 4 | 19 | 2 | 107.05 | 967.85 | 968.00 | 968.16 | 340.92 | 344.92 | 348.93 | 0.14286 |
| 9 | 4 | 21 | 2 | 105.71 | 986.25 | 986.41 | 986.58 | 131.96 | 136.98 | 142.01 | 0.15789 |
| 9 | 4 | 22 | 2 | 104.81 | 998.57 | 998.74 | 998.92 | 153.79 | 159.49 | 165.19 | 0.16541 |
| 9 | 4 | 15 | 2 | 103.99 | 1009.72 | 1009.84 | 1009.96 | 160.11 | 164.17 | 168.23 | 0.11278 |
| 9 | 4 | 15 | 2 | 103.95 | 1010.20 | 1010.32 | 1010.45 | 176.35 | 180.41 | 184.47 | 0.11278 |
| 9 | 4 | 22 | 2 | 103.06 | 1022.28 | 1022.45 | 1022.63 | 228.72 | 234.88 | 241.04 | 0.16541 |
| 10 | 4 | 19 | 2 | 100.65 | 1055.67 | 1055.82 | 1055.98 | 355.33 | 1.09 | 6.86 | 0.16379 |
| 10 | 4 | 18 | 2 | 99.93 | 1065.90 | 1066.04 | 1066.19 | 25.43 | 31.05 | 36.67 | 0.15517 |
| 10 | 4 | 17 | 2 | 99.24 | 1075.77 | 1075.91 | 1076.05 | 53.36 | 58.82 | 64.29 | 0.14655 |
| 10 | 4 | 16 | 2 | 98.49 | 1086.69 | 1086.82 | 1086.95 | 136.31 | 141.64 | 146.97 | 0.13793 |
| 10 | 4 | 16 | 2 | 97.86 | 1095.87 | 1096.00 | 1096.13 | 161.46 | 166.96 | 172.46 | 0.13793 |
| 10 | 4 | 15 | 2 | 97.25 | 1104.83 | 1104.95 | 1105.07 | 189.04 | 194.35 | 199.67 | 0.12931 |
| 10 | 4 | 15 | 2 | 96.67 | 1113.54 | 1113.66 | 1113.78 | 216.45 | 221.91 | 227.36 | 0.12931 |
| 11 | 4 | 14 | 2 | 94.51 | 1146.47 | 1146.58 | 1146.69 | 330.15 | 335.72 | 341.28 | 0.14000 |
| 11 | 4 | 13 | 2 | 94.01 | 1154.30 | 1154.40 | 1154.51 | 359.93 | 5.20 | 10.48 | 0.13000 |
| 11 | 4 | 13 | 2 | 93.36 | 1164.51 | 1164.62 | 1164.72 | 160.07 | 165.48 | 170.88 | 0.13000 |
| 11 | 4 | 13 | 2 | 92.89 | 1171.95 | 1172.05 | 1172.16 | 186.41 | 191.90 | 197.40 | 0.13000 |
| 11 | 4 | 12 | 2 | 92.43 | 1179.31 | 1179.41 | 1179.50 | 215.65 | 220.81 | 225.97 | 0.12000 |
| 11 | 4 | 12 | 2 | 91.99 | 1186.46 | 1186.56 | 1186.66 | 240.56 | 245.81 | 251.07 | 0.12000 |
| 11 | 4 | 12 | 2 | 91.57 | 1193.47 | 1193.57 | 1193.66 | 264.49 | 269.84 | 275.19 | 0.12000 |
| 11 | 4 | 11 | 2 | 91.15 | 1200.36 | 1200.45 | 1200.53 | 288.22 | 293.21 | 298.20 | 0.11000 |
| 12 | 4 | 10 | 2 | 89.43 | 1229.18 | 1229.26 | 1229.34 | 147.99 | 152.71 | 157.43 | 0.06024 |
| 12 | 4 | 10 | 2 | 89.08 | 1235.24 | 1235.32 | 1235.40 | 142.57 | 147.29 | 152.02 | 0.06024 |
| 12 | 4 | 21 | 2 | 88.72 | 1241.32 | 1241.49 | 1241.66 | 138.55 | 148.46 | 158.36 | 0.12651 |
| 12 | 4 | 10 | 2 | 88.37 | 1247.48 | 1247.56 | 1247.64 | 138.38 | 143.07 | 147.78 | 0.06024 |
| 12 | 4 | 10 | 2 | 88.02 | 1253.57 | 1253.66 | 1253.74 | 132.85 | 137.53 | 142.20 | 0.06024 |
| 12 | 4 | 21 | 2 | 87.66 | 1259.69 | 1259.86 | 1260.03 | 126.41 | 136.18 | 145.95 | 0.12651 |
| 12 | 4 | 21 | 2 | 87.32 | 1265.69 | 1265.86 | 1266.03 | 111.74 | 121.46 | 131.19 | 0.12651 |
| 12 | 4 | 21 | 2 | 86.99 | 1271.61 | 1271.78 | 1271.95 | 91.51 | 101.20 | 110.89 | 0.12651 |
| 12 | 4 | 21 | 2 | 86.66 | 1277.48 | 1277.65 | 1277.82 | 66.50 | 76.15 | 85.82 | 0.12651 |
| 12 | 4 | 21 | 2 | 86.34 | 1283.32 | 1283.49 | 1283.66 | 39.64 | 49.28 | 58.92 | 0.12651 |
| 13 | 4 | 20 | 2 | 84.35 | 1320.07 | 1320.23 | 1320.39 | 317.20 | 326.28 | 335.36 | 0.09756 |
| 13 | 4 | 22 | 2 | 83.83 | 1329.83 | 1330.01 | 1330.18 | 145.78 | 155.73 | 165.67 | 0.10732 |
| 13 | 4 | 20 | 2 | 83.51 | 1336.07 | 1336.23 | 1336.39 | 134.86 | 143.85 | 152.83 | 0.09756 |
| 13 | 4 | 20 | 2 | 83.21 | 1341.79 | 1341.95 | 1342.11 | 92.58 | 101.51 | 110.43 | 0.09756 |
| 13 | 4 | 20 | 2 | 82.91 | 1347.50 | 1347.66 | 1347.82 | 47.31 | 56.16 | 65.00 | 0.09756 |
| 13 | 4 | 20 | 2 | 82.62 | 1353.21 | 1353.38 | 1353.54 | 359.24 | 7.99 | 16.73 | 0.09756 |
| 13 | 4 | 21 | 2 | 82.32 | 1358.97 | 1359.14 | 1359.31 | 309.38 | 318.44 | 327.49 | 0.10244 |
| 13 | 4 | 20 | 2 | 82.03 | 1364.75 | 1364.91 | 1365.07 | 255.78 | 264.26 | 272.74 | 0.09756 |
| 13 | 4 | 21 | 2 | 81.74 | 1370.54 | 1370.71 | 1370.87 | 198.04 | 206.80 | 215.55 | 0.10244 |
| 13 | 4 | 21 | 2 | 81.45 | 1376.36 | 1376.53 | 1376.70 | 136.82 | 145.43 | 154.04 | 0.10244 |
| 14 | 4 | 13 | 2 | 79.65 | 1413.03 | 1413.14 | 1413.24 | 131.61 | 136.62 | 141.63 | 0.06047 |
| 14 | 4 | 22 | 2 | 79.36 | 1419.07 | 1419.25 | 1419.43 | 59.77 | 68.22 | 76.67 | 0.10233 |
| 14 | 4 | 22 | 2 | 79.07 | 1425.19 | 1425.37 | 1425.55 | 351.05 | 359.48 | 7.91 | 0.10233 |
| 14 | 4 | 22 | 2 | 78.78 | 1431.40 | 1431.57 | 1431.75 | 285.47 | 293.88 | 302.29 | 0.10233 |
| 14 | 4 | 22 | 2 | 78.49 | 1437.64 | 1437.81 | 1437.99 | 220.71 | 229.09 | 237.46 | 0.10233 |
| 14 | 4 | 22 | 2 | 78.19 | 1443.91 | 1444.09 | 1444.27 | 155.91 | 164.23 | 172.54 | 0.10233 |
| 14 | 4 | 23 | 2 | 77.90 | 1450.22 | 1450.41 | 1450.60 | 90.77 | 99.39 | 108.01 | 0.10698 |
| 14 | 4 | 23 | 2 | 77.61 | 1456.60 | 1456.79 | 1456.97 | 25.76 | 34.29 | 42.82 | 0.10698 |
| 14 | 4 | 23 | 2 | 77.31 | 1463.07 | 1463.26 | 1463.44 | 321.40 | 329.83 | 338.26 | 0.10698 |
| 14 | 4 | 23 | 2 | 77.02 | 1469.58 | 1469.76 | 1469.95 | 255.11 | 263.43 | 271.74 | 0.10698 |
| 15 | 4 | 16 | 2 | 74.17 | 1535.16 | 1535.29 | 1535.42 | 148.55 | 153.76 | 158.98 | 0.14953 |
| 15 | 4 | 13 | 2 | 73.88 | 1542.25 | 1542.35 | 1542.46 | 75.30 | 79.54 | 83.77 | 0.12150 |
| 15 | 4 | 13 | 2 | 73.57 | 1549.63 | 1549.74 | 1549.84 | 11.57 | 15.72 | 19.87 | 0.12150 |
| 15 | 4 | 13 | 2 | 73.27 | 1557.05 | 1557.15 | 1557.26 | 302.07 | 306.12 | 310.17 | 0.12150 |
| 15 | 4 | 13 | 2 | 72.96 | 1564.54 | 1564.64 | 1564.75 | 226.24 | 230.14 | 234.04 | 0.12150 |
| 15 | 4 | 13 | 2 | 72.65 | 1572.12 | 1572.22 | 1572.33 | 144.98 | 148.82 | 152.66 | 0.12150 |
| 15 | 4 | 13 | 2 | 72.34 | 1579.82 | 1579.92 | 1580.03 | 63.12 | 66.84 | 70.55 | 0.12150 |
| 15 | 4 | 13 | 2 | 72.03 | 1587.61 | 1587.71 | 1587.82 | 338.52 | 342.19 | 345.86 | 0.12150 |
| 16 | 4 | 14 | 2 | 70.69 | 1622.29 | 1622.41 | 1622.52 | 118.28 | 122.24 | 126.20 | 0.13462 |
| 16 | 4 | 15 | 2 | 70.37 | 1630.66 | 1630.79 | 1630.91 | 52.11 | 56.34 | 60.57 | 0.14423 |
| 16 | 4 | 15 | 2 | 70.06 | 1639.16 | 1639.29 | 1639.41 | 348.54 | 352.69 | 356.85 | 0.14423 |
| 16 | 4 | 15 | 2 | 69.74 | 1647.74 | 1647.86 | 1647.98 | 283.62 | 287.81 | 292.00 | 0.14423 |
| 16 | 4 | 15 | 2 | 69.43 | 1656.44 | 1656.56 | 1656.68 | 225.68 | 229.82 | 233.96 | 0.14423 |
| 16 | 4 | 15 | 2 | 69.11 | 1665.13 | 1665.25 | 1665.37 | 150.51 | 154.43 | 158.36 | 0.14423 |
| 16 | 4 | 15 | 2 | 68.80 | 1673.94 | 1674.07 | 1674.19 | 76.16 | 80.09 | 84.02 | 0.14423 |
| 17 | 4 | 17 | 2 | 67.21 | 1720.11 | 1720.25 | 1720.39 | 142.22 | 146.58 | 150.94 | 0.14286 |
| 17 | 4 | 16 | 2 | 66.90 | 1729.51 | 1729.64 | 1729.76 | 72.19 | 76.06 | 79.93 | 0.13445 |
| 17 | 4 | 17 | 2 | 66.58 | 1739.01 | 1739.14 | 1739.28 | 0.35 | 4.61 | 8.87 | 0.14286 |
| 17 | 4 | 17 | 2 | 66.28 | 1748.53 | 1748.67 | 1748.80 | 294.62 | 298.75 | 302.87 | 0.14286 |
| 17 | 4 | 17 | 2 | 65.97 | 1758.26 | 1758.39 | 1758.53 | 220.95 | 224.94 | 228.93 | 0.14286 |
| 17 | 4 | 17 | 2 | 65.66 | 1768.10 | 1768.24 | 1768.38 | 145.33 | 149.24 | 153.14 | 0.14286 |
| 17 | 4 | 18 | 2 | 65.35 | 1778.05 | 1778.20 | 1778.34 | 67.62 | 71.78 | 75.94 | 0.15126 |
| 18 | 4 | 18 | 2 | 64.38 | 1810.44 | 1810.59 | 1810.73 | 282.55 | 286.66 | 290.77 | 0.15929 |
| 18 | 4 | 18 | 2 | 64.07 | 1820.79 | 1820.93 | 1821.08 | 213.16 | 217.26 | 221.36 | 0.15929 |
| 18 | 4 | 19 | 2 | 63.77 | 1831.34 | 1831.50 | 1831.65 | 153.58 | 157.93 | 162.27 | 0.16814 |
| 18 | 4 | 19 | 2 | 63.47 | 1842.06 | 1842.21 | 1842.37 | 95.90 | 100.19 | 104.47 | 0.16814 |
| 18 | 4 | 19 | 2 | 63.16 | 1852.98 | 1853.13 | 1853.28 | 35.69 | 39.83 | 43.96 | 0.16814 |
| 18 | 4 | 20 | 2 | 62.86 | 1864.13 | 1864.29 | 1864.45 | 342.21 | 346.68 | 351.16 | 0.17699 |
| 22 | 4 | 23 | 2 | 58.97 | 2018.51 | 2018.70 | 2018.89 | 133.29 | 137.88 | 142.48 | 0.20536 |
| 22 | 4 | 22 | 2 | 58.68 | 2031.04 | 2031.21 | 2031.39 | 84.51 | 88.96 | 93.42 | 0.19643 |
| 22 | 4 | 22 | 2 | 58.39 | 2043.54 | 2043.71 | 2043.89 | 42.82 | 47.39 | 51.96 | 0.19643 |
| 22 | 4 | 22 | 2 | 58.10 | 2056.04 | 2056.21 | 2056.39 | 356.81 | 1.36 | 5.90 | 0.19643 |
| 22 | 4 | 23 | 2 | 57.82 | 2068.56 | 2068.74 | 2068.93 | 315.36 | 320.03 | 324.70 | 0.20536 |
| 23 | 4 | 23 | 2 | 56.97 | 2106.88 | 2107.06 | 2107.25 | 159.83 | 164.14 | 168.44 | 0.19658 |
| 23 | 4 | 23 | 2 | 56.68 | 2120.33 | 2120.51 | 2120.70 | 112.16 | 116.52 | 120.88 | 0.19658 |
| 23 | 4 | 23 | 2 | 56.41 | 2133.10 | 2133.29 | 2133.47 | 51.90 | 56.22 | 60.52 | 0.19658 |
| 23 | 4 | 24 | 2 | 56.07 | 2149.08 | 2149.27 | 2149.46 | 62.02 | 66.40 | 70.77 | 0.20513 |
| 23 | 4 | 24 | 2 | 55.80 | 2162.12 | 2162.32 | 2162.51 | 346.44 | 350.60 | 354.75 | 0.20513 |
| 24 | 4 | 25 | 2 | 54.42 | 2229.27 | 2229.47 | 2229.67 | 288.90 | 292.68 | 296.46 | 0.19084 |
| 24 | 4 | 26 | 2 | 54.14 | 2243.38 | 2243.59 | 2243.80 | 196.31 | 200.29 | 204.27 | 0.19847 |
| 24 | 4 | 26 | 2 | 53.86 | 2257.67 | 2257.88 | 2258.09 | 102.55 | 106.46 | 110.38 | 0.19847 |
| 24 | 4 | 27 | 2 | 53.58 | 2272.33 | 2272.55 | 2272.77 | 17.21 | 21.29 | 25.36 | 0.20611 |
| 24 | 4 | 27 | 2 | 53.29 | 2287.36 | 2287.58 | 2287.80 | 302.62 | 306.66 | 310.69 | 0.20611 |
| 25 | 4 | 29 | 2 | 52.43 | 2332.74 | 2332.98 | 2333.21 | 20.81 | 24.97 | 29.14 | 0.23770 |
| 25 | 4 | 30 | 2 | 52.12 | 2348.92 | 2349.16 | 2349.40 | 300.84 | 304.92 | 309.00 | 0.24590 |
| 25 | 4 | 31 | 2 | 51.82 | 2365.44 | 2365.69 | 2365.94 | 212.51 | 216.52 | 220.53 | 0.25410 |
| 25 | 4 | 32 | 2 | 51.51 | 2382.40 | 2382.66 | 2382.92 | 118.52 | 122.47 | 126.42 | 0.26230 |
| 26 | 4 | 32 | 2 | 50.90 | 2416.68 | 2416.94 | 2417.20 | 296.94 | 301.06 | 305.19 | 0.24806 |
| 26 | 4 | 33 | 2 | 50.58 | 2434.35 | 2434.62 | 2434.88 | 220.17 | 224.44 | 228.71 | 0.25581 |
| 26 | 4 | 32 | 2 | 50.27 | 2452.02 | 2452.28 | 2452.54 | 140.46 | 144.45 | 148.44 | 0.24806 |
| 26 | 4 | 32 | 2 | 49.97 | 2469.66 | 2469.92 | 2470.18 | 43.88 | 47.65 | 51.42 | 0.24806 |
| 30 | 4 | 64 | 2 | 47.33 | 2628.83 | 2629.34 | 2629.86 | 23.19 | 31.43 | 39.67 | 0.32990 |
| 30 | 4 | 64 | 2 | 47.06 | 2645.97 | 2646.49 | 2647.00 | 291.92 | 299.87 | 307.80 | 0.32990 |
| 30 | 4 | 33 | 2 | 46.80 | 2663.22 | 2663.49 | 2663.75 | 193.89 | 197.89 | 201.88 | 0.17010 |
| 30 | 4 | 33 | 2 | 46.52 | 2680.75 | 2681.02 | 2681.28 | 91.83 | 95.71 | 99.57 | 0.17010 |
| 31 | 4 | 68 | 2 | 46.13 | 2706.29 | 2706.83 | 2707.38 | 120.93 | 130.12 | 139.36 | 0.31481 |
| 31 | 4 | 67 | 2 | 45.86 | 2724.21 | 2724.75 | 2725.29 | 58.80 | 67.52 | 76.22 | 0.31019 |
| 31 | 4 | 81 | 2 | 45.46 | 2750.65 | 2751.31 | 2751.96 | 116.35 | 126.75 | 137.15 | 0.37500 |
| 32 | 4 | 63 | 2 | 44.71 | 2801.17 | 2801.68 | 2802.19 | 243.34 | 252.15 | 260.95 | 0.25200 |
| 32 | 4 | 63 | 2 | 44.47 | 2817.84 | 2818.35 | 2818.85 | 166.10 | 174.40 | 182.69 | 0.25200 |
| 32 | 4 | 61 | 2 | 44.24 | 2834.05 | 2834.55 | 2835.04 | 74.50 | 82.68 | 90.85 | 0.24400 |
| 32 | 4 | 63 | 2 | 44.00 | 2850.84 | 2851.34 | 2851.85 | 349.15 | 357.21 | 5.26 | 0.25200 |
| 33 | 4 | 95 | 2 | 43.19 | 2907.89 | 2908.65 | 2909.42 | 160.16 | 172.47 | 184.81 | 0.42035 |
| 33 | 4 | 67 | 2 | 42.92 | 2927.73 | 2928.27 | 2928.81 | 129.69 | 138.82 | 147.94 | 0.29646 |
| 33 | 4 | 64 | 2 | 42.68 | 2945.23 | 2945.75 | 2946.26 | 50.83 | 58.72 | 66.61 | 0.28319 |
| 34 | 4 | 67 | 2 | 42.20 | 2980.49 | 2981.03 | 2981.57 | 213.27 | 221.34 | 229.44 | 0.25869 |
| 34 | 4 | 67 | 2 | 41.96 | 2998.44 | 2998.98 | 2999.52 | 139.19 | 148.16 | 157.15 | 0.25869 |
| 34 | 4 | 65 | 2 | 41.73 | 3015.85 | 3016.38 | 3016.90 | 66.58 | 75.10 | 83.63 | 0.25097 |
| 34 | 4 | 60 | 2 | 41.47 | 3034.97 | 3035.46 | 3035.94 | 13.03 | 20.68 | 28.32 | 0.23166 |
| 35 | 4 | 62 | 2 | 40.35 | 3122.10 | 3122.60 | 3123.10 | 268.96 | 276.30 | 283.62 | 0.24900 |
| 35 | 4 | 62 | 2 | 40.14 | 3138.58 | 3139.08 | 3139.58 | 151.30 | 158.72 | 166.15 | 0.24900 |
| 35 | 4 | 62 | 2 | 39.93 | 3155.16 | 3155.66 | 3156.16 | 15.80 | 22.08 | 28.39 | 0.24900 |
| 35 | 4 | 63 | 2 | 39.72 | 3171.83 | 3172.33 | 3172.84 | 242.92 | 249.98 | 257.02 | 0.25301 |
| 36 | 4 | 63 | 2 | 39.30 | 3205.52 | 3206.02 | 3206.53 | 343.75 | 350.88 | 358.01 | 0.24609 |
| 36 | 4 | 64 | 2 | 39.09 | 3222.48 | 3223.00 | 3223.52 | 220.69 | 227.88 | 235.07 | 0.25000 |
| 36 | 4 | 64 | 2 | 38.88 | 3239.59 | 3240.10 | 3240.62 | 99.35 | 106.48 | 113.62 | 0.25000 |
| 36 | 4 | 65 | 2 | 38.68 | 3256.79 | 3257.31 | 3257.84 | 327.42 | 333.86 | 340.29 | 0.25391 |
| 37 | 4 | 66 | 2 | 38.05 | 3309.03 | 3309.56 | 3310.09 | 328.33 | 335.55 | 342.76 | 0.24812 |
| 37 | 4 | 66 | 2 | 37.84 | 3326.66 | 3327.19 | 3327.72 | 203.08 | 210.11 | 217.13 | 0.24812 |
| 37 | 4 | 67 | 2 | 37.64 | 3344.38 | 3344.92 | 3345.46 | 81.79 | 89.25 | 96.72 | 0.25188 |
| 37 | 4 | 67 | 2 | 37.43 | 3362.23 | 3362.77 | 3363.31 | 339.01 | 347.32 | 355.66 | 0.25188 |
| 38 | 4 | 68 | 2 | 37.12 | 3388.90 | 3389.44 | 3389.99 | 31.50 | 40.01 | 48.52 | 0.24818 |
| 38 | 4 | 68 | 2 | 36.91 | 3407.02 | 3407.57 | 3408.12 | 307.39 | 315.30 | 323.18 | 0.24818 |
| 38 | 4 | 69 | 2 | 36.70 | 3425.26 | 3425.82 | 3426.38 | 197.84 | 205.15 | 212.48 | 0.25182 |
| 38 | 4 | 69 | 2 | 36.49 | 3443.62 | 3444.17 | 3444.73 | 80.73 | 88.04 | 95.33 | 0.25182 |
| 39 | 4 | 70 | 2 | 36.08 | 3480.65 | 3481.21 | 3481.78 | 224.18 | 232.42 | 240.67 | 0.24648 |
| 39 | 4 | 71 | 2 | 35.89 | 3497.67 | 3498.25 | 3498.82 | 118.31 | 127.05 | 135.80 | 0.25000 |
| 39 | 4 | 71 | 2 | 35.68 | 3516.48 | 3517.05 | 3517.62 | 37.66 | 45.73 | 53.74 | 0.25000 |
| 39 | 4 | 72 | 2 | 35.43 | 3539.52 | 3540.10 | 3540.68 | 336.71 | 343.77 | 350.84 | 0.25352 |
| 40 | 4 | 72 | 2 | 35.01 | 3577.77 | 3578.35 | 3578.93 | 112.07 | 119.53 | 126.98 | 0.24658 |
| 40 | 4 | 73 | 2 | 34.80 | 3597.06 | 3597.65 | 3598.24 | 14.00 | 22.65 | 31.32 | 0.25000 |
| 40 | 4 | 73 | 2 | 34.60 | 3616.50 | 3617.09 | 3617.68 | 303.76 | 312.11 | 320.44 | 0.25000 |
| 40 | 4 | 74 | 2 | 34.39 | 3636.05 | 3636.64 | 3637.24 | 228.97 | 237.86 | 246.74 | 0.25342 |
| 41 | 4 | 74 | 2 | 34.02 | 3671.42 | 3672.02 | 3672.61 | 17.18 | 25.86 | 34.48 | 0.24585 |
| 41 | 4 | 75 | 2 | 33.81 | 3691.32 | 3691.92 | 3692.52 | 291.57 | 300.09 | 308.62 | 0.24917 |
| 41 | 4 | 76 | 2 | 33.60 | 3711.32 | 3711.94 | 3712.55 | 215.18 | 223.83 | 232.46 | 0.25249 |
| 41 | 4 | 76 | 2 | 33.40 | 3730.65 | 3731.26 | 3731.88 | 125.07 | 133.84 | 142.63 | 0.25249 |
| 42 | 4 | 78 | 2 | 32.57 | 3812.47 | 3813.10 | 3813.73 | 143.42 | 150.27 | 157.09 | 0.33191 |
| 42 | 4 | 78 | 2 | 32.36 | 3833.25 | 3833.88 | 3834.51 | 36.09 | 44.48 | 52.82 | 0.33191 |
| 42 | 4 | 79 | 2 | 32.15 | 3854.15 | 3854.79 | 3855.42 | 304.07 | 312.23 | 320.38 | 0.33617 |
| 43 | 4 | 80 | 2 | 31.74 | 3896.35 | 3897.00 | 3897.64 | 44.07 | 50.28 | 56.59 | 0.33195 |
| 43 | 4 | 80 | 2 | 31.53 | 3917.64 | 3918.29 | 3918.93 | 297.85 | 307.73 | 317.78 | 0.33195 |
| 43 | 4 | 81 | 2 | 31.32 | 3939.07 | 3939.73 | 3940.38 | 259.76 | 268.75 | 277.66 | 0.33610 |
| 44 | 4 | 82 | 2 | 30.90 | 3982.32 | 3982.98 | 3983.64 | 174.17 | 184.11 | 194.10 | 0.33065 |
| 44 | 4 | 83 | 2 | 30.71 | 4002.64 | 4003.31 | 4003.98 | 154.68 | 166.56 | 178.47 | 0.33468 |
| 44 | 4 | 83 | 2 | 30.50 | 4024.60 | 4025.27 | 4025.94 | 169.43 | 180.19 | 190.92 | 0.33468 |
| 45 | 4 | 84 | 2 | 30.08 | 4068.92 | 4069.60 | 4070.28 | 94.56 | 103.80 | 113.04 | 0.33071 |
| 45 | 4 | 85 | 2 | 29.88 | 4091.29 | 4091.97 | 4092.66 | 43.54 | 53.04 | 62.49 | 0.33465 |
| 45 | 4 | 85 | 2 | 29.67 | 4113.79 | 4114.48 | 4115.16 | 345.78 | 354.81 | 3.81 | 0.33465 |
| 46 | 4 | 86 | 2 | 29.25 | 4159.25 | 4159.94 | 4160.64 | 150.42 | 156.81 | 163.22 | 0.33077 |
| 46 | 4 | 87 | 2 | 29.04 | 4182.18 | 4182.88 | 4183.58 | 30.07 | 38.35 | 46.67 | 0.33462 |
| 46 | 4 | 87 | 2 | 28.83 | 4205.27 | 4205.97 | 4206.67 | 315.01 | 323.76 | 332.53 | 0.33462 |
| 47 | 4 | 88 | 2 | 28.41 | 4251.86 | 4252.57 | 4253.28 | 184.14 | 192.76 | 201.35 | 0.32959 |
| 47 | 4 | 89 | 2 | 28.21 | 4275.39 | 4276.11 | 4276.83 | 97.64 | 106.03 | 114.43 | 0.33333 |
| 47 | 4 | 90 | 2 | 28.00 | 4299.05 | 4299.78 | 4300.50 | 19.58 | 27.97 | 36.22 | 0.33708 |
| 48 | 4 | 91 | 2 | 27.58 | 4346.82 | 4347.55 | 4348.29 | 135.68 | 143.82 | 152.00 | 0.33212 |
| 48 | 4 | 91 | 2 | 27.37 | 4370.93 | 4371.66 | 4372.40 | 35.01 | 42.62 | 50.21 | 0.33212 |
| 48 | 4 | 92 | 2 | 27.16 | 4395.19 | 4395.93 | 4396.67 | 280.91 | 288.77 | 296.65 | 0.33577 |
| 50 | 4 | 128 | 2 | 23.31 | 4866.34 | 4867.37 | 4868.40 | 257.99 | 269.31 | 280.71 | 0.49805 |
| 50 | 4 | 129 | 2 | 23.05 | 4899.97 | 4901.01 | 4902.05 | 247.38 | 257.82 | 268.32 | 0.50195 |
| 51 | 4 | 132 | 2 | 20.74 | 5208.94 | 5210.01 | 5211.07 | 178.70 | 189.62 | 200.49 | 0.49811 |
| 51 | 4 | 133 | 2 | 20.48 | 5243.80 | 5244.88 | 5245.95 | 152.00 | 161.42 | 170.84 | 0.50189 |
| 54 | 4 | 131 | 2 | 17.08 | 5736.83 | 5737.88 | 5738.94 | 303.53 | 313.52 | 323.53 | 0.51172 |
| 54 | 4 | 125 | 2 | 16.86 | 5769.73 | 5770.73 | 5771.74 | 251.06 | 260.46 | 269.85 | 0.48828 |
| 55 | 4 | 117 | 2 | 14.87 | 6078.19 | 6079.14 | 6080.08 | 294.88 | 303.61 | 312.35 | 0.49787 |
| 55 | 4 | 118 | 2 | 14.68 | 6109.21 | 6110.16 | 6111.11 | 225.15 | 234.27 | 243.42 | 0.50213 |
| 56 | 4 | 124 | 2 | 12.72 | 6428.57 | 6429.57 | 6430.57 | 272.99 | 282.11 | 291.23 | 0.49799 |
| 56 | 4 | 125 | 2 | 12.53 | 6461.37 | 6462.37 | 6463.38 | 200.74 | 209.05 | 217.31 | 0.50201 |
| 57 | 4 | 131 | 2 | 10.76 | 6764.38 | 6765.44 | 6766.49 | 161.61 | 170.54 | 179.46 | 0.50000 |
| 57 | 4 | 131 | 2 | 10.56 | 6798.75 | 6799.81 | 6800.86 | 75.79 | 83.69 | 91.55 | 0.50000 |
| 58 | 4 | 136 | 2 | 8.99 | 7079.76 | 7080.86 | 7081.96 | 198.31 | 206.37 | 214.43 | 0.50000 |
| 58 | 4 | 136 | 2 | 8.79 | 7115.64 | 7116.74 | 7117.83 | 100.61 | 109.41 | 118.28 | 0.50000 |
| 59 | 4 | 141 | 2 | 7.21 | 7407.85 | 7408.99 | 7410.13 | 117.70 | 125.77 | 133.86 | 0.49823 |
| 59 | 4 | 142 | 2 | 7.01 | 7445.01 | 7446.15 | 7447.30 | 28.87 | 37.41 | 45.94 | 0.50177 |
| 60 | 4 | 146 | 2 | 5.44 | 7747.67 | 7748.84 | 7750.02 | 107.25 | 115.09 | 122.93 | 0.50000 |
| 60 | 4 | 146 | 2 | 5.24 | 7785.99 | 7787.16 | 7788.34 | 4.66 | 12.67 | 20.67 | 0.50000 |
| 1 | 5 | 1 | 10 | 142.31 | 144.42 | 144.46 | 144.50 | 177.61 | 176.42 | 175.22 | 0.03030 |
| 1 | 5 | 1 | 10 | 142.31 | 144.58 | 144.62 | 144.66 | 172.84 | 171.64 | 170.45 | 0.03030 |
| 1 | 5 | 1 | 10 | 142.30 | 144.74 | 144.78 | 144.82 | 168.07 | 166.87 | 165.68 | 0.03030 |
| 1 | 5 | 1 | 10 | 142.29 | 144.90 | 144.95 | 144.99 | 163.30 | 162.10 | 160.91 | 0.03030 |
| 1 | 5 | 1 | 12 | 141.84 | 153.55 | 153.60 | 153.65 | 270.49 | 269.09 | 267.69 | 0.03030 |
| 1 | 5 | 1 | 12 | 141.83 | 153.74 | 153.79 | 153.84 | 264.90 | 263.50 | 262.10 | 0.03030 |
| 1 | 5 | 1 | 12 | 141.82 | 153.94 | 153.99 | 154.03 | 259.30 | 257.91 | 256.51 | 0.03030 |
| 1 | 5 | 1 | 12 | 141.81 | 154.13 | 154.18 | 154.23 | 253.71 | 252.32 | 250.92 | 0.03030 |
| 1 | 5 | 1 | 10 | 141.35 | 163.02 | 163.06 | 163.10 | 359.90 | 358.76 | 357.62 | 0.03030 |
| 1 | 5 | 1 | 10 | 141.34 | 163.18 | 163.22 | 163.26 | 355.34 | 354.20 | 353.06 | 0.03030 |
| 1 | 5 | 1 | 10 | 141.33 | 163.34 | 163.38 | 163.42 | 350.78 | 349.65 | 348.51 | 0.03030 |
| 1 | 5 | 1 | 10 | 141.32 | 163.50 | 163.54 | 163.58 | 346.23 | 345.09 | 343.96 | 0.03030 |
| 1 | 5 | 1 | 10 | 141.32 | 163.66 | 163.70 | 163.74 | 341.68 | 340.54 | 339.40 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.85 | 172.84 | 172.88 | 172.92 | 85.71 | 84.60 | 83.49 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.84 | 173.00 | 173.04 | 173.08 | 81.26 | 80.15 | 79.04 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.83 | 173.16 | 173.20 | 173.24 | 76.82 | 75.71 | 74.59 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.82 | 173.32 | 173.36 | 173.40 | 72.37 | 71.26 | 70.15 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.81 | 173.48 | 173.52 | 173.56 | 67.93 | 66.82 | 65.71 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.33 | 182.95 | 182.99 | 183.03 | 170.22 | 169.14 | 168.06 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.32 | 183.11 | 183.15 | 183.19 | 165.89 | 164.80 | 163.72 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.31 | 183.27 | 183.31 | 183.35 | 161.55 | 160.46 | 159.38 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.31 | 183.43 | 183.48 | 183.52 | 157.21 | 156.13 | 155.04 | 0.03030 |
| 1 | 5 | 1 | 10 | 140.30 | 183.60 | 183.64 | 183.68 | 152.87 | 151.79 | 150.71 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.81 | 193.29 | 193.33 | 193.37 | 255.83 | 254.78 | 253.72 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.80 | 193.45 | 193.49 | 193.53 | 251.61 | 250.56 | 249.50 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.79 | 193.61 | 193.65 | 193.69 | 247.39 | 246.34 | 245.29 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.78 | 193.77 | 193.81 | 193.85 | 243.18 | 242.12 | 241.07 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.77 | 193.93 | 193.97 | 194.01 | 238.96 | 237.91 | 236.86 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.27 | 204.03 | 204.07 | 204.11 | 339.15 | 338.13 | 337.11 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.26 | 204.19 | 204.23 | 204.27 | 335.06 | 334.04 | 333.02 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.25 | 204.35 | 204.39 | 204.43 | 330.98 | 329.96 | 328.94 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.24 | 204.51 | 204.55 | 204.59 | 326.90 | 325.88 | 324.86 | 0.03030 |
| 1 | 5 | 1 | 10 | 139.24 | 204.68 | 204.72 | 204.76 | 322.82 | 321.80 | 320.78 | 0.03030 |
| 2 | 5 | 16 | 2 | 135.29 | 285.60 | 285.73 | 285.86 | 307.91 | 305.36 | 302.81 | 0.12500 |
| 2 | 5 | 16 | 2 | 135.27 | 286.12 | 286.25 | 286.37 | 297.72 | 295.17 | 292.63 | 0.12500 |
| 2 | 5 | 3 | 2 | 134.91 | 293.87 | 293.90 | 293.92 | 147.26 | 146.80 | 146.34 | 0.02344 |
| 2 | 5 | 19 | 2 | 134.03 | 312.53 | 312.69 | 312.84 | 168.02 | 165.38 | 162.74 | 0.14844 |
| 2 | 5 | 19 | 2 | 134.00 | 313.15 | 313.30 | 313.45 | 157.47 | 154.83 | 152.20 | 0.14844 |
| 2 | 5 | 19 | 2 | 133.97 | 313.76 | 313.91 | 314.06 | 146.95 | 144.33 | 141.71 | 0.14844 |
| 2 | 5 | 18 | 2 | 133.17 | 331.07 | 331.22 | 331.36 | 225.38 | 223.15 | 220.92 | 0.14063 |
| 2 | 5 | 18 | 2 | 133.14 | 331.65 | 331.80 | 331.94 | 216.46 | 214.24 | 212.02 | 0.14063 |
| 3 | 5 | 23 | 2 | 127.93 | 449.12 | 449.31 | 449.50 | 59.10 | 57.94 | 56.78 | 0.09091 |
| 3 | 5 | 23 | 2 | 127.90 | 449.87 | 450.05 | 450.24 | 54.47 | 53.32 | 52.17 | 0.09091 |
| 3 | 5 | 23 | 2 | 127.87 | 450.61 | 450.79 | 450.98 | 49.88 | 48.74 | 47.61 | 0.09091 |
| 3 | 5 | 21 | 2 | 126.93 | 472.87 | 473.03 | 473.20 | 292.69 | 291.94 | 291.19 | 0.08300 |
| 3 | 5 | 21 | 2 | 126.90 | 473.54 | 473.71 | 473.88 | 289.71 | 288.97 | 288.23 | 0.08300 |
| 3 | 5 | 21 | 2 | 126.87 | 474.22 | 474.39 | 474.56 | 286.76 | 286.04 | 285.31 | 0.08300 |
| 3 | 5 | 21 | 2 | 126.84 | 474.90 | 475.07 | 475.24 | 283.87 | 283.16 | 282.44 | 0.08300 |
| 3 | 5 | 20 | 2 | 125.57 | 505.56 | 505.72 | 505.88 | 186.77 | 186.42 | 186.08 | 0.07905 |
| 3 | 5 | 20 | 2 | 125.54 | 506.20 | 506.36 | 506.52 | 185.41 | 185.07 | 184.74 | 0.07905 |
| 3 | 5 | 20 | 2 | 125.52 | 506.85 | 507.01 | 507.17 | 184.07 | 183.74 | 183.41 | 0.07905 |
| 3 | 5 | 20 | 2 | 125.49 | 507.49 | 507.65 | 507.81 | 182.75 | 182.42 | 182.10 | 0.07905 |
| 3 | 5 | 20 | 2 | 125.46 | 508.14 | 508.30 | 508.46 | 181.46 | 181.14 | 180.83 | 0.07905 |
| 4 | 5 | 23 | 2 | 124.07 | 542.35 | 542.53 | 542.72 | 149.16 | 149.17 | 149.18 | 0.09312 |
| 4 | 5 | 23 | 2 | 124.04 | 543.09 | 543.27 | 543.46 | 149.21 | 149.23 | 149.25 | 0.09312 |
| 4 | 5 | 23 | 2 | 124.01 | 543.83 | 544.02 | 544.20 | 149.30 | 149.33 | 149.36 | 0.09312 |
| 4 | 5 | 23 | 2 | 123.98 | 544.57 | 544.76 | 544.94 | 149.41 | 149.44 | 149.48 | 0.09312 |
| 4 | 5 | 23 | 2 | 123.95 | 545.32 | 545.50 | 545.69 | 149.54 | 149.58 | 149.62 | 0.09312 |
| 4 | 5 | 22 | 2 | 122.36 | 585.13 | 585.31 | 585.49 | 204.05 | 204.50 | 204.96 | 0.08907 |
| 4 | 5 | 22 | 2 | 122.33 | 585.84 | 586.02 | 586.20 | 205.87 | 206.33 | 206.79 | 0.08907 |
| 4 | 5 | 22 | 2 | 122.30 | 586.55 | 586.73 | 586.91 | 207.72 | 208.19 | 208.65 | 0.08907 |
| 4 | 5 | 22 | 2 | 122.28 | 587.26 | 587.44 | 587.62 | 209.59 | 210.07 | 210.54 | 0.08907 |
| 4 | 5 | 22 | 2 | 122.25 | 587.97 | 588.15 | 588.33 | 211.49 | 211.98 | 212.46 | 0.08907 |
| 4 | 5 | 22 | 2 | 122.22 | 588.68 | 588.86 | 589.04 | 213.43 | 213.92 | 214.41 | 0.08907 |
| 8 | 5 | 22 | 2 | 112.58 | 850.67 | 850.84 | 851.02 | 197.22 | 200.06 | 202.90 | 0.11828 |
| 8 | 5 | 22 | 2 | 112.56 | 851.38 | 851.55 | 851.73 | 208.59 | 211.43 | 214.28 | 0.11828 |
| 8 | 5 | 22 | 2 | 112.54 | 852.09 | 852.26 | 852.44 | 219.97 | 222.83 | 225.68 | 0.11828 |
| 8 | 5 | 21 | 2 | 111.80 | 873.78 | 873.95 | 874.12 | 219.30 | 222.20 | 225.09 | 0.11290 |
| 8 | 5 | 21 | 2 | 111.78 | 874.46 | 874.62 | 874.79 | 230.87 | 233.77 | 236.67 | 0.11290 |
| 8 | 5 | 21 | 2 | 111.76 | 875.13 | 875.30 | 875.47 | 242.46 | 245.36 | 248.27 | 0.11290 |
| 8 | 5 | 19 | 2 | 111.08 | 895.46 | 895.62 | 895.77 | 239.88 | 242.65 | 245.41 | 0.10215 |
| 8 | 5 | 19 | 2 | 111.06 | 896.08 | 896.23 | 896.38 | 250.95 | 253.71 | 256.48 | 0.10215 |
| 8 | 5 | 19 | 2 | 111.04 | 896.69 | 896.84 | 896.99 | 262.03 | 264.80 | 267.57 | 0.10215 |
| 9 | 5 | 19 | 2 | 106.95 | 969.26 | 969.42 | 969.57 | 17.98 | 22.03 | 26.08 | 0.14394 |
| 9 | 5 | 19 | 2 | 106.90 | 969.88 | 970.03 | 970.18 | 34.20 | 38.27 | 42.34 | 0.14394 |
| 9 | 5 | 21 | 2 | 105.61 | 987.70 | 987.87 | 988.04 | 175.46 | 180.54 | 185.63 | 0.15909 |
| 9 | 5 | 21 | 2 | 104.69 | 1000.08 | 1000.25 | 1000.42 | 202.63 | 208.10 | 213.58 | 0.15909 |
| 9 | 5 | 15 | 2 | 103.87 | 1011.35 | 1011.47 | 1011.59 | 214.84 | 218.92 | 222.99 | 0.11364 |
| 9 | 5 | 15 | 2 | 103.83 | 1011.83 | 1011.95 | 1012.07 | 231.15 | 235.23 | 239.31 | 0.11364 |
| 9 | 5 | 22 | 2 | 102.95 | 1023.81 | 1023.98 | 1024.16 | 282.01 | 288.19 | 294.38 | 0.16667 |
| 10 | 5 | 19 | 2 | 100.56 | 1056.99 | 1057.14 | 1057.29 | 44.85 | 50.64 | 56.43 | 0.16379 |
| 10 | 5 | 18 | 2 | 99.84 | 1067.17 | 1067.32 | 1067.46 | 74.82 | 80.45 | 86.09 | 0.15517 |
| 10 | 5 | 17 | 2 | 99.15 | 1077.01 | 1077.15 | 1077.28 | 102.62 | 108.10 | 113.59 | 0.14655 |
| 10 | 5 | 16 | 2 | 98.40 | 1087.86 | 1087.99 | 1088.12 | 184.70 | 190.05 | 195.40 | 0.13793 |
| 10 | 5 | 16 | 2 | 97.78 | 1097.01 | 1097.14 | 1097.27 | 210.02 | 215.54 | 221.06 | 0.13793 |
| 10 | 5 | 15 | 2 | 97.18 | 1105.93 | 1106.05 | 1106.18 | 237.60 | 242.94 | 248.27 | 0.12931 |
| 10 | 5 | 15 | 2 | 96.59 | 1114.63 | 1114.75 | 1114.87 | 265.59 | 271.06 | 276.53 | 0.12931 |
| 11 | 5 | 14 | 2 | 94.96 | 1139.51 | 1139.62 | 1139.73 | 350.25 | 355.71 | 1.18 | 0.13592 |
| 11 | 5 | 14 | 2 | 94.44 | 1147.45 | 1147.56 | 1147.68 | 18.68 | 24.26 | 29.84 | 0.13592 |
| 11 | 5 | 13 | 2 | 93.94 | 1155.27 | 1155.38 | 1155.48 | 49.06 | 54.34 | 59.63 | 0.12621 |
| 11 | 5 | 13 | 2 | 93.30 | 1165.46 | 1165.57 | 1165.67 | 209.16 | 214.57 | 219.99 | 0.12621 |
| 11 | 5 | 13 | 2 | 92.83 | 1172.88 | 1172.98 | 1173.08 | 235.07 | 240.58 | 246.09 | 0.12621 |
| 11 | 5 | 12 | 2 | 92.38 | 1180.22 | 1180.32 | 1180.42 | 264.31 | 269.48 | 274.66 | 0.11650 |
| 11 | 5 | 12 | 2 | 91.94 | 1187.36 | 1187.45 | 1187.55 | 289.24 | 294.51 | 299.78 | 0.11650 |
| 11 | 5 | 12 | 2 | 91.51 | 1194.34 | 1194.44 | 1194.54 | 312.68 | 318.03 | 323.40 | 0.11650 |
| 12 | 5 | 10 | 2 | 89.38 | 1229.96 | 1230.04 | 1230.12 | 193.27 | 197.99 | 202.71 | 0.05650 |
| 12 | 5 | 21 | 2 | 89.02 | 1236.08 | 1236.25 | 1236.41 | 191.70 | 201.62 | 211.54 | 0.11864 |
| 12 | 5 | 21 | 2 | 88.67 | 1242.06 | 1242.23 | 1242.40 | 181.93 | 191.84 | 201.74 | 0.11864 |
| 12 | 5 | 10 | 2 | 88.32 | 1248.24 | 1248.33 | 1248.41 | 183.02 | 187.71 | 192.41 | 0.05650 |
| 12 | 5 | 10 | 2 | 87.97 | 1254.34 | 1254.42 | 1254.50 | 177.27 | 181.94 | 186.61 | 0.05650 |
| 12 | 5 | 21 | 2 | 87.62 | 1260.44 | 1260.61 | 1260.78 | 169.65 | 179.42 | 189.19 | 0.11864 |
| 12 | 5 | 21 | 2 | 87.28 | 1266.44 | 1266.61 | 1266.78 | 154.81 | 164.52 | 174.25 | 0.11864 |
| 12 | 5 | 21 | 2 | 86.95 | 1272.36 | 1272.53 | 1272.69 | 133.94 | 143.63 | 153.32 | 0.11864 |
| 12 | 5 | 21 | 2 | 86.62 | 1278.21 | 1278.38 | 1278.55 | 108.34 | 118.00 | 127.66 | 0.11864 |
| 12 | 5 | 21 | 2 | 86.30 | 1284.06 | 1284.23 | 1284.40 | 81.39 | 91.02 | 100.65 | 0.11864 |
| 13 | 5 | 20 | 2 | 84.31 | 1320.79 | 1320.95 | 1321.11 | 357.61 | 6.69 | 15.77 | 0.09434 |
| 13 | 5 | 7 | 2 | 84.13 | 1324.28 | 1324.34 | 1324.39 | 194.14 | 197.32 | 200.49 | 0.03302 |
| 13 | 5 | 22 | 2 | 83.79 | 1330.61 | 1330.79 | 1330.97 | 189.63 | 199.57 | 209.51 | 0.10377 |
| 13 | 5 | 20 | 2 | 83.47 | 1336.79 | 1336.95 | 1337.11 | 174.86 | 183.85 | 192.83 | 0.09434 |
| 13 | 5 | 20 | 2 | 83.17 | 1342.50 | 1342.66 | 1342.82 | 131.86 | 140.77 | 149.69 | 0.09434 |
| 13 | 5 | 20 | 2 | 82.88 | 1348.21 | 1348.38 | 1348.54 | 86.65 | 95.48 | 104.31 | 0.09434 |
| 13 | 5 | 20 | 2 | 82.58 | 1353.93 | 1354.09 | 1354.25 | 38.13 | 46.86 | 55.58 | 0.09434 |
| 13 | 5 | 20 | 2 | 82.29 | 1359.70 | 1359.86 | 1360.02 | 348.17 | 356.78 | 5.38 | 0.09434 |
| 13 | 5 | 21 | 2 | 82.00 | 1365.46 | 1365.63 | 1365.80 | 293.51 | 302.41 | 311.30 | 0.09906 |
| 13 | 5 | 21 | 2 | 81.70 | 1371.26 | 1371.43 | 1371.60 | 235.55 | 244.29 | 253.02 | 0.09906 |
| 13 | 5 | 21 | 2 | 81.41 | 1377.09 | 1377.26 | 1377.43 | 174.11 | 182.70 | 191.30 | 0.09906 |
| 14 | 5 | 12 | 2 | 79.61 | 1413.97 | 1414.07 | 1414.16 | 176.29 | 180.91 | 185.53 | 0.05607 |
| 14 | 5 | 22 | 2 | 79.33 | 1419.84 | 1420.02 | 1420.19 | 96.26 | 104.71 | 113.16 | 0.10280 |
| 14 | 5 | 22 | 2 | 79.03 | 1425.96 | 1426.14 | 1426.32 | 27.45 | 35.88 | 44.32 | 0.10280 |
| 14 | 5 | 22 | 2 | 78.74 | 1432.17 | 1432.35 | 1432.52 | 322.17 | 330.57 | 338.97 | 0.10280 |
| 14 | 5 | 22 | 2 | 78.45 | 1438.42 | 1438.60 | 1438.77 | 257.62 | 265.99 | 274.36 | 0.10280 |
| 14 | 5 | 23 | 2 | 78.16 | 1444.69 | 1444.88 | 1445.06 | 192.57 | 201.26 | 209.94 | 0.10748 |
| 14 | 5 | 22 | 2 | 77.86 | 1451.02 | 1451.19 | 1451.37 | 127.48 | 135.71 | 143.94 | 0.10280 |
| 14 | 5 | 23 | 2 | 77.57 | 1457.40 | 1457.59 | 1457.77 | 62.46 | 70.98 | 79.50 | 0.10748 |
| 14 | 5 | 23 | 2 | 77.28 | 1463.88 | 1464.06 | 1464.25 | 358.03 | 6.44 | 14.86 | 0.10748 |
| 14 | 5 | 23 | 2 | 76.98 | 1470.39 | 1470.58 | 1470.76 | 291.58 | 299.88 | 308.18 | 0.10748 |
| 15 | 5 | 16 | 2 | 74.12 | 1536.27 | 1536.40 | 1536.53 | 193.55 | 198.77 | 203.99 | 0.15094 |
| 15 | 5 | 12 | 2 | 73.84 | 1543.18 | 1543.28 | 1543.38 | 113.07 | 116.97 | 120.87 | 0.11321 |
| 15 | 5 | 13 | 2 | 73.53 | 1550.56 | 1550.67 | 1550.77 | 48.28 | 52.42 | 56.57 | 0.12264 |
| 15 | 5 | 13 | 2 | 73.23 | 1557.97 | 1558.08 | 1558.18 | 337.80 | 341.83 | 345.86 | 0.12264 |
| 15 | 5 | 13 | 2 | 72.92 | 1565.47 | 1565.58 | 1565.68 | 261.01 | 264.89 | 268.78 | 0.12264 |
| 15 | 5 | 13 | 2 | 72.61 | 1573.07 | 1573.17 | 1573.28 | 179.79 | 183.63 | 187.47 | 0.12264 |
| 15 | 5 | 13 | 2 | 72.30 | 1580.78 | 1580.88 | 1580.99 | 97.17 | 100.90 | 104.62 | 0.12264 |
| 15 | 5 | 13 | 2 | 71.99 | 1588.58 | 1588.69 | 1588.79 | 12.63 | 16.29 | 19.95 | 0.12264 |
| 16 | 5 | 14 | 2 | 70.65 | 1623.33 | 1623.44 | 1623.55 | 154.51 | 158.48 | 162.45 | 0.13462 |
| 16 | 5 | 15 | 2 | 70.33 | 1631.72 | 1631.84 | 1631.96 | 89.06 | 93.29 | 97.52 | 0.14423 |
| 16 | 5 | 15 | 2 | 70.02 | 1640.24 | 1640.36 | 1640.48 | 25.33 | 29.49 | 33.64 | 0.14423 |
| 16 | 5 | 15 | 2 | 69.70 | 1648.81 | 1648.93 | 1649.05 | 320.83 | 325.04 | 329.24 | 0.14423 |
| 16 | 5 | 15 | 2 | 69.39 | 1657.53 | 1657.65 | 1657.77 | 262.81 | 266.92 | 271.03 | 0.14423 |
| 16 | 5 | 15 | 2 | 69.07 | 1666.22 | 1666.34 | 1666.46 | 185.83 | 189.75 | 193.67 | 0.14423 |
| 16 | 5 | 15 | 2 | 68.76 | 1675.05 | 1675.17 | 1675.29 | 112.04 | 115.97 | 119.90 | 0.14423 |
| 17 | 5 | 16 | 2 | 67.48 | 1711.94 | 1712.06 | 1712.19 | 238.79 | 242.97 | 247.15 | 0.13793 |
| 17 | 5 | 16 | 2 | 67.17 | 1721.31 | 1721.43 | 1721.56 | 180.12 | 184.21 | 188.29 | 0.13793 |
| 17 | 5 | 16 | 2 | 66.86 | 1730.68 | 1730.81 | 1730.93 | 107.34 | 111.24 | 115.12 | 0.13793 |
| 17 | 5 | 17 | 2 | 66.55 | 1740.20 | 1740.34 | 1740.47 | 37.48 | 41.74 | 46.00 | 0.14655 |
| 17 | 5 | 17 | 2 | 66.24 | 1749.72 | 1749.86 | 1750.00 | 330.36 | 334.43 | 338.51 | 0.14655 |
| 17 | 5 | 17 | 2 | 65.93 | 1759.46 | 1759.60 | 1759.74 | 256.14 | 260.13 | 264.11 | 0.14655 |
| 17 | 5 | 17 | 2 | 65.62 | 1769.33 | 1769.47 | 1769.61 | 180.35 | 184.24 | 188.11 | 0.14655 |
| 18 | 5 | 18 | 2 | 64.34 | 1811.74 | 1811.88 | 1812.03 | 319.20 | 323.28 | 327.35 | 0.15929 |
| 18 | 5 | 18 | 2 | 64.04 | 1822.08 | 1822.22 | 1822.37 | 249.62 | 253.72 | 257.83 | 0.15929 |
| 18 | 5 | 19 | 2 | 63.73 | 1832.66 | 1832.81 | 1832.96 | 190.81 | 195.15 | 199.48 | 0.16814 |
| 18 | 5 | 19 | 2 | 63.43 | 1843.40 | 1843.55 | 1843.70 | 133.25 | 137.52 | 141.78 | 0.16814 |
| 18 | 5 | 19 | 2 | 63.13 | 1854.34 | 1854.49 | 1854.64 | 72.50 | 76.66 | 80.82 | 0.16814 |
| 18 | 5 | 20 | 2 | 62.82 | 1865.52 | 1865.68 | 1865.84 | 20.74 | 25.23 | 29.72 | 0.17699 |
| 22 | 5 | 26 | 2 | 59.25 | 2006.62 | 2006.83 | 2007.04 | 192.71 | 198.24 | 203.76 | 0.22034 |
| 22 | 5 | 23 | 2 | 58.93 | 2020.08 | 2020.26 | 2020.45 | 171.97 | 176.55 | 181.13 | 0.19492 |
| 22 | 5 | 23 | 2 | 58.64 | 2032.58 | 2032.76 | 2032.95 | 123.22 | 127.89 | 132.57 | 0.19492 |
| 22 | 5 | 23 | 2 | 58.35 | 2045.08 | 2045.26 | 2045.45 | 82.21 | 86.92 | 91.62 | 0.19492 |
| 22 | 5 | 23 | 2 | 58.07 | 2057.58 | 2057.76 | 2057.95 | 36.33 | 41.10 | 45.87 | 0.19492 |
| 23 | 5 | 23 | 2 | 57.22 | 2095.71 | 2095.90 | 2096.08 | 255.04 | 259.58 | 264.12 | 0.19828 |
| 23 | 5 | 23 | 2 | 56.94 | 2108.48 | 2108.66 | 2108.85 | 196.85 | 201.14 | 205.44 | 0.19828 |
| 23 | 5 | 23 | 2 | 56.65 | 2121.93 | 2122.11 | 2122.30 | 149.80 | 154.19 | 158.57 | 0.19828 |
| 23 | 5 | 23 | 2 | 56.37 | 2134.70 | 2134.88 | 2135.07 | 89.02 | 93.35 | 97.67 | 0.19828 |
| 23 | 5 | 24 | 2 | 56.03 | 2150.71 | 2150.90 | 2151.09 | 98.64 | 102.94 | 107.25 | 0.20690 |
| 24 | 5 | 25 | 2 | 54.39 | 2231.01 | 2231.22 | 2231.42 | 321.64 | 325.44 | 329.24 | 0.15823 |
| 24 | 5 | 26 | 2 | 54.11 | 2245.17 | 2245.38 | 2245.59 | 230.23 | 234.20 | 238.16 | 0.16456 |
| 24 | 5 | 26 | 2 | 53.83 | 2259.46 | 2259.66 | 2259.87 | 135.89 | 139.83 | 143.77 | 0.16456 |
| 24 | 5 | 27 | 2 | 53.54 | 2274.16 | 2274.38 | 2274.60 | 51.47 | 55.55 | 59.65 | 0.17089 |
| 24 | 5 | 27 | 2 | 53.25 | 2289.24 | 2289.45 | 2289.67 | 337.20 | 341.21 | 345.21 | 0.17089 |
| 24 | 5 | 27 | 2 | 53.07 | 2298.60 | 2298.82 | 2299.03 | 148.02 | 151.97 | 155.93 | 0.17089 |
| 25 | 5 | 29 | 2 | 52.39 | 2334.72 | 2334.95 | 2335.19 | 55.95 | 60.10 | 64.22 | 0.23967 |
| 25 | 5 | 30 | 2 | 52.09 | 2350.95 | 2351.19 | 2351.44 | 335.00 | 339.05 | 343.10 | 0.24793 |
| 25 | 5 | 31 | 2 | 51.78 | 2367.53 | 2367.78 | 2368.03 | 245.91 | 249.89 | 253.87 | 0.25620 |
| 25 | 5 | 31 | 2 | 51.47 | 2384.56 | 2384.81 | 2385.06 | 151.51 | 155.31 | 159.12 | 0.25620 |
| 26 | 5 | 32 | 2 | 50.86 | 2418.83 | 2419.09 | 2419.34 | 331.24 | 335.36 | 339.48 | 0.24615 |
| 26 | 5 | 32 | 2 | 50.54 | 2436.58 | 2436.84 | 2437.09 | 255.85 | 259.98 | 264.11 | 0.24615 |
| 26 | 5 | 33 | 2 | 50.24 | 2454.21 | 2454.47 | 2454.74 | 174.16 | 178.24 | 182.32 | 0.25385 |
| 26 | 5 | 33 | 2 | 49.93 | 2471.85 | 2472.12 | 2472.38 | 75.95 | 79.87 | 83.80 | 0.25385 |
| 30 | 5 | 64 | 2 | 47.30 | 2630.97 | 2631.49 | 2632.00 | 57.39 | 65.58 | 73.77 | 0.24521 |
| 30 | 5 | 65 | 2 | 47.03 | 2648.11 | 2648.63 | 2649.16 | 324.77 | 332.80 | 340.82 | 0.24904 |
| 30 | 5 | 66 | 2 | 46.75 | 2665.70 | 2666.24 | 2666.77 | 231.08 | 239.00 | 246.92 | 0.25287 |
| 30 | 5 | 66 | 2 | 46.48 | 2683.24 | 2683.77 | 2684.30 | 127.83 | 135.48 | 143.14 | 0.25287 |
| 31 | 5 | 69 | 2 | 46.09 | 2708.55 | 2709.11 | 2709.67 | 159.18 | 168.64 | 178.13 | 0.31944 |
| 31 | 5 | 66 | 2 | 45.82 | 2726.42 | 2726.95 | 2727.49 | 94.36 | 102.89 | 111.40 | 0.30556 |
| 31 | 5 | 81 | 2 | 45.42 | 2753.32 | 2753.97 | 2754.62 | 158.79 | 169.24 | 179.70 | 0.37500 |
| 32 | 5 | 63 | 2 | 44.68 | 2803.25 | 2803.76 | 2804.27 | 279.35 | 288.12 | 296.89 | 0.25200 |
| 32 | 5 | 63 | 2 | 44.44 | 2819.92 | 2820.43 | 2820.94 | 200.02 | 208.32 | 216.62 | 0.25200 |
| 32 | 5 | 61 | 2 | 44.21 | 2836.09 | 2836.58 | 2837.07 | 108.22 | 116.41 | 124.58 | 0.24400 |
| 32 | 5 | 63 | 2 | 43.97 | 2852.92 | 2853.42 | 2853.93 | 22.09 | 30.11 | 38.11 | 0.25200 |
| 33 | 5 | 84 | 2 | 43.50 | 2886.10 | 2886.77 | 2887.45 | 179.05 | 189.56 | 200.09 | 0.34146 |
| 33 | 5 | 96 | 2 | 43.15 | 2911.01 | 2911.79 | 2912.56 | 210.64 | 223.20 | 235.82 | 0.39024 |
| 33 | 5 | 66 | 2 | 42.89 | 2929.94 | 2930.47 | 2931.00 | 167.00 | 175.96 | 184.90 | 0.26829 |
| 34 | 5 | 67 | 2 | 42.17 | 2982.70 | 2983.24 | 2983.78 | 246.51 | 254.73 | 263.01 | 0.26070 |
| 34 | 5 | 66 | 2 | 41.93 | 3000.65 | 3001.18 | 3001.71 | 175.97 | 184.87 | 193.77 | 0.25681 |
| 34 | 5 | 65 | 2 | 41.70 | 3018.00 | 3018.52 | 3019.05 | 101.45 | 109.93 | 118.43 | 0.25292 |
| 34 | 5 | 59 | 2 | 41.45 | 3036.96 | 3037.43 | 3037.91 | 44.38 | 51.90 | 59.42 | 0.22957 |
| 35 | 5 | 61 | 2 | 40.32 | 3124.15 | 3124.64 | 3125.14 | 299.01 | 306.23 | 313.46 | 0.24597 |
| 35 | 5 | 62 | 2 | 40.11 | 3140.63 | 3141.13 | 3141.63 | 181.75 | 189.07 | 196.39 | 0.25000 |
| 35 | 5 | 62 | 2 | 39.90 | 3157.23 | 3157.73 | 3158.23 | 42.02 | 48.42 | 54.85 | 0.25000 |
| 35 | 5 | 63 | 2 | 39.69 | 3173.91 | 3174.42 | 3174.93 | 271.79 | 278.76 | 285.72 | 0.25403 |
| 36 | 5 | 64 | 2 | 39.28 | 3207.61 | 3208.13 | 3208.65 | 13.19 | 20.43 | 27.68 | 0.24903 |
| 36 | 5 | 64 | 2 | 39.07 | 3224.60 | 3225.12 | 3225.64 | 250.25 | 257.45 | 264.65 | 0.24903 |
| 36 | 5 | 64 | 2 | 38.86 | 3241.72 | 3242.24 | 3242.76 | 128.91 | 136.05 | 143.17 | 0.24903 |
| 36 | 5 | 65 | 2 | 38.65 | 3258.93 | 3259.46 | 3259.98 | 353.72 | 0.14 | 6.54 | 0.25292 |
| 37 | 5 | 66 | 2 | 38.03 | 3311.21 | 3311.75 | 3312.28 | 357.89 | 5.05 | 12.19 | 0.24812 |
| 37 | 5 | 66 | 2 | 37.82 | 3328.86 | 3329.39 | 3329.92 | 232.20 | 239.27 | 246.36 | 0.24812 |
| 37 | 5 | 67 | 2 | 37.61 | 3346.60 | 3347.14 | 3347.68 | 112.48 | 119.99 | 127.50 | 0.25188 |
| 37 | 5 | 67 | 2 | 37.40 | 3364.46 | 3365.00 | 3365.54 | 13.38 | 21.75 | 30.12 | 0.25188 |
| 38 | 5 | 68 | 2 | 37.09 | 3391.15 | 3391.69 | 3392.24 | 66.36 | 74.82 | 83.29 | 0.24818 |
| 38 | 5 | 68 | 2 | 36.88 | 3409.29 | 3409.84 | 3410.39 | 339.92 | 347.73 | 355.52 | 0.24818 |
| 38 | 5 | 69 | 2 | 36.68 | 3427.55 | 3428.10 | 3428.66 | 227.88 | 235.22 | 242.57 | 0.25182 |
| 38 | 5 | 69 | 2 | 36.47 | 3445.92 | 3446.47 | 3447.03 | 110.80 | 118.04 | 125.28 | 0.25182 |
| 39 | 5 | 70 | 2 | 36.05 | 3482.96 | 3483.53 | 3484.09 | 258.02 | 266.29 | 274.58 | 0.24823 |
| 39 | 5 | 70 | 2 | 35.86 | 3500.01 | 3500.58 | 3501.14 | 154.09 | 162.72 | 171.34 | 0.24823 |
| 39 | 5 | 71 | 2 | 35.66 | 3518.83 | 3519.40 | 3519.97 | 70.43 | 78.31 | 86.14 | 0.25177 |
| 39 | 5 | 71 | 2 | 35.40 | 3541.89 | 3542.46 | 3543.03 | 5.55 | 12.54 | 19.56 | 0.25177 |
| 40 | 5 | 73 | 2 | 34.99 | 3580.16 | 3580.75 | 3581.34 | 142.81 | 150.37 | 157.94 | 0.24915 |
| 40 | 5 | 73 | 2 | 34.78 | 3599.48 | 3600.06 | 3600.65 | 49.50 | 58.19 | 66.89 | 0.24915 |
| 40 | 5 | 73 | 2 | 34.57 | 3618.93 | 3619.51 | 3620.10 | 338.09 | 346.38 | 354.66 | 0.24915 |
| 40 | 5 | 74 | 2 | 34.36 | 3638.49 | 3639.09 | 3639.68 | 265.28 | 274.10 | 282.93 | 0.25256 |
| 41 | 5 | 75 | 2 | 33.99 | 3673.88 | 3674.49 | 3675.09 | 52.62 | 61.23 | 69.77 | 0.24917 |
| 41 | 5 | 75 | 2 | 33.78 | 3693.80 | 3694.40 | 3695.01 | 326.62 | 335.20 | 343.77 | 0.24917 |
| 41 | 5 | 75 | 2 | 33.57 | 3713.82 | 3714.43 | 3715.03 | 250.33 | 258.80 | 267.25 | 0.24917 |
| 41 | 5 | 76 | 2 | 33.37 | 3733.17 | 3733.78 | 3734.39 | 161.13 | 169.90 | 178.68 | 0.25249 |
| 42 | 5 | 78 | 2 | 32.54 | 3815.05 | 3815.68 | 3816.31 | 171.45 | 178.30 | 185.16 | 0.33051 |
| 42 | 5 | 79 | 2 | 32.33 | 3835.84 | 3836.47 | 3837.11 | 70.27 | 78.55 | 86.78 | 0.33475 |
| 42 | 5 | 79 | 2 | 32.13 | 3856.76 | 3857.40 | 3858.04 | 337.48 | 345.59 | 353.68 | 0.33475 |
| 43 | 5 | 80 | 2 | 31.71 | 3898.99 | 3899.63 | 3900.28 | 70.16 | 76.85 | 83.65 | 0.33058 |
| 43 | 5 | 81 | 2 | 31.50 | 3920.29 | 3920.95 | 3921.60 | 339.58 | 350.17 | 0.85 | 0.33471 |
| 43 | 5 | 81 | 2 | 31.29 | 3941.74 | 3942.39 | 3943.05 | 295.95 | 304.62 | 313.28 | 0.33471 |
| 44 | 5 | 82 | 2 | 30.88 | 3985.03 | 3985.69 | 3986.35 | 215.26 | 225.46 | 235.77 | 0.33065 |
| 44 | 5 | 83 | 2 | 30.68 | 4005.36 | 4006.03 | 4006.70 | 203.08 | 214.94 | 226.71 | 0.33468 |
| 44 | 5 | 83 | 2 | 30.47 | 4027.34 | 4028.01 | 4028.68 | 213.27 | 223.87 | 234.43 | 0.33468 |
| 45 | 5 | 84 | 2 | 30.06 | 4071.71 | 4072.38 | 4073.06 | 132.51 | 141.75 | 150.99 | 0.33071 |
| 45 | 5 | 85 | 2 | 29.85 | 4094.08 | 4094.76 | 4095.45 | 81.94 | 91.30 | 100.63 | 0.33465 |
| 45 | 5 | 85 | 2 | 29.64 | 4116.61 | 4117.29 | 4117.98 | 22.69 | 31.60 | 40.47 | 0.33465 |
| 46 | 5 | 86 | 2 | 29.22 | 4162.10 | 4162.79 | 4163.48 | 176.73 | 183.19 | 189.66 | 0.33077 |
| 46 | 5 | 87 | 2 | 29.02 | 4185.04 | 4185.74 | 4186.44 | 64.01 | 72.36 | 80.72 | 0.33462 |
| 46 | 5 | 87 | 2 | 28.81 | 4208.14 | 4208.85 | 4209.55 | 351.01 | 359.84 | 8.67 | 0.33462 |
| 47 | 5 | 89 | 2 | 28.39 | 4254.77 | 4255.49 | 4256.21 | 219.33 | 227.93 | 236.49 | 0.33209 |
| 47 | 5 | 89 | 2 | 28.18 | 4278.33 | 4279.04 | 4279.76 | 132.08 | 140.50 | 148.96 | 0.33209 |
| 47 | 5 | 90 | 2 | 27.97 | 4302.00 | 4302.73 | 4303.45 | 52.87 | 60.72 | 68.42 | 0.33582 |
| 48 | 5 | 91 | 2 | 27.55 | 4349.81 | 4350.54 | 4351.28 | 168.97 | 177.17 | 185.35 | 0.33212 |
| 48 | 5 | 91 | 2 | 27.34 | 4373.94 | 4374.67 | 4375.40 | 66.06 | 73.56 | 81.06 | 0.33212 |
| 48 | 5 | 92 | 2 | 27.13 | 4398.22 | 4398.96 | 4399.70 | 313.14 | 321.09 | 329.07 | 0.33577 |
| 50 | 5 | 128 | 2 | 23.28 | 4870.54 | 4871.57 | 4872.60 | 304.49 | 316.03 | 327.57 | 0.49805 |
| 50 | 5 | 129 | 2 | 23.02 | 4904.19 | 4905.23 | 4906.27 | 290.02 | 300.62 | 311.26 | 0.50195 |
| 51 | 5 | 132 | 2 | 20.71 | 5213.27 | 5214.33 | 5215.39 | 222.71 | 233.43 | 244.12 | 0.49811 |
| 51 | 5 | 133 | 2 | 20.45 | 5248.14 | 5249.21 | 5250.29 | 190.03 | 199.43 | 208.84 | 0.50189 |
| 54 | 5 | 131 | 2 | 17.05 | 5741.12 | 5742.17 | 5743.23 | 344.12 | 354.10 | 4.09 | 0.51172 |
| 54 | 5 | 125 | 2 | 16.83 | 5773.81 | 5774.82 | 5775.82 | 289.05 | 298.43 | 307.81 | 0.48828 |
| 55 | 5 | 118 | 2 | 14.85 | 6082.04 | 6082.99 | 6083.94 | 330.58 | 339.46 | 348.33 | 0.50000 |
| 55 | 5 | 118 | 2 | 14.65 | 6113.06 | 6114.01 | 6114.96 | 262.35 | 271.65 | 280.98 | 0.50000 |
| 56 | 5 | 124 | 2 | 12.70 | 6432.64 | 6433.64 | 6434.64 | 310.05 | 319.09 | 328.13 | 0.49799 |
| 56 | 5 | 125 | 2 | 12.50 | 6465.46 | 6466.47 | 6467.48 | 234.10 | 242.13 | 250.09 | 0.50201 |
| 57 | 5 | 131 | 2 | 10.73 | 6768.66 | 6769.72 | 6770.78 | 197.68 | 206.47 | 215.23 | 0.50000 |
| 57 | 5 | 131 | 2 | 10.54 | 6803.04 | 6804.10 | 6805.15 | 107.57 | 115.29 | 122.99 | 0.50000 |
| 58 | 5 | 136 | 2 | 8.96 | 7084.20 | 7085.29 | 7086.39 | 230.80 | 238.80 | 246.79 | 0.50000 |
| 58 | 5 | 136 | 2 | 8.76 | 7120.10 | 7121.20 | 7122.29 | 136.77 | 145.77 | 154.79 | 0.50000 |
| 59 | 5 | 141 | 2 | 7.19 | 7412.48 | 7413.61 | 7414.75 | 150.51 | 158.60 | 166.67 | 0.49823 |
| 59 | 5 | 142 | 2 | 6.99 | 7449.64 | 7450.79 | 7451.93 | 63.35 | 71.85 | 80.42 | 0.50177 |
| 60 | 5 | 146 | 2 | 5.41 | 7752.43 | 7753.61 | 7754.79 | 138.93 | 146.71 | 154.49 | 0.50000 |
| 60 | 5 | 146 | 2 | 5.22 | 7790.77 | 7791.94 | 7793.12 | 37.07 | 44.99 | 52.89 | 0.50000 |
| 1 | 6 | 1 | 10 | 142.26 | 145.54 | 145.58 | 145.62 | 144.48 | 143.29 | 142.10 | 0.03030 |
| 1 | 6 | 1 | 10 | 142.25 | 145.70 | 145.74 | 145.78 | 139.72 | 138.53 | 137.34 | 0.03030 |
| 1 | 6 | 1 | 10 | 142.24 | 145.86 | 145.90 | 145.95 | 134.96 | 133.78 | 132.59 | 0.03030 |
| 1 | 6 | 1 | 10 | 142.23 | 146.03 | 146.07 | 146.11 | 130.21 | 129.02 | 127.83 | 0.03030 |
| 1 | 6 | 1 | 12 | 141.78 | 154.70 | 154.75 | 154.80 | 237.19 | 235.79 | 234.40 | 0.03030 |
| 1 | 6 | 1 | 12 | 141.77 | 154.90 | 154.94 | 154.99 | 231.61 | 230.21 | 228.82 | 0.03030 |
| 1 | 6 | 1 | 12 | 141.76 | 155.09 | 155.14 | 155.19 | 226.03 | 224.64 | 223.24 | 0.03030 |
| 1 | 6 | 1 | 12 | 141.75 | 155.28 | 155.33 | 155.38 | 220.46 | 219.06 | 217.67 | 0.03030 |
| 1 | 6 | 1 | 10 | 141.29 | 164.21 | 164.25 | 164.29 | 326.22 | 325.08 | 323.94 | 0.03030 |
| 1 | 6 | 1 | 10 | 141.28 | 164.37 | 164.41 | 164.45 | 321.67 | 320.54 | 319.40 | 0.03030 |
| 1 | 6 | 1 | 10 | 141.27 | 164.53 | 164.57 | 164.61 | 317.13 | 316.00 | 314.86 | 0.03030 |
| 1 | 6 | 1 | 10 | 141.26 | 164.69 | 164.73 | 164.77 | 312.59 | 311.46 | 310.32 | 0.03030 |
| 1 | 6 | 1 | 10 | 141.26 | 164.86 | 164.90 | 164.94 | 308.06 | 306.92 | 305.79 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.78 | 174.07 | 174.11 | 174.15 | 51.72 | 50.62 | 49.51 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.77 | 174.23 | 174.27 | 174.31 | 47.29 | 46.18 | 45.07 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.77 | 174.39 | 174.44 | 174.48 | 42.86 | 41.75 | 40.64 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.76 | 174.56 | 174.60 | 174.64 | 38.43 | 37.32 | 36.21 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.75 | 174.72 | 174.76 | 174.80 | 34.00 | 32.89 | 31.79 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.26 | 184.22 | 184.27 | 184.31 | 135.98 | 134.90 | 133.82 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.26 | 184.39 | 184.43 | 184.47 | 131.66 | 130.58 | 129.50 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.25 | 184.55 | 184.59 | 184.63 | 127.34 | 126.26 | 125.18 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.24 | 184.71 | 184.75 | 184.79 | 123.01 | 121.93 | 120.86 | 0.03030 |
| 1 | 6 | 1 | 10 | 140.23 | 184.87 | 184.91 | 184.95 | 118.70 | 117.62 | 116.54 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.74 | 194.58 | 194.62 | 194.66 | 222.13 | 221.07 | 220.02 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.73 | 194.74 | 194.78 | 194.82 | 217.92 | 216.87 | 215.82 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.73 | 194.90 | 194.94 | 194.98 | 213.72 | 212.67 | 211.62 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.72 | 195.06 | 195.10 | 195.14 | 209.52 | 208.47 | 207.42 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.71 | 195.22 | 195.26 | 195.30 | 205.32 | 204.27 | 203.22 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.20 | 205.38 | 205.42 | 205.46 | 305.11 | 304.09 | 303.07 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.19 | 205.54 | 205.58 | 205.62 | 301.04 | 300.02 | 299.01 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.18 | 205.70 | 205.74 | 205.78 | 296.97 | 295.96 | 294.94 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.18 | 205.86 | 205.90 | 205.94 | 292.91 | 291.90 | 290.88 | 0.03030 |
| 1 | 6 | 1 | 10 | 139.17 | 206.02 | 206.06 | 206.10 | 288.85 | 287.84 | 286.82 | 0.03030 |
| 2 | 6 | 16 | 2 | 135.21 | 287.26 | 287.39 | 287.52 | 275.18 | 272.65 | 270.12 | 0.09412 |
| 2 | 6 | 16 | 2 | 135.19 | 287.78 | 287.91 | 288.04 | 265.06 | 262.53 | 260.00 | 0.09412 |
| 2 | 6 | 21 | 2 | 134.88 | 294.21 | 294.38 | 294.55 | 140.81 | 137.59 | 134.37 | 0.12353 |
| 2 | 6 | 21 | 2 | 134.85 | 294.89 | 295.06 | 295.23 | 127.94 | 124.73 | 121.52 | 0.12353 |
| 2 | 6 | 20 | 2 | 133.90 | 315.19 | 315.35 | 315.51 | 122.58 | 119.84 | 117.10 | 0.11765 |
| 2 | 6 | 20 | 2 | 133.87 | 315.83 | 315.99 | 316.15 | 111.63 | 108.89 | 106.17 | 0.11765 |
| 2 | 6 | 20 | 2 | 133.84 | 316.48 | 316.64 | 316.80 | 100.72 | 98.00 | 95.28 | 0.11765 |
| 2 | 6 | 18 | 2 | 133.08 | 332.99 | 333.14 | 333.28 | 196.04 | 193.83 | 191.63 | 0.10588 |
| 2 | 6 | 18 | 2 | 133.05 | 333.57 | 333.72 | 333.86 | 187.23 | 185.04 | 182.84 | 0.10588 |
| 3 | 6 | 23 | 2 | 127.80 | 452.25 | 452.44 | 452.62 | 39.87 | 38.76 | 37.64 | 0.09237 |
| 3 | 6 | 23 | 2 | 127.77 | 453.00 | 453.18 | 453.37 | 35.43 | 34.32 | 33.22 | 0.09237 |
| 3 | 6 | 23 | 2 | 127.74 | 453.74 | 453.92 | 454.11 | 31.03 | 29.94 | 28.85 | 0.09237 |
| 3 | 6 | 21 | 2 | 126.78 | 476.43 | 476.60 | 476.77 | 277.47 | 276.77 | 276.07 | 0.08434 |
| 3 | 6 | 21 | 2 | 126.75 | 477.11 | 477.28 | 477.45 | 274.68 | 273.99 | 273.29 | 0.08434 |
| 3 | 6 | 21 | 2 | 126.72 | 477.78 | 477.95 | 478.12 | 271.92 | 271.23 | 270.55 | 0.08434 |
| 3 | 6 | 21 | 2 | 126.69 | 478.46 | 478.63 | 478.80 | 269.19 | 268.51 | 267.84 | 0.08434 |
| 3 | 6 | 24 | 2 | 125.39 | 509.77 | 509.97 | 510.16 | 178.33 | 177.97 | 177.62 | 0.09639 |
| 3 | 6 | 24 | 2 | 125.36 | 510.55 | 510.74 | 510.94 | 176.91 | 176.56 | 176.21 | 0.09639 |
| 3 | 6 | 24 | 2 | 125.33 | 511.32 | 511.52 | 511.71 | 175.53 | 175.19 | 174.85 | 0.09639 |
| 3 | 6 | 24 | 2 | 125.30 | 512.10 | 512.29 | 512.48 | 174.18 | 173.85 | 173.52 | 0.09639 |
| 4 | 6 | 23 | 2 | 123.88 | 547.03 | 547.22 | 547.40 | 149.97 | 150.03 | 150.09 | 0.09312 |
| 4 | 6 | 23 | 2 | 123.85 | 547.77 | 547.96 | 548.15 | 150.22 | 150.29 | 150.35 | 0.09312 |
| 4 | 6 | 23 | 2 | 123.82 | 548.52 | 548.70 | 548.89 | 150.50 | 150.57 | 150.65 | 0.09312 |
| 4 | 6 | 23 | 2 | 123.79 | 549.26 | 549.44 | 549.63 | 150.81 | 150.90 | 150.98 | 0.09312 |
| 4 | 6 | 23 | 2 | 123.76 | 550.00 | 550.19 | 550.37 | 151.16 | 151.25 | 151.34 | 0.09312 |
| 4 | 6 | 22 | 2 | 122.15 | 590.49 | 590.66 | 590.84 | 218.51 | 219.01 | 219.53 | 0.08907 |
| 4 | 6 | 22 | 2 | 122.12 | 591.20 | 591.37 | 591.55 | 220.55 | 221.07 | 221.59 | 0.08907 |
| 4 | 6 | 22 | 2 | 122.09 | 591.91 | 592.08 | 592.26 | 222.64 | 223.16 | 223.69 | 0.08907 |
| 4 | 6 | 22 | 2 | 122.07 | 592.62 | 592.79 | 592.97 | 224.74 | 225.27 | 225.80 | 0.08907 |
| 4 | 6 | 22 | 2 | 122.04 | 593.33 | 593.50 | 593.68 | 226.87 | 227.41 | 227.95 | 0.08907 |
| 4 | 6 | 22 | 2 | 122.01 | 594.04 | 594.21 | 594.39 | 229.03 | 229.58 | 230.13 | 0.08907 |
| 8 | 6 | 22 | 2 | 112.48 | 853.73 | 853.91 | 854.09 | 246.47 | 249.33 | 252.20 | 0.11828 |
| 8 | 6 | 22 | 2 | 112.46 | 854.44 | 854.62 | 854.80 | 257.93 | 260.80 | 263.68 | 0.11828 |
| 8 | 6 | 22 | 2 | 112.43 | 855.15 | 855.33 | 855.51 | 269.42 | 272.30 | 275.18 | 0.11828 |
| 8 | 6 | 21 | 2 | 111.70 | 876.67 | 876.83 | 877.00 | 268.75 | 271.66 | 274.58 | 0.11290 |
| 8 | 6 | 21 | 2 | 111.68 | 877.34 | 877.51 | 877.68 | 280.41 | 283.33 | 286.25 | 0.11290 |
| 8 | 6 | 21 | 2 | 111.66 | 878.02 | 878.19 | 878.36 | 292.09 | 295.01 | 297.93 | 0.11290 |
| 8 | 6 | 19 | 2 | 110.99 | 898.14 | 898.29 | 898.45 | 288.34 | 291.12 | 293.90 | 0.10215 |
| 8 | 6 | 19 | 2 | 110.97 | 898.75 | 898.91 | 899.06 | 299.48 | 302.26 | 305.05 | 0.10215 |
| 8 | 6 | 19 | 2 | 110.95 | 899.37 | 899.52 | 899.67 | 310.63 | 313.42 | 316.21 | 0.10215 |
| 9 | 6 | 19 | 2 | 106.80 | 971.29 | 971.45 | 971.60 | 72.05 | 76.16 | 80.27 | 0.12925 |
| 9 | 6 | 19 | 2 | 106.76 | 971.91 | 972.06 | 972.21 | 88.52 | 92.64 | 96.77 | 0.12925 |
| 9 | 6 | 15 | 2 | 106.22 | 979.34 | 979.46 | 979.58 | 292.94 | 296.34 | 299.75 | 0.10204 |
| 9 | 6 | 21 | 2 | 105.50 | 989.17 | 989.34 | 989.51 | 219.71 | 224.86 | 230.00 | 0.14286 |
| 9 | 6 | 21 | 2 | 104.58 | 1001.57 | 1001.74 | 1001.91 | 251.01 | 256.52 | 262.04 | 0.14286 |
| 9 | 6 | 15 | 2 | 103.75 | 1012.99 | 1013.11 | 1013.23 | 270.12 | 274.21 | 278.31 | 0.10204 |
| 9 | 6 | 15 | 2 | 103.71 | 1013.47 | 1013.59 | 1013.71 | 286.51 | 290.61 | 294.72 | 0.10204 |
| 9 | 6 | 22 | 2 | 102.84 | 1025.34 | 1025.52 | 1025.69 | 335.48 | 341.69 | 347.90 | 0.14966 |
| 10 | 6 | 19 | 2 | 100.47 | 1058.30 | 1058.45 | 1058.61 | 94.55 | 100.36 | 106.17 | 0.16239 |
| 10 | 6 | 18 | 2 | 99.75 | 1068.44 | 1068.58 | 1068.73 | 124.06 | 129.72 | 135.38 | 0.15385 |
| 10 | 6 | 17 | 2 | 99.07 | 1078.23 | 1078.37 | 1078.51 | 151.74 | 157.24 | 162.76 | 0.14530 |
| 10 | 6 | 17 | 2 | 98.33 | 1089.02 | 1089.15 | 1089.29 | 232.63 | 238.34 | 244.06 | 0.14530 |
| 10 | 6 | 16 | 2 | 97.70 | 1098.14 | 1098.27 | 1098.40 | 258.76 | 264.30 | 269.84 | 0.13675 |
| 10 | 6 | 15 | 2 | 97.10 | 1107.04 | 1107.16 | 1107.28 | 286.34 | 291.69 | 297.04 | 0.12821 |
| 10 | 6 | 15 | 2 | 96.52 | 1115.71 | 1115.83 | 1115.95 | 314.51 | 319.99 | 325.48 | 0.12821 |
| 11 | 6 | 14 | 2 | 94.89 | 1140.51 | 1140.62 | 1140.73 | 38.70 | 44.18 | 49.66 | 0.13462 |
| 11 | 6 | 14 | 2 | 94.38 | 1148.44 | 1148.55 | 1148.67 | 67.75 | 73.34 | 78.93 | 0.13462 |
| 11 | 6 | 13 | 2 | 93.88 | 1156.25 | 1156.35 | 1156.46 | 98.31 | 103.62 | 108.92 | 0.12500 |
| 11 | 6 | 13 | 2 | 93.24 | 1166.41 | 1166.51 | 1166.62 | 257.93 | 263.36 | 268.79 | 0.12500 |
| 11 | 6 | 13 | 2 | 92.78 | 1173.79 | 1173.90 | 1174.00 | 283.41 | 288.93 | 294.45 | 0.12500 |
| 11 | 6 | 13 | 2 | 92.32 | 1181.12 | 1181.22 | 1181.33 | 312.21 | 317.84 | 323.46 | 0.12500 |
| 11 | 6 | 12 | 2 | 91.88 | 1188.25 | 1188.34 | 1188.44 | 337.59 | 342.87 | 348.14 | 0.11538 |
| 11 | 6 | 12 | 2 | 91.46 | 1195.21 | 1195.31 | 1195.41 | 0.96 | 6.33 | 11.70 | 0.11538 |
| 12 | 6 | 10 | 2 | 89.34 | 1230.73 | 1230.81 | 1230.89 | 238.58 | 243.31 | 248.02 | 0.06024 |
| 12 | 6 | 10 | 2 | 88.99 | 1236.75 | 1236.83 | 1236.91 | 230.90 | 235.63 | 240.35 | 0.06024 |
| 12 | 6 | 21 | 2 | 88.63 | 1242.81 | 1242.98 | 1243.15 | 225.77 | 235.68 | 245.57 | 0.12651 |
| 12 | 6 | 10 | 2 | 88.28 | 1249.02 | 1249.10 | 1249.18 | 228.09 | 232.78 | 237.47 | 0.06024 |
| 12 | 6 | 10 | 2 | 87.93 | 1255.11 | 1255.19 | 1255.27 | 221.64 | 226.31 | 230.98 | 0.06024 |
| 12 | 6 | 21 | 2 | 87.58 | 1261.19 | 1261.36 | 1261.53 | 212.89 | 222.65 | 232.41 | 0.12651 |
| 12 | 6 | 21 | 2 | 87.24 | 1267.19 | 1267.36 | 1267.53 | 197.84 | 207.55 | 217.28 | 0.12651 |
| 12 | 6 | 21 | 2 | 86.91 | 1273.10 | 1273.27 | 1273.44 | 176.38 | 186.06 | 195.75 | 0.12651 |
| 12 | 6 | 21 | 2 | 86.58 | 1278.94 | 1279.11 | 1279.28 | 150.19 | 159.86 | 169.51 | 0.12651 |
| 12 | 6 | 21 | 2 | 86.26 | 1284.78 | 1284.95 | 1285.12 | 122.67 | 132.30 | 141.92 | 0.12651 |
| 13 | 6 | 20 | 2 | 84.27 | 1321.50 | 1321.66 | 1321.82 | 37.57 | 46.64 | 55.72 | 0.08889 |
| 13 | 6 | 20 | 2 | 84.10 | 1324.68 | 1324.84 | 1325.00 | 216.37 | 225.44 | 234.51 | 0.08889 |
| 13 | 6 | 22 | 2 | 83.75 | 1331.39 | 1331.57 | 1331.75 | 233.46 | 243.40 | 253.33 | 0.09778 |
| 13 | 6 | 20 | 2 | 83.43 | 1337.50 | 1337.66 | 1337.82 | 214.40 | 223.37 | 232.34 | 0.08889 |
| 13 | 6 | 20 | 2 | 83.13 | 1343.22 | 1343.38 | 1343.54 | 171.54 | 180.45 | 189.36 | 0.08889 |
| 13 | 6 | 20 | 2 | 82.84 | 1348.93 | 1349.09 | 1349.26 | 125.95 | 134.76 | 143.59 | 0.08889 |
| 13 | 6 | 20 | 2 | 82.55 | 1354.64 | 1354.80 | 1354.96 | 76.52 | 85.23 | 93.94 | 0.08889 |
| 13 | 6 | 21 | 2 | 82.25 | 1360.42 | 1360.58 | 1360.75 | 26.44 | 35.47 | 44.49 | 0.09333 |
| 13 | 6 | 20 | 2 | 81.96 | 1366.19 | 1366.35 | 1366.51 | 331.59 | 340.04 | 348.49 | 0.08889 |
| 13 | 6 | 21 | 2 | 81.67 | 1372.00 | 1372.17 | 1372.33 | 273.38 | 282.10 | 290.82 | 0.09333 |
| 13 | 6 | 21 | 2 | 81.38 | 1377.82 | 1377.99 | 1378.16 | 210.91 | 219.49 | 228.07 | 0.09333 |
| 14 | 6 | 12 | 2 | 79.57 | 1414.89 | 1414.99 | 1415.08 | 220.19 | 224.80 | 229.42 | 0.05607 |
| 14 | 6 | 22 | 2 | 79.29 | 1420.61 | 1420.78 | 1420.96 | 132.75 | 141.19 | 149.64 | 0.10280 |
| 14 | 6 | 22 | 2 | 79.00 | 1426.73 | 1426.90 | 1427.08 | 63.86 | 72.28 | 80.71 | 0.10280 |
| 14 | 6 | 22 | 2 | 78.71 | 1432.94 | 1433.12 | 1433.30 | 358.82 | 7.22 | 15.62 | 0.10280 |
| 14 | 6 | 22 | 2 | 78.41 | 1439.20 | 1439.38 | 1439.56 | 294.51 | 302.87 | 311.23 | 0.10280 |
| 14 | 6 | 23 | 2 | 78.12 | 1445.48 | 1445.67 | 1445.85 | 229.56 | 238.24 | 246.91 | 0.10748 |
| 14 | 6 | 22 | 2 | 77.83 | 1451.81 | 1451.98 | 1452.16 | 164.15 | 172.37 | 180.59 | 0.10280 |
| 14 | 6 | 23 | 2 | 77.53 | 1458.20 | 1458.39 | 1458.57 | 99.10 | 107.61 | 116.11 | 0.10748 |
| 14 | 6 | 23 | 2 | 77.24 | 1464.68 | 1464.87 | 1465.05 | 34.58 | 42.99 | 51.38 | 0.10748 |
| 14 | 6 | 23 | 2 | 76.94 | 1471.21 | 1471.39 | 1471.58 | 328.00 | 336.29 | 344.57 | 0.10748 |
| 15 | 6 | 16 | 2 | 74.08 | 1537.39 | 1537.52 | 1537.65 | 238.89 | 244.11 | 249.32 | 0.14953 |
| 15 | 6 | 13 | 2 | 73.80 | 1544.08 | 1544.18 | 1544.29 | 149.15 | 153.38 | 157.61 | 0.12150 |
| 15 | 6 | 13 | 2 | 73.49 | 1551.49 | 1551.59 | 1551.70 | 84.90 | 89.03 | 93.16 | 0.12150 |
| 15 | 6 | 13 | 2 | 73.19 | 1558.90 | 1559.01 | 1559.11 | 13.38 | 17.39 | 21.40 | 0.12150 |
| 15 | 6 | 13 | 2 | 72.88 | 1566.42 | 1566.52 | 1566.63 | 295.94 | 299.82 | 303.70 | 0.12150 |
| 15 | 6 | 13 | 2 | 72.58 | 1574.02 | 1574.13 | 1574.23 | 214.62 | 218.44 | 222.26 | 0.12150 |
| 15 | 6 | 13 | 2 | 72.27 | 1581.74 | 1581.84 | 1581.95 | 131.21 | 134.93 | 138.64 | 0.12150 |
| 15 | 6 | 13 | 2 | 71.96 | 1589.56 | 1589.66 | 1589.77 | 46.64 | 50.29 | 53.94 | 0.12150 |
| 16 | 6 | 15 | 2 | 70.61 | 1624.35 | 1624.47 | 1624.59 | 190.52 | 194.78 | 199.03 | 0.14286 |
| 16 | 6 | 15 | 2 | 70.29 | 1632.78 | 1632.90 | 1633.02 | 126.04 | 130.28 | 134.52 | 0.14286 |
| 16 | 6 | 15 | 2 | 69.98 | 1641.31 | 1641.43 | 1641.55 | 62.12 | 66.27 | 70.41 | 0.14286 |
| 16 | 6 | 15 | 2 | 69.66 | 1649.88 | 1650.00 | 1650.12 | 358.16 | 2.38 | 6.60 | 0.14286 |
| 16 | 6 | 15 | 2 | 69.35 | 1658.61 | 1658.73 | 1658.85 | 299.15 | 303.17 | 307.19 | 0.14286 |
| 16 | 6 | 15 | 2 | 69.04 | 1667.31 | 1667.43 | 1667.55 | 221.10 | 225.02 | 228.94 | 0.14286 |
| 16 | 6 | 15 | 2 | 68.72 | 1676.15 | 1676.27 | 1676.40 | 147.95 | 151.88 | 155.82 | 0.14286 |
| 17 | 6 | 17 | 2 | 67.44 | 1713.08 | 1713.22 | 1713.35 | 275.93 | 280.37 | 284.82 | 0.14407 |
| 17 | 6 | 16 | 2 | 67.13 | 1722.48 | 1722.60 | 1722.73 | 217.07 | 221.13 | 225.18 | 0.13559 |
| 17 | 6 | 17 | 2 | 66.82 | 1731.83 | 1731.97 | 1732.10 | 142.12 | 146.26 | 150.39 | 0.14407 |
| 17 | 6 | 17 | 2 | 66.51 | 1741.39 | 1741.52 | 1741.66 | 74.31 | 78.57 | 82.84 | 0.14407 |
| 17 | 6 | 17 | 2 | 66.20 | 1750.91 | 1751.05 | 1751.18 | 5.56 | 9.62 | 13.67 | 0.14407 |
| 17 | 6 | 17 | 2 | 65.89 | 1760.67 | 1760.81 | 1760.95 | 291.28 | 295.25 | 299.23 | 0.14407 |
| 17 | 6 | 17 | 2 | 65.58 | 1770.56 | 1770.70 | 1770.83 | 215.03 | 218.91 | 222.79 | 0.14407 |
| 18 | 6 | 18 | 2 | 64.61 | 1802.68 | 1802.83 | 1802.97 | 61.22 | 65.32 | 69.44 | 0.16216 |
| 18 | 6 | 18 | 2 | 64.30 | 1813.03 | 1813.17 | 1813.32 | 355.40 | 359.47 | 3.53 | 0.16216 |
| 18 | 6 | 18 | 2 | 64.00 | 1823.37 | 1823.52 | 1823.66 | 286.41 | 290.53 | 294.65 | 0.16216 |
| 18 | 6 | 19 | 2 | 63.70 | 1833.97 | 1834.12 | 1834.28 | 227.99 | 232.32 | 236.65 | 0.17117 |
| 18 | 6 | 19 | 2 | 63.39 | 1844.74 | 1844.89 | 1845.04 | 170.47 | 174.72 | 178.97 | 0.17117 |
| 18 | 6 | 19 | 2 | 63.09 | 1855.70 | 1855.85 | 1856.01 | 109.59 | 113.78 | 117.96 | 0.17117 |
| 22 | 6 | 26 | 2 | 59.21 | 2008.39 | 2008.60 | 2008.80 | 239.13 | 244.59 | 250.04 | 0.22034 |
| 22 | 6 | 23 | 2 | 58.90 | 2021.64 | 2021.83 | 2022.01 | 210.60 | 215.18 | 219.75 | 0.19492 |
| 22 | 6 | 23 | 2 | 58.61 | 2034.14 | 2034.33 | 2034.51 | 162.71 | 167.40 | 172.08 | 0.19492 |
| 22 | 6 | 23 | 2 | 58.32 | 2046.64 | 2046.83 | 2047.01 | 121.67 | 126.30 | 130.93 | 0.19492 |
| 22 | 6 | 23 | 2 | 58.03 | 2059.14 | 2059.33 | 2059.51 | 76.42 | 81.17 | 85.91 | 0.19492 |
| 23 | 6 | 23 | 2 | 57.18 | 2097.31 | 2097.49 | 2097.68 | 293.80 | 298.27 | 302.72 | 0.19828 |
| 23 | 6 | 23 | 2 | 56.90 | 2110.07 | 2110.26 | 2110.44 | 233.82 | 238.13 | 242.43 | 0.19828 |
| 23 | 6 | 23 | 2 | 56.61 | 2123.52 | 2123.71 | 2123.89 | 187.58 | 191.97 | 196.37 | 0.19828 |
| 23 | 6 | 23 | 2 | 56.34 | 2136.30 | 2136.48 | 2136.67 | 126.23 | 130.56 | 134.88 | 0.19828 |
| 23 | 6 | 24 | 2 | 56.00 | 2152.33 | 2152.53 | 2152.72 | 134.67 | 138.93 | 143.17 | 0.20690 |
| 24 | 6 | 25 | 2 | 54.35 | 2232.76 | 2232.97 | 2233.17 | 354.73 | 358.55 | 2.37 | 0.15723 |
| 24 | 6 | 26 | 2 | 54.07 | 2246.96 | 2247.17 | 2247.38 | 263.81 | 267.75 | 271.68 | 0.16352 |
| 24 | 6 | 26 | 2 | 53.79 | 2261.25 | 2261.46 | 2261.66 | 169.51 | 173.45 | 177.40 | 0.16352 |
| 24 | 6 | 27 | 2 | 53.50 | 2275.99 | 2276.21 | 2276.43 | 85.94 | 90.05 | 94.19 | 0.16981 |
| 24 | 6 | 27 | 2 | 53.21 | 2291.11 | 2291.33 | 2291.55 | 11.70 | 15.69 | 19.67 | 0.16981 |
| 24 | 6 | 28 | 2 | 53.04 | 2300.45 | 2300.68 | 2300.90 | 181.66 | 185.73 | 189.81 | 0.17610 |
| 25 | 6 | 29 | 2 | 52.65 | 2320.91 | 2321.14 | 2321.37 | 175.07 | 179.01 | 182.96 | 0.24370 |
| 25 | 6 | 29 | 2 | 52.35 | 2336.70 | 2336.93 | 2337.16 | 90.75 | 94.83 | 98.91 | 0.24370 |
| 25 | 6 | 30 | 2 | 52.05 | 2352.98 | 2353.22 | 2353.46 | 8.79 | 12.82 | 16.84 | 0.25210 |
| 25 | 6 | 31 | 2 | 51.74 | 2369.61 | 2369.86 | 2370.11 | 279.01 | 282.97 | 286.94 | 0.26050 |
| 26 | 6 | 32 | 2 | 50.82 | 2420.97 | 2421.23 | 2421.49 | 5.52 | 9.65 | 13.77 | 0.24615 |
| 26 | 6 | 33 | 2 | 50.51 | 2438.76 | 2439.03 | 2439.30 | 290.82 | 295.09 | 299.34 | 0.25385 |
| 26 | 6 | 32 | 2 | 50.20 | 2456.43 | 2456.69 | 2456.95 | 208.17 | 212.10 | 216.02 | 0.24615 |
| 26 | 6 | 33 | 2 | 49.89 | 2474.06 | 2474.33 | 2474.59 | 108.74 | 112.71 | 116.67 | 0.25385 |
| 30 | 6 | 64 | 2 | 47.54 | 2615.97 | 2616.49 | 2617.01 | 180.96 | 188.92 | 196.88 | 0.28194 |
| 30 | 6 | 65 | 2 | 47.26 | 2633.11 | 2633.63 | 2634.16 | 91.22 | 99.47 | 107.71 | 0.28634 |
| 30 | 6 | 65 | 2 | 47.00 | 2650.25 | 2650.78 | 2651.30 | 357.57 | 5.58 | 13.57 | 0.28634 |
| 30 | 6 | 33 | 2 | 46.73 | 2667.63 | 2667.90 | 2668.16 | 259.74 | 263.69 | 267.64 | 0.14537 |
| 31 | 6 | 69 | 2 | 46.06 | 2710.83 | 2711.38 | 2711.94 | 197.91 | 207.39 | 216.87 | 0.31797 |
| 31 | 6 | 67 | 2 | 45.79 | 2728.62 | 2729.16 | 2729.70 | 129.53 | 138.12 | 146.69 | 0.30876 |
| 31 | 6 | 81 | 2 | 45.38 | 2755.98 | 2756.63 | 2757.28 | 201.44 | 211.95 | 222.48 | 0.37327 |
| 32 | 6 | 63 | 2 | 44.65 | 2805.33 | 2805.84 | 2806.35 | 315.31 | 324.10 | 332.87 | 0.25200 |
| 32 | 6 | 63 | 2 | 44.41 | 2822.00 | 2822.51 | 2823.02 | 234.03 | 242.37 | 250.73 | 0.25200 |
| 32 | 6 | 61 | 2 | 44.18 | 2838.11 | 2838.60 | 2839.10 | 141.89 | 150.07 | 158.24 | 0.24400 |
| 32 | 6 | 63 | 2 | 43.94 | 2855.01 | 2855.51 | 2856.02 | 55.01 | 62.99 | 70.97 | 0.25200 |
| 33 | 6 | 84 | 2 | 43.46 | 2888.88 | 2889.56 | 2890.23 | 222.26 | 232.78 | 243.32 | 0.33871 |
| 33 | 6 | 97 | 2 | 43.10 | 2914.18 | 2914.96 | 2915.74 | 262.22 | 275.11 | 288.07 | 0.39113 |
| 33 | 6 | 67 | 2 | 42.86 | 2932.14 | 2932.68 | 2933.22 | 203.93 | 212.90 | 221.84 | 0.27016 |
| 34 | 6 | 66 | 2 | 42.14 | 2984.91 | 2985.45 | 2985.98 | 280.47 | 288.78 | 297.11 | 0.25681 |
| 34 | 6 | 67 | 2 | 41.90 | 3002.85 | 3003.39 | 3003.93 | 212.74 | 221.73 | 230.71 | 0.26070 |
| 34 | 6 | 64 | 2 | 41.67 | 3020.15 | 3020.66 | 3021.18 | 136.15 | 144.49 | 152.83 | 0.24903 |
| 34 | 6 | 60 | 2 | 41.42 | 3038.93 | 3039.42 | 3039.90 | 75.62 | 83.27 | 90.93 | 0.23346 |
| 35 | 6 | 62 | 2 | 40.29 | 3126.19 | 3126.69 | 3127.19 | 328.99 | 336.32 | 343.66 | 0.24900 |
| 35 | 6 | 62 | 2 | 40.08 | 3142.70 | 3143.20 | 3143.70 | 211.81 | 218.99 | 226.12 | 0.24900 |
| 35 | 6 | 62 | 2 | 39.88 | 3159.30 | 3159.80 | 3160.30 | 68.85 | 75.44 | 82.08 | 0.24900 |
| 35 | 6 | 63 | 2 | 39.67 | 3176.00 | 3176.50 | 3177.01 | 300.29 | 307.22 | 314.15 | 0.25301 |
| 36 | 6 | 64 | 2 | 39.25 | 3209.73 | 3210.24 | 3210.76 | 42.86 | 50.09 | 57.33 | 0.24903 |
| 36 | 6 | 64 | 2 | 39.04 | 3226.73 | 3227.24 | 3227.76 | 279.87 | 287.08 | 294.32 | 0.24903 |
| 36 | 6 | 64 | 2 | 38.83 | 3243.86 | 3244.38 | 3244.89 | 158.40 | 165.49 | 172.57 | 0.24903 |
| 36 | 6 | 65 | 2 | 38.62 | 3261.08 | 3261.60 | 3262.13 | 19.95 | 26.35 | 32.76 | 0.25292 |
| 37 | 6 | 65 | 2 | 38.21 | 3295.89 | 3296.42 | 3296.94 | 137.09 | 145.10 | 153.11 | 0.24621 |
| 37 | 6 | 66 | 2 | 38.00 | 3313.41 | 3313.94 | 3314.47 | 27.30 | 34.40 | 41.51 | 0.25000 |
| 37 | 6 | 66 | 2 | 37.79 | 3331.06 | 3331.59 | 3332.12 | 261.52 | 268.64 | 275.75 | 0.25000 |
| 37 | 6 | 67 | 2 | 37.58 | 3348.82 | 3349.36 | 3349.90 | 143.35 | 150.90 | 158.46 | 0.25379 |
| 38 | 6 | 68 | 2 | 37.07 | 3393.40 | 3393.94 | 3394.49 | 101.15 | 109.66 | 118.16 | 0.24818 |
| 38 | 6 | 68 | 2 | 36.86 | 3411.56 | 3412.10 | 3412.65 | 12.07 | 19.79 | 27.48 | 0.24818 |
| 38 | 6 | 69 | 2 | 36.65 | 3429.82 | 3430.38 | 3430.93 | 257.94 | 265.32 | 272.71 | 0.25182 |
| 38 | 6 | 69 | 2 | 36.44 | 3448.21 | 3448.76 | 3449.32 | 140.53 | 147.73 | 154.93 | 0.25182 |
| 39 | 6 | 70 | 2 | 36.03 | 3485.29 | 3485.85 | 3486.42 | 292.21 | 300.60 | 308.98 | 0.24648 |
| 39 | 6 | 71 | 2 | 35.84 | 3502.34 | 3502.91 | 3503.49 | 189.68 | 198.38 | 207.08 | 0.25000 |
| 39 | 6 | 71 | 2 | 35.63 | 3521.17 | 3521.74 | 3522.32 | 102.43 | 110.12 | 117.78 | 0.25000 |
| 39 | 6 | 72 | 2 | 35.38 | 3544.25 | 3544.83 | 3545.41 | 34.52 | 41.73 | 48.96 | 0.25352 |
| 40 | 6 | 72 | 2 | 34.96 | 3582.57 | 3583.15 | 3583.73 | 173.73 | 181.24 | 188.77 | 0.24658 |
| 40 | 6 | 73 | 2 | 34.75 | 3601.89 | 3602.48 | 3603.06 | 85.12 | 93.91 | 102.73 | 0.25000 |
| 40 | 6 | 73 | 2 | 34.55 | 3621.35 | 3621.94 | 3622.53 | 12.29 | 20.60 | 28.95 | 0.25000 |
| 40 | 6 | 74 | 2 | 34.34 | 3640.93 | 3641.53 | 3642.13 | 301.40 | 310.22 | 319.03 | 0.25342 |
| 41 | 6 | 74 | 2 | 34.17 | 3656.60 | 3657.20 | 3657.79 | 165.90 | 174.13 | 182.37 | 0.24667 |
| 41 | 6 | 75 | 2 | 33.96 | 3676.35 | 3676.95 | 3677.56 | 87.40 | 95.77 | 104.08 | 0.25000 |
| 41 | 6 | 75 | 2 | 33.76 | 3696.28 | 3696.89 | 3697.49 | 1.89 | 10.50 | 19.11 | 0.25000 |
| 41 | 6 | 76 | 2 | 33.55 | 3716.31 | 3716.93 | 3717.54 | 285.14 | 293.65 | 302.15 | 0.25333 |
| 42 | 6 | 77 | 2 | 32.72 | 3796.99 | 3797.61 | 3798.24 | 319.41 | 327.28 | 335.14 | 0.32906 |
| 42 | 6 | 78 | 2 | 32.52 | 3817.62 | 3818.25 | 3818.88 | 199.61 | 206.58 | 213.62 | 0.33333 |
| 42 | 6 | 79 | 2 | 32.31 | 3838.43 | 3839.07 | 3839.71 | 103.75 | 111.93 | 120.09 | 0.33761 |
| 43 | 6 | 80 | 2 | 31.68 | 3901.63 | 3902.28 | 3902.92 | 98.15 | 105.17 | 112.26 | 0.33195 |
| 43 | 6 | 80 | 2 | 31.48 | 3922.95 | 3923.60 | 3924.24 | 22.94 | 33.39 | 43.71 | 0.33195 |
| 43 | 6 | 81 | 2 | 31.27 | 3944.42 | 3945.07 | 3945.73 | 331.30 | 339.91 | 348.48 | 0.33610 |
| 44 | 6 | 82 | 2 | 30.85 | 3987.74 | 3988.40 | 3989.06 | 257.86 | 268.58 | 279.42 | 0.33198 |
| 44 | 6 | 82 | 2 | 30.66 | 4008.09 | 4008.75 | 4009.41 | 251.02 | 262.55 | 274.05 | 0.33198 |
| 44 | 6 | 83 | 2 | 30.45 | 4030.08 | 4030.75 | 4031.42 | 256.50 | 266.99 | 277.44 | 0.33603 |
| 45 | 6 | 84 | 2 | 30.03 | 4074.48 | 4075.16 | 4075.84 | 170.33 | 179.59 | 188.87 | 0.33071 |
| 45 | 6 | 85 | 2 | 29.82 | 4096.87 | 4097.55 | 4098.24 | 119.89 | 129.17 | 138.42 | 0.33465 |
| 45 | 6 | 85 | 2 | 29.62 | 4119.41 | 4120.10 | 4120.78 | 58.95 | 67.73 | 76.49 | 0.33465 |
| 46 | 6 | 86 | 2 | 29.20 | 4164.94 | 4165.64 | 4166.33 | 203.38 | 209.97 | 216.61 | 0.33077 |
| 46 | 6 | 87 | 2 | 28.99 | 4187.90 | 4188.60 | 4189.31 | 98.26 | 106.79 | 115.36 | 0.33462 |
| 46 | 6 | 87 | 2 | 28.78 | 4211.02 | 4211.72 | 4212.43 | 27.33 | 36.24 | 45.16 | 0.33462 |
| 47 | 6 | 89 | 2 | 28.36 | 4257.69 | 4258.41 | 4259.13 | 254.08 | 262.51 | 270.90 | 0.33209 |
| 47 | 6 | 89 | 2 | 28.15 | 4281.26 | 4281.98 | 4282.70 | 166.64 | 175.11 | 183.57 | 0.33209 |
| 47 | 6 | 90 | 2 | 27.94 | 4304.95 | 4305.68 | 4306.40 | 83.86 | 91.09 | 98.23 | 0.33582 |
| 48 | 6 | 91 | 2 | 27.53 | 4352.79 | 4353.53 | 4354.26 | 202.23 | 210.34 | 218.44 | 0.33212 |
| 48 | 6 | 91 | 2 | 27.32 | 4376.94 | 4377.68 | 4378.41 | 96.67 | 104.06 | 111.42 | 0.33212 |
| 48 | 6 | 92 | 2 | 27.11 | 4401.24 | 4401.98 | 4402.73 | 345.65 | 353.66 | 1.66 | 0.33577 |
| 50 | 6 | 128 | 2 | 23.24 | 4874.74 | 4875.77 | 4876.81 | 351.40 | 2.85 | 14.21 | 0.49805 |
| 50 | 6 | 129 | 2 | 22.98 | 4908.40 | 4909.44 | 4910.48 | 333.07 | 343.72 | 354.38 | 0.50195 |
| 51 | 6 | 132 | 2 | 20.67 | 5217.59 | 5218.65 | 5219.72 | 265.96 | 276.50 | 287.01 | 0.49811 |
| 51 | 6 | 133 | 2 | 20.42 | 5252.49 | 5253.56 | 5254.63 | 228.20 | 237.66 | 247.14 | 0.50189 |
| 54 | 6 | 131 | 2 | 17.02 | 5745.41 | 5746.46 | 5747.52 | 24.63 | 34.60 | 44.54 | 0.51172 |
| 54 | 6 | 125 | 2 | 16.80 | 5777.89 | 5778.90 | 5779.90 | 327.06 | 336.47 | 345.88 | 0.48828 |
| 55 | 6 | 117 | 2 | 14.82 | 6085.89 | 6086.84 | 6087.78 | 6.60 | 15.44 | 24.29 | 0.49787 |
| 55 | 6 | 118 | 2 | 14.63 | 6116.92 | 6117.88 | 6118.83 | 300.30 | 309.73 | 319.17 | 0.50213 |
| 56 | 6 | 124 | 2 | 12.67 | 6436.71 | 6437.71 | 6438.71 | 346.77 | 355.74 | 4.67 | 0.49799 |
| 56 | 6 | 125 | 2 | 12.48 | 6469.56 | 6470.57 | 6471.57 | 266.25 | 273.97 | 281.62 | 0.50201 |
| 57 | 6 | 130 | 2 | 10.90 | 6738.68 | 6739.73 | 6740.78 | 309.27 | 317.71 | 326.18 | 0.49808 |
| 57 | 6 | 131 | 2 | 10.71 | 6772.95 | 6774.00 | 6775.06 | 233.10 | 241.68 | 250.26 | 0.50192 |
| 58 | 6 | 136 | 2 | 8.94 | 7088.64 | 7089.74 | 7090.83 | 262.99 | 270.83 | 278.65 | 0.50000 |
| 58 | 6 | 136 | 2 | 8.74 | 7124.56 | 7125.65 | 7126.75 | 173.53 | 182.56 | 191.60 | 0.50000 |
| 59 | 6 | 141 | 2 | 7.16 | 7417.09 | 7418.23 | 7419.36 | 183.33 | 191.45 | 199.58 | 0.49823 |
| 59 | 6 | 142 | 2 | 6.97 | 7454.28 | 7455.42 | 7456.57 | 97.99 | 106.59 | 115.19 | 0.50177 |
| 60 | 6 | 146 | 2 | 5.39 | 7757.20 | 7758.37 | 7759.55 | 170.42 | 178.23 | 186.02 | 0.50000 |
| 60 | 6 | 146 | 2 | 5.19 | 7795.54 | 7796.72 | 7797.90 | 69.03 | 76.84 | 84.61 | 0.50000 |
| 1 | 7 | 1 | 10 | 142.17 | 147.11 | 147.15 | 147.19 | 98.39 | 97.21 | 96.02 | 0.07143 |
| 1 | 7 | 1 | 10 | 142.17 | 147.27 | 147.31 | 147.35 | 93.65 | 92.47 | 91.28 | 0.07143 |
| 1 | 7 | 1 | 10 | 141.69 | 156.38 | 156.42 | 156.46 | 188.92 | 187.76 | 186.60 | 0.07143 |
| 1 | 7 | 1 | 10 | 141.68 | 156.54 | 156.58 | 156.62 | 184.29 | 183.13 | 181.97 | 0.07143 |
| 1 | 7 | 1 | 10 | 141.20 | 165.96 | 166.00 | 166.04 | 277.03 | 275.90 | 274.77 | 0.07143 |
| 1 | 7 | 1 | 10 | 141.19 | 166.12 | 166.16 | 166.20 | 272.50 | 271.37 | 270.24 | 0.07143 |
| 1 | 7 | 1 | 10 | 140.69 | 175.86 | 175.90 | 175.94 | 2.61 | 1.51 | 0.40 | 0.07143 |
| 1 | 7 | 1 | 10 | 140.68 | 176.02 | 176.06 | 176.10 | 358.20 | 357.09 | 355.99 | 0.07143 |
| 1 | 7 | 1 | 12 | 140.17 | 186.03 | 186.08 | 186.13 | 87.65 | 86.36 | 85.07 | 0.07143 |
| 1 | 7 | 1 | 12 | 140.16 | 186.22 | 186.27 | 186.32 | 82.49 | 81.20 | 79.91 | 0.07143 |
| 1 | 7 | 1 | 12 | 139.65 | 196.41 | 196.46 | 196.51 | 174.52 | 173.26 | 172.01 | 0.07143 |
| 1 | 7 | 1 | 12 | 139.64 | 196.60 | 196.65 | 196.70 | 169.50 | 168.25 | 166.99 | 0.07143 |
| 1 | 7 | 1 | 12 | 139.11 | 207.26 | 207.30 | 207.35 | 257.86 | 256.64 | 255.43 | 0.07143 |
| 1 | 7 | 1 | 12 | 139.10 | 207.45 | 207.50 | 207.55 | 253.01 | 251.79 | 250.58 | 0.07143 |
| 2 | 7 | 13 | 2 | 135.10 | 289.65 | 289.75 | 289.86 | 228.54 | 226.51 | 224.47 | 0.19403 |
| 2 | 7 | 17 | 2 | 134.74 | 297.37 | 297.51 | 297.65 | 81.12 | 78.55 | 75.98 | 0.25373 |
| 2 | 7 | 22 | 2 | 133.72 | 319.19 | 319.36 | 319.54 | 55.39 | 52.45 | 49.51 | 0.32836 |
| 2 | 7 | 15 | 2 | 132.96 | 335.72 | 335.85 | 335.97 | 154.90 | 153.10 | 151.30 | 0.22388 |
| 3 | 7 | 16 | 2 | 127.61 | 456.74 | 456.87 | 457.00 | 13.70 | 12.97 | 12.24 | 0.14035 |
| 3 | 7 | 16 | 2 | 127.59 | 457.25 | 457.38 | 457.51 | 10.79 | 10.07 | 9.35 | 0.14035 |
| 3 | 7 | 19 | 2 | 126.56 | 481.59 | 481.74 | 481.90 | 257.06 | 256.48 | 255.91 | 0.16667 |
| 3 | 7 | 19 | 2 | 126.54 | 482.20 | 482.36 | 482.51 | 254.77 | 254.20 | 253.64 | 0.16667 |
| 3 | 7 | 22 | 2 | 125.15 | 515.75 | 515.93 | 516.10 | 168.28 | 168.01 | 167.75 | 0.19298 |
| 3 | 7 | 22 | 2 | 125.12 | 516.46 | 516.64 | 516.81 | 167.23 | 166.98 | 166.73 | 0.19298 |
| 4 | 7 | 19 | 2 | 123.61 | 553.76 | 553.91 | 554.06 | 153.36 | 153.47 | 153.59 | 0.15833 |
| 4 | 7 | 19 | 2 | 123.59 | 554.37 | 554.52 | 554.68 | 153.81 | 153.92 | 154.03 | 0.15833 |
| 4 | 7 | 19 | 2 | 123.56 | 554.98 | 555.14 | 555.29 | 154.26 | 154.38 | 154.49 | 0.15833 |
| 4 | 7 | 21 | 2 | 121.85 | 598.20 | 598.37 | 598.54 | 242.35 | 242.91 | 243.47 | 0.17500 |
| 4 | 7 | 21 | 2 | 121.82 | 598.88 | 599.05 | 599.22 | 244.61 | 245.18 | 245.75 | 0.17500 |
| 4 | 7 | 21 | 2 | 121.79 | 599.56 | 599.73 | 599.90 | 246.90 | 247.47 | 248.05 | 0.17500 |
| 8 | 7 | 15 | 2 | 113.21 | 832.58 | 832.70 | 832.82 | 275.31 | 277.14 | 278.98 | 0.16304 |
| 8 | 7 | 15 | 2 | 113.19 | 833.06 | 833.18 | 833.30 | 282.65 | 284.48 | 286.32 | 0.16304 |
| 8 | 7 | 16 | 2 | 112.33 | 858.10 | 858.23 | 858.36 | 317.49 | 319.60 | 321.71 | 0.17391 |
| 8 | 7 | 16 | 2 | 112.32 | 858.62 | 858.75 | 858.88 | 325.93 | 328.05 | 330.16 | 0.17391 |
| 8 | 7 | 15 | 2 | 111.57 | 880.79 | 880.91 | 881.03 | 339.94 | 342.04 | 344.15 | 0.16304 |
| 8 | 7 | 15 | 2 | 111.55 | 881.27 | 881.39 | 881.51 | 348.36 | 350.46 | 352.57 | 0.16304 |
| 9 | 7 | 16 | 2 | 106.60 | 974.17 | 974.29 | 974.42 | 149.66 | 153.18 | 156.70 | 0.20779 |
| 9 | 7 | 15 | 2 | 106.08 | 981.26 | 981.38 | 981.50 | 347.26 | 350.71 | 354.16 | 0.19481 |
| 9 | 7 | 11 | 2 | 105.35 | 991.20 | 991.29 | 991.37 | 281.58 | 284.32 | 287.05 | 0.14286 |
| 9 | 7 | 11 | 2 | 104.44 | 1003.62 | 1003.71 | 1003.80 | 317.93 | 320.84 | 323.75 | 0.14286 |
| 9 | 7 | 12 | 2 | 103.58 | 1015.31 | 1015.40 | 1015.50 | 349.02 | 352.32 | 355.62 | 0.15584 |
| 9 | 7 | 12 | 2 | 102.69 | 1027.45 | 1027.55 | 1027.65 | 49.56 | 52.96 | 56.37 | 0.15584 |
| 10 | 7 | 21 | 2 | 101.08 | 1049.70 | 1049.86 | 1050.03 | 132.25 | 138.52 | 144.80 | 0.15672 |
| 10 | 7 | 21 | 2 | 100.33 | 1060.22 | 1060.39 | 1060.56 | 167.46 | 173.92 | 180.37 | 0.15672 |
| 10 | 7 | 20 | 2 | 99.62 | 1070.28 | 1070.44 | 1070.61 | 196.27 | 202.60 | 208.93 | 0.14925 |
| 10 | 7 | 19 | 2 | 98.94 | 1080.01 | 1080.16 | 1080.31 | 223.24 | 229.43 | 235.63 | 0.14179 |
| 10 | 7 | 18 | 2 | 98.21 | 1090.73 | 1090.87 | 1091.02 | 304.04 | 310.12 | 316.21 | 0.13433 |
| 10 | 7 | 18 | 2 | 97.59 | 1099.80 | 1099.94 | 1100.09 | 329.96 | 336.23 | 342.50 | 0.13433 |
| 10 | 7 | 17 | 2 | 96.99 | 1108.63 | 1108.76 | 1108.90 | 356.72 | 2.82 | 8.91 | 0.12687 |
| 11 | 7 | 15 | 2 | 94.80 | 1141.96 | 1142.08 | 1142.20 | 109.25 | 115.15 | 121.04 | 0.11719 |
| 11 | 7 | 15 | 2 | 94.29 | 1149.88 | 1150.00 | 1150.12 | 138.98 | 145.00 | 151.02 | 0.11719 |
| 11 | 7 | 15 | 2 | 93.79 | 1157.64 | 1157.76 | 1157.88 | 168.92 | 175.06 | 181.19 | 0.11719 |
| 11 | 7 | 14 | 2 | 93.61 | 1160.50 | 1160.62 | 1160.73 | 314.55 | 320.31 | 326.07 | 0.10938 |
| 11 | 7 | 14 | 2 | 93.15 | 1167.79 | 1167.90 | 1168.01 | 329.39 | 335.25 | 341.11 | 0.10938 |
| 11 | 7 | 14 | 2 | 92.69 | 1175.15 | 1175.26 | 1175.38 | 354.85 | 0.81 | 6.77 | 0.10938 |
| 11 | 7 | 14 | 2 | 92.24 | 1182.45 | 1182.56 | 1182.67 | 23.63 | 29.70 | 35.77 | 0.10938 |
| 11 | 7 | 14 | 2 | 91.80 | 1189.54 | 1189.65 | 1189.76 | 48.06 | 54.23 | 60.41 | 0.10938 |
| 11 | 7 | 13 | 2 | 91.38 | 1196.48 | 1196.58 | 1196.69 | 71.33 | 77.17 | 83.01 | 0.10156 |
| 12 | 7 | 12 | 2 | 89.27 | 1231.84 | 1231.94 | 1232.04 | 303.76 | 309.42 | 315.09 | 0.10435 |
| 12 | 7 | 11 | 2 | 88.92 | 1237.83 | 1237.92 | 1238.01 | 294.66 | 299.85 | 305.05 | 0.09565 |
| 12 | 7 | 11 | 2 | 88.58 | 1243.80 | 1243.89 | 1243.98 | 283.74 | 288.93 | 294.11 | 0.09565 |
| 12 | 7 | 11 | 2 | 88.21 | 1250.13 | 1250.22 | 1250.31 | 292.81 | 297.97 | 303.13 | 0.09565 |
| 12 | 7 | 11 | 2 | 87.86 | 1256.21 | 1256.30 | 1256.39 | 285.58 | 290.71 | 295.84 | 0.09565 |
| 12 | 7 | 11 | 2 | 87.52 | 1262.19 | 1262.28 | 1262.36 | 270.49 | 275.60 | 280.71 | 0.09565 |
| 12 | 7 | 11 | 2 | 87.19 | 1268.19 | 1268.28 | 1268.36 | 255.19 | 260.28 | 265.37 | 0.09565 |
| 12 | 7 | 11 | 2 | 86.86 | 1274.08 | 1274.17 | 1274.26 | 232.62 | 237.69 | 242.77 | 0.09565 |
| 12 | 7 | 11 | 2 | 86.53 | 1279.93 | 1280.02 | 1280.11 | 206.28 | 211.34 | 216.40 | 0.09565 |
| 12 | 7 | 11 | 2 | 86.21 | 1285.75 | 1285.84 | 1285.93 | 177.65 | 182.70 | 187.74 | 0.09565 |
| 12 | 7 | 4 | 2 | 86.11 | 1287.60 | 1287.63 | 1287.66 | 282.58 | 284.41 | 286.24 | 0.03478 |
| 13 | 7 | 11 | 2 | 84.22 | 1322.45 | 1322.54 | 1322.63 | 91.11 | 96.11 | 101.11 | 0.09322 |
| 13 | 7 | 10 | 2 | 84.06 | 1325.63 | 1325.71 | 1325.79 | 269.87 | 274.40 | 278.93 | 0.08475 |
| 13 | 7 | 12 | 2 | 83.70 | 1332.43 | 1332.52 | 1332.62 | 291.26 | 296.66 | 302.08 | 0.10169 |
| 13 | 7 | 10 | 2 | 83.39 | 1338.45 | 1338.53 | 1338.61 | 267.33 | 271.82 | 276.30 | 0.08475 |
| 13 | 7 | 10 | 2 | 83.09 | 1344.17 | 1344.25 | 1344.33 | 224.08 | 228.53 | 232.97 | 0.08475 |
| 13 | 7 | 11 | 2 | 82.79 | 1349.88 | 1349.96 | 1350.05 | 177.50 | 182.35 | 187.20 | 0.09322 |
| 13 | 7 | 10 | 2 | 82.50 | 1355.59 | 1355.67 | 1355.75 | 127.90 | 132.25 | 136.60 | 0.08475 |
| 13 | 7 | 11 | 2 | 82.21 | 1361.38 | 1361.46 | 1361.55 | 77.51 | 82.23 | 86.94 | 0.09322 |
| 13 | 7 | 11 | 2 | 81.91 | 1367.15 | 1367.24 | 1367.33 | 21.84 | 26.47 | 31.11 | 0.09322 |
| 13 | 7 | 11 | 2 | 81.62 | 1372.96 | 1373.05 | 1373.14 | 323.16 | 327.73 | 332.28 | 0.09322 |
| 13 | 7 | 11 | 2 | 81.33 | 1378.79 | 1378.87 | 1378.96 | 259.87 | 264.35 | 268.84 | 0.09322 |
| 14 | 7 | 14 | 2 | 79.85 | 1409.06 | 1409.17 | 1409.28 | 301.51 | 306.91 | 312.31 | 0.11570 |
| 14 | 7 | 14 | 2 | 79.50 | 1416.20 | 1416.32 | 1416.43 | 282.92 | 288.31 | 293.68 | 0.11570 |
| 14 | 7 | 11 | 2 | 79.24 | 1421.63 | 1421.72 | 1421.81 | 181.49 | 185.72 | 189.94 | 0.09091 |
| 14 | 7 | 11 | 2 | 78.95 | 1427.75 | 1427.84 | 1427.93 | 112.49 | 116.70 | 120.92 | 0.09091 |
| 14 | 7 | 11 | 2 | 78.66 | 1433.98 | 1434.06 | 1434.15 | 47.66 | 51.86 | 56.06 | 0.09091 |
| 14 | 7 | 12 | 2 | 78.37 | 1440.24 | 1440.34 | 1440.44 | 343.50 | 348.06 | 352.61 | 0.09917 |
| 14 | 7 | 12 | 2 | 78.07 | 1446.54 | 1446.64 | 1446.73 | 278.96 | 283.48 | 288.00 | 0.09917 |
| 14 | 7 | 12 | 2 | 77.78 | 1452.85 | 1452.95 | 1453.05 | 212.70 | 217.17 | 221.65 | 0.09917 |
| 14 | 7 | 12 | 2 | 77.49 | 1459.26 | 1459.36 | 1459.46 | 147.88 | 152.30 | 156.73 | 0.09917 |
| 14 | 7 | 12 | 2 | 77.20 | 1465.76 | 1465.85 | 1465.95 | 83.11 | 87.49 | 91.85 | 0.09917 |
| 15 | 7 | 17 | 2 | 74.01 | 1539.02 | 1539.16 | 1539.30 | 304.77 | 310.31 | 315.86 | 0.14530 |
| 15 | 7 | 14 | 2 | 73.74 | 1545.41 | 1545.52 | 1545.64 | 202.75 | 207.29 | 211.83 | 0.11966 |
| 15 | 7 | 14 | 2 | 73.44 | 1552.84 | 1552.96 | 1553.07 | 138.22 | 142.65 | 147.08 | 0.11966 |
| 15 | 7 | 14 | 2 | 73.13 | 1560.25 | 1560.36 | 1560.47 | 64.78 | 69.07 | 73.36 | 0.11966 |
| 15 | 7 | 14 | 2 | 72.83 | 1567.78 | 1567.89 | 1568.00 | 346.22 | 350.37 | 354.52 | 0.11966 |
| 15 | 7 | 14 | 2 | 72.52 | 1575.39 | 1575.50 | 1575.62 | 264.53 | 268.63 | 272.73 | 0.11966 |
| 15 | 7 | 15 | 2 | 72.21 | 1583.12 | 1583.25 | 1583.37 | 180.39 | 184.68 | 188.97 | 0.12821 |
| 15 | 7 | 15 | 2 | 71.90 | 1590.95 | 1591.08 | 1591.20 | 95.23 | 99.45 | 103.66 | 0.12821 |
| 16 | 7 | 16 | 2 | 70.55 | 1625.87 | 1626.00 | 1626.13 | 244.14 | 248.68 | 253.21 | 0.13675 |
| 16 | 7 | 16 | 2 | 70.24 | 1634.32 | 1634.45 | 1634.58 | 180.00 | 184.52 | 189.03 | 0.13675 |
| 16 | 7 | 17 | 2 | 69.92 | 1642.87 | 1643.00 | 1643.14 | 115.54 | 120.26 | 124.98 | 0.14530 |
| 16 | 7 | 17 | 2 | 69.61 | 1651.44 | 1651.57 | 1651.71 | 52.40 | 57.19 | 61.98 | 0.14530 |
| 16 | 7 | 17 | 2 | 69.29 | 1660.18 | 1660.32 | 1660.45 | 350.82 | 355.30 | 359.75 | 0.14530 |
| 16 | 7 | 17 | 2 | 68.98 | 1668.87 | 1669.01 | 1669.15 | 271.79 | 276.23 | 280.67 | 0.14530 |
| 16 | 7 | 17 | 2 | 68.66 | 1677.75 | 1677.89 | 1678.02 | 199.88 | 204.34 | 208.81 | 0.14530 |
| 17 | 7 | 18 | 2 | 67.39 | 1714.79 | 1714.94 | 1715.08 | 331.29 | 335.99 | 340.69 | 0.13846 |
| 17 | 7 | 18 | 2 | 67.07 | 1724.17 | 1724.31 | 1724.46 | 270.05 | 274.55 | 279.04 | 0.13846 |
| 17 | 7 | 18 | 2 | 66.76 | 1733.54 | 1733.68 | 1733.83 | 193.58 | 197.94 | 202.30 | 0.13846 |
| 17 | 7 | 19 | 2 | 66.45 | 1743.12 | 1743.27 | 1743.43 | 128.11 | 132.86 | 137.60 | 0.14615 |
| 17 | 7 | 19 | 2 | 66.14 | 1752.64 | 1752.80 | 1752.95 | 56.65 | 61.16 | 65.66 | 0.14615 |
| 17 | 7 | 19 | 2 | 65.84 | 1762.42 | 1762.58 | 1762.73 | 342.02 | 346.46 | 350.90 | 0.14615 |
| 17 | 7 | 19 | 2 | 65.53 | 1772.33 | 1772.49 | 1772.64 | 265.23 | 269.58 | 273.92 | 0.14615 |
| 18 | 7 | 20 | 2 | 64.55 | 1804.55 | 1804.71 | 1804.88 | 114.28 | 118.88 | 123.47 | 0.16000 |
| 18 | 7 | 20 | 2 | 64.24 | 1814.90 | 1815.06 | 1815.22 | 47.69 | 52.20 | 56.70 | 0.16000 |
| 18 | 7 | 21 | 2 | 63.94 | 1825.24 | 1825.41 | 1825.58 | 339.46 | 344.31 | 349.15 | 0.16800 |
| 18 | 7 | 21 | 2 | 63.64 | 1835.89 | 1836.06 | 1836.23 | 282.20 | 286.98 | 291.76 | 0.16800 |
| 18 | 7 | 21 | 2 | 63.34 | 1846.68 | 1846.85 | 1847.02 | 224.27 | 228.94 | 233.61 | 0.16800 |
| 18 | 7 | 22 | 2 | 63.03 | 1857.67 | 1857.85 | 1858.02 | 163.53 | 168.41 | 173.30 | 0.17600 |
| 22 | 7 | 29 | 2 | 59.15 | 2010.97 | 2011.20 | 2011.43 | 305.51 | 311.43 | 317.32 | 0.22481 |
| 22 | 7 | 25 | 2 | 58.84 | 2023.92 | 2024.12 | 2024.32 | 266.89 | 271.89 | 276.89 | 0.19380 |
| 22 | 7 | 25 | 2 | 58.55 | 2036.42 | 2036.63 | 2036.83 | 220.49 | 225.61 | 230.72 | 0.19380 |
| 22 | 7 | 25 | 2 | 58.27 | 2048.92 | 2049.12 | 2049.33 | 178.26 | 183.22 | 188.20 | 0.19380 |
| 22 | 7 | 25 | 2 | 57.98 | 2061.42 | 2061.62 | 2061.83 | 134.66 | 139.80 | 144.93 | 0.19380 |
| 23 | 7 | 26 | 2 | 57.13 | 2099.62 | 2099.83 | 2100.04 | 349.16 | 354.14 | 359.12 | 0.20000 |
| 23 | 7 | 26 | 2 | 56.85 | 2112.39 | 2112.60 | 2112.81 | 287.51 | 292.37 | 297.24 | 0.20000 |
| 23 | 7 | 26 | 2 | 56.56 | 2125.84 | 2126.05 | 2126.26 | 242.30 | 247.26 | 252.20 | 0.20000 |
| 23 | 7 | 26 | 2 | 56.29 | 2138.61 | 2138.82 | 2139.03 | 180.22 | 185.11 | 190.00 | 0.20000 |
| 23 | 7 | 26 | 2 | 55.95 | 2154.72 | 2154.93 | 2155.14 | 186.70 | 191.24 | 195.79 | 0.20000 |
| 24 | 7 | 28 | 2 | 54.30 | 2235.29 | 2235.51 | 2235.74 | 42.66 | 46.94 | 51.23 | 0.15819 |
| 24 | 7 | 29 | 2 | 54.02 | 2249.55 | 2249.79 | 2250.02 | 312.25 | 316.59 | 320.92 | 0.16384 |
| 24 | 7 | 29 | 2 | 53.74 | 2263.83 | 2264.07 | 2264.30 | 218.17 | 222.57 | 226.96 | 0.16384 |
| 24 | 7 | 30 | 2 | 53.45 | 2278.66 | 2278.91 | 2279.15 | 136.81 | 141.47 | 146.12 | 0.16949 |
| 24 | 7 | 30 | 2 | 53.16 | 2293.83 | 2294.07 | 2294.32 | 61.30 | 65.71 | 70.12 | 0.16949 |
| 24 | 7 | 31 | 2 | 52.98 | 2303.25 | 2303.50 | 2303.75 | 231.89 | 236.35 | 240.80 | 0.17514 |
| 25 | 7 | 32 | 2 | 52.59 | 2323.78 | 2324.03 | 2324.29 | 223.69 | 228.07 | 232.48 | 0.24427 |
| 25 | 7 | 32 | 2 | 52.30 | 2339.57 | 2339.82 | 2340.08 | 140.71 | 145.19 | 149.67 | 0.24427 |
| 25 | 7 | 33 | 2 | 51.99 | 2355.93 | 2356.19 | 2356.46 | 57.71 | 62.10 | 66.48 | 0.25191 |
| 25 | 7 | 34 | 2 | 51.69 | 2372.65 | 2372.93 | 2373.20 | 326.99 | 331.30 | 335.60 | 0.25954 |
| 26 | 7 | 35 | 2 | 50.76 | 2424.09 | 2424.38 | 2424.66 | 55.48 | 60.01 | 64.52 | 0.24476 |
| 26 | 7 | 36 | 2 | 50.45 | 2441.98 | 2442.27 | 2442.56 | 342.37 | 347.02 | 351.66 | 0.25175 |
| 26 | 7 | 36 | 2 | 50.14 | 2459.63 | 2459.92 | 2460.21 | 256.62 | 260.99 | 265.34 | 0.25175 |
| 26 | 7 | 36 | 2 | 49.84 | 2477.28 | 2477.57 | 2477.86 | 156.87 | 161.22 | 165.57 | 0.25175 |
| 30 | 7 | 35 | 2 | 47.50 | 2618.83 | 2619.11 | 2619.39 | 225.04 | 229.43 | 233.81 | 0.24823 |
| 30 | 7 | 35 | 2 | 47.22 | 2635.97 | 2636.25 | 2636.54 | 136.26 | 140.69 | 145.13 | 0.24823 |
| 30 | 7 | 35 | 2 | 46.95 | 2653.12 | 2653.40 | 2653.68 | 41.17 | 45.45 | 49.74 | 0.24823 |
| 30 | 7 | 36 | 2 | 46.68 | 2670.85 | 2671.14 | 2671.43 | 307.27 | 311.52 | 315.79 | 0.25532 |
| 31 | 7 | 37 | 2 | 46.02 | 2713.86 | 2714.16 | 2714.46 | 249.23 | 254.22 | 259.19 | 0.23871 |
| 31 | 7 | 36 | 2 | 45.75 | 2731.57 | 2731.86 | 2732.15 | 176.18 | 180.77 | 185.36 | 0.23226 |
| 31 | 7 | 38 | 2 | 45.66 | 2737.24 | 2737.55 | 2737.86 | 265.67 | 270.51 | 275.34 | 0.24516 |
| 31 | 7 | 44 | 2 | 45.33 | 2759.52 | 2759.87 | 2760.23 | 258.66 | 264.42 | 270.19 | 0.28387 |
| 32 | 7 | 34 | 2 | 44.62 | 2808.11 | 2808.39 | 2808.66 | 3.24 | 7.94 | 12.65 | 0.25185 |
| 32 | 7 | 34 | 2 | 44.37 | 2824.78 | 2825.06 | 2825.33 | 279.92 | 284.46 | 289.02 | 0.25185 |
| 32 | 7 | 33 | 2 | 44.14 | 2840.81 | 2841.08 | 2841.34 | 186.73 | 191.13 | 195.52 | 0.24444 |
| 32 | 7 | 34 | 2 | 43.90 | 2857.78 | 2858.05 | 2858.33 | 98.46 | 102.75 | 107.01 | 0.25185 |
| 33 | 7 | 46 | 2 | 43.41 | 2892.60 | 2892.97 | 2893.34 | 280.09 | 285.84 | 291.60 | 0.34074 |
| 33 | 7 | 53 | 2 | 43.05 | 2918.43 | 2918.86 | 2919.29 | 333.11 | 340.30 | 347.50 | 0.39259 |
| 33 | 7 | 36 | 2 | 42.82 | 2935.08 | 2935.38 | 2935.67 | 252.34 | 257.05 | 261.73 | 0.26667 |
| 34 | 7 | 36 | 2 | 42.10 | 2987.86 | 2988.15 | 2988.44 | 326.86 | 331.49 | 336.13 | 0.25899 |
| 34 | 7 | 36 | 2 | 41.86 | 3005.79 | 3006.08 | 3006.37 | 261.58 | 266.37 | 271.17 | 0.25899 |
| 34 | 7 | 35 | 2 | 41.63 | 3023.00 | 3023.28 | 3023.56 | 182.22 | 186.76 | 191.32 | 0.25180 |
| 34 | 7 | 32 | 2 | 41.39 | 3041.58 | 3041.84 | 3042.09 | 117.49 | 121.57 | 125.66 | 0.23022 |
| 35 | 7 | 33 | 2 | 40.26 | 3128.92 | 3129.18 | 3129.45 | 9.01 | 12.92 | 16.83 | 0.24444 |
| 35 | 7 | 34 | 2 | 40.05 | 3145.44 | 3145.72 | 3145.99 | 250.53 | 254.31 | 258.08 | 0.25185 |
| 35 | 7 | 34 | 2 | 39.84 | 3162.06 | 3162.33 | 3162.61 | 105.91 | 109.67 | 113.43 | 0.25185 |
| 35 | 7 | 34 | 2 | 39.64 | 3178.78 | 3179.05 | 3179.33 | 338.17 | 341.89 | 345.62 | 0.25185 |
| 36 | 7 | 35 | 2 | 39.22 | 3212.54 | 3212.82 | 3213.11 | 82.27 | 86.20 | 90.14 | 0.25000 |
| 36 | 7 | 35 | 2 | 39.01 | 3229.56 | 3229.84 | 3230.12 | 319.54 | 323.49 | 327.46 | 0.25000 |
| 36 | 7 | 35 | 2 | 38.80 | 3246.71 | 3246.99 | 3247.27 | 197.31 | 201.13 | 204.94 | 0.25000 |
| 36 | 7 | 35 | 2 | 38.59 | 3263.95 | 3264.23 | 3264.51 | 55.24 | 58.75 | 62.25 | 0.25000 |
| 37 | 7 | 36 | 2 | 38.18 | 3298.79 | 3299.08 | 3299.37 | 181.27 | 185.67 | 190.06 | 0.25000 |
| 37 | 7 | 36 | 2 | 37.97 | 3316.33 | 3316.62 | 3316.91 | 66.32 | 70.21 | 74.11 | 0.25000 |
| 37 | 7 | 36 | 2 | 37.76 | 3333.99 | 3334.28 | 3334.57 | 300.82 | 304.71 | 308.59 | 0.25000 |
| 37 | 7 | 36 | 2 | 37.55 | 3351.78 | 3352.07 | 3352.36 | 184.92 | 189.03 | 193.16 | 0.25000 |
| 38 | 7 | 37 | 2 | 37.03 | 3396.40 | 3396.69 | 3396.99 | 147.69 | 152.32 | 156.94 | 0.24832 |
| 38 | 7 | 37 | 2 | 36.83 | 3414.57 | 3414.87 | 3415.17 | 54.23 | 58.37 | 62.49 | 0.24832 |
| 38 | 7 | 37 | 2 | 36.62 | 3432.87 | 3433.17 | 3433.47 | 298.43 | 302.40 | 306.35 | 0.24832 |
| 38 | 7 | 38 | 2 | 36.41 | 3451.25 | 3451.56 | 3451.87 | 180.01 | 183.99 | 187.97 | 0.25503 |
| 39 | 7 | 38 | 2 | 36.00 | 3488.38 | 3488.68 | 3488.99 | 338.21 | 342.80 | 347.38 | 0.24675 |
| 39 | 7 | 38 | 2 | 35.81 | 3505.46 | 3505.76 | 3506.07 | 236.71 | 241.28 | 245.85 | 0.24675 |
| 39 | 7 | 39 | 2 | 35.60 | 3524.31 | 3524.62 | 3524.94 | 144.18 | 148.32 | 152.45 | 0.25325 |
| 39 | 7 | 39 | 2 | 35.34 | 3547.40 | 3547.72 | 3548.03 | 74.14 | 78.17 | 82.22 | 0.25325 |
| 40 | 7 | 39 | 2 | 35.14 | 3566.58 | 3566.90 | 3567.21 | 328.11 | 332.16 | 336.20 | 0.24684 |
| 40 | 7 | 39 | 2 | 34.93 | 3585.76 | 3586.07 | 3586.39 | 215.38 | 219.53 | 223.71 | 0.24684 |
| 40 | 7 | 40 | 2 | 34.72 | 3605.10 | 3605.43 | 3605.75 | 133.63 | 138.54 | 143.48 | 0.25316 |
| 40 | 7 | 40 | 2 | 34.51 | 3624.58 | 3624.90 | 3625.22 | 58.19 | 62.84 | 67.50 | 0.25316 |
| 41 | 7 | 41 | 2 | 34.14 | 3659.87 | 3660.20 | 3660.53 | 211.01 | 215.60 | 220.20 | 0.25000 |
| 41 | 7 | 41 | 2 | 33.93 | 3679.64 | 3679.97 | 3680.30 | 132.40 | 136.86 | 141.32 | 0.25000 |
| 41 | 7 | 41 | 2 | 33.72 | 3699.59 | 3699.92 | 3700.25 | 48.96 | 53.66 | 58.35 | 0.25000 |
| 41 | 7 | 41 | 2 | 33.52 | 3719.65 | 3719.98 | 3720.31 | 331.37 | 335.94 | 340.50 | 0.25000 |
| 42 | 7 | 42 | 2 | 32.69 | 3800.41 | 3800.75 | 3801.09 | 2.48 | 6.70 | 10.90 | 0.32813 |
| 42 | 7 | 43 | 2 | 32.49 | 3821.06 | 3821.40 | 3821.75 | 238.50 | 242.57 | 246.65 | 0.33594 |
| 42 | 7 | 43 | 2 | 32.28 | 3841.89 | 3842.23 | 3842.58 | 147.97 | 152.40 | 156.82 | 0.33594 |
| 43 | 7 | 44 | 2 | 31.65 | 3905.16 | 3905.51 | 3905.87 | 137.36 | 141.43 | 145.52 | 0.33333 |
| 43 | 7 | 44 | 2 | 31.44 | 3926.49 | 3926.85 | 3927.20 | 78.41 | 83.74 | 89.03 | 0.33333 |
| 43 | 7 | 44 | 2 | 31.24 | 3947.99 | 3948.35 | 3948.70 | 18.43 | 23.16 | 27.93 | 0.33333 |
| 44 | 7 | 45 | 2 | 30.82 | 3991.34 | 3991.70 | 3992.07 | 317.64 | 323.82 | 330.03 | 0.33333 |
| 44 | 7 | 45 | 2 | 30.63 | 4011.72 | 4012.08 | 4012.44 | 314.08 | 320.36 | 326.65 | 0.33333 |
| 44 | 7 | 45 | 2 | 30.42 | 4033.73 | 4034.10 | 4034.46 | 313.36 | 318.96 | 324.55 | 0.33333 |
| 45 | 7 | 46 | 2 | 30.00 | 4078.17 | 4078.55 | 4078.92 | 220.91 | 226.01 | 231.11 | 0.33094 |
| 45 | 7 | 46 | 2 | 29.79 | 4100.59 | 4100.96 | 4101.33 | 170.14 | 175.12 | 180.10 | 0.33094 |
| 45 | 7 | 47 | 2 | 29.58 | 4123.16 | 4123.53 | 4123.91 | 106.61 | 111.42 | 116.22 | 0.33813 |
| 46 | 7 | 47 | 2 | 29.17 | 4168.74 | 4169.12 | 4169.50 | 240.39 | 244.24 | 248.13 | 0.33099 |
| 46 | 7 | 47 | 2 | 28.96 | 4191.73 | 4192.10 | 4192.48 | 145.28 | 150.02 | 154.74 | 0.33099 |
| 46 | 7 | 48 | 2 | 28.75 | 4214.86 | 4215.25 | 4215.63 | 76.41 | 81.44 | 86.46 | 0.33803 |
| 47 | 7 | 48 | 2 | 28.33 | 4261.58 | 4261.96 | 4262.35 | 299.26 | 303.68 | 308.12 | 0.32877 |
| 47 | 7 | 49 | 2 | 28.12 | 4285.17 | 4285.57 | 4285.96 | 212.82 | 217.47 | 222.15 | 0.33562 |
| 47 | 7 | 49 | 2 | 27.91 | 4308.89 | 4309.28 | 4309.68 | 121.98 | 125.68 | 129.36 | 0.33562 |
| 48 | 7 | 50 | 2 | 27.49 | 4356.78 | 4357.19 | 4357.59 | 245.95 | 250.28 | 254.61 | 0.33333 |
| 48 | 7 | 50 | 2 | 27.28 | 4380.95 | 4381.36 | 4381.76 | 136.74 | 140.75 | 144.74 | 0.33333 |
| 48 | 7 | 50 | 2 | 27.07 | 4405.28 | 4405.68 | 4406.09 | 29.28 | 33.63 | 37.99 | 0.33333 |
| 50 | 7 | 70 | 2 | 23.20 | 4880.35 | 4880.91 | 4881.47 | 52.41 | 58.36 | 64.27 | 0.49645 |
| 50 | 7 | 71 | 2 | 22.94 | 4914.02 | 4914.59 | 4915.17 | 30.45 | 36.25 | 42.03 | 0.50355 |
| 51 | 7 | 72 | 2 | 20.63 | 5223.35 | 5223.93 | 5224.52 | 322.70 | 328.36 | 334.01 | 0.49655 |
| 51 | 7 | 73 | 2 | 20.38 | 5258.27 | 5258.86 | 5259.45 | 279.18 | 284.36 | 289.52 | 0.50345 |
| 54 | 7 | 72 | 2 | 16.98 | 5751.12 | 5751.70 | 5752.28 | 78.31 | 83.72 | 89.12 | 0.51429 |
| 54 | 7 | 68 | 2 | 16.77 | 5783.34 | 5783.89 | 5784.44 | 18.13 | 23.28 | 28.45 | 0.48571 |
| 55 | 7 | 64 | 2 | 14.79 | 6091.02 | 6091.54 | 6092.05 | 54.60 | 59.44 | 64.27 | 0.49612 |
| 55 | 7 | 65 | 2 | 14.60 | 6122.07 | 6122.59 | 6123.12 | 351.51 | 356.73 | 1.97 | 0.50388 |
| 56 | 7 | 68 | 2 | 12.84 | 6409.51 | 6410.05 | 6410.60 | 94.99 | 100.12 | 105.25 | 0.50000 |
| 56 | 7 | 68 | 2 | 12.64 | 6442.15 | 6442.70 | 6443.25 | 35.15 | 39.98 | 44.80 | 0.50000 |
| 57 | 7 | 72 | 2 | 10.87 | 6744.36 | 6744.94 | 6745.52 | 355.12 | 359.82 | 4.53 | 0.50000 |
| 57 | 7 | 72 | 2 | 10.68 | 6778.65 | 6779.23 | 6779.82 | 279.21 | 283.86 | 288.50 | 0.50000 |
| 58 | 7 | 75 | 2 | 8.91 | 7094.57 | 7095.17 | 7095.78 | 305.09 | 309.38 | 313.65 | 0.50000 |
| 58 | 7 | 75 | 2 | 8.71 | 7130.50 | 7131.11 | 7131.71 | 222.31 | 227.22 | 232.11 | 0.50000 |
| 59 | 7 | 77 | 2 | 7.33 | 7386.26 | 7386.88 | 7387.50 | 321.25 | 325.82 | 330.38 | 0.49677 |
| 59 | 7 | 78 | 2 | 7.13 | 7423.24 | 7423.87 | 7424.50 | 227.48 | 232.03 | 236.61 | 0.50323 |
| 60 | 7 | 80 | 2 | 5.56 | 7725.33 | 7725.97 | 7726.62 | 313.17 | 317.71 | 322.28 | 0.50000 |
| 60 | 7 | 80 | 2 | 5.36 | 7763.56 | 7764.21 | 7764.85 | 212.63 | 216.93 | 221.23 | 0.50000 |
| 1 | 8 | 1 | 10 | 142.14 | 147.78 | 147.82 | 147.86 | 78.49 | 77.31 | 76.13 | 0.03030 |
| 1 | 8 | 1 | 10 | 142.13 | 147.94 | 147.99 | 148.03 | 73.76 | 72.58 | 71.39 | 0.03030 |
| 1 | 8 | 1 | 10 | 142.12 | 148.11 | 148.15 | 148.19 | 69.03 | 67.85 | 66.66 | 0.03030 |
| 1 | 8 | 1 | 10 | 142.11 | 148.27 | 148.31 | 148.35 | 64.30 | 63.12 | 61.93 | 0.03030 |
| 1 | 8 | 1 | 12 | 141.66 | 157.01 | 157.06 | 157.11 | 170.87 | 169.48 | 168.09 | 0.03030 |
| 1 | 8 | 1 | 12 | 141.65 | 157.20 | 157.25 | 157.30 | 165.32 | 163.93 | 162.55 | 0.03030 |
| 1 | 8 | 1 | 12 | 141.64 | 157.40 | 157.44 | 157.49 | 159.77 | 158.39 | 157.00 | 0.03030 |
| 1 | 8 | 1 | 12 | 141.63 | 157.59 | 157.64 | 157.69 | 154.23 | 152.85 | 151.46 | 0.03030 |
| 1 | 8 | 1 | 10 | 141.17 | 166.59 | 166.63 | 166.67 | 259.40 | 258.27 | 257.14 | 0.03030 |
| 1 | 8 | 1 | 10 | 141.16 | 166.75 | 166.79 | 166.83 | 254.88 | 253.76 | 252.63 | 0.03030 |
| 1 | 8 | 1 | 10 | 141.15 | 166.91 | 166.95 | 166.99 | 250.37 | 249.24 | 248.12 | 0.03030 |
| 1 | 8 | 1 | 10 | 141.14 | 167.07 | 167.11 | 167.15 | 245.86 | 244.73 | 243.60 | 0.03030 |
| 1 | 8 | 1 | 10 | 141.13 | 167.23 | 167.27 | 167.31 | 241.35 | 240.22 | 239.10 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.66 | 176.53 | 176.57 | 176.61 | 344.30 | 343.20 | 342.10 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.65 | 176.69 | 176.73 | 176.77 | 339.89 | 338.79 | 337.69 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.64 | 176.85 | 176.89 | 176.93 | 335.49 | 334.39 | 333.29 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.63 | 177.02 | 177.06 | 177.10 | 331.08 | 329.98 | 328.88 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.62 | 177.18 | 177.22 | 177.26 | 326.68 | 325.58 | 324.48 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.14 | 186.77 | 186.81 | 186.85 | 68.08 | 67.01 | 65.94 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.13 | 186.93 | 186.97 | 187.01 | 63.79 | 62.71 | 61.64 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.12 | 187.09 | 187.13 | 187.17 | 59.49 | 58.42 | 57.35 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.11 | 187.25 | 187.29 | 187.33 | 55.20 | 54.13 | 53.06 | 0.03030 |
| 1 | 8 | 1 | 10 | 140.10 | 187.41 | 187.45 | 187.49 | 50.91 | 49.84 | 48.77 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.61 | 197.17 | 197.21 | 197.25 | 154.88 | 153.84 | 152.80 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.60 | 197.33 | 197.37 | 197.41 | 150.71 | 149.67 | 148.63 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.60 | 197.49 | 197.53 | 197.57 | 146.54 | 145.50 | 144.46 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.59 | 197.65 | 197.69 | 197.73 | 142.38 | 141.34 | 140.29 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.58 | 197.81 | 197.85 | 197.89 | 138.21 | 137.17 | 136.13 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.07 | 208.05 | 208.09 | 208.13 | 237.87 | 236.86 | 235.85 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.06 | 208.22 | 208.26 | 208.30 | 233.84 | 232.83 | 231.82 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.05 | 208.38 | 208.42 | 208.46 | 229.81 | 228.80 | 227.79 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.04 | 208.54 | 208.58 | 208.62 | 225.78 | 224.77 | 223.77 | 0.03030 |
| 1 | 8 | 1 | 10 | 139.03 | 208.70 | 208.74 | 208.78 | 221.75 | 220.75 | 219.74 | 0.03030 |
| 2 | 8 | 16 | 2 | 135.06 | 290.60 | 290.73 | 290.86 | 210.09 | 207.59 | 205.10 | 0.09412 |
| 2 | 8 | 16 | 2 | 135.03 | 291.12 | 291.25 | 291.37 | 200.11 | 197.62 | 195.14 | 0.09412 |
| 2 | 8 | 21 | 2 | 134.68 | 298.65 | 298.82 | 298.99 | 57.16 | 54.01 | 50.86 | 0.12353 |
| 2 | 8 | 21 | 2 | 134.64 | 299.33 | 299.50 | 299.67 | 44.57 | 41.43 | 38.29 | 0.12353 |
| 2 | 8 | 20 | 2 | 133.65 | 320.74 | 320.90 | 321.06 | 29.69 | 27.05 | 24.40 | 0.11765 |
| 2 | 8 | 20 | 2 | 133.62 | 321.39 | 321.55 | 321.71 | 19.12 | 16.48 | 13.85 | 0.11765 |
| 2 | 8 | 20 | 2 | 133.59 | 322.03 | 322.19 | 322.35 | 8.59 | 5.96 | 3.34 | 0.11765 |
| 2 | 8 | 18 | 2 | 132.90 | 336.84 | 336.98 | 337.13 | 138.37 | 136.22 | 134.08 | 0.10588 |
| 2 | 8 | 18 | 2 | 132.88 | 337.42 | 337.56 | 337.71 | 129.79 | 127.65 | 125.51 | 0.10588 |
| 3 | 8 | 23 | 2 | 127.53 | 458.50 | 458.69 | 458.87 | 3.83 | 2.81 | 1.79 | 0.09237 |
| 3 | 8 | 23 | 2 | 127.50 | 459.25 | 459.43 | 459.62 | 359.75 | 358.74 | 357.73 | 0.09237 |
| 3 | 8 | 23 | 2 | 127.47 | 459.99 | 460.17 | 460.36 | 355.71 | 354.71 | 353.72 | 0.09237 |
| 3 | 8 | 21 | 2 | 126.48 | 483.57 | 483.74 | 483.91 | 249.75 | 249.14 | 248.53 | 0.08434 |
| 3 | 8 | 21 | 2 | 126.45 | 484.25 | 484.42 | 484.59 | 247.32 | 246.71 | 246.11 | 0.08434 |
| 3 | 8 | 21 | 2 | 126.42 | 484.93 | 485.10 | 485.27 | 244.91 | 244.31 | 243.72 | 0.08434 |
| 3 | 8 | 21 | 2 | 126.40 | 485.61 | 485.78 | 485.94 | 242.53 | 241.94 | 241.35 | 0.08434 |
| 3 | 8 | 24 | 2 | 125.05 | 518.10 | 518.30 | 518.49 | 164.95 | 164.69 | 164.43 | 0.09639 |
| 3 | 8 | 24 | 2 | 125.02 | 518.88 | 519.07 | 519.27 | 163.92 | 163.67 | 163.41 | 0.09639 |
| 3 | 8 | 24 | 2 | 124.99 | 519.65 | 519.85 | 520.04 | 162.91 | 162.66 | 162.42 | 0.09639 |
| 3 | 8 | 24 | 2 | 124.96 | 520.43 | 520.62 | 520.81 | 161.94 | 161.71 | 161.48 | 0.09639 |
| 4 | 8 | 23 | 2 | 123.50 | 556.41 | 556.60 | 556.78 | 155.39 | 155.55 | 155.71 | 0.09312 |
| 4 | 8 | 23 | 2 | 123.47 | 557.15 | 557.34 | 557.52 | 156.04 | 156.20 | 156.37 | 0.09312 |
| 4 | 8 | 23 | 2 | 123.44 | 557.89 | 558.08 | 558.27 | 156.72 | 156.89 | 157.07 | 0.09312 |
| 4 | 8 | 23 | 2 | 123.41 | 558.64 | 558.82 | 559.01 | 157.43 | 157.61 | 157.80 | 0.09312 |
| 4 | 8 | 23 | 2 | 123.38 | 559.38 | 559.56 | 559.75 | 158.18 | 158.37 | 158.57 | 0.09312 |
| 4 | 8 | 22 | 2 | 121.73 | 601.20 | 601.38 | 601.56 | 252.56 | 253.18 | 253.80 | 0.08907 |
| 4 | 8 | 22 | 2 | 121.70 | 601.91 | 602.09 | 602.27 | 255.05 | 255.67 | 256.30 | 0.08907 |
| 4 | 8 | 22 | 2 | 121.67 | 602.62 | 602.80 | 602.98 | 257.57 | 258.20 | 258.84 | 0.08907 |
| 4 | 8 | 22 | 2 | 121.65 | 603.33 | 603.51 | 603.69 | 260.12 | 260.76 | 261.41 | 0.08907 |
| 4 | 8 | 22 | 2 | 121.62 | 604.04 | 604.22 | 604.40 | 262.71 | 263.35 | 264.01 | 0.08907 |
| 4 | 8 | 22 | 2 | 121.59 | 604.75 | 604.93 | 605.11 | 265.31 | 265.97 | 266.63 | 0.08907 |
| 8 | 8 | 21 | 2 | 113.15 | 834.22 | 834.39 | 834.55 | 300.18 | 302.76 | 305.34 | 0.10938 |
| 8 | 8 | 21 | 2 | 113.13 | 834.89 | 835.06 | 835.23 | 310.51 | 313.09 | 315.68 | 0.10938 |
| 8 | 8 | 21 | 2 | 113.10 | 835.57 | 835.74 | 835.91 | 320.86 | 323.45 | 326.04 | 0.10938 |
| 8 | 8 | 22 | 2 | 112.27 | 859.88 | 860.05 | 860.23 | 346.55 | 349.46 | 352.38 | 0.11458 |
| 8 | 8 | 22 | 2 | 112.25 | 860.59 | 860.76 | 860.94 | 358.22 | 1.14 | 4.06 | 0.11458 |
| 8 | 8 | 22 | 2 | 112.22 | 861.30 | 861.47 | 861.65 | 9.92 | 12.85 | 15.78 | 0.11458 |
| 8 | 8 | 21 | 2 | 111.51 | 882.43 | 882.60 | 882.77 | 8.60 | 11.55 | 14.51 | 0.10938 |
| 8 | 8 | 21 | 2 | 111.49 | 883.11 | 883.28 | 883.45 | 20.43 | 23.39 | 26.35 | 0.10938 |
| 8 | 8 | 21 | 2 | 111.47 | 883.79 | 883.96 | 884.12 | 32.28 | 35.25 | 38.22 | 0.10938 |
| 9 | 8 | 20 | 2 | 107.67 | 959.04 | 959.20 | 959.36 | 119.92 | 123.81 | 127.69 | 0.10753 |
| 9 | 8 | 20 | 2 | 107.62 | 959.68 | 959.84 | 960.00 | 135.48 | 139.38 | 143.29 | 0.10753 |
| 9 | 8 | 19 | 2 | 106.51 | 975.34 | 975.50 | 975.65 | 181.87 | 186.07 | 190.28 | 0.10215 |
| 9 | 8 | 19 | 2 | 106.46 | 975.96 | 976.11 | 976.26 | 198.72 | 202.94 | 207.17 | 0.10215 |
| 9 | 8 | 18 | 2 | 106.00 | 982.37 | 982.52 | 982.66 | 19.10 | 23.27 | 27.44 | 0.09677 |
| 9 | 8 | 18 | 2 | 105.96 | 982.96 | 983.10 | 983.25 | 35.81 | 40.00 | 44.19 | 0.09677 |
| 9 | 8 | 21 | 2 | 105.28 | 992.08 | 992.24 | 992.41 | 308.73 | 313.98 | 319.24 | 0.11290 |
| 9 | 8 | 21 | 2 | 104.37 | 1004.53 | 1004.70 | 1004.86 | 347.60 | 353.18 | 358.76 | 0.11290 |
| 9 | 8 | 15 | 2 | 103.51 | 1016.24 | 1016.36 | 1016.49 | 20.94 | 25.08 | 29.21 | 0.08065 |
| 9 | 8 | 15 | 2 | 103.47 | 1016.73 | 1016.85 | 1016.97 | 37.49 | 41.63 | 45.76 | 0.08065 |
| 10 | 8 | 19 | 2 | 101.03 | 1050.43 | 1050.58 | 1050.74 | 159.46 | 165.15 | 170.84 | 0.13869 |
| 10 | 8 | 19 | 2 | 100.28 | 1060.94 | 1061.09 | 1061.24 | 194.83 | 200.68 | 206.54 | 0.13869 |
| 10 | 8 | 18 | 2 | 99.57 | 1070.98 | 1071.12 | 1071.27 | 223.49 | 229.20 | 234.90 | 0.13139 |
| 10 | 8 | 17 | 2 | 98.90 | 1080.69 | 1080.83 | 1080.97 | 250.95 | 256.50 | 262.06 | 0.12409 |
| 10 | 8 | 17 | 2 | 98.73 | 1083.17 | 1083.31 | 1083.44 | 351.58 | 357.19 | 2.79 | 0.12409 |
| 10 | 8 | 16 | 2 | 98.16 | 1091.38 | 1091.51 | 1091.64 | 331.44 | 336.85 | 342.27 | 0.11679 |
| 10 | 8 | 16 | 2 | 97.55 | 1100.42 | 1100.55 | 1100.68 | 356.80 | 2.39 | 7.97 | 0.11679 |
| 10 | 8 | 15 | 2 | 96.95 | 1109.24 | 1109.36 | 1109.48 | 23.97 | 29.35 | 34.74 | 0.10949 |
| 11 | 8 | 14 | 2 | 94.76 | 1142.51 | 1142.62 | 1142.73 | 135.99 | 141.50 | 147.01 | 0.11966 |
| 11 | 8 | 14 | 2 | 94.25 | 1150.41 | 1150.52 | 1150.63 | 165.47 | 171.08 | 176.70 | 0.11966 |
| 11 | 8 | 13 | 2 | 93.76 | 1158.19 | 1158.30 | 1158.40 | 196.74 | 202.06 | 207.39 | 0.11111 |
| 11 | 8 | 13 | 2 | 93.58 | 1161.03 | 1161.13 | 1161.24 | 341.32 | 346.68 | 352.04 | 0.11111 |
| 11 | 8 | 13 | 2 | 93.12 | 1168.31 | 1168.42 | 1168.52 | 356.60 | 2.05 | 7.51 | 0.11111 |
| 11 | 8 | 13 | 2 | 92.66 | 1175.65 | 1175.75 | 1175.86 | 21.26 | 26.81 | 32.35 | 0.11111 |
| 11 | 8 | 13 | 2 | 92.21 | 1182.95 | 1183.05 | 1183.16 | 50.52 | 56.16 | 61.81 | 0.11111 |
| 11 | 8 | 12 | 2 | 91.77 | 1190.04 | 1190.13 | 1190.23 | 75.42 | 80.72 | 86.03 | 0.10256 |
| 11 | 8 | 12 | 2 | 91.35 | 1196.95 | 1197.05 | 1197.15 | 97.83 | 103.22 | 108.61 | 0.10256 |
| 12 | 8 | 10 | 2 | 89.25 | 1232.29 | 1232.37 | 1232.45 | 329.73 | 334.45 | 339.17 | 0.05051 |
| 12 | 8 | 21 | 2 | 88.89 | 1238.33 | 1238.50 | 1238.67 | 323.94 | 333.86 | 343.77 | 0.10606 |
| 12 | 8 | 21 | 2 | 88.54 | 1244.30 | 1244.47 | 1244.64 | 312.95 | 322.84 | 332.73 | 0.10606 |
| 12 | 8 | 10 | 2 | 88.19 | 1250.57 | 1250.65 | 1250.73 | 318.14 | 322.83 | 327.52 | 0.05051 |
| 12 | 8 | 10 | 2 | 87.84 | 1256.64 | 1256.72 | 1256.80 | 310.30 | 314.96 | 319.63 | 0.05051 |
| 12 | 8 | 21 | 2 | 87.49 | 1262.69 | 1262.86 | 1263.03 | 299.28 | 309.03 | 318.78 | 0.10606 |
| 12 | 8 | 21 | 2 | 87.15 | 1268.69 | 1268.86 | 1269.03 | 283.86 | 293.57 | 303.28 | 0.10606 |
| 12 | 8 | 21 | 2 | 86.82 | 1274.58 | 1274.75 | 1274.92 | 261.20 | 270.88 | 280.56 | 0.10606 |
| 12 | 8 | 21 | 2 | 86.50 | 1280.42 | 1280.59 | 1280.76 | 234.32 | 243.97 | 253.63 | 0.10606 |
| 12 | 8 | 21 | 2 | 86.18 | 1286.24 | 1286.41 | 1286.58 | 205.62 | 215.24 | 224.87 | 0.10606 |
| 12 | 8 | 21 | 2 | 86.08 | 1287.92 | 1288.09 | 1288.26 | 300.90 | 310.51 | 320.13 | 0.10606 |
| 13 | 8 | 20 | 2 | 84.19 | 1322.93 | 1323.10 | 1323.26 | 118.37 | 127.44 | 136.52 | 0.08889 |
| 13 | 8 | 20 | 2 | 84.03 | 1326.10 | 1326.26 | 1326.43 | 296.62 | 305.68 | 314.74 | 0.08889 |
| 13 | 8 | 22 | 2 | 83.67 | 1332.95 | 1333.13 | 1333.30 | 320.58 | 330.50 | 340.41 | 0.09778 |
| 13 | 8 | 20 | 2 | 83.36 | 1338.93 | 1339.09 | 1339.26 | 294.24 | 303.20 | 312.15 | 0.08889 |
| 13 | 8 | 20 | 2 | 83.06 | 1344.64 | 1344.80 | 1344.96 | 250.32 | 259.21 | 268.10 | 0.08889 |
| 13 | 8 | 20 | 2 | 82.76 | 1350.36 | 1350.52 | 1350.68 | 203.93 | 212.73 | 221.52 | 0.08889 |
| 13 | 8 | 20 | 2 | 82.47 | 1356.07 | 1356.23 | 1356.39 | 153.56 | 162.25 | 170.93 | 0.08889 |
| 13 | 8 | 21 | 2 | 82.18 | 1361.86 | 1362.03 | 1362.20 | 103.22 | 112.21 | 121.20 | 0.09333 |
| 13 | 8 | 20 | 2 | 81.89 | 1367.63 | 1367.79 | 1367.96 | 47.13 | 55.54 | 63.95 | 0.08889 |
| 13 | 8 | 21 | 2 | 81.59 | 1373.45 | 1373.62 | 1373.79 | 348.02 | 356.71 | 5.41 | 0.09333 |
| 13 | 8 | 21 | 2 | 81.30 | 1379.28 | 1379.45 | 1379.62 | 284.72 | 293.27 | 301.81 | 0.09333 |
| 14 | 8 | 13 | 2 | 79.82 | 1409.58 | 1409.69 | 1409.79 | 326.60 | 331.63 | 336.65 | 0.06341 |
| 14 | 8 | 12 | 2 | 79.48 | 1416.72 | 1416.82 | 1416.91 | 307.53 | 312.15 | 316.76 | 0.05854 |
| 14 | 8 | 22 | 2 | 79.22 | 1422.13 | 1422.31 | 1422.48 | 205.30 | 213.74 | 222.18 | 0.10732 |
| 14 | 8 | 22 | 2 | 78.93 | 1428.26 | 1428.44 | 1428.61 | 136.63 | 145.05 | 153.47 | 0.10732 |
| 14 | 8 | 22 | 2 | 78.63 | 1434.49 | 1434.67 | 1434.85 | 72.09 | 80.48 | 88.87 | 0.10732 |
| 14 | 8 | 22 | 2 | 78.34 | 1440.77 | 1440.94 | 1441.12 | 8.18 | 16.52 | 24.86 | 0.10732 |
| 14 | 8 | 23 | 2 | 78.05 | 1447.06 | 1447.25 | 1447.43 | 303.44 | 312.11 | 320.77 | 0.11220 |
| 14 | 8 | 23 | 2 | 77.75 | 1453.38 | 1453.56 | 1453.75 | 236.96 | 245.53 | 254.11 | 0.11220 |
| 14 | 8 | 23 | 2 | 77.46 | 1459.80 | 1459.98 | 1460.17 | 172.22 | 180.71 | 189.18 | 0.11220 |
| 14 | 8 | 23 | 2 | 77.17 | 1466.30 | 1466.48 | 1466.67 | 107.52 | 115.89 | 124.26 | 0.11220 |
| 15 | 8 | 16 | 2 | 73.98 | 1539.63 | 1539.76 | 1539.89 | 329.23 | 334.45 | 339.67 | 0.14953 |
| 15 | 8 | 13 | 2 | 73.72 | 1545.91 | 1546.01 | 1546.12 | 222.87 | 227.09 | 231.30 | 0.12150 |
| 15 | 8 | 13 | 2 | 73.42 | 1553.34 | 1553.44 | 1553.55 | 157.52 | 161.63 | 165.74 | 0.12150 |
| 15 | 8 | 13 | 2 | 73.11 | 1560.75 | 1560.85 | 1560.96 | 83.78 | 87.76 | 91.73 | 0.12150 |
| 15 | 8 | 13 | 2 | 72.81 | 1568.29 | 1568.39 | 1568.50 | 4.89 | 8.74 | 12.58 | 0.12150 |
| 15 | 8 | 13 | 2 | 72.50 | 1575.92 | 1576.02 | 1576.13 | 283.56 | 287.36 | 291.15 | 0.12150 |
| 15 | 8 | 13 | 2 | 72.19 | 1583.66 | 1583.77 | 1583.87 | 199.55 | 203.27 | 206.99 | 0.12150 |
| 15 | 8 | 13 | 2 | 71.88 | 1591.50 | 1591.61 | 1591.71 | 114.34 | 118.00 | 121.65 | 0.12150 |
| 16 | 8 | 14 | 2 | 70.53 | 1626.46 | 1626.57 | 1626.68 | 264.54 | 268.50 | 272.47 | 0.13462 |
| 16 | 8 | 15 | 2 | 70.22 | 1634.89 | 1635.01 | 1635.13 | 200.04 | 204.28 | 208.52 | 0.14423 |
| 16 | 8 | 15 | 2 | 69.90 | 1643.45 | 1643.58 | 1643.70 | 135.81 | 139.98 | 144.15 | 0.14423 |
| 16 | 8 | 15 | 2 | 69.59 | 1652.03 | 1652.15 | 1652.27 | 72.95 | 77.16 | 81.36 | 0.14423 |
| 16 | 8 | 15 | 2 | 69.27 | 1660.78 | 1660.90 | 1661.03 | 10.46 | 14.38 | 18.29 | 0.14423 |
| 16 | 8 | 15 | 2 | 68.96 | 1669.48 | 1669.60 | 1669.72 | 291.39 | 295.31 | 299.23 | 0.14423 |
| 16 | 8 | 15 | 2 | 68.64 | 1678.36 | 1678.48 | 1678.61 | 219.83 | 223.77 | 227.71 | 0.14423 |
| 17 | 8 | 16 | 2 | 67.36 | 1715.44 | 1715.57 | 1715.70 | 352.43 | 356.61 | 0.78 | 0.13675 |
| 17 | 8 | 17 | 2 | 67.05 | 1724.80 | 1724.93 | 1725.07 | 289.51 | 293.73 | 297.95 | 0.14530 |
| 17 | 8 | 16 | 2 | 66.74 | 1734.19 | 1734.32 | 1734.45 | 213.23 | 217.12 | 221.01 | 0.13675 |
| 17 | 8 | 17 | 2 | 66.43 | 1743.77 | 1743.91 | 1744.05 | 148.32 | 152.57 | 156.81 | 0.14530 |
| 17 | 8 | 17 | 2 | 66.12 | 1753.30 | 1753.43 | 1753.57 | 75.83 | 79.85 | 83.88 | 0.14530 |
| 17 | 8 | 17 | 2 | 65.82 | 1763.09 | 1763.23 | 1763.37 | 1.41 | 5.38 | 9.35 | 0.14530 |
| 17 | 8 | 17 | 2 | 65.51 | 1773.02 | 1773.16 | 1773.29 | 284.66 | 288.55 | 292.44 | 0.14530 |
| 18 | 8 | 18 | 2 | 64.53 | 1805.27 | 1805.42 | 1805.56 | 134.76 | 138.91 | 143.06 | 0.16216 |
| 18 | 8 | 18 | 2 | 64.22 | 1815.62 | 1815.76 | 1815.91 | 67.74 | 71.80 | 75.86 | 0.16216 |
| 18 | 8 | 18 | 2 | 63.92 | 1825.96 | 1826.11 | 1826.25 | 0.24 | 4.39 | 8.54 | 0.16216 |
| 18 | 8 | 19 | 2 | 63.62 | 1836.61 | 1836.76 | 1836.91 | 302.46 | 306.78 | 311.11 | 0.17117 |
| 18 | 8 | 19 | 2 | 63.32 | 1847.41 | 1847.57 | 1847.72 | 244.50 | 248.72 | 252.94 | 0.17117 |
| 18 | 8 | 19 | 2 | 63.01 | 1858.43 | 1858.58 | 1858.73 | 184.41 | 188.63 | 192.86 | 0.17117 |
| 22 | 8 | 26 | 2 | 59.12 | 2011.94 | 2012.15 | 2012.36 | 330.07 | 335.32 | 340.57 | 0.22609 |
| 22 | 8 | 23 | 2 | 58.82 | 2024.76 | 2024.95 | 2025.13 | 287.91 | 292.52 | 297.14 | 0.20000 |
| 22 | 8 | 22 | 2 | 58.53 | 2037.29 | 2037.46 | 2037.64 | 242.40 | 246.91 | 251.42 | 0.19130 |
| 22 | 8 | 22 | 2 | 58.25 | 2049.79 | 2049.96 | 2050.14 | 199.51 | 203.87 | 208.26 | 0.19130 |
| 22 | 8 | 22 | 2 | 57.96 | 2062.29 | 2062.46 | 2062.64 | 156.63 | 161.14 | 165.66 | 0.19130 |
| 23 | 8 | 23 | 2 | 57.11 | 2100.50 | 2100.69 | 2100.87 | 10.03 | 14.43 | 18.83 | 0.16667 |
| 23 | 8 | 23 | 2 | 56.84 | 2113.26 | 2113.45 | 2113.64 | 307.90 | 312.20 | 316.51 | 0.16667 |
| 23 | 8 | 23 | 2 | 56.55 | 2126.72 | 2126.90 | 2127.09 | 263.01 | 267.36 | 271.71 | 0.16667 |
| 23 | 8 | 23 | 2 | 56.27 | 2139.49 | 2139.67 | 2139.86 | 200.71 | 205.04 | 209.36 | 0.16667 |
| 23 | 8 | 22 | 2 | 56.20 | 2142.67 | 2142.84 | 2143.02 | 274.48 | 278.60 | 282.71 | 0.15942 |
| 23 | 8 | 24 | 2 | 55.93 | 2155.60 | 2155.79 | 2155.99 | 205.72 | 209.88 | 214.05 | 0.17391 |
| 24 | 8 | 25 | 2 | 54.28 | 2236.25 | 2236.45 | 2236.65 | 60.89 | 64.70 | 68.52 | 0.15723 |
| 24 | 8 | 26 | 2 | 54.00 | 2250.53 | 2250.74 | 2250.95 | 330.33 | 334.20 | 338.08 | 0.16352 |
| 24 | 8 | 26 | 2 | 53.72 | 2264.82 | 2265.03 | 2265.24 | 236.65 | 240.60 | 244.53 | 0.16352 |
| 24 | 8 | 27 | 2 | 53.43 | 2279.66 | 2279.87 | 2280.09 | 155.92 | 160.12 | 164.33 | 0.16981 |
| 24 | 8 | 27 | 2 | 53.14 | 2294.86 | 2295.08 | 2295.30 | 80.12 | 84.08 | 88.06 | 0.16981 |
| 24 | 8 | 28 | 2 | 52.96 | 2304.30 | 2304.52 | 2304.75 | 250.54 | 254.53 | 258.52 | 0.17610 |
| 25 | 8 | 29 | 2 | 52.57 | 2324.85 | 2325.08 | 2325.32 | 242.00 | 246.00 | 250.01 | 0.24370 |
| 25 | 8 | 29 | 2 | 52.28 | 2340.64 | 2340.87 | 2341.11 | 159.31 | 163.37 | 167.41 | 0.24370 |
| 25 | 8 | 30 | 2 | 51.97 | 2357.03 | 2357.27 | 2357.52 | 75.89 | 79.87 | 83.84 | 0.25210 |
| 25 | 8 | 31 | 2 | 51.67 | 2373.78 | 2374.03 | 2374.28 | 344.70 | 348.61 | 352.53 | 0.26050 |
| 26 | 8 | 32 | 2 | 50.74 | 2425.26 | 2425.51 | 2425.77 | 74.06 | 78.19 | 82.32 | 0.24615 |
| 26 | 8 | 33 | 2 | 50.43 | 2443.17 | 2443.44 | 2443.71 | 1.46 | 5.71 | 9.96 | 0.25385 |
| 26 | 8 | 32 | 2 | 50.12 | 2460.84 | 2461.10 | 2461.36 | 274.75 | 278.59 | 282.43 | 0.24615 |
| 26 | 8 | 33 | 2 | 49.82 | 2478.47 | 2478.74 | 2479.00 | 174.79 | 178.79 | 182.80 | 0.25385 |
| 30 | 8 | 65 | 2 | 47.47 | 2620.26 | 2620.78 | 2621.30 | 247.24 | 255.42 | 263.62 | 0.25000 |
| 30 | 8 | 64 | 2 | 47.20 | 2637.40 | 2637.92 | 2638.43 | 158.67 | 166.77 | 174.85 | 0.24615 |
| 30 | 8 | 65 | 2 | 46.93 | 2654.54 | 2655.06 | 2655.58 | 62.70 | 70.66 | 78.61 | 0.25000 |
| 30 | 8 | 66 | 2 | 46.65 | 2672.32 | 2672.86 | 2673.39 | 328.91 | 336.71 | 344.49 | 0.25385 |
| 31 | 8 | 68 | 2 | 45.99 | 2715.37 | 2715.92 | 2716.47 | 274.45 | 283.49 | 292.54 | 0.23529 |
| 31 | 8 | 67 | 2 | 45.72 | 2733.03 | 2733.57 | 2734.12 | 199.35 | 207.87 | 216.37 | 0.23183 |
| 31 | 8 | 74 | 2 | 45.64 | 2738.82 | 2739.41 | 2740.01 | 290.41 | 299.75 | 309.07 | 0.25606 |
| 31 | 8 | 80 | 2 | 45.30 | 2761.30 | 2761.94 | 2762.59 | 287.69 | 298.24 | 308.82 | 0.27682 |
| 32 | 8 | 63 | 2 | 44.59 | 2809.50 | 2810.01 | 2810.52 | 26.99 | 35.67 | 44.32 | 0.25200 |
| 32 | 8 | 63 | 2 | 44.35 | 2826.17 | 2826.68 | 2827.18 | 302.98 | 311.43 | 319.90 | 0.25200 |
| 32 | 8 | 61 | 2 | 44.12 | 2842.17 | 2842.66 | 2843.15 | 209.11 | 217.21 | 225.29 | 0.24400 |
| 32 | 8 | 63 | 2 | 43.88 | 2859.17 | 2859.67 | 2860.18 | 120.03 | 127.89 | 135.73 | 0.25200 |
| 33 | 8 | 84 | 2 | 43.38 | 2894.45 | 2895.13 | 2895.81 | 308.85 | 319.36 | 329.88 | 0.42424 |
| 33 | 8 | 48 | 2 | 43.02 | 2920.19 | 2920.58 | 2920.97 | 2.75 | 9.27 | 15.81 | 0.24242 |
| 33 | 8 | 66 | 2 | 42.80 | 2936.56 | 2937.09 | 2937.63 | 276.11 | 284.60 | 293.04 | 0.33333 |
| 34 | 8 | 66 | 2 | 42.08 | 2989.32 | 2989.86 | 2990.39 | 350.36 | 358.94 | 7.54 | 0.24719 |
| 34 | 8 | 67 | 2 | 41.84 | 3007.26 | 3007.80 | 3008.34 | 285.79 | 294.66 | 303.55 | 0.25094 |
| 34 | 8 | 65 | 2 | 41.61 | 3024.43 | 3024.95 | 3025.48 | 205.18 | 213.60 | 222.00 | 0.24345 |
| 34 | 8 | 9 | 2 | 41.57 | 3028.27 | 3028.35 | 3028.42 | 266.78 | 267.93 | 269.09 | 0.03371 |
| 34 | 8 | 60 | 2 | 41.37 | 3042.89 | 3043.38 | 3043.86 | 138.31 | 145.96 | 153.62 | 0.22472 |
| 35 | 8 | 62 | 2 | 40.24 | 3130.28 | 3130.78 | 3131.28 | 29.02 | 36.36 | 43.68 | 0.24900 |
| 35 | 8 | 62 | 2 | 40.03 | 3146.81 | 3147.31 | 3147.81 | 269.23 | 275.92 | 282.54 | 0.24900 |
| 35 | 8 | 62 | 2 | 39.82 | 3163.45 | 3163.95 | 3164.45 | 125.01 | 131.97 | 138.94 | 0.24900 |
| 35 | 8 | 63 | 2 | 39.62 | 3180.16 | 3180.67 | 3181.18 | 357.01 | 3.91 | 10.80 | 0.25301 |
| 36 | 8 | 63 | 2 | 39.20 | 3213.95 | 3214.46 | 3214.97 | 101.93 | 109.02 | 116.09 | 0.24609 |
| 36 | 8 | 64 | 2 | 38.99 | 3230.97 | 3231.48 | 3232.00 | 339.34 | 346.59 | 353.83 | 0.25000 |
| 36 | 8 | 64 | 2 | 38.78 | 3248.14 | 3248.65 | 3249.17 | 216.55 | 223.45 | 230.32 | 0.25000 |
| 36 | 8 | 65 | 2 | 38.57 | 3265.38 | 3265.90 | 3266.42 | 73.05 | 79.63 | 86.22 | 0.25391 |
| 37 | 8 | 66 | 2 | 38.16 | 3300.25 | 3300.78 | 3301.31 | 203.32 | 211.29 | 219.22 | 0.24906 |
| 37 | 8 | 66 | 2 | 37.95 | 3317.79 | 3318.32 | 3318.85 | 85.90 | 93.02 | 100.11 | 0.24906 |
| 37 | 8 | 66 | 2 | 37.74 | 3335.46 | 3335.99 | 3336.53 | 320.52 | 327.67 | 334.83 | 0.24906 |
| 37 | 8 | 67 | 2 | 37.53 | 3353.24 | 3353.78 | 3354.32 | 205.78 | 213.51 | 221.26 | 0.25283 |
| 38 | 8 | 68 | 2 | 37.01 | 3397.90 | 3398.44 | 3398.99 | 170.93 | 179.39 | 187.82 | 0.24818 |
| 38 | 8 | 68 | 2 | 36.81 | 3416.08 | 3416.63 | 3417.18 | 75.04 | 82.56 | 90.05 | 0.24818 |
| 38 | 8 | 69 | 2 | 36.60 | 3434.39 | 3434.94 | 3435.50 | 318.57 | 325.96 | 333.35 | 0.25182 |
| 38 | 8 | 69 | 2 | 36.39 | 3452.79 | 3453.34 | 3453.90 | 199.92 | 207.18 | 214.44 | 0.25182 |
| 39 | 8 | 70 | 2 | 35.98 | 3489.92 | 3490.49 | 3491.05 | 1.41 | 9.89 | 18.37 | 0.20772 |
| 39 | 8 | 70 | 2 | 35.79 | 3507.01 | 3507.58 | 3508.14 | 259.87 | 268.21 | 276.54 | 0.20772 |
| 39 | 8 | 71 | 2 | 35.58 | 3525.87 | 3526.45 | 3527.02 | 164.67 | 172.12 | 179.55 | 0.21068 |
| 39 | 8 | 54 | 2 | 35.53 | 3530.46 | 3530.90 | 3531.33 | 223.85 | 229.41 | 234.96 | 0.16024 |
| 39 | 8 | 72 | 2 | 35.32 | 3548.98 | 3549.56 | 3550.15 | 94.57 | 102.16 | 109.81 | 0.21365 |
| 40 | 8 | 72 | 2 | 35.12 | 3568.17 | 3568.75 | 3569.33 | 348.52 | 355.98 | 3.43 | 0.24828 |
| 40 | 8 | 72 | 2 | 34.91 | 3587.36 | 3587.94 | 3588.52 | 236.65 | 244.48 | 252.36 | 0.24828 |
| 40 | 8 | 73 | 2 | 34.70 | 3606.71 | 3607.30 | 3607.89 | 158.15 | 167.17 | 176.20 | 0.25172 |
| 40 | 8 | 73 | 2 | 34.49 | 3626.20 | 3626.79 | 3627.38 | 81.69 | 90.31 | 98.99 | 0.25172 |
| 41 | 8 | 74 | 2 | 34.12 | 3661.51 | 3662.11 | 3662.70 | 233.95 | 242.32 | 250.75 | 0.24667 |
| 41 | 8 | 75 | 2 | 33.91 | 3681.28 | 3681.89 | 3682.49 | 154.58 | 162.70 | 170.84 | 0.25000 |
| 41 | 8 | 75 | 2 | 33.70 | 3701.25 | 3701.85 | 3702.46 | 72.52 | 81.07 | 89.63 | 0.25000 |
| 41 | 8 | 76 | 2 | 33.50 | 3721.31 | 3721.93 | 3722.54 | 354.31 | 2.76 | 11.24 | 0.25333 |
| 42 | 8 | 78 | 2 | 32.67 | 3802.12 | 3802.75 | 3803.38 | 23.68 | 31.39 | 39.08 | 0.33191 |
| 42 | 8 | 78 | 2 | 32.46 | 3822.77 | 3823.40 | 3824.03 | 258.95 | 266.65 | 274.47 | 0.33191 |
| 42 | 8 | 79 | 2 | 32.26 | 3843.62 | 3844.26 | 3844.89 | 170.08 | 178.19 | 186.30 | 0.33617 |
| 43 | 8 | 80 | 2 | 31.84 | 3885.77 | 3886.42 | 3887.06 | 308.18 | 314.03 | 319.86 | 0.33195 |
| 43 | 8 | 80 | 2 | 31.63 | 3906.92 | 3907.56 | 3908.21 | 157.76 | 165.39 | 173.11 | 0.33195 |
| 43 | 8 | 81 | 2 | 31.42 | 3928.27 | 3928.92 | 3929.57 | 104.71 | 114.35 | 123.97 | 0.33610 |
| 44 | 8 | 82 | 2 | 30.80 | 3993.16 | 3993.82 | 3994.48 | 348.74 | 0.19 | 11.67 | 0.33065 |
| 44 | 8 | 83 | 2 | 30.61 | 4013.53 | 4014.20 | 4014.87 | 345.44 | 356.96 | 8.42 | 0.33468 |
| 44 | 8 | 83 | 2 | 30.40 | 4035.56 | 4036.23 | 4036.90 | 341.52 | 351.75 | 1.95 | 0.33468 |
| 45 | 8 | 84 | 2 | 29.98 | 4080.03 | 4080.71 | 4081.38 | 246.42 | 255.79 | 265.18 | 0.33071 |
| 45 | 8 | 85 | 2 | 29.77 | 4102.45 | 4103.13 | 4103.82 | 195.01 | 204.18 | 213.32 | 0.33465 |
| 45 | 8 | 85 | 2 | 29.56 | 4125.03 | 4125.72 | 4126.41 | 130.37 | 139.00 | 147.61 | 0.33465 |
| 46 | 8 | 86 | 2 | 29.15 | 4170.64 | 4171.33 | 4172.02 | 259.95 | 267.28 | 274.70 | 0.33077 |
| 46 | 8 | 87 | 2 | 28.94 | 4193.63 | 4194.33 | 4195.03 | 169.10 | 177.92 | 186.76 | 0.33462 |
| 46 | 8 | 87 | 2 | 28.73 | 4216.79 | 4217.49 | 4218.19 | 101.57 | 110.84 | 120.16 | 0.33462 |
| 47 | 8 | 89 | 2 | 28.31 | 4263.52 | 4264.24 | 4264.96 | 321.49 | 329.67 | 337.84 | 0.33209 |
| 47 | 8 | 89 | 2 | 28.10 | 4287.13 | 4287.85 | 4288.57 | 235.94 | 244.40 | 252.89 | 0.33209 |
| 47 | 8 | 90 | 2 | 27.89 | 4310.86 | 4311.58 | 4312.31 | 140.26 | 146.93 | 153.54 | 0.33582 |
| 48 | 8 | 91 | 2 | 27.47 | 4358.78 | 4359.51 | 4360.24 | 267.29 | 275.13 | 282.94 | 0.33212 |
| 48 | 8 | 91 | 2 | 27.26 | 4382.96 | 4383.69 | 4384.43 | 156.68 | 163.96 | 171.27 | 0.33212 |
| 48 | 8 | 92 | 2 | 27.05 | 4407.30 | 4408.04 | 4408.78 | 51.02 | 58.99 | 66.95 | 0.33577 |
| 50 | 8 | 128 | 2 | 23.18 | 4883.14 | 4884.18 | 4885.21 | 81.51 | 91.99 | 102.39 | 0.49805 |
| 50 | 8 | 129 | 2 | 22.92 | 4916.83 | 4917.87 | 4918.91 | 58.74 | 69.14 | 79.60 | 0.50195 |
| 51 | 8 | 132 | 2 | 20.61 | 5226.23 | 5227.30 | 5228.36 | 350.61 | 0.83 | 11.01 | 0.49811 |
| 51 | 8 | 133 | 2 | 20.36 | 5261.17 | 5262.24 | 5263.31 | 304.58 | 313.96 | 323.36 | 0.50189 |
| 54 | 8 | 131 | 2 | 16.96 | 5753.99 | 5755.04 | 5756.10 | 104.95 | 114.76 | 124.57 | 0.51172 |
| 54 | 8 | 125 | 2 | 16.75 | 5786.06 | 5787.06 | 5788.07 | 43.81 | 53.39 | 63.01 | 0.48828 |
| 55 | 8 | 118 | 2 | 14.78 | 6093.59 | 6094.54 | 6095.49 | 78.56 | 87.43 | 96.28 | 0.50000 |
| 55 | 8 | 118 | 2 | 14.58 | 6124.64 | 6125.59 | 6126.54 | 17.16 | 26.64 | 36.12 | 0.50000 |
| 56 | 8 | 124 | 2 | 12.82 | 6412.21 | 6413.21 | 6414.21 | 120.41 | 129.84 | 139.28 | 0.50000 |
| 56 | 8 | 124 | 2 | 12.62 | 6444.86 | 6445.86 | 6446.86 | 58.95 | 67.66 | 76.40 | 0.50000 |
| 57 | 8 | 130 | 2 | 10.86 | 6747.21 | 6748.25 | 6749.30 | 18.20 | 26.73 | 35.29 | 0.49808 |
| 57 | 8 | 131 | 2 | 10.66 | 6781.51 | 6782.57 | 6783.62 | 302.05 | 310.45 | 318.80 | 0.50192 |
| 58 | 8 | 136 | 2 | 8.89 | 7097.53 | 7098.62 | 7099.72 | 326.02 | 333.79 | 341.56 | 0.50000 |
| 58 | 8 | 136 | 2 | 8.69 | 7133.48 | 7134.57 | 7135.67 | 246.31 | 255.10 | 263.87 | 0.50000 |
| 59 | 8 | 141 | 2 | 7.31 | 7389.32 | 7390.46 | 7391.60 | 343.73 | 352.02 | 0.36 | 0.50000 |
| 59 | 8 | 141 | 2 | 7.11 | 7426.32 | 7427.46 | 7428.60 | 249.96 | 258.29 | 266.64 | 0.50000 |
| 60 | 8 | 146 | 2 | 5.54 | 7728.50 | 7729.67 | 7730.85 | 335.54 | 343.78 | 352.03 | 0.50000 |
| 60 | 8 | 146 | 2 | 5.34 | 7766.73 | 7767.91 | 7769.08 | 233.81 | 241.71 | 249.66 | 0.50000 |

For the ULIS, the instrument collected data in 4 regions as given below:

ULIS Azimuth for Each Region

| cycle | region | samp.t. | # reads | altitude | mt start | mt center | mt end | az start | az center | az end | weights |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 8 | 1 | 10 | 142.14 | 147.78 | 147.82 | 147.86 | 78.49 | 77.31 | 76.13 | 0.01515 |
| 1 | 8 | 1 | 10 | 142.13 | 147.94 | 147.99 | 148.03 | 73.76 | 72.58 | 71.39 | 0.01515 |
| 1 | 8 | 1 | 10 | 142.12 | 148.11 | 148.15 | 148.19 | 69.03 | 67.85 | 66.66 | 0.01515 |
| 1 | 8 | 1 | 10 | 142.11 | 148.27 | 148.31 | 148.35 | 64.30 | 63.12 | 61.93 | 0.01515 |
| 1 | 1 | 1 | 10 | 142.08 | 148.90 | 148.94 | 148.99 | 45.63 | 44.45 | 43.27 | 0.01515 |
| 1 | 1 | 1 | 10 | 142.07 | 149.07 | 149.11 | 149.15 | 40.91 | 39.73 | 38.55 | 0.01515 |
| 1 | 1 | 1 | 10 | 142.06 | 149.23 | 149.27 | 149.31 | 36.19 | 35.01 | 33.84 | 0.01515 |
| 1 | 1 | 1 | 10 | 142.05 | 149.39 | 149.43 | 149.47 | 31.48 | 30.30 | 29.12 | 0.01515 |
| 1 | 8 | 1 | 12 | 141.66 | 157.01 | 157.06 | 157.11 | 170.87 | 169.48 | 168.09 | 0.01515 |
| 1 | 8 | 1 | 12 | 141.65 | 157.20 | 157.25 | 157.30 | 165.32 | 163.93 | 162.55 | 0.01515 |
| 1 | 8 | 1 | 12 | 141.64 | 157.40 | 157.44 | 157.49 | 159.77 | 158.39 | 157.00 | 0.01515 |
| 1 | 8 | 1 | 12 | 141.63 | 157.59 | 157.64 | 157.69 | 154.23 | 152.85 | 151.46 | 0.01515 |
| 1 | 1 | 1 | 12 | 141.60 | 158.16 | 158.21 | 158.26 | 137.85 | 136.47 | 135.08 | 0.01515 |
| 1 | 1 | 1 | 12 | 141.59 | 158.36 | 158.40 | 158.45 | 132.32 | 130.94 | 129.55 | 0.01515 |
| 1 | 1 | 1 | 12 | 141.58 | 158.55 | 158.60 | 158.65 | 126.79 | 125.41 | 124.03 | 0.01515 |
| 1 | 1 | 1 | 12 | 141.57 | 158.74 | 158.79 | 158.84 | 121.26 | 119.88 | 118.50 | 0.01515 |
| 1 | 8 | 1 | 10 | 141.17 | 166.59 | 166.63 | 166.67 | 259.40 | 258.27 | 257.14 | 0.01515 |
| 1 | 8 | 1 | 10 | 141.16 | 166.75 | 166.79 | 166.83 | 254.88 | 253.76 | 252.63 | 0.01515 |
| 1 | 8 | 1 | 10 | 141.15 | 166.91 | 166.95 | 166.99 | 250.37 | 249.24 | 248.12 | 0.01515 |
| 1 | 8 | 1 | 10 | 141.14 | 167.07 | 167.11 | 167.15 | 245.86 | 244.73 | 243.60 | 0.01515 |
| 1 | 8 | 1 | 10 | 141.13 | 167.23 | 167.27 | 167.31 | 241.35 | 240.22 | 239.10 | 0.01515 |
| 1 | 1 | 1 | 10 | 141.10 | 167.78 | 167.82 | 167.86 | 226.03 | 224.90 | 223.78 | 0.01515 |
| 1 | 1 | 1 | 10 | 141.10 | 167.94 | 167.98 | 168.02 | 221.53 | 220.40 | 219.28 | 0.01515 |
| 1 | 1 | 1 | 10 | 141.09 | 168.11 | 168.15 | 168.19 | 217.03 | 215.90 | 214.78 | 0.01515 |
| 1 | 1 | 1 | 10 | 141.08 | 168.27 | 168.31 | 168.35 | 212.53 | 211.41 | 210.28 | 0.01515 |
| 1 | 1 | 1 | 10 | 141.07 | 168.43 | 168.47 | 168.51 | 208.03 | 206.91 | 205.78 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.66 | 176.53 | 176.57 | 176.61 | 344.30 | 343.20 | 342.10 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.65 | 176.69 | 176.73 | 176.77 | 339.89 | 338.79 | 337.69 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.64 | 176.85 | 176.89 | 176.93 | 335.49 | 334.39 | 333.29 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.63 | 177.02 | 177.06 | 177.10 | 331.08 | 329.98 | 328.88 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.62 | 177.18 | 177.22 | 177.26 | 326.68 | 325.58 | 324.48 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.59 | 177.76 | 177.80 | 177.84 | 310.84 | 309.74 | 308.65 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.59 | 177.92 | 177.96 | 178.00 | 306.45 | 305.35 | 304.25 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.58 | 178.08 | 178.12 | 178.16 | 302.05 | 300.96 | 299.86 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.57 | 178.24 | 178.28 | 178.32 | 297.66 | 296.56 | 295.47 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.56 | 178.40 | 178.44 | 178.48 | 293.27 | 292.17 | 291.08 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.14 | 186.77 | 186.81 | 186.85 | 68.08 | 67.01 | 65.94 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.13 | 186.93 | 186.97 | 187.01 | 63.79 | 62.71 | 61.64 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.12 | 187.09 | 187.13 | 187.17 | 59.49 | 58.42 | 57.35 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.11 | 187.25 | 187.29 | 187.33 | 55.20 | 54.13 | 53.06 | 0.01515 |
| 1 | 8 | 1 | 10 | 140.10 | 187.41 | 187.45 | 187.49 | 50.91 | 49.84 | 48.77 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.07 | 188.04 | 188.08 | 188.12 | 34.20 | 33.13 | 32.06 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.06 | 188.20 | 188.24 | 188.28 | 29.92 | 28.85 | 27.78 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.06 | 188.36 | 188.40 | 188.44 | 25.64 | 24.57 | 23.50 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.05 | 188.52 | 188.56 | 188.60 | 21.37 | 20.30 | 19.23 | 0.01515 |
| 1 | 1 | 1 | 10 | 140.04 | 188.68 | 188.72 | 188.77 | 17.09 | 16.02 | 14.95 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.61 | 197.17 | 197.21 | 197.25 | 154.88 | 153.84 | 152.80 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.60 | 197.33 | 197.37 | 197.41 | 150.71 | 149.67 | 148.63 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.60 | 197.49 | 197.53 | 197.57 | 146.54 | 145.50 | 144.46 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.59 | 197.65 | 197.69 | 197.73 | 142.38 | 141.34 | 140.29 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.58 | 197.81 | 197.85 | 197.89 | 138.21 | 137.17 | 136.13 | 0.01515 |
| 1 | 1 | 1 | 10 | 139.55 | 198.46 | 198.50 | 198.54 | 121.57 | 120.53 | 119.49 | 0.01515 |
| 1 | 1 | 1 | 10 | 139.54 | 198.62 | 198.66 | 198.70 | 117.41 | 116.37 | 115.33 | 0.01515 |
| 1 | 1 | 1 | 10 | 139.53 | 198.78 | 198.82 | 198.86 | 113.26 | 112.22 | 111.18 | 0.01515 |
| 1 | 1 | 1 | 10 | 139.52 | 198.94 | 198.98 | 199.02 | 109.10 | 108.07 | 107.03 | 0.01515 |
| 1 | 1 | 1 | 10 | 139.51 | 199.10 | 199.14 | 199.18 | 104.95 | 103.92 | 102.88 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.07 | 208.05 | 208.09 | 208.13 | 237.87 | 236.86 | 235.85 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.06 | 208.22 | 208.26 | 208.30 | 233.84 | 232.83 | 231.82 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.05 | 208.38 | 208.42 | 208.46 | 229.81 | 228.80 | 227.79 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.04 | 208.54 | 208.58 | 208.62 | 225.78 | 224.77 | 223.77 | 0.01515 |
| 1 | 8 | 1 | 10 | 139.03 | 208.70 | 208.74 | 208.78 | 221.75 | 220.75 | 219.74 | 0.01515 |
| 1 | 1 | 1 | 10 | 139.00 | 209.39 | 209.43 | 209.47 | 204.46 | 203.46 | 202.46 | 0.01515 |
| 1 | 1 | 1 | 10 | 138.99 | 209.55 | 209.59 | 209.63 | 200.45 | 199.45 | 198.44 | 0.01515 |
| 1 | 1 | 1 | 10 | 138.98 | 209.72 | 209.76 | 209.80 | 196.44 | 195.43 | 194.43 | 0.01515 |
| 1 | 1 | 1 | 10 | 138.98 | 209.88 | 209.92 | 209.96 | 192.43 | 191.42 | 190.42 | 0.01515 |
| 1 | 1 | 1 | 10 | 138.97 | 210.04 | 210.08 | 210.12 | 188.42 | 187.42 | 186.42 | 0.01515 |
| 2 | 8 | 16 | 2 | 135.06 | 290.60 | 290.73 | 290.86 | 210.09 | 207.59 | 205.10 | 0.04665 |
| 2 | 8 | 16 | 2 | 135.03 | 291.12 | 291.25 | 291.37 | 200.11 | 197.62 | 195.14 | 0.04665 |
| 2 | 1 | 16 | 2 | 134.98 | 292.26 | 292.39 | 292.52 | 178.07 | 175.59 | 173.12 | 0.04665 |
| 2 | 1 | 16 | 2 | 134.95 | 292.78 | 292.91 | 293.04 | 168.17 | 165.70 | 163.23 | 0.04665 |
| 2 | 8 | 21 | 2 | 134.68 | 298.65 | 298.82 | 298.99 | 57.16 | 54.01 | 50.86 | 0.06122 |
| 2 | 8 | 21 | 2 | 134.64 | 299.33 | 299.50 | 299.67 | 44.57 | 41.43 | 38.29 | 0.06122 |
| 2 | 1 | 21 | 2 | 134.57 | 300.88 | 301.05 | 301.22 | 15.94 | 12.83 | 9.72 | 0.06122 |
| 2 | 1 | 21 | 2 | 134.54 | 301.56 | 301.73 | 301.90 | 3.50 | 0.39 | 357.29 | 0.06122 |
| 2 | 8 | 20 | 2 | 133.65 | 320.74 | 320.90 | 321.06 | 29.69 | 27.05 | 24.40 | 0.05831 |
| 2 | 8 | 20 | 2 | 133.62 | 321.39 | 321.55 | 321.71 | 19.12 | 16.48 | 13.85 | 0.05831 |
| 2 | 8 | 20 | 2 | 133.59 | 322.03 | 322.19 | 322.35 | 8.59 | 5.96 | 3.34 | 0.05831 |
| 2 | 1 | 21 | 2 | 133.52 | 323.54 | 323.71 | 323.88 | 344.13 | 341.40 | 338.67 | 0.06122 |
| 2 | 1 | 21 | 2 | 133.48 | 324.22 | 324.39 | 324.56 | 333.22 | 330.49 | 327.78 | 0.06122 |
| 2 | 1 | 21 | 2 | 133.45 | 324.89 | 325.06 | 325.23 | 322.35 | 319.64 | 316.93 | 0.06122 |
| 2 | 8 | 18 | 2 | 132.90 | 336.84 | 336.98 | 337.13 | 138.37 | 136.22 | 134.08 | 0.05248 |
| 2 | 8 | 18 | 2 | 132.88 | 337.42 | 337.56 | 337.71 | 129.79 | 127.65 | 125.51 | 0.05248 |
| 2 | 1 | 18 | 2 | 132.82 | 338.76 | 338.91 | 339.05 | 110.00 | 107.88 | 105.76 | 0.05248 |
| 2 | 1 | 18 | 2 | 132.79 | 339.35 | 339.49 | 339.64 | 101.53 | 99.42 | 97.31 | 0.05248 |
| 3 | 1 | 23 | 2 | 128.47 | 436.63 | 436.81 | 437.00 | 143.69 | 142.34 | 141.00 | 0.04883 |
| 3 | 1 | 23 | 2 | 128.43 | 437.37 | 437.55 | 437.74 | 138.32 | 136.99 | 135.65 | 0.04883 |
| 3 | 1 | 23 | 2 | 128.40 | 438.11 | 438.29 | 438.48 | 132.99 | 131.67 | 130.35 | 0.04883 |
| 3 | 8 | 23 | 2 | 127.53 | 458.50 | 458.69 | 458.87 | 3.83 | 2.81 | 1.79 | 0.04883 |
| 3 | 8 | 23 | 2 | 127.50 | 459.25 | 459.43 | 459.62 | 359.75 | 358.74 | 357.73 | 0.04883 |
| 3 | 8 | 23 | 2 | 127.47 | 459.99 | 460.17 | 460.36 | 355.71 | 354.71 | 353.72 | 0.04883 |
| 3 | 1 | 23 | 2 | 127.40 | 461.62 | 461.81 | 461.99 | 346.99 | 346.02 | 345.05 | 0.04883 |
| 3 | 1 | 23 | 2 | 127.37 | 462.37 | 462.55 | 462.74 | 343.11 | 342.15 | 341.18 | 0.04883 |
| 3 | 1 | 23 | 2 | 127.34 | 463.11 | 463.29 | 463.48 | 339.27 | 338.31 | 337.36 | 0.04883 |
| 3 | 8 | 21 | 2 | 126.48 | 483.57 | 483.74 | 483.91 | 249.75 | 249.14 | 248.53 | 0.04459 |
| 3 | 8 | 21 | 2 | 126.45 | 484.25 | 484.42 | 484.59 | 247.32 | 246.71 | 246.11 | 0.04459 |
| 3 | 8 | 21 | 2 | 126.42 | 484.93 | 485.10 | 485.27 | 244.91 | 244.31 | 243.72 | 0.04459 |
| 3 | 8 | 21 | 2 | 126.40 | 485.61 | 485.78 | 485.94 | 242.53 | 241.94 | 241.35 | 0.04459 |
| 3 | 1 | 21 | 2 | 126.33 | 487.15 | 487.32 | 487.49 | 237.25 | 236.68 | 236.11 | 0.04459 |
| 3 | 1 | 21 | 2 | 126.30 | 487.82 | 487.99 | 488.16 | 234.98 | 234.42 | 233.86 | 0.04459 |
| 3 | 1 | 21 | 2 | 126.27 | 488.50 | 488.67 | 488.84 | 232.75 | 232.19 | 231.64 | 0.04459 |
| 3 | 1 | 21 | 2 | 126.25 | 489.18 | 489.35 | 489.52 | 230.54 | 230.00 | 229.45 | 0.04459 |
| 3 | 8 | 24 | 2 | 125.05 | 518.10 | 518.30 | 518.49 | 164.95 | 164.69 | 164.43 | 0.05096 |
| 3 | 8 | 24 | 2 | 125.02 | 518.88 | 519.07 | 519.27 | 163.92 | 163.67 | 163.41 | 0.05096 |
| 3 | 8 | 24 | 2 | 124.99 | 519.65 | 519.85 | 520.04 | 162.91 | 162.66 | 162.42 | 0.05096 |
| 3 | 8 | 24 | 2 | 124.96 | 520.43 | 520.62 | 520.81 | 161.94 | 161.71 | 161.48 | 0.05096 |
| 4 | 8 | 23 | 2 | 123.50 | 556.41 | 556.60 | 556.78 | 155.39 | 155.55 | 155.71 | 0.04656 |
| 4 | 8 | 23 | 2 | 123.47 | 557.15 | 557.34 | 557.52 | 156.04 | 156.20 | 156.37 | 0.04656 |
| 4 | 8 | 23 | 2 | 123.44 | 557.89 | 558.08 | 558.27 | 156.72 | 156.89 | 157.07 | 0.04656 |
| 4 | 8 | 23 | 2 | 123.41 | 558.64 | 558.82 | 559.01 | 157.43 | 157.61 | 157.80 | 0.04656 |
| 4 | 8 | 23 | 2 | 123.38 | 559.38 | 559.56 | 559.75 | 158.18 | 158.37 | 158.57 | 0.04656 |
| 4 | 1 | 23 | 2 | 123.31 | 561.10 | 561.28 | 561.47 | 160.03 | 160.24 | 160.45 | 0.04656 |
| 4 | 1 | 23 | 2 | 123.28 | 561.84 | 562.02 | 562.21 | 160.88 | 161.09 | 161.31 | 0.04656 |
| 4 | 1 | 23 | 2 | 123.26 | 562.58 | 562.77 | 562.95 | 161.76 | 161.99 | 162.21 | 0.04656 |
| 4 | 1 | 23 | 2 | 123.23 | 563.32 | 563.51 | 563.69 | 162.68 | 162.91 | 163.15 | 0.04656 |
| 4 | 1 | 23 | 2 | 123.20 | 564.06 | 564.25 | 564.43 | 163.63 | 163.87 | 164.12 | 0.04656 |
| 4 | 8 | 22 | 2 | 121.73 | 601.20 | 601.38 | 601.56 | 252.56 | 253.18 | 253.80 | 0.04453 |
| 4 | 8 | 22 | 2 | 121.70 | 601.91 | 602.09 | 602.27 | 255.05 | 255.67 | 256.30 | 0.04453 |
| 4 | 8 | 22 | 2 | 121.67 | 602.62 | 602.80 | 602.98 | 257.57 | 258.20 | 258.84 | 0.04453 |
| 4 | 8 | 22 | 2 | 121.65 | 603.33 | 603.51 | 603.69 | 260.12 | 260.76 | 261.41 | 0.04453 |
| 4 | 8 | 22 | 2 | 121.62 | 604.04 | 604.22 | 604.40 | 262.71 | 263.35 | 264.01 | 0.04453 |
| 4 | 8 | 22 | 2 | 121.59 | 604.75 | 604.93 | 605.11 | 265.31 | 265.97 | 266.63 | 0.04453 |
| 4 | 1 | 22 | 2 | 121.52 | 606.56 | 606.74 | 606.91 | 272.08 | 272.75 | 273.42 | 0.04453 |
| 4 | 1 | 22 | 2 | 121.49 | 607.27 | 607.45 | 607.62 | 274.78 | 275.46 | 276.14 | 0.04453 |
| 4 | 1 | 22 | 2 | 121.46 | 607.98 | 608.16 | 608.33 | 277.51 | 278.20 | 278.89 | 0.04453 |
| 4 | 1 | 22 | 2 | 121.44 | 608.69 | 608.87 | 609.04 | 280.27 | 280.96 | 281.66 | 0.04453 |
| 4 | 1 | 22 | 2 | 121.41 | 609.40 | 609.57 | 609.75 | 283.06 | 283.76 | 284.46 | 0.04453 |
| 4 | 1 | 22 | 2 | 121.38 | 610.11 | 610.28 | 610.46 | 285.87 | 286.58 | 287.29 | 0.04453 |
| 8 | 8 | 21 | 2 | 113.15 | 834.22 | 834.39 | 834.55 | 300.18 | 302.76 | 305.34 | 0.05512 |
| 8 | 8 | 21 | 2 | 113.13 | 834.89 | 835.06 | 835.23 | 310.51 | 313.09 | 315.68 | 0.05512 |
| 8 | 8 | 21 | 2 | 113.10 | 835.57 | 835.74 | 835.91 | 320.86 | 323.45 | 326.04 | 0.05512 |
| 8 | 1 | 20 | 2 | 113.05 | 837.13 | 837.29 | 837.45 | 344.72 | 347.20 | 349.69 | 0.05249 |
| 8 | 1 | 20 | 2 | 113.03 | 837.77 | 837.93 | 838.09 | 354.66 | 357.14 | 359.63 | 0.05249 |
| 8 | 1 | 20 | 2 | 113.01 | 838.42 | 838.58 | 838.74 | 4.61 | 7.10 | 9.59 | 0.05249 |
| 8 | 8 | 22 | 2 | 112.27 | 859.88 | 860.05 | 860.23 | 346.55 | 349.46 | 352.38 | 0.05774 |
| 8 | 8 | 22 | 2 | 112.25 | 860.59 | 860.76 | 860.94 | 358.22 | 1.14 | 4.06 | 0.05774 |
| 8 | 8 | 22 | 2 | 112.22 | 861.30 | 861.47 | 861.65 | 9.92 | 12.85 | 15.78 | 0.05774 |
| 8 | 1 | 22 | 2 | 112.17 | 862.94 | 863.12 | 863.29 | 37.11 | 40.05 | 42.99 | 0.05774 |
| 8 | 1 | 22 | 2 | 112.14 | 863.65 | 863.83 | 864.00 | 48.87 | 51.82 | 54.76 | 0.05774 |
| 8 | 1 | 22 | 2 | 112.12 | 864.36 | 864.54 | 864.71 | 60.66 | 63.61 | 66.56 | 0.05774 |
| 8 | 8 | 21 | 2 | 111.51 | 882.43 | 882.60 | 882.77 | 8.60 | 11.55 | 14.51 | 0.05512 |
| 8 | 8 | 21 | 2 | 111.49 | 883.11 | 883.28 | 883.45 | 20.43 | 23.39 | 26.35 | 0.05512 |
| 8 | 8 | 21 | 2 | 111.47 | 883.79 | 883.96 | 884.12 | 32.28 | 35.25 | 38.22 | 0.05512 |
| 8 | 1 | 21 | 2 | 111.41 | 885.32 | 885.49 | 885.66 | 59.17 | 62.14 | 65.12 | 0.05512 |
| 8 | 1 | 21 | 2 | 111.39 | 886.00 | 886.16 | 886.33 | 71.09 | 74.07 | 77.06 | 0.05512 |
| 8 | 1 | 21 | 2 | 111.37 | 886.67 | 886.84 | 887.01 | 83.03 | 86.03 | 89.02 | 0.05512 |
| 9 | 8 | 20 | 2 | 107.67 | 959.04 | 959.20 | 959.36 | 119.92 | 123.81 | 127.69 | 0.05391 |
| 9 | 8 | 20 | 2 | 107.62 | 959.68 | 959.84 | 960.00 | 135.48 | 139.38 | 143.29 | 0.05391 |
| 9 | 1 | 20 | 2 | 107.52 | 961.12 | 961.28 | 961.44 | 170.43 | 174.38 | 178.34 | 0.05391 |
| 9 | 1 | 20 | 2 | 107.48 | 961.76 | 961.92 | 962.09 | 186.28 | 190.26 | 194.24 | 0.05391 |
| 9 | 8 | 19 | 2 | 106.51 | 975.34 | 975.50 | 975.65 | 181.87 | 186.07 | 190.28 | 0.05121 |
| 9 | 8 | 19 | 2 | 106.46 | 975.96 | 976.11 | 976.26 | 198.72 | 202.94 | 207.17 | 0.05121 |
| 9 | 1 | 19 | 2 | 106.36 | 977.37 | 977.53 | 977.68 | 237.97 | 242.22 | 246.48 | 0.05121 |
| 9 | 1 | 19 | 2 | 106.32 | 977.99 | 978.14 | 978.29 | 255.01 | 259.28 | 263.56 | 0.05121 |
| 9 | 8 | 18 | 2 | 106.00 | 982.37 | 982.52 | 982.66 | 19.10 | 23.27 | 27.44 | 0.04852 |
| 9 | 8 | 18 | 2 | 105.96 | 982.96 | 983.10 | 983.25 | 35.81 | 40.00 | 44.19 | 0.04852 |
| 9 | 1 | 18 | 2 | 105.86 | 984.30 | 984.45 | 984.59 | 74.85 | 79.08 | 83.32 | 0.04852 |
| 9 | 1 | 18 | 2 | 105.81 | 984.88 | 985.03 | 985.17 | 91.82 | 96.07 | 100.33 | 0.04852 |
| 9 | 8 | 21 | 2 | 105.28 | 992.08 | 992.24 | 992.41 | 308.73 | 313.98 | 319.24 | 0.05660 |
| 9 | 1 | 20 | 2 | 105.18 | 993.53 | 993.69 | 993.85 | 353.95 | 359.00 | 4.06 | 0.05391 |
| 9 | 8 | 21 | 2 | 104.37 | 1004.53 | 1004.70 | 1004.86 | 347.60 | 353.18 | 358.76 | 0.05660 |
| 9 | 1 | 21 | 2 | 104.26 | 1006.00 | 1006.17 | 1006.34 | 36.32 | 41.92 | 47.53 | 0.05660 |
| 9 | 8 | 15 | 2 | 103.51 | 1016.24 | 1016.36 | 1016.49 | 20.94 | 25.08 | 29.21 | 0.04043 |
| 9 | 8 | 15 | 2 | 103.47 | 1016.73 | 1016.85 | 1016.97 | 37.49 | 41.63 | 45.76 | 0.04043 |
| 9 | 1 | 15 | 2 | 103.39 | 1017.87 | 1017.99 | 1018.11 | 76.72 | 80.87 | 85.03 | 0.04043 |
| 9 | 1 | 15 | 2 | 103.35 | 1018.36 | 1018.48 | 1018.60 | 93.34 | 97.49 | 101.65 | 0.04043 |
| 10 | 8 | 19 | 2 | 101.03 | 1050.43 | 1050.58 | 1050.74 | 159.46 | 165.15 | 170.84 | 0.06909 |
| 10 | 1 | 19 | 2 | 100.93 | 1051.77 | 1051.92 | 1052.07 | 209.23 | 214.94 | 220.65 | 0.06909 |
| 10 | 8 | 19 | 2 | 100.28 | 1060.94 | 1061.09 | 1061.24 | 194.83 | 200.68 | 206.54 | 0.06909 |
| 10 | 1 | 19 | 2 | 100.19 | 1062.25 | 1062.40 | 1062.56 | 245.09 | 250.97 | 256.84 | 0.06909 |
| 10 | 8 | 18 | 2 | 99.57 | 1070.98 | 1071.12 | 1071.27 | 223.49 | 229.20 | 234.90 | 0.06545 |
| 10 | 1 | 18 | 2 | 99.49 | 1072.25 | 1072.40 | 1072.54 | 273.65 | 279.38 | 285.11 | 0.06545 |
| 10 | 8 | 17 | 2 | 98.90 | 1080.69 | 1080.83 | 1080.97 | 250.95 | 256.50 | 262.06 | 0.06182 |
| 10 | 1 | 17 | 2 | 98.81 | 1081.93 | 1082.06 | 1082.20 | 300.99 | 306.56 | 312.14 | 0.06182 |
| 10 | 8 | 17 | 2 | 98.73 | 1083.17 | 1083.31 | 1083.44 | 351.58 | 357.19 | 2.79 | 0.06182 |
| 10 | 1 | 17 | 2 | 98.65 | 1084.36 | 1084.50 | 1084.64 | 40.42 | 46.04 | 51.66 | 0.06182 |
| 10 | 8 | 16 | 2 | 98.16 | 1091.38 | 1091.51 | 1091.64 | 331.44 | 336.85 | 342.27 | 0.05818 |
| 10 | 1 | 17 | 2 | 98.08 | 1092.53 | 1092.67 | 1092.81 | 19.94 | 25.72 | 31.50 | 0.06182 |
| 10 | 8 | 16 | 2 | 97.55 | 1100.42 | 1100.55 | 1100.68 | 356.80 | 2.39 | 7.97 | 0.05818 |
| 10 | 1 | 16 | 2 | 97.47 | 1101.56 | 1101.68 | 1101.81 | 46.11 | 51.71 | 57.32 | 0.05818 |
| 10 | 8 | 15 | 2 | 96.95 | 1109.24 | 1109.36 | 1109.48 | 23.97 | 29.35 | 34.74 | 0.05455 |
| 10 | 1 | 15 | 2 | 96.88 | 1110.34 | 1110.47 | 1110.59 | 73.20 | 78.60 | 84.01 | 0.05455 |
| 11 | 8 | 14 | 2 | 94.76 | 1142.51 | 1142.62 | 1142.73 | 135.99 | 141.50 | 147.01 | 0.06009 |
| 11 | 1 | 14 | 2 | 94.70 | 1143.51 | 1143.62 | 1143.73 | 184.84 | 190.36 | 195.89 | 0.06009 |
| 11 | 8 | 14 | 2 | 94.25 | 1150.41 | 1150.52 | 1150.63 | 165.47 | 171.08 | 176.70 | 0.06009 |
| 11 | 1 | 14 | 2 | 94.19 | 1151.40 | 1151.51 | 1151.63 | 214.91 | 220.55 | 226.18 | 0.06009 |
| 11 | 8 | 13 | 2 | 93.76 | 1158.19 | 1158.30 | 1158.40 | 196.74 | 202.06 | 207.39 | 0.05579 |
| 11 | 1 | 13 | 2 | 93.70 | 1159.17 | 1159.27 | 1159.38 | 246.33 | 251.66 | 257.00 | 0.05579 |
| 11 | 8 | 13 | 2 | 93.58 | 1161.03 | 1161.13 | 1161.24 | 341.32 | 346.68 | 352.04 | 0.05579 |
| 11 | 1 | 13 | 2 | 93.52 | 1161.97 | 1162.08 | 1162.18 | 29.59 | 34.96 | 40.33 | 0.05579 |
| 11 | 8 | 13 | 2 | 93.12 | 1168.31 | 1168.42 | 1168.52 | 356.60 | 2.05 | 7.51 | 0.05579 |
| 11 | 1 | 13 | 2 | 93.06 | 1169.26 | 1169.37 | 1169.47 | 46.11 | 51.58 | 57.05 | 0.05579 |
| 11 | 8 | 13 | 2 | 92.66 | 1175.65 | 1175.75 | 1175.86 | 21.26 | 26.81 | 32.35 | 0.05579 |
| 11 | 1 | 13 | 2 | 92.60 | 1176.58 | 1176.68 | 1176.79 | 70.35 | 75.90 | 81.46 | 0.05579 |
| 11 | 8 | 13 | 2 | 92.21 | 1182.95 | 1183.05 | 1183.16 | 50.52 | 56.16 | 61.81 | 0.05579 |
| 11 | 1 | 12 | 2 | 92.15 | 1183.88 | 1183.98 | 1184.08 | 100.91 | 106.13 | 111.35 | 0.05150 |
| 11 | 8 | 12 | 2 | 91.77 | 1190.04 | 1190.13 | 1190.23 | 75.42 | 80.72 | 86.03 | 0.05150 |
| 11 | 1 | 12 | 2 | 91.72 | 1190.92 | 1191.02 | 1191.12 | 124.08 | 129.40 | 134.71 | 0.05150 |
| 11 | 8 | 12 | 2 | 91.35 | 1196.95 | 1197.05 | 1197.15 | 97.83 | 103.22 | 108.61 | 0.05150 |
| 11 | 1 | 12 | 2 | 91.30 | 1197.83 | 1197.93 | 1198.03 | 146.86 | 152.26 | 157.66 | 0.05150 |
| 12 | 8 | 10 | 2 | 89.25 | 1232.29 | 1232.37 | 1232.45 | 329.73 | 334.45 | 339.17 | 0.02525 |
| 12 | 1 | 10 | 2 | 89.20 | 1233.06 | 1233.14 | 1233.22 | 15.06 | 19.78 | 24.51 | 0.02525 |
| 12 | 8 | 21 | 2 | 88.89 | 1238.33 | 1238.50 | 1238.67 | 323.94 | 333.86 | 343.77 | 0.05303 |
| 12 | 1 | 21 | 2 | 88.85 | 1239.08 | 1239.25 | 1239.42 | 7.84 | 17.76 | 27.67 | 0.05303 |
| 12 | 8 | 21 | 2 | 88.54 | 1244.30 | 1244.47 | 1244.64 | 312.95 | 322.84 | 332.73 | 0.05303 |
| 12 | 1 | 21 | 2 | 88.50 | 1245.05 | 1245.22 | 1245.39 | 356.75 | 6.64 | 16.53 | 0.05303 |
| 12 | 8 | 10 | 2 | 88.19 | 1250.57 | 1250.65 | 1250.73 | 318.14 | 322.83 | 327.52 | 0.02525 |
| 12 | 1 | 10 | 2 | 88.14 | 1251.34 | 1251.42 | 1251.50 | 3.14 | 7.83 | 12.51 | 0.02525 |
| 12 | 8 | 10 | 2 | 87.84 | 1256.64 | 1256.72 | 1256.80 | 310.30 | 314.96 | 319.63 | 0.02525 |
| 12 | 1 | 10 | 2 | 87.80 | 1257.41 | 1257.49 | 1257.57 | 354.61 | 359.27 | 3.93 | 0.02525 |
| 12 | 8 | 21 | 2 | 87.49 | 1262.69 | 1262.86 | 1263.03 | 299.28 | 309.03 | 318.78 | 0.05303 |
| 12 | 1 | 21 | 2 | 87.45 | 1263.44 | 1263.61 | 1263.78 | 342.43 | 352.17 | 1.92 | 0.05303 |
| 12 | 8 | 21 | 2 | 87.15 | 1268.69 | 1268.86 | 1269.03 | 283.86 | 293.57 | 303.28 | 0.05303 |
| 12 | 1 | 21 | 2 | 87.11 | 1269.44 | 1269.61 | 1269.78 | 326.84 | 336.54 | 346.25 | 0.05303 |
| 12 | 8 | 21 | 2 | 86.82 | 1274.58 | 1274.75 | 1274.92 | 261.20 | 270.88 | 280.56 | 0.05303 |
| 12 | 1 | 21 | 2 | 86.78 | 1275.32 | 1275.49 | 1275.66 | 303.59 | 313.26 | 322.93 | 0.05303 |
| 12 | 8 | 21 | 2 | 86.50 | 1280.42 | 1280.59 | 1280.76 | 234.32 | 243.97 | 253.63 | 0.05303 |
| 12 | 1 | 21 | 2 | 86.46 | 1281.15 | 1281.32 | 1281.49 | 276.14 | 285.78 | 295.43 | 0.05303 |
| 12 | 8 | 21 | 2 | 86.18 | 1286.24 | 1286.41 | 1286.58 | 205.62 | 215.24 | 224.87 | 0.05303 |
| 12 | 1 | 21 | 2 | 86.14 | 1286.97 | 1287.14 | 1287.31 | 246.84 | 256.46 | 266.09 | 0.05303 |
| 12 | 8 | 21 | 2 | 86.08 | 1287.92 | 1288.09 | 1288.26 | 300.90 | 310.51 | 320.13 | 0.05303 |
| 12 | 1 | 21 | 2 | 86.04 | 1288.65 | 1288.81 | 1288.98 | 342.10 | 351.71 | 1.33 | 0.05303 |
| 13 | 8 | 20 | 2 | 84.19 | 1322.93 | 1323.10 | 1323.26 | 118.37 | 127.44 | 136.52 | 0.04454 |
| 13 | 1 | 20 | 2 | 84.16 | 1323.64 | 1323.81 | 1323.97 | 158.30 | 167.38 | 176.45 | 0.04454 |
| 13 | 8 | 20 | 2 | 84.03 | 1326.10 | 1326.26 | 1326.43 | 296.62 | 305.68 | 314.74 | 0.04454 |
| 13 | 1 | 20 | 2 | 83.99 | 1326.82 | 1326.98 | 1327.14 | 336.95 | 346.01 | 355.06 | 0.04454 |
| 13 | 8 | 22 | 2 | 83.67 | 1332.95 | 1333.13 | 1333.30 | 320.58 | 330.50 | 340.41 | 0.04900 |
| 13 | 1 | 22 | 2 | 83.63 | 1333.72 | 1333.90 | 1334.08 | 3.86 | 13.77 | 23.68 | 0.04900 |
| 13 | 8 | 20 | 2 | 83.36 | 1338.93 | 1339.09 | 1339.26 | 294.24 | 303.20 | 312.15 | 0.04454 |
| 13 | 1 | 20 | 2 | 83.32 | 1339.64 | 1339.80 | 1339.97 | 333.65 | 342.60 | 351.55 | 0.04454 |
| 13 | 8 | 20 | 2 | 83.06 | 1344.64 | 1344.80 | 1344.96 | 250.32 | 259.21 | 268.10 | 0.04454 |
| 13 | 1 | 20 | 2 | 83.02 | 1345.36 | 1345.52 | 1345.68 | 289.87 | 298.75 | 307.62 | 0.04454 |
| 13 | 8 | 20 | 2 | 82.76 | 1350.36 | 1350.52 | 1350.68 | 203.93 | 212.73 | 221.52 | 0.04454 |
| 13 | 1 | 20 | 2 | 82.73 | 1351.07 | 1351.23 | 1351.39 | 242.62 | 251.41 | 260.19 | 0.04454 |
| 13 | 8 | 20 | 2 | 82.47 | 1356.07 | 1356.23 | 1356.39 | 153.56 | 162.25 | 170.93 | 0.04454 |
| 13 | 1 | 20 | 2 | 82.44 | 1356.79 | 1356.95 | 1357.11 | 192.19 | 200.86 | 209.53 | 0.04454 |
| 13 | 8 | 21 | 2 | 82.18 | 1361.86 | 1362.03 | 1362.20 | 103.22 | 112.21 | 121.20 | 0.04677 |
| 13 | 1 | 20 | 2 | 82.14 | 1362.58 | 1362.75 | 1362.91 | 141.70 | 150.23 | 158.76 | 0.04454 |
| 13 | 8 | 20 | 2 | 81.89 | 1367.63 | 1367.79 | 1367.96 | 47.13 | 55.54 | 63.95 | 0.04454 |
| 13 | 1 | 20 | 2 | 81.85 | 1368.35 | 1368.51 | 1368.67 | 84.56 | 92.96 | 101.35 | 0.04454 |
| 13 | 8 | 21 | 2 | 81.59 | 1373.45 | 1373.62 | 1373.79 | 348.02 | 356.71 | 5.41 | 0.04677 |
| 13 | 1 | 21 | 2 | 81.56 | 1374.17 | 1374.34 | 1374.51 | 25.24 | 33.91 | 42.58 | 0.04677 |
| 13 | 8 | 21 | 2 | 81.30 | 1379.28 | 1379.45 | 1379.62 | 284.72 | 293.27 | 301.81 | 0.04677 |
| 13 | 1 | 21 | 2 | 81.27 | 1380.00 | 1380.17 | 1380.34 | 321.31 | 329.84 | 338.37 | 0.04677 |
| 14 | 8 | 13 | 2 | 79.82 | 1409.58 | 1409.69 | 1409.79 | 326.60 | 331.63 | 336.65 | 0.03163 |
| 14 | 1 | 13 | 2 | 79.77 | 1410.54 | 1410.65 | 1410.75 | 12.52 | 17.54 | 22.55 | 0.03163 |
| 14 | 8 | 12 | 2 | 79.48 | 1416.72 | 1416.82 | 1416.91 | 307.53 | 312.15 | 316.76 | 0.02920 |
| 14 | 1 | 13 | 2 | 79.44 | 1417.61 | 1417.72 | 1417.82 | 350.21 | 355.21 | 0.20 | 0.03163 |
| 14 | 8 | 22 | 2 | 79.22 | 1422.13 | 1422.31 | 1422.48 | 205.30 | 213.74 | 222.18 | 0.05353 |
| 14 | 1 | 22 | 2 | 79.18 | 1422.90 | 1423.07 | 1423.25 | 241.75 | 250.19 | 258.63 | 0.05353 |
| 14 | 8 | 22 | 2 | 78.93 | 1428.26 | 1428.44 | 1428.61 | 136.63 | 145.05 | 153.47 | 0.05353 |
| 14 | 1 | 22 | 2 | 78.89 | 1429.02 | 1429.19 | 1429.37 | 172.61 | 181.02 | 189.44 | 0.05353 |
| 14 | 8 | 22 | 2 | 78.63 | 1434.49 | 1434.67 | 1434.85 | 72.09 | 80.48 | 88.87 | 0.05353 |
| 14 | 1 | 22 | 2 | 78.60 | 1435.27 | 1435.44 | 1435.62 | 108.70 | 117.08 | 125.47 | 0.05353 |
| 14 | 8 | 22 | 2 | 78.34 | 1440.77 | 1440.94 | 1441.12 | 8.18 | 16.52 | 24.86 | 0.05353 |
| 14 | 1 | 22 | 2 | 78.30 | 1441.55 | 1441.73 | 1441.90 | 44.96 | 53.30 | 61.63 | 0.05353 |
| 14 | 8 | 23 | 2 | 78.05 | 1447.06 | 1447.25 | 1447.43 | 303.44 | 312.11 | 320.77 | 0.05596 |
| 14 | 1 | 23 | 2 | 78.01 | 1447.85 | 1448.04 | 1448.22 | 340.33 | 348.99 | 357.64 | 0.05596 |
| 14 | 8 | 23 | 2 | 77.75 | 1453.38 | 1453.56 | 1453.75 | 236.96 | 245.53 | 254.11 | 0.05596 |
| 14 | 1 | 23 | 2 | 77.72 | 1454.17 | 1454.35 | 1454.54 | 273.48 | 282.05 | 290.62 | 0.05596 |
| 14 | 8 | 23 | 2 | 77.46 | 1459.80 | 1459.98 | 1460.17 | 172.22 | 180.71 | 189.18 | 0.05596 |
| 14 | 1 | 23 | 2 | 77.42 | 1460.59 | 1460.78 | 1460.97 | 208.68 | 217.15 | 225.60 | 0.05596 |
| 14 | 8 | 23 | 2 | 77.17 | 1466.30 | 1466.48 | 1466.67 | 107.52 | 115.89 | 124.26 | 0.05596 |
| 14 | 1 | 23 | 2 | 77.13 | 1467.10 | 1467.29 | 1467.47 | 143.88 | 152.23 | 160.58 | 0.05596 |
| 15 | 8 | 16 | 2 | 73.98 | 1539.63 | 1539.76 | 1539.89 | 329.23 | 334.45 | 339.67 | 0.07512 |
| 15 | 1 | 16 | 2 | 73.94 | 1540.75 | 1540.88 | 1541.01 | 14.59 | 19.80 | 25.03 | 0.07512 |
| 15 | 8 | 13 | 2 | 73.72 | 1545.91 | 1546.01 | 1546.12 | 222.87 | 227.09 | 231.30 | 0.06103 |
| 15 | 1 | 12 | 2 | 73.69 | 1546.84 | 1546.94 | 1547.04 | 260.39 | 264.26 | 268.13 | 0.05634 |
| 15 | 8 | 13 | 2 | 73.42 | 1553.34 | 1553.44 | 1553.55 | 157.52 | 161.63 | 165.74 | 0.06103 |
| 15 | 1 | 13 | 2 | 73.38 | 1554.26 | 1554.37 | 1554.47 | 193.84 | 197.94 | 202.04 | 0.06103 |
| 15 | 8 | 13 | 2 | 73.11 | 1560.75 | 1560.85 | 1560.96 | 83.78 | 87.76 | 91.73 | 0.06103 |
| 15 | 1 | 13 | 2 | 73.08 | 1561.67 | 1561.78 | 1561.88 | 118.89 | 122.86 | 126.82 | 0.06103 |
| 15 | 8 | 13 | 2 | 72.81 | 1568.29 | 1568.39 | 1568.50 | 4.89 | 8.74 | 12.58 | 0.06103 |
| 15 | 1 | 13 | 2 | 72.77 | 1569.23 | 1569.34 | 1569.44 | 39.45 | 43.29 | 47.12 | 0.06103 |
| 15 | 8 | 13 | 2 | 72.50 | 1575.92 | 1576.02 | 1576.13 | 283.56 | 287.36 | 291.15 | 0.06103 |
| 15 | 1 | 13 | 2 | 72.46 | 1576.87 | 1576.97 | 1577.08 | 317.86 | 321.63 | 325.40 | 0.06103 |
| 15 | 8 | 13 | 2 | 72.19 | 1583.66 | 1583.77 | 1583.87 | 199.55 | 203.27 | 206.99 | 0.06103 |
| 15 | 1 | 13 | 2 | 72.15 | 1584.62 | 1584.73 | 1584.83 | 233.52 | 237.23 | 240.93 | 0.06103 |
| 15 | 8 | 13 | 2 | 71.88 | 1591.50 | 1591.61 | 1591.71 | 114.34 | 118.00 | 121.65 | 0.06103 |
| 15 | 1 | 13 | 2 | 71.84 | 1592.48 | 1592.58 | 1592.69 | 148.34 | 152.00 | 155.65 | 0.06103 |
| 16 | 8 | 14 | 2 | 70.53 | 1626.46 | 1626.57 | 1626.68 | 264.54 | 268.50 | 272.47 | 0.06731 |
| 16 | 1 | 14 | 2 | 70.49 | 1627.50 | 1627.61 | 1627.72 | 301.05 | 305.01 | 308.98 | 0.06731 |
| 16 | 8 | 15 | 2 | 70.22 | 1634.89 | 1635.01 | 1635.13 | 200.04 | 204.28 | 208.52 | 0.07212 |
| 16 | 1 | 15 | 2 | 70.18 | 1635.95 | 1636.07 | 1636.19 | 237.02 | 241.25 | 245.47 | 0.07212 |
| 16 | 8 | 15 | 2 | 69.90 | 1643.45 | 1643.58 | 1643.70 | 135.81 | 139.98 | 144.15 | 0.07212 |
| 16 | 1 | 15 | 2 | 69.86 | 1644.53 | 1644.65 | 1644.77 | 172.77 | 176.93 | 181.12 | 0.07212 |
| 16 | 8 | 15 | 2 | 69.59 | 1652.03 | 1652.15 | 1652.27 | 72.95 | 77.16 | 81.36 | 0.07212 |
| 16 | 1 | 15 | 2 | 69.55 | 1653.10 | 1653.22 | 1653.34 | 110.18 | 114.37 | 118.56 | 0.07212 |
| 16 | 8 | 15 | 2 | 69.27 | 1660.78 | 1660.90 | 1661.03 | 10.46 | 14.38 | 18.29 | 0.07212 |
| 16 | 1 | 15 | 2 | 69.23 | 1661.87 | 1661.99 | 1662.11 | 45.46 | 49.33 | 53.22 | 0.07212 |
| 16 | 8 | 15 | 2 | 68.96 | 1669.48 | 1669.60 | 1669.72 | 291.39 | 295.31 | 299.23 | 0.07212 |
| 16 | 1 | 15 | 2 | 68.92 | 1670.57 | 1670.69 | 1670.81 | 326.66 | 330.58 | 334.50 | 0.07212 |
| 16 | 8 | 15 | 2 | 68.64 | 1678.36 | 1678.48 | 1678.61 | 219.83 | 223.77 | 227.71 | 0.07212 |
| 16 | 1 | 15 | 2 | 68.61 | 1679.46 | 1679.58 | 1679.70 | 255.56 | 259.51 | 263.45 | 0.07212 |
| 17 | 8 | 16 | 2 | 67.36 | 1715.44 | 1715.57 | 1715.70 | 352.43 | 356.61 | 0.78 | 0.06809 |
| 17 | 1 | 17 | 2 | 67.32 | 1716.60 | 1716.73 | 1716.87 | 29.67 | 34.10 | 38.51 | 0.07234 |
| 17 | 8 | 17 | 2 | 67.05 | 1724.80 | 1724.93 | 1725.07 | 289.51 | 293.73 | 297.95 | 0.07234 |
| 17 | 1 | 16 | 2 | 67.01 | 1725.99 | 1726.12 | 1726.25 | 326.10 | 330.04 | 333.97 | 0.06809 |
| 17 | 8 | 16 | 2 | 66.74 | 1734.19 | 1734.32 | 1734.45 | 213.23 | 217.12 | 221.01 | 0.06809 |
| 17 | 1 | 17 | 2 | 66.70 | 1735.35 | 1735.48 | 1735.62 | 248.08 | 252.23 | 256.38 | 0.07234 |
| 17 | 8 | 17 | 2 | 66.43 | 1743.77 | 1743.91 | 1744.05 | 148.32 | 152.57 | 156.81 | 0.07234 |
| 17 | 1 | 17 | 2 | 66.39 | 1744.96 | 1745.09 | 1745.23 | 185.04 | 189.29 | 193.54 | 0.07234 |
| 17 | 8 | 17 | 2 | 66.12 | 1753.30 | 1753.43 | 1753.57 | 75.83 | 79.85 | 83.88 | 0.07234 |
| 17 | 1 | 17 | 2 | 66.09 | 1754.48 | 1754.62 | 1754.76 | 110.60 | 114.62 | 118.65 | 0.07234 |
| 17 | 8 | 17 | 2 | 65.82 | 1763.09 | 1763.23 | 1763.37 | 1.41 | 5.38 | 9.35 | 0.07234 |
| 17 | 1 | 17 | 2 | 65.78 | 1764.30 | 1764.44 | 1764.58 | 36.39 | 40.35 | 44.31 | 0.07234 |
| 17 | 8 | 17 | 2 | 65.51 | 1773.02 | 1773.16 | 1773.29 | 284.66 | 288.55 | 292.44 | 0.07234 |
| 17 | 1 | 17 | 2 | 65.47 | 1774.25 | 1774.38 | 1774.52 | 319.46 | 323.35 | 327.24 | 0.07234 |
| 18 | 8 | 18 | 2 | 64.53 | 1805.27 | 1805.42 | 1805.56 | 134.76 | 138.91 | 143.06 | 0.08108 |
| 18 | 1 | 18 | 2 | 64.49 | 1806.56 | 1806.71 | 1806.85 | 171.69 | 175.85 | 180.01 | 0.08108 |
| 18 | 8 | 18 | 2 | 64.22 | 1815.62 | 1815.76 | 1815.91 | 67.74 | 71.80 | 75.86 | 0.08108 |
| 18 | 1 | 18 | 2 | 64.19 | 1816.91 | 1817.05 | 1817.20 | 103.90 | 107.98 | 112.06 | 0.08108 |
| 18 | 8 | 18 | 2 | 63.92 | 1825.96 | 1826.11 | 1826.25 | 0.24 | 4.39 | 8.54 | 0.08108 |
| 18 | 1 | 18 | 2 | 63.89 | 1827.25 | 1827.40 | 1827.54 | 37.14 | 41.29 | 45.44 | 0.08108 |
| 18 | 8 | 19 | 2 | 63.62 | 1836.61 | 1836.76 | 1836.91 | 302.46 | 306.78 | 311.11 | 0.08559 |
| 18 | 1 | 19 | 2 | 63.58 | 1837.92 | 1838.08 | 1838.23 | 339.58 | 343.91 | 348.23 | 0.08559 |
| 18 | 8 | 19 | 2 | 63.32 | 1847.41 | 1847.57 | 1847.72 | 244.50 | 248.72 | 252.94 | 0.08559 |
| 18 | 1 | 19 | 2 | 63.28 | 1848.75 | 1848.91 | 1849.06 | 281.25 | 285.44 | 289.63 | 0.08559 |
| 18 | 8 | 19 | 2 | 63.01 | 1858.43 | 1858.58 | 1858.73 | 184.41 | 188.63 | 192.86 | 0.08559 |
| 18 | 1 | 19 | 2 | 62.98 | 1859.79 | 1859.94 | 1860.10 | 222.02 | 226.26 | 230.50 | 0.08559 |
| 22 | 8 | 26 | 2 | 59.12 | 2011.94 | 2012.15 | 2012.36 | 330.07 | 335.32 | 340.57 | 0.11159 |
| 22 | 1 | 26 | 2 | 59.08 | 2013.72 | 2013.93 | 2014.14 | 14.32 | 19.51 | 24.71 | 0.11159 |
| 22 | 8 | 23 | 2 | 58.82 | 2024.76 | 2024.95 | 2025.13 | 287.91 | 292.52 | 297.14 | 0.09871 |
| 22 | 1 | 23 | 2 | 58.79 | 2026.33 | 2026.51 | 2026.70 | 326.83 | 331.45 | 336.07 | 0.09871 |
| 22 | 8 | 22 | 2 | 58.53 | 2037.29 | 2037.46 | 2037.64 | 242.40 | 246.91 | 251.42 | 0.09442 |
| 22 | 1 | 23 | 2 | 58.50 | 2038.83 | 2039.01 | 2039.20 | 281.63 | 286.37 | 291.12 | 0.09871 |
| 22 | 8 | 22 | 2 | 58.25 | 2049.79 | 2049.96 | 2050.14 | 199.51 | 203.87 | 208.26 | 0.09442 |
| 22 | 1 | 23 | 2 | 58.21 | 2051.33 | 2051.51 | 2051.70 | 237.64 | 242.25 | 246.85 | 0.09871 |
| 22 | 8 | 22 | 2 | 57.96 | 2062.29 | 2062.46 | 2062.64 | 156.63 | 161.14 | 165.66 | 0.09442 |
| 22 | 1 | 23 | 2 | 57.93 | 2063.83 | 2064.01 | 2064.20 | 195.80 | 200.51 | 205.22 | 0.09871 |
| 23 | 8 | 23 | 2 | 57.11 | 2100.50 | 2100.69 | 2100.87 | 10.03 | 14.43 | 18.83 | 0.08303 |
| 23 | 1 | 23 | 2 | 57.08 | 2102.09 | 2102.27 | 2102.46 | 47.59 | 51.97 | 56.34 | 0.08303 |
| 23 | 8 | 23 | 2 | 56.84 | 2113.26 | 2113.45 | 2113.64 | 307.90 | 312.20 | 316.51 | 0.08303 |
| 23 | 1 | 23 | 2 | 56.80 | 2114.86 | 2115.05 | 2115.23 | 344.90 | 349.20 | 353.50 | 0.08303 |
| 23 | 8 | 23 | 2 | 56.55 | 2126.72 | 2126.90 | 2127.09 | 263.01 | 267.36 | 271.71 | 0.08303 |
| 23 | 1 | 23 | 2 | 56.51 | 2128.31 | 2128.50 | 2128.68 | 300.42 | 304.76 | 309.09 | 0.08303 |
| 23 | 8 | 23 | 2 | 56.27 | 2139.49 | 2139.67 | 2139.86 | 200.71 | 205.04 | 209.36 | 0.08303 |
| 23 | 1 | 23 | 2 | 56.24 | 2141.08 | 2141.26 | 2141.45 | 237.67 | 241.97 | 246.28 | 0.08303 |
| 23 | 8 | 22 | 2 | 56.20 | 2142.67 | 2142.84 | 2143.02 | 274.48 | 278.60 | 282.71 | 0.07942 |
| 23 | 1 | 23 | 2 | 56.17 | 2144.20 | 2144.38 | 2144.57 | 309.90 | 314.20 | 318.49 | 0.08303 |
| 23 | 8 | 24 | 2 | 55.93 | 2155.60 | 2155.79 | 2155.99 | 205.72 | 209.88 | 214.05 | 0.08664 |
| 23 | 1 | 24 | 2 | 55.90 | 2157.23 | 2157.42 | 2157.62 | 240.88 | 245.06 | 249.24 | 0.08664 |
| 24 | 8 | 25 | 2 | 54.28 | 2236.25 | 2236.45 | 2236.65 | 60.89 | 64.70 | 68.52 | 0.07862 |
| 24 | 1 | 25 | 2 | 54.25 | 2237.99 | 2238.19 | 2238.39 | 93.90 | 97.71 | 101.54 | 0.07862 |
| 24 | 8 | 26 | 2 | 54.00 | 2250.53 | 2250.74 | 2250.95 | 330.33 | 334.20 | 338.08 | 0.08176 |
| 24 | 1 | 26 | 2 | 53.97 | 2252.31 | 2252.52 | 2252.73 | 3.21 | 7.08 | 10.95 | 0.08176 |
| 24 | 8 | 26 | 2 | 53.72 | 2264.82 | 2265.03 | 2265.24 | 236.65 | 240.60 | 244.53 | 0.08176 |
| 24 | 1 | 26 | 2 | 53.69 | 2266.60 | 2266.81 | 2267.02 | 270.12 | 274.05 | 277.98 | 0.08176 |
| 24 | 8 | 27 | 2 | 53.43 | 2279.66 | 2279.87 | 2280.09 | 155.92 | 160.12 | 164.33 | 0.08491 |
| 24 | 1 | 27 | 2 | 53.40 | 2281.49 | 2281.70 | 2281.92 | 191.20 | 195.39 | 199.58 | 0.08491 |
| 24 | 8 | 27 | 2 | 53.14 | 2294.86 | 2295.08 | 2295.30 | 80.12 | 84.08 | 88.06 | 0.08491 |
| 24 | 1 | 27 | 2 | 53.11 | 2296.74 | 2296.95 | 2297.17 | 114.17 | 118.13 | 122.09 | 0.08491 |
| 24 | 8 | 28 | 2 | 52.96 | 2304.30 | 2304.52 | 2304.75 | 250.54 | 254.53 | 258.52 | 0.08805 |
| 24 | 1 | 28 | 2 | 52.92 | 2306.23 | 2306.45 | 2306.68 | 284.48 | 288.41 | 292.35 | 0.08805 |
| 25 | 8 | 29 | 2 | 52.57 | 2324.85 | 2325.08 | 2325.32 | 242.00 | 246.00 | 250.01 | 0.12185 |
| 25 | 1 | 29 | 2 | 52.54 | 2326.82 | 2327.06 | 2327.29 | 276.06 | 280.13 | 284.21 | 0.12185 |
| 25 | 8 | 29 | 2 | 52.28 | 2340.64 | 2340.87 | 2341.11 | 159.31 | 163.37 | 167.41 | 0.12185 |
| 25 | 1 | 29 | 2 | 52.24 | 2342.61 | 2342.85 | 2343.08 | 193.43 | 197.45 | 201.47 | 0.12185 |
| 25 | 8 | 30 | 2 | 51.97 | 2357.03 | 2357.27 | 2357.52 | 75.89 | 79.87 | 83.84 | 0.12605 |
| 25 | 1 | 30 | 2 | 51.94 | 2359.06 | 2359.30 | 2359.54 | 109.06 | 113.03 | 116.98 | 0.12605 |
| 25 | 8 | 31 | 2 | 51.67 | 2373.78 | 2374.03 | 2374.28 | 344.70 | 348.61 | 352.53 | 0.13025 |
| 25 | 1 | 31 | 2 | 51.63 | 2375.86 | 2376.11 | 2376.36 | 17.22 | 21.12 | 25.02 | 0.13025 |
| 26 | 1 | 31 | 2 | 51.01 | 2410.28 | 2410.53 | 2410.78 | 194.68 | 198.67 | 202.66 | 0.12016 |
| 26 | 8 | 32 | 2 | 50.74 | 2425.26 | 2425.51 | 2425.77 | 74.06 | 78.19 | 82.32 | 0.12403 |
| 26 | 1 | 32 | 2 | 50.71 | 2427.40 | 2427.66 | 2427.92 | 108.45 | 112.59 | 116.75 | 0.12403 |
| 26 | 8 | 33 | 2 | 50.43 | 2443.17 | 2443.44 | 2443.71 | 1.46 | 5.71 | 9.96 | 0.12791 |
| 26 | 1 | 32 | 2 | 50.39 | 2445.40 | 2445.66 | 2445.92 | 36.90 | 40.99 | 45.08 | 0.12403 |
| 26 | 8 | 32 | 2 | 50.12 | 2460.84 | 2461.10 | 2461.36 | 274.75 | 278.59 | 282.43 | 0.12403 |
| 26 | 1 | 33 | 2 | 50.08 | 2463.03 | 2463.29 | 2463.56 | 307.13 | 311.03 | 314.93 | 0.12791 |
| 26 | 8 | 33 | 2 | 49.82 | 2478.47 | 2478.74 | 2479.00 | 174.79 | 178.79 | 182.80 | 0.12791 |
| 30 | 8 | 65 | 2 | 47.47 | 2620.26 | 2620.78 | 2621.30 | 247.24 | 255.42 | 263.62 | 0.15439 |
| 30 | 1 | 64 | 2 | 47.43 | 2622.40 | 2622.92 | 2623.43 | 280.85 | 289.00 | 297.17 | 0.15202 |
| 30 | 8 | 64 | 2 | 47.20 | 2637.40 | 2637.92 | 2638.43 | 158.67 | 166.77 | 174.85 | 0.15202 |
| 30 | 1 | 32 | 2 | 47.17 | 2639.27 | 2639.53 | 2639.79 | 187.95 | 191.98 | 196.01 | 0.07601 |
| 30 | 8 | 65 | 2 | 46.93 | 2654.54 | 2655.06 | 2655.58 | 62.70 | 70.66 | 78.61 | 0.15439 |
| 30 | 1 | 32 | 2 | 46.90 | 2656.41 | 2656.67 | 2656.93 | 91.21 | 95.12 | 99.02 | 0.07601 |
| 30 | 8 | 66 | 2 | 46.65 | 2672.32 | 2672.86 | 2673.39 | 328.91 | 336.71 | 344.49 | 0.15677 |
| 30 | 1 | 33 | 2 | 46.63 | 2674.25 | 2674.52 | 2674.78 | 357.09 | 0.97 | 4.86 | 0.07838 |
| 31 | 8 | 68 | 2 | 45.99 | 2715.37 | 2715.92 | 2716.47 | 274.45 | 283.49 | 292.54 | 0.11765 |
| 31 | 1 | 69 | 2 | 45.96 | 2717.64 | 2718.20 | 2718.75 | 311.79 | 320.93 | 330.05 | 0.11938 |
| 31 | 8 | 67 | 2 | 45.72 | 2733.03 | 2733.57 | 2734.12 | 199.35 | 207.87 | 216.37 | 0.11592 |
| 31 | 1 | 66 | 2 | 45.69 | 2735.24 | 2735.78 | 2736.31 | 234.16 | 242.54 | 250.93 | 0.11419 |
| 31 | 8 | 74 | 2 | 45.64 | 2738.82 | 2739.41 | 2740.01 | 290.41 | 299.75 | 309.07 | 0.12803 |
| 31 | 1 | 73 | 2 | 45.60 | 2741.24 | 2741.83 | 2742.42 | 328.35 | 337.54 | 346.73 | 0.12630 |
| 31 | 8 | 80 | 2 | 45.30 | 2761.30 | 2761.94 | 2762.59 | 287.69 | 298.24 | 308.82 | 0.13841 |
| 31 | 1 | 81 | 2 | 45.26 | 2763.95 | 2764.60 | 2765.26 | 331.23 | 342.02 | 352.80 | 0.14014 |
| 32 | 1 | 64 | 2 | 44.81 | 2794.90 | 2795.41 | 2795.93 | 133.95 | 142.95 | 151.97 | 0.12774 |
| 32 | 8 | 63 | 2 | 44.59 | 2809.50 | 2810.01 | 2810.52 | 26.99 | 35.67 | 44.32 | 0.12575 |
| 32 | 1 | 63 | 2 | 44.56 | 2811.59 | 2812.10 | 2812.60 | 62.50 | 71.08 | 79.63 | 0.12575 |
| 32 | 8 | 63 | 2 | 44.35 | 2826.17 | 2826.68 | 2827.18 | 302.98 | 311.43 | 319.90 | 0.12575 |
| 32 | 1 | 63 | 2 | 44.32 | 2828.25 | 2828.76 | 2829.26 | 337.66 | 346.14 | 354.63 | 0.12575 |
| 32 | 8 | 61 | 2 | 44.12 | 2842.17 | 2842.66 | 2843.15 | 209.11 | 217.21 | 225.29 | 0.12176 |
| 32 | 1 | 61 | 2 | 44.09 | 2844.19 | 2844.68 | 2845.18 | 242.31 | 250.32 | 258.31 | 0.12176 |
| 32 | 8 | 63 | 2 | 43.88 | 2859.17 | 2859.67 | 2860.18 | 120.03 | 127.89 | 135.73 | 0.12575 |
| 33 | 8 | 84 | 2 | 43.38 | 2894.45 | 2895.13 | 2895.81 | 308.85 | 319.36 | 329.88 | 0.21053 |
| 33 | 1 | 85 | 2 | 43.34 | 2897.23 | 2897.92 | 2898.60 | 352.03 | 2.69 | 13.36 | 0.21303 |
| 33 | 8 | 48 | 2 | 43.02 | 2920.19 | 2920.58 | 2920.97 | 2.75 | 9.27 | 15.81 | 0.12030 |
| 33 | 1 | 49 | 2 | 42.98 | 2923.42 | 2923.81 | 2924.21 | 57.13 | 63.79 | 70.43 | 0.12281 |
| 33 | 8 | 66 | 2 | 42.80 | 2936.56 | 2937.09 | 2937.63 | 276.11 | 284.60 | 293.04 | 0.16541 |
| 33 | 1 | 67 | 2 | 42.77 | 2938.76 | 2939.30 | 2939.84 | 310.94 | 319.39 | 327.83 | 0.16792 |
| 34 | 8 | 66 | 2 | 42.08 | 2989.32 | 2989.86 | 2990.39 | 350.36 | 358.94 | 7.54 | 0.11301 |
| 34 | 1 | 67 | 2 | 42.05 | 2991.53 | 2992.07 | 2992.61 | 25.97 | 34.77 | 43.55 | 0.11473 |
| 34 | 8 | 67 | 2 | 41.84 | 3007.26 | 3007.80 | 3008.34 | 285.79 | 294.66 | 303.55 | 0.11473 |
| 34 | 1 | 67 | 2 | 41.81 | 3009.47 | 3010.01 | 3010.55 | 322.11 | 330.98 | 339.85 | 0.11473 |
| 34 | 8 | 65 | 2 | 41.61 | 3024.43 | 3024.95 | 3025.48 | 205.18 | 213.60 | 222.00 | 0.11130 |
| 34 | 1 | 64 | 2 | 41.58 | 3026.57 | 3027.09 | 3027.60 | 239.56 | 247.82 | 256.07 | 0.10959 |
| 34 | 8 | 9 | 2 | 41.57 | 3028.27 | 3028.35 | 3028.42 | 266.78 | 267.93 | 269.09 | 0.01541 |
| 34 | 1 | 60 | 2 | 41.55 | 3029.03 | 3029.52 | 3030.00 | 278.87 | 286.58 | 294.29 | 0.10274 |
| 34 | 8 | 60 | 2 | 41.37 | 3042.89 | 3043.38 | 3043.86 | 138.31 | 145.96 | 153.62 | 0.10274 |
| 34 | 1 | 59 | 2 | 41.34 | 3044.88 | 3045.35 | 3045.83 | 169.69 | 177.23 | 184.78 | 0.10103 |
| 35 | 1 | 61 | 2 | 40.42 | 3115.97 | 3116.47 | 3116.96 | 179.57 | 186.72 | 193.89 | 0.12323 |
| 35 | 8 | 62 | 2 | 40.24 | 3130.28 | 3130.78 | 3131.28 | 29.02 | 36.36 | 43.68 | 0.12525 |
| 35 | 1 | 61 | 2 | 40.22 | 3132.33 | 3132.82 | 3133.31 | 59.02 | 66.23 | 73.43 | 0.12323 |
| 35 | 8 | 62 | 2 | 40.03 | 3146.81 | 3147.31 | 3147.81 | 269.23 | 275.92 | 282.54 | 0.12525 |
| 35 | 1 | 62 | 2 | 40.01 | 3148.87 | 3149.37 | 3149.87 | 296.33 | 302.78 | 309.18 | 0.12525 |
| 35 | 8 | 62 | 2 | 39.82 | 3163.45 | 3163.95 | 3164.45 | 125.01 | 131.97 | 138.94 | 0.12525 |
| 35 | 1 | 62 | 2 | 39.80 | 3165.51 | 3166.01 | 3166.51 | 153.89 | 160.94 | 168.00 | 0.12525 |
| 35 | 8 | 63 | 2 | 39.62 | 3180.16 | 3180.67 | 3181.18 | 357.01 | 3.91 | 10.80 | 0.12727 |
| 36 | 8 | 63 | 2 | 39.20 | 3213.95 | 3214.46 | 3214.97 | 101.93 | 109.02 | 116.09 | 0.12281 |
| 36 | 1 | 64 | 2 | 39.17 | 3216.06 | 3216.57 | 3217.09 | 131.25 | 138.43 | 145.62 | 0.12476 |
| 36 | 8 | 64 | 2 | 38.99 | 3230.97 | 3231.48 | 3232.00 | 339.34 | 346.59 | 353.83 | 0.12476 |
| 36 | 1 | 64 | 2 | 38.96 | 3233.09 | 3233.60 | 3234.12 | 9.09 | 16.31 | 23.53 | 0.12476 |
| 36 | 8 | 64 | 2 | 38.78 | 3248.14 | 3248.65 | 3249.17 | 216.55 | 223.45 | 230.32 | 0.12476 |
| 36 | 1 | 64 | 2 | 38.75 | 3250.27 | 3250.79 | 3251.30 | 244.92 | 251.68 | 258.40 | 0.12476 |
| 36 | 8 | 65 | 2 | 38.57 | 3265.38 | 3265.90 | 3266.42 | 73.05 | 79.63 | 86.22 | 0.12671 |
| 36 | 1 | 65 | 2 | 38.55 | 3267.53 | 3268.05 | 3268.58 | 100.10 | 106.69 | 113.25 | 0.12671 |
| 37 | 8 | 66 | 2 | 38.16 | 3300.25 | 3300.78 | 3301.31 | 203.32 | 211.29 | 219.22 | 0.12453 |
| 37 | 1 | 66 | 2 | 38.13 | 3302.42 | 3302.96 | 3303.49 | 235.64 | 243.39 | 251.07 | 0.12453 |
| 37 | 8 | 66 | 2 | 37.95 | 3317.79 | 3318.32 | 3318.85 | 85.90 | 93.02 | 100.11 | 0.12453 |
| 37 | 1 | 66 | 2 | 37.92 | 3319.98 | 3320.51 | 3321.04 | 115.10 | 122.16 | 129.19 | 0.12453 |
| 37 | 8 | 66 | 2 | 37.74 | 3335.46 | 3335.99 | 3336.53 | 320.52 | 327.67 | 334.83 | 0.12453 |
| 37 | 1 | 66 | 2 | 37.71 | 3337.66 | 3338.20 | 3338.73 | 350.15 | 357.32 | 4.52 | 0.12453 |
| 37 | 8 | 67 | 2 | 37.53 | 3353.24 | 3353.78 | 3354.32 | 205.78 | 213.51 | 221.26 | 0.12642 |
| 37 | 1 | 67 | 2 | 37.51 | 3355.46 | 3356.00 | 3356.54 | 237.67 | 245.50 | 253.39 | 0.12642 |
| 38 | 8 | 68 | 2 | 37.01 | 3397.90 | 3398.44 | 3398.99 | 170.93 | 179.39 | 187.82 | 0.12409 |
| 38 | 1 | 68 | 2 | 36.99 | 3400.15 | 3400.69 | 3401.24 | 205.47 | 213.81 | 222.11 | 0.12409 |
| 38 | 8 | 68 | 2 | 36.81 | 3416.08 | 3416.63 | 3417.18 | 75.04 | 82.56 | 90.05 | 0.12409 |
| 38 | 1 | 68 | 2 | 36.78 | 3418.35 | 3418.89 | 3419.44 | 105.95 | 113.35 | 120.71 | 0.12409 |
| 38 | 8 | 69 | 2 | 36.60 | 3434.39 | 3434.94 | 3435.50 | 318.57 | 325.96 | 333.35 | 0.12591 |
| 38 | 1 | 69 | 2 | 36.57 | 3436.66 | 3437.22 | 3437.77 | 348.76 | 356.15 | 3.53 | 0.12591 |
| 38 | 8 | 69 | 2 | 36.39 | 3452.79 | 3453.34 | 3453.90 | 199.92 | 207.18 | 214.44 | 0.12591 |
| 38 | 1 | 69 | 2 | 36.37 | 3455.09 | 3455.64 | 3456.20 | 229.98 | 237.30 | 244.65 | 0.12591 |
| 39 | 8 | 70 | 2 | 35.98 | 3489.92 | 3490.49 | 3491.05 | 1.41 | 9.89 | 18.37 | 0.10116 |
| 39 | 1 | 70 | 2 | 35.95 | 3492.24 | 3492.80 | 3493.37 | 36.20 | 44.70 | 53.21 | 0.10116 |
| 39 | 8 | 70 | 2 | 35.79 | 3507.01 | 3507.58 | 3508.14 | 259.87 | 268.21 | 276.54 | 0.10116 |
| 39 | 1 | 71 | 2 | 35.76 | 3509.34 | 3509.91 | 3510.49 | 294.26 | 302.69 | 311.11 | 0.10260 |
| 39 | 8 | 71 | 2 | 35.58 | 3525.87 | 3526.45 | 3527.02 | 164.67 | 172.12 | 179.55 | 0.10260 |
| 39 | 1 | 71 | 2 | 35.55 | 3528.22 | 3528.79 | 3529.36 | 195.07 | 202.45 | 209.80 | 0.10260 |
| 39 | 8 | 54 | 2 | 35.53 | 3530.46 | 3530.90 | 3531.33 | 223.85 | 229.41 | 234.96 | 0.07803 |
| 39 | 1 | 71 | 2 | 35.51 | 3532.41 | 3532.99 | 3533.56 | 248.71 | 255.98 | 263.23 | 0.10260 |
| 39 | 8 | 72 | 2 | 35.32 | 3548.98 | 3549.56 | 3550.15 | 94.57 | 102.16 | 109.81 | 0.10405 |
| 39 | 1 | 72 | 2 | 35.30 | 3551.36 | 3551.94 | 3552.52 | 125.89 | 133.67 | 141.47 | 0.10405 |
| 40 | 8 | 72 | 2 | 35.12 | 3568.17 | 3568.75 | 3569.33 | 348.52 | 355.98 | 3.43 | 0.12414 |
| 40 | 1 | 72 | 2 | 35.09 | 3570.56 | 3571.14 | 3571.72 | 19.15 | 26.60 | 34.06 | 0.12414 |
| 40 | 8 | 72 | 2 | 34.91 | 3587.36 | 3587.94 | 3588.52 | 236.65 | 244.48 | 252.36 | 0.12414 |
| 40 | 1 | 72 | 2 | 34.88 | 3589.75 | 3590.33 | 3590.91 | 269.29 | 277.35 | 285.46 | 0.12414 |
| 40 | 8 | 73 | 2 | 34.70 | 3606.71 | 3607.30 | 3607.89 | 158.15 | 167.17 | 176.20 | 0.12586 |
| 40 | 1 | 73 | 2 | 34.68 | 3609.12 | 3609.71 | 3610.30 | 195.03 | 203.98 | 212.90 | 0.12586 |
| 40 | 8 | 73 | 2 | 34.49 | 3626.20 | 3626.79 | 3627.38 | 81.69 | 90.31 | 98.99 | 0.12586 |
| 40 | 1 | 73 | 2 | 34.47 | 3628.63 | 3629.22 | 3629.80 | 117.54 | 126.33 | 135.16 | 0.12586 |
| 41 | 8 | 74 | 2 | 34.12 | 3661.51 | 3662.11 | 3662.70 | 233.95 | 242.32 | 250.75 | 0.12333 |
| 41 | 1 | 74 | 2 | 34.09 | 3663.96 | 3664.56 | 3665.16 | 268.63 | 277.19 | 285.80 | 0.12333 |
| 41 | 8 | 75 | 2 | 33.91 | 3681.28 | 3681.89 | 3682.49 | 154.58 | 162.70 | 170.84 | 0.12500 |
| 41 | 1 | 75 | 2 | 33.89 | 3683.75 | 3684.36 | 3684.96 | 187.75 | 195.90 | 204.05 | 0.12500 |
| 41 | 8 | 75 | 2 | 33.70 | 3701.25 | 3701.85 | 3702.46 | 72.52 | 81.07 | 89.63 | 0.12500 |
| 41 | 1 | 75 | 2 | 33.68 | 3703.73 | 3704.34 | 3704.94 | 107.68 | 116.26 | 124.83 | 0.12500 |
| 41 | 8 | 76 | 2 | 33.50 | 3721.31 | 3721.93 | 3722.54 | 354.31 | 2.76 | 11.24 | 0.12667 |
| 41 | 1 | 76 | 2 | 33.47 | 3723.81 | 3724.43 | 3725.04 | 28.88 | 37.39 | 45.91 | 0.12667 |
| 42 | 8 | 78 | 2 | 32.67 | 3802.12 | 3802.75 | 3803.38 | 23.68 | 31.39 | 39.08 | 0.16596 |
| 42 | 1 | 78 | 2 | 32.65 | 3804.69 | 3805.31 | 3805.94 | 54.90 | 62.44 | 69.93 | 0.16596 |
| 42 | 8 | 78 | 2 | 32.46 | 3822.77 | 3823.40 | 3824.03 | 258.95 | 266.65 | 274.47 | 0.16596 |
| 42 | 1 | 78 | 2 | 32.44 | 3825.35 | 3825.98 | 3826.61 | 291.17 | 299.26 | 307.42 | 0.16596 |
| 42 | 8 | 79 | 2 | 32.26 | 3843.62 | 3844.26 | 3844.89 | 170.08 | 178.19 | 186.30 | 0.16809 |
| 42 | 1 | 79 | 2 | 32.23 | 3846.22 | 3846.85 | 3847.49 | 203.12 | 211.21 | 219.29 | 0.16809 |
| 43 | 8 | 80 | 2 | 31.84 | 3885.77 | 3886.42 | 3887.06 | 308.18 | 314.03 | 319.86 | 0.16632 |
| 43 | 1 | 80 | 2 | 31.81 | 3888.40 | 3889.05 | 3889.69 | 331.84 | 337.60 | 343.35 | 0.16632 |
| 43 | 8 | 80 | 2 | 31.63 | 3906.92 | 3907.56 | 3908.21 | 157.76 | 165.39 | 173.11 | 0.16632 |
| 43 | 1 | 80 | 2 | 31.61 | 3909.56 | 3910.21 | 3910.85 | 189.59 | 197.57 | 205.62 | 0.16632 |
| 43 | 8 | 81 | 2 | 31.42 | 3928.27 | 3928.92 | 3929.57 | 104.71 | 114.35 | 123.97 | 0.16840 |
| 43 | 1 | 80 | 2 | 31.40 | 3930.93 | 3931.58 | 3932.22 | 143.81 | 153.20 | 162.57 | 0.16632 |
| 44 | 8 | 82 | 2 | 30.80 | 3993.16 | 3993.82 | 3994.48 | 348.74 | 0.19 | 11.67 | 0.16532 |
| 44 | 1 | 82 | 2 | 30.77 | 3995.86 | 3996.52 | 3997.18 | 35.72 | 47.25 | 58.75 | 0.16532 |
| 44 | 8 | 83 | 2 | 30.61 | 4013.53 | 4014.20 | 4014.87 | 345.44 | 356.96 | 8.42 | 0.16734 |
| 44 | 1 | 83 | 2 | 30.58 | 4016.26 | 4016.93 | 4017.60 | 32.04 | 43.34 | 54.59 | 0.16734 |
| 44 | 8 | 83 | 2 | 30.40 | 4035.56 | 4036.23 | 4036.90 | 341.52 | 351.75 | 1.95 | 0.16734 |
| 44 | 1 | 83 | 2 | 30.37 | 4038.31 | 4038.98 | 4039.64 | 23.19 | 33.27 | 43.32 | 0.16734 |
| 45 | 8 | 84 | 2 | 29.98 | 4080.03 | 4080.71 | 4081.38 | 246.42 | 255.79 | 265.18 | 0.16535 |
| 45 | 1 | 84 | 2 | 29.95 | 4082.81 | 4083.49 | 4084.17 | 285.04 | 294.50 | 303.96 | 0.16535 |
| 45 | 8 | 85 | 2 | 29.77 | 4102.45 | 4103.13 | 4103.82 | 195.01 | 204.18 | 213.32 | 0.16732 |
| 45 | 1 | 85 | 2 | 29.75 | 4105.25 | 4105.93 | 4106.62 | 232.35 | 241.46 | 250.59 | 0.16732 |
| 45 | 8 | 85 | 2 | 29.56 | 4125.03 | 4125.72 | 4126.41 | 130.37 | 139.00 | 147.61 | 0.16732 |
| 45 | 1 | 85 | 2 | 29.54 | 4127.84 | 4128.53 | 4129.21 | 165.53 | 174.03 | 182.50 | 0.16732 |
| 46 | 8 | 86 | 2 | 29.15 | 4170.64 | 4171.33 | 4172.02 | 259.95 | 267.28 | 274.70 | 0.16538 |
| 46 | 1 | 86 | 2 | 29.12 | 4173.48 | 4174.18 | 4174.87 | 290.53 | 298.17 | 305.85 | 0.16538 |
| 46 | 8 | 87 | 2 | 28.94 | 4193.63 | 4194.33 | 4195.03 | 169.10 | 177.92 | 186.76 | 0.16731 |
| 46 | 1 | 87 | 2 | 28.91 | 4196.49 | 4197.19 | 4197.89 | 205.17 | 213.99 | 222.84 | 0.16731 |
| 46 | 8 | 87 | 2 | 28.73 | 4216.79 | 4217.49 | 4218.19 | 101.57 | 110.84 | 120.16 | 0.16731 |
| 46 | 1 | 87 | 2 | 28.70 | 4219.67 | 4220.37 | 4221.07 | 139.86 | 149.25 | 158.63 | 0.16731 |
| 47 | 8 | 89 | 2 | 28.31 | 4263.52 | 4264.24 | 4264.96 | 321.49 | 329.67 | 337.84 | 0.16604 |
| 47 | 1 | 89 | 2 | 28.28 | 4266.44 | 4267.16 | 4267.88 | 354.70 | 2.86 | 11.00 | 0.16604 |
| 47 | 8 | 89 | 2 | 28.10 | 4287.13 | 4287.85 | 4288.57 | 235.94 | 244.40 | 252.89 | 0.16604 |
| 47 | 1 | 89 | 2 | 28.08 | 4290.07 | 4290.78 | 4291.50 | 270.74 | 279.40 | 288.15 | 0.16604 |
| 47 | 8 | 90 | 2 | 27.89 | 4310.86 | 4311.58 | 4312.31 | 140.26 | 146.93 | 153.54 | 0.16791 |
| 47 | 1 | 90 | 2 | 27.87 | 4313.81 | 4314.53 | 4315.26 | 167.14 | 173.71 | 180.28 | 0.16791 |
| 48 | 1 | 90 | 2 | 27.66 | 4337.84 | 4338.57 | 4339.29 | 39.87 | 47.37 | 54.89 | 0.16484 |
| 48 | 8 | 91 | 2 | 27.47 | 4358.78 | 4359.51 | 4360.24 | 267.29 | 275.13 | 282.94 | 0.16667 |
| 48 | 1 | 91 | 2 | 27.45 | 4361.77 | 4362.50 | 4363.24 | 299.08 | 306.84 | 314.60 | 0.16667 |
| 48 | 8 | 91 | 2 | 27.26 | 4382.96 | 4383.69 | 4384.43 | 156.68 | 163.96 | 171.27 | 0.16667 |
| 48 | 1 | 91 | 2 | 27.24 | 4385.97 | 4386.70 | 4387.44 | 186.59 | 193.92 | 201.28 | 0.16667 |
| 48 | 8 | 92 | 2 | 27.05 | 4407.30 | 4408.04 | 4408.78 | 51.02 | 58.99 | 66.95 | 0.16850 |
| 50 | 8 | 128 | 2 | 23.18 | 4883.14 | 4884.18 | 4885.21 | 81.51 | 91.99 | 102.39 | 0.24854 |
| 50 | 1 | 129 | 2 | 23.15 | 4887.35 | 4888.39 | 4889.43 | 123.64 | 133.86 | 144.05 | 0.25049 |
| 50 | 8 | 129 | 2 | 22.92 | 4916.83 | 4917.87 | 4918.91 | 58.74 | 69.14 | 79.60 | 0.25049 |
| 50 | 1 | 129 | 2 | 22.89 | 4921.04 | 4922.08 | 4923.13 | 101.29 | 111.97 | 122.61 | 0.25049 |
| 51 | 8 | 132 | 2 | 20.61 | 5226.23 | 5227.30 | 5228.36 | 350.61 | 0.83 | 11.01 | 0.24906 |
| 51 | 1 | 132 | 2 | 20.58 | 5230.56 | 5231.62 | 5232.68 | 31.82 | 41.83 | 51.78 | 0.24906 |
| 51 | 8 | 133 | 2 | 20.36 | 5261.17 | 5262.24 | 5263.31 | 304.58 | 313.96 | 323.36 | 0.25094 |
| 51 | 1 | 133 | 2 | 20.33 | 5265.50 | 5266.58 | 5267.65 | 342.61 | 352.06 | 1.54 | 0.25094 |
| 54 | 8 | 131 | 2 | 16.96 | 5753.99 | 5755.04 | 5756.10 | 104.95 | 114.76 | 124.57 | 0.25586 |
| 54 | 1 | 131 | 2 | 16.93 | 5758.28 | 5759.33 | 5760.39 | 144.74 | 154.54 | 164.34 | 0.25586 |
| 54 | 8 | 125 | 2 | 16.75 | 5786.06 | 5787.06 | 5788.07 | 43.81 | 53.39 | 63.01 | 0.24414 |
| 54 | 1 | 125 | 2 | 16.72 | 5790.14 | 5791.15 | 5792.15 | 82.83 | 92.54 | 102.27 | 0.24414 |
| 55 | 8 | 118 | 2 | 14.78 | 6093.59 | 6094.54 | 6095.49 | 78.56 | 87.43 | 96.28 | 0.25053 |
| 55 | 1 | 117 | 2 | 14.75 | 6097.44 | 6098.38 | 6099.33 | 114.43 | 123.21 | 132.02 | 0.24841 |
| 55 | 8 | 118 | 2 | 14.58 | 6124.64 | 6125.59 | 6126.54 | 17.16 | 26.64 | 36.12 | 0.25053 |
| 55 | 1 | 118 | 2 | 14.56 | 6128.50 | 6129.46 | 6130.41 | 55.66 | 65.07 | 74.51 | 0.25053 |
| 56 | 8 | 124 | 2 | 12.82 | 6412.21 | 6413.21 | 6414.21 | 120.41 | 129.84 | 139.28 | 0.25000 |
| 56 | 1 | 124 | 2 | 12.80 | 6416.28 | 6417.28 | 6418.28 | 158.75 | 168.21 | 177.67 | 0.25000 |
| 56 | 8 | 124 | 2 | 12.62 | 6444.86 | 6445.86 | 6446.86 | 58.95 | 67.66 | 76.40 | 0.25000 |
| 56 | 1 | 124 | 2 | 12.60 | 6448.93 | 6449.93 | 6450.93 | 94.48 | 103.21 | 111.91 | 0.25000 |
| 57 | 8 | 130 | 2 | 10.86 | 6747.21 | 6748.25 | 6749.30 | 18.20 | 26.73 | 35.29 | 0.24904 |
| 57 | 1 | 130 | 2 | 10.83 | 6751.47 | 6752.52 | 6753.57 | 53.06 | 61.70 | 70.37 | 0.24904 |
| 57 | 8 | 131 | 2 | 10.66 | 6781.51 | 6782.57 | 6783.62 | 302.05 | 310.45 | 318.80 | 0.25096 |
| 57 | 1 | 131 | 2 | 10.63 | 6785.79 | 6786.85 | 6787.90 | 335.93 | 344.22 | 352.49 | 0.25096 |
| 58 | 1 | 135 | 2 | 9.06 | 7066.43 | 7067.51 | 7068.60 | 99.79 | 107.89 | 115.96 | 0.24862 |
| 58 | 8 | 136 | 2 | 8.89 | 7097.53 | 7098.62 | 7099.72 | 326.02 | 333.79 | 341.56 | 0.25046 |
| 58 | 1 | 136 | 2 | 8.86 | 7101.97 | 7103.07 | 7104.16 | 357.64 | 5.52 | 13.45 | 0.25046 |
| 58 | 8 | 136 | 2 | 8.69 | 7133.48 | 7134.57 | 7135.67 | 246.31 | 255.10 | 263.87 | 0.25046 |
| 59 | 8 | 141 | 2 | 7.31 | 7389.32 | 7390.46 | 7391.60 | 343.73 | 352.02 | 0.36 | 0.25000 |
| 59 | 1 | 141 | 2 | 7.28 | 7393.94 | 7395.07 | 7396.21 | 17.44 | 25.71 | 33.96 | 0.25000 |
| 59 | 8 | 141 | 2 | 7.11 | 7426.32 | 7427.46 | 7428.60 | 249.96 | 258.29 | 266.64 | 0.25000 |
| 59 | 1 | 141 | 2 | 7.09 | 7430.93 | 7432.07 | 7433.21 | 283.94 | 292.38 | 300.82 | 0.25000 |
| 60 | 8 | 146 | 2 | 5.54 | 7728.50 | 7729.67 | 7730.85 | 335.54 | 343.78 | 352.03 | 0.25000 |
| 60 | 1 | 146 | 2 | 5.51 | 7733.25 | 7734.43 | 7735.61 | 8.84 | 17.05 | 25.20 | 0.25000 |
| 60 | 8 | 146 | 2 | 5.34 | 7766.73 | 7767.91 | 7769.08 | 233.81 | 241.71 | 249.66 | 0.25000 |
| 60 | 1 | 146 | 2 | 5.32 | 7771.50 | 7772.68 | 7773.86 | 266.06 | 274.03 | 282.04 | 0.25000 |
| 1 | 7 | 1 | 10 | 142.17 | 147.11 | 147.15 | 147.19 | 98.39 | 97.21 | 96.02 | 0.03571 |
| 1 | 7 | 1 | 10 | 142.17 | 147.27 | 147.31 | 147.35 | 93.65 | 92.47 | 91.28 | 0.03571 |
| 1 | 2 | 1 | 10 | 142.02 | 150.10 | 150.14 | 150.18 | 10.74 | 9.57 | 8.39 | 0.03571 |
| 1 | 2 | 1 | 10 | 142.01 | 150.26 | 150.30 | 150.34 | 6.04 | 4.86 | 3.69 | 0.03571 |
| 1 | 7 | 1 | 10 | 141.69 | 156.38 | 156.42 | 156.46 | 188.92 | 187.76 | 186.60 | 0.03571 |
| 1 | 7 | 1 | 10 | 141.68 | 156.54 | 156.58 | 156.62 | 184.29 | 183.13 | 181.97 | 0.03571 |
| 1 | 2 | 1 | 10 | 141.53 | 159.48 | 159.52 | 159.56 | 100.34 | 99.19 | 98.04 | 0.03571 |
| 1 | 2 | 1 | 10 | 141.52 | 159.64 | 159.68 | 159.72 | 95.74 | 94.59 | 93.44 | 0.03571 |
| 1 | 7 | 1 | 10 | 141.20 | 165.96 | 166.00 | 166.04 | 277.03 | 275.90 | 274.77 | 0.03571 |
| 1 | 7 | 1 | 10 | 141.19 | 166.12 | 166.16 | 166.20 | 272.50 | 271.37 | 270.24 | 0.03571 |
| 1 | 2 | 1 | 10 | 141.03 | 169.17 | 169.21 | 169.25 | 187.37 | 186.25 | 185.12 | 0.03571 |
| 1 | 2 | 1 | 10 | 141.03 | 169.33 | 169.37 | 169.41 | 182.88 | 181.76 | 180.64 | 0.03571 |
| 1 | 7 | 1 | 10 | 140.69 | 175.86 | 175.90 | 175.94 | 2.61 | 1.51 | 0.40 | 0.03571 |
| 1 | 7 | 1 | 10 | 140.68 | 176.02 | 176.06 | 176.10 | 358.20 | 357.09 | 355.99 | 0.03571 |
| 1 | 2 | 1 | 12 | 140.52 | 179.11 | 179.16 | 179.21 | 273.97 | 272.65 | 271.34 | 0.03571 |
| 1 | 2 | 1 | 12 | 140.51 | 179.31 | 179.35 | 179.40 | 268.71 | 267.40 | 266.08 | 0.03571 |
| 1 | 7 | 1 | 12 | 140.17 | 186.03 | 186.08 | 186.13 | 87.65 | 86.36 | 85.07 | 0.03571 |
| 1 | 7 | 1 | 12 | 140.16 | 186.22 | 186.27 | 186.32 | 82.49 | 81.20 | 79.91 | 0.03571 |
| 1 | 2 | 1 | 12 | 140.00 | 189.38 | 189.43 | 189.47 | 358.73 | 357.45 | 356.17 | 0.03571 |
| 1 | 2 | 1 | 12 | 139.99 | 189.57 | 189.62 | 189.67 | 353.61 | 352.33 | 351.05 | 0.03571 |
| 1 | 7 | 1 | 12 | 139.65 | 196.41 | 196.46 | 196.51 | 174.52 | 173.26 | 172.01 | 0.03571 |
| 1 | 7 | 1 | 12 | 139.64 | 196.60 | 196.65 | 196.70 | 169.50 | 168.25 | 166.99 | 0.03571 |
| 1 | 2 | 1 | 12 | 139.47 | 199.97 | 200.01 | 200.06 | 82.77 | 81.53 | 80.29 | 0.03571 |
| 1 | 2 | 1 | 12 | 139.46 | 200.16 | 200.21 | 200.26 | 77.81 | 76.57 | 75.33 | 0.03571 |
| 1 | 7 | 1 | 12 | 139.11 | 207.26 | 207.30 | 207.35 | 257.86 | 256.64 | 255.43 | 0.03571 |
| 1 | 7 | 1 | 12 | 139.10 | 207.45 | 207.50 | 207.55 | 253.01 | 251.79 | 250.58 | 0.03571 |
| 1 | 2 | 1 | 12 | 138.93 | 210.87 | 210.92 | 210.97 | 167.81 | 166.61 | 165.42 | 0.03571 |
| 1 | 2 | 1 | 12 | 138.92 | 211.06 | 211.11 | 211.16 | 163.02 | 161.82 | 160.62 | 0.03571 |
| 2 | 7 | 13 | 2 | 135.10 | 289.65 | 289.75 | 289.86 | 228.54 | 226.51 | 224.47 | 0.12621 |
| 2 | 7 | 17 | 2 | 134.74 | 297.37 | 297.51 | 297.65 | 81.12 | 78.55 | 75.98 | 0.16505 |
| 2 | 2 | 20 | 2 | 134.39 | 304.73 | 304.89 | 305.05 | 305.83 | 302.93 | 300.02 | 0.19417 |
| 2 | 7 | 22 | 2 | 133.72 | 319.19 | 319.36 | 319.54 | 55.39 | 52.45 | 49.51 | 0.21359 |
| 2 | 7 | 15 | 2 | 132.96 | 335.72 | 335.85 | 335.97 | 154.90 | 153.10 | 151.30 | 0.14563 |
| 2 | 2 | 16 | 2 | 132.71 | 341.00 | 341.13 | 341.26 | 77.60 | 75.75 | 73.90 | 0.15534 |
| 3 | 2 | 16 | 2 | 128.34 | 439.74 | 439.87 | 440.00 | 121.46 | 120.56 | 119.65 | 0.08421 |
| 3 | 2 | 16 | 2 | 128.31 | 440.25 | 440.38 | 440.51 | 117.85 | 116.95 | 116.06 | 0.08421 |
| 3 | 7 | 16 | 2 | 127.61 | 456.74 | 456.87 | 457.00 | 13.70 | 12.97 | 12.24 | 0.08421 |
| 3 | 7 | 16 | 2 | 127.59 | 457.25 | 457.38 | 457.51 | 10.79 | 10.07 | 9.35 | 0.08421 |
| 3 | 7 | 19 | 2 | 126.56 | 481.59 | 481.74 | 481.90 | 257.06 | 256.48 | 255.91 | 0.10000 |
| 3 | 7 | 19 | 2 | 126.54 | 482.20 | 482.36 | 482.51 | 254.77 | 254.20 | 253.64 | 0.10000 |
| 3 | 2 | 22 | 2 | 126.08 | 493.07 | 493.25 | 493.43 | 218.50 | 217.98 | 217.46 | 0.11579 |
| 3 | 2 | 22 | 2 | 126.05 | 493.78 | 493.96 | 494.14 | 216.42 | 215.90 | 215.39 | 0.11579 |
| 3 | 7 | 22 | 2 | 125.15 | 515.75 | 515.93 | 516.10 | 168.28 | 168.01 | 167.75 | 0.11579 |
| 3 | 7 | 22 | 2 | 125.12 | 516.46 | 516.64 | 516.81 | 167.23 | 166.98 | 166.73 | 0.11579 |
| 4 | 2 | 18 | 2 | 124.64 | 528.30 | 528.45 | 528.59 | 154.21 | 154.10 | 153.99 | 0.07595 |
| 4 | 2 | 18 | 2 | 124.62 | 528.88 | 529.03 | 529.17 | 153.78 | 153.67 | 153.57 | 0.07595 |
| 4 | 2 | 18 | 2 | 124.59 | 529.46 | 529.61 | 529.75 | 153.37 | 153.27 | 153.17 | 0.07595 |
| 4 | 7 | 19 | 2 | 123.61 | 553.76 | 553.91 | 554.06 | 153.36 | 153.47 | 153.59 | 0.08017 |
| 4 | 7 | 19 | 2 | 123.59 | 554.37 | 554.52 | 554.68 | 153.81 | 153.92 | 154.03 | 0.08017 |
| 4 | 7 | 19 | 2 | 123.56 | 554.98 | 555.14 | 555.29 | 154.26 | 154.38 | 154.49 | 0.08017 |
| 4 | 2 | 21 | 2 | 123.00 | 569.05 | 569.21 | 569.38 | 170.79 | 171.06 | 171.33 | 0.08861 |
| 4 | 2 | 21 | 2 | 122.97 | 569.72 | 569.89 | 570.06 | 171.88 | 172.15 | 172.43 | 0.08861 |
| 4 | 2 | 21 | 2 | 122.94 | 570.40 | 570.57 | 570.74 | 173.00 | 173.28 | 173.57 | 0.08861 |
| 4 | 7 | 21 | 2 | 121.85 | 598.20 | 598.37 | 598.54 | 242.35 | 242.91 | 243.47 | 0.08861 |
| 4 | 7 | 21 | 2 | 121.82 | 598.88 | 599.05 | 599.22 | 244.61 | 245.18 | 245.75 | 0.08861 |
| 4 | 7 | 21 | 2 | 121.79 | 599.56 | 599.73 | 599.90 | 246.90 | 247.47 | 248.05 | 0.08861 |
| 8 | 7 | 15 | 2 | 113.21 | 832.58 | 832.70 | 832.82 | 275.31 | 277.14 | 278.98 | 0.09091 |
| 8 | 7 | 15 | 2 | 113.19 | 833.06 | 833.18 | 833.30 | 282.65 | 284.48 | 286.32 | 0.09091 |
| 8 | 2 | 16 | 2 | 112.90 | 841.39 | 841.52 | 841.65 | 50.74 | 52.75 | 54.75 | 0.09697 |
| 8 | 2 | 16 | 2 | 112.89 | 841.91 | 842.04 | 842.17 | 58.78 | 60.79 | 62.80 | 0.09697 |
| 8 | 7 | 16 | 2 | 112.33 | 858.10 | 858.23 | 858.36 | 317.49 | 319.60 | 321.71 | 0.09697 |
| 8 | 7 | 16 | 2 | 112.32 | 858.62 | 858.75 | 858.88 | 325.93 | 328.05 | 330.16 | 0.09697 |
| 8 | 2 | 22 | 2 | 112.06 | 865.96 | 866.14 | 866.32 | 87.42 | 90.39 | 93.35 | 0.13333 |
| 8 | 7 | 15 | 2 | 111.57 | 880.79 | 880.91 | 881.03 | 339.94 | 342.04 | 344.15 | 0.09091 |
| 8 | 7 | 15 | 2 | 111.55 | 881.27 | 881.39 | 881.51 | 348.36 | 350.46 | 352.57 | 0.09091 |
| 8 | 2 | 19 | 2 | 111.32 | 888.20 | 888.36 | 888.51 | 110.15 | 112.86 | 115.58 | 0.11515 |
| 9 | 2 | 15 | 2 | 107.38 | 963.17 | 963.30 | 963.42 | 221.33 | 224.35 | 227.38 | 0.12821 |
| 9 | 7 | 16 | 2 | 106.60 | 974.17 | 974.29 | 974.42 | 149.66 | 153.18 | 156.70 | 0.13675 |
| 9 | 7 | 15 | 2 | 106.08 | 981.26 | 981.38 | 981.50 | 347.26 | 350.71 | 354.16 | 0.12821 |
| 9 | 7 | 11 | 2 | 105.35 | 991.20 | 991.29 | 991.37 | 281.58 | 284.32 | 287.05 | 0.09402 |
| 9 | 2 | 11 | 2 | 105.04 | 995.49 | 995.58 | 995.67 | 55.85 | 58.66 | 61.47 | 0.09402 |
| 9 | 7 | 11 | 2 | 104.44 | 1003.62 | 1003.71 | 1003.80 | 317.93 | 320.84 | 323.75 | 0.09402 |
| 9 | 7 | 12 | 2 | 103.58 | 1015.31 | 1015.40 | 1015.50 | 349.02 | 352.32 | 355.62 | 0.10256 |
| 9 | 2 | 14 | 2 | 103.27 | 1019.48 | 1019.59 | 1019.70 | 131.91 | 135.80 | 139.69 | 0.11966 |
| 9 | 7 | 12 | 2 | 102.69 | 1027.45 | 1027.55 | 1027.65 | 49.56 | 52.96 | 56.37 | 0.10256 |
| 10 | 7 | 21 | 2 | 101.08 | 1049.70 | 1049.86 | 1050.03 | 132.25 | 138.52 | 144.80 | 0.09677 |
| 10 | 2 | 18 | 2 | 100.84 | 1053.11 | 1053.26 | 1053.40 | 259.47 | 264.90 | 270.33 | 0.08295 |
| 10 | 7 | 21 | 2 | 100.33 | 1060.22 | 1060.39 | 1060.56 | 167.46 | 173.92 | 180.37 | 0.09677 |
| 10 | 2 | 12 | 2 | 100.10 | 1063.55 | 1063.65 | 1063.74 | 294.91 | 298.62 | 302.34 | 0.05530 |
| 10 | 7 | 20 | 2 | 99.62 | 1070.28 | 1070.44 | 1070.61 | 196.27 | 202.60 | 208.93 | 0.09217 |
| 10 | 2 | 12 | 2 | 99.40 | 1073.49 | 1073.59 | 1073.69 | 322.72 | 326.54 | 330.38 | 0.05530 |
| 10 | 7 | 19 | 2 | 98.94 | 1080.01 | 1080.16 | 1080.31 | 223.24 | 229.43 | 235.63 | 0.08756 |
| 10 | 7 | 18 | 2 | 98.21 | 1090.73 | 1090.87 | 1091.02 | 304.04 | 310.12 | 316.21 | 0.08295 |
| 10 | 2 | 13 | 2 | 98.00 | 1093.72 | 1093.83 | 1093.93 | 70.35 | 74.79 | 79.23 | 0.05991 |
| 10 | 7 | 18 | 2 | 97.59 | 1099.80 | 1099.94 | 1100.09 | 329.96 | 336.23 | 342.50 | 0.08295 |
| 10 | 2 | 14 | 2 | 97.39 | 1102.70 | 1102.81 | 1102.93 | 95.94 | 100.86 | 105.79 | 0.06452 |
| 10 | 7 | 17 | 2 | 96.99 | 1108.63 | 1108.76 | 1108.90 | 356.72 | 2.82 | 8.91 | 0.07834 |
| 10 | 2 | 14 | 2 | 96.81 | 1111.43 | 1111.55 | 1111.66 | 121.88 | 126.94 | 132.00 | 0.06452 |
| 11 | 7 | 15 | 2 | 94.80 | 1141.96 | 1142.08 | 1142.20 | 109.25 | 115.15 | 121.04 | 0.07143 |
| 11 | 2 | 14 | 2 | 94.63 | 1144.53 | 1144.64 | 1144.76 | 235.01 | 240.55 | 246.09 | 0.06667 |
| 11 | 7 | 15 | 2 | 94.29 | 1149.88 | 1150.00 | 1150.12 | 138.98 | 145.00 | 151.02 | 0.07143 |
| 11 | 2 | 12 | 2 | 94.13 | 1152.41 | 1152.51 | 1152.60 | 265.29 | 270.12 | 274.97 | 0.05714 |
| 11 | 7 | 15 | 2 | 93.79 | 1157.64 | 1157.76 | 1157.88 | 168.92 | 175.06 | 181.19 | 0.07143 |
| 11 | 7 | 14 | 2 | 93.61 | 1160.50 | 1160.62 | 1160.73 | 314.55 | 320.31 | 326.07 | 0.06667 |
| 11 | 2 | 4 | 2 | 93.47 | 1162.87 | 1162.90 | 1162.93 | 75.49 | 77.14 | 78.80 | 0.01905 |
| 11 | 7 | 14 | 2 | 93.15 | 1167.79 | 1167.90 | 1168.01 | 329.39 | 335.25 | 341.11 | 0.06667 |
| 11 | 2 | 10 | 2 | 93.00 | 1170.21 | 1170.29 | 1170.37 | 95.33 | 99.53 | 103.74 | 0.04762 |
| 11 | 7 | 14 | 2 | 92.69 | 1175.15 | 1175.26 | 1175.38 | 354.85 | 0.81 | 6.77 | 0.06667 |
| 11 | 2 | 12 | 2 | 92.54 | 1177.53 | 1177.63 | 1177.72 | 120.81 | 125.95 | 131.09 | 0.05714 |
| 11 | 7 | 14 | 2 | 92.24 | 1182.45 | 1182.56 | 1182.67 | 23.63 | 29.70 | 35.77 | 0.06667 |
| 11 | 2 | 10 | 2 | 92.10 | 1184.76 | 1184.84 | 1184.92 | 148.38 | 152.74 | 157.09 | 0.04762 |
| 11 | 7 | 14 | 2 | 91.80 | 1189.54 | 1189.65 | 1189.76 | 48.06 | 54.23 | 60.41 | 0.06667 |
| 11 | 2 | 10 | 2 | 91.67 | 1191.81 | 1191.89 | 1191.97 | 172.85 | 177.30 | 181.73 | 0.04762 |
| 11 | 7 | 13 | 2 | 91.38 | 1196.48 | 1196.58 | 1196.69 | 71.33 | 77.17 | 83.01 | 0.06190 |
| 11 | 2 | 10 | 2 | 91.25 | 1198.71 | 1198.79 | 1198.87 | 195.99 | 200.50 | 205.02 | 0.04762 |
| 12 | 7 | 12 | 2 | 89.27 | 1231.84 | 1231.94 | 1232.04 | 303.76 | 309.42 | 315.09 | 0.05797 |
| 12 | 2 | 7 | 2 | 89.16 | 1233.82 | 1233.87 | 1233.93 | 59.44 | 62.75 | 66.06 | 0.03382 |
| 12 | 7 | 11 | 2 | 88.92 | 1237.83 | 1237.92 | 1238.01 | 294.66 | 299.85 | 305.05 | 0.05314 |
| 12 | 2 | 10 | 2 | 88.81 | 1239.76 | 1239.84 | 1239.92 | 47.49 | 52.21 | 56.92 | 0.04831 |
| 12 | 7 | 11 | 2 | 88.58 | 1243.80 | 1243.89 | 1243.98 | 283.74 | 288.93 | 294.11 | 0.05314 |
| 12 | 2 | 11 | 2 | 88.46 | 1245.93 | 1246.02 | 1246.11 | 48.05 | 53.23 | 58.41 | 0.05314 |
| 12 | 7 | 11 | 2 | 88.21 | 1250.13 | 1250.22 | 1250.31 | 292.81 | 297.97 | 303.13 | 0.05314 |
| 12 | 2 | 9 | 2 | 88.10 | 1252.10 | 1252.17 | 1252.24 | 47.20 | 51.42 | 55.64 | 0.04348 |
| 12 | 7 | 11 | 2 | 87.86 | 1256.21 | 1256.30 | 1256.39 | 285.58 | 290.71 | 295.84 | 0.05314 |
| 12 | 2 | 9 | 2 | 87.75 | 1258.16 | 1258.24 | 1258.31 | 38.42 | 42.61 | 46.81 | 0.04348 |
| 12 | 7 | 11 | 2 | 87.52 | 1262.19 | 1262.28 | 1262.36 | 270.49 | 275.60 | 280.71 | 0.05314 |
| 12 | 2 | 10 | 2 | 87.42 | 1264.12 | 1264.20 | 1264.28 | 21.85 | 26.48 | 31.12 | 0.04831 |
| 12 | 7 | 11 | 2 | 87.19 | 1268.19 | 1268.28 | 1268.36 | 255.19 | 260.28 | 265.37 | 0.05314 |
| 12 | 2 | 8 | 2 | 87.08 | 1270.11 | 1270.17 | 1270.24 | 5.18 | 8.87 | 12.57 | 0.03865 |
| 12 | 7 | 11 | 2 | 86.86 | 1274.08 | 1274.17 | 1274.26 | 232.62 | 237.69 | 242.77 | 0.05314 |
| 12 | 2 | 8 | 2 | 86.75 | 1275.99 | 1276.06 | 1276.12 | 341.81 | 345.49 | 349.17 | 0.03865 |
| 12 | 7 | 11 | 2 | 86.53 | 1279.93 | 1280.02 | 1280.11 | 206.28 | 211.34 | 216.40 | 0.05314 |
| 12 | 2 | 9 | 2 | 86.43 | 1281.83 | 1281.90 | 1281.98 | 314.71 | 318.85 | 322.98 | 0.04348 |
| 12 | 7 | 11 | 2 | 86.21 | 1285.75 | 1285.84 | 1285.93 | 177.65 | 182.70 | 187.74 | 0.05314 |
| 12 | 7 | 4 | 2 | 86.11 | 1287.60 | 1287.63 | 1287.66 | 282.58 | 284.41 | 286.24 | 0.01932 |
| 12 | 2 | 11 | 2 | 85.99 | 1289.77 | 1289.85 | 1289.94 | 45.69 | 50.73 | 55.76 | 0.05314 |
| 13 | 2 | 10 | 2 | 84.43 | 1318.58 | 1318.66 | 1318.74 | 233.14 | 237.68 | 242.22 | 0.04566 |
| 13 | 7 | 11 | 2 | 84.22 | 1322.45 | 1322.54 | 1322.63 | 91.11 | 96.11 | 101.11 | 0.05023 |
| 13 | 7 | 10 | 2 | 84.06 | 1325.63 | 1325.71 | 1325.79 | 269.87 | 274.40 | 278.93 | 0.04566 |
| 13 | 2 | 11 | 2 | 83.92 | 1328.18 | 1328.26 | 1328.35 | 53.03 | 58.01 | 62.99 | 0.05023 |
| 13 | 7 | 12 | 2 | 83.70 | 1332.43 | 1332.52 | 1332.62 | 291.26 | 296.66 | 302.08 | 0.05479 |
| 13 | 2 | 10 | 2 | 83.59 | 1334.57 | 1334.65 | 1334.73 | 51.16 | 55.66 | 60.16 | 0.04566 |
| 13 | 7 | 10 | 2 | 83.39 | 1338.45 | 1338.53 | 1338.61 | 267.33 | 271.82 | 276.30 | 0.04566 |
| 13 | 2 | 10 | 2 | 83.29 | 1340.30 | 1340.38 | 1340.46 | 9.91 | 14.38 | 18.85 | 0.04566 |
| 13 | 7 | 10 | 2 | 83.09 | 1344.17 | 1344.25 | 1344.33 | 224.08 | 228.53 | 232.97 | 0.04566 |
| 13 | 2 | 10 | 2 | 82.99 | 1346.01 | 1346.09 | 1346.17 | 325.82 | 330.25 | 334.69 | 0.04566 |
| 13 | 7 | 11 | 2 | 82.79 | 1349.88 | 1349.96 | 1350.05 | 177.50 | 182.35 | 187.20 | 0.05023 |
| 13 | 2 | 10 | 2 | 82.70 | 1351.72 | 1351.80 | 1351.88 | 278.20 | 282.58 | 286.97 | 0.04566 |
| 13 | 7 | 10 | 2 | 82.50 | 1355.59 | 1355.67 | 1355.75 | 127.90 | 132.25 | 136.60 | 0.04566 |
| 13 | 2 | 10 | 2 | 82.41 | 1357.46 | 1357.54 | 1357.62 | 228.15 | 232.48 | 236.81 | 0.04566 |
| 13 | 7 | 11 | 2 | 82.21 | 1361.38 | 1361.46 | 1361.55 | 77.51 | 82.23 | 86.94 | 0.05023 |
| 13 | 2 | 10 | 2 | 82.11 | 1363.24 | 1363.32 | 1363.40 | 176.25 | 180.51 | 184.78 | 0.04566 |
| 13 | 7 | 11 | 2 | 81.91 | 1367.15 | 1367.24 | 1367.33 | 21.84 | 26.47 | 31.11 | 0.05023 |
| 13 | 2 | 10 | 2 | 81.82 | 1369.01 | 1369.09 | 1369.17 | 118.97 | 123.16 | 127.35 | 0.04566 |
| 13 | 7 | 11 | 2 | 81.62 | 1372.96 | 1373.05 | 1373.14 | 323.16 | 327.73 | 332.28 | 0.05023 |
| 13 | 2 | 10 | 2 | 81.53 | 1374.85 | 1374.93 | 1375.01 | 59.89 | 64.02 | 68.13 | 0.04566 |
| 13 | 7 | 11 | 2 | 81.33 | 1378.79 | 1378.87 | 1378.96 | 259.87 | 264.35 | 268.84 | 0.05023 |
| 14 | 7 | 14 | 2 | 79.85 | 1409.06 | 1409.17 | 1409.28 | 301.51 | 306.91 | 312.31 | 0.06512 |
| 14 | 2 | 5 | 2 | 79.73 | 1411.44 | 1411.48 | 1411.52 | 55.33 | 57.26 | 59.18 | 0.02326 |
| 14 | 7 | 14 | 2 | 79.50 | 1416.20 | 1416.32 | 1416.43 | 282.92 | 288.31 | 293.68 | 0.06512 |
| 14 | 7 | 11 | 2 | 79.24 | 1421.63 | 1421.72 | 1421.81 | 181.49 | 185.72 | 189.94 | 0.05116 |
| 14 | 2 | 10 | 2 | 79.15 | 1423.61 | 1423.69 | 1423.77 | 275.51 | 279.34 | 283.17 | 0.04651 |
| 14 | 7 | 11 | 2 | 78.95 | 1427.75 | 1427.84 | 1427.93 | 112.49 | 116.70 | 120.92 | 0.05116 |
| 14 | 2 | 11 | 2 | 78.86 | 1429.77 | 1429.85 | 1429.94 | 208.18 | 212.40 | 216.61 | 0.05116 |
| 14 | 7 | 11 | 2 | 78.66 | 1433.98 | 1434.06 | 1434.15 | 47.66 | 51.86 | 56.06 | 0.05116 |
| 14 | 2 | 11 | 2 | 78.57 | 1436.00 | 1436.09 | 1436.18 | 143.38 | 147.57 | 151.76 | 0.05116 |
| 14 | 7 | 12 | 2 | 78.37 | 1440.24 | 1440.34 | 1440.44 | 343.50 | 348.06 | 352.61 | 0.05581 |
| 14 | 2 | 11 | 2 | 78.27 | 1442.26 | 1442.35 | 1442.44 | 78.31 | 82.48 | 86.65 | 0.05116 |
| 14 | 7 | 12 | 2 | 78.07 | 1446.54 | 1446.64 | 1446.73 | 278.96 | 283.48 | 288.00 | 0.05581 |
| 14 | 2 | 10 | 2 | 77.98 | 1448.59 | 1448.67 | 1448.75 | 14.55 | 18.31 | 22.07 | 0.04651 |
| 14 | 7 | 12 | 2 | 77.78 | 1452.85 | 1452.95 | 1453.05 | 212.70 | 217.17 | 221.65 | 0.05581 |
| 14 | 2 | 12 | 2 | 77.69 | 1454.93 | 1455.02 | 1455.12 | 308.48 | 312.94 | 317.41 | 0.05581 |
| 14 | 7 | 12 | 2 | 77.49 | 1459.26 | 1459.36 | 1459.46 | 147.88 | 152.30 | 156.73 | 0.05581 |
| 14 | 2 | 12 | 2 | 77.39 | 1461.37 | 1461.47 | 1461.56 | 243.98 | 248.39 | 252.80 | 0.05581 |
| 14 | 7 | 12 | 2 | 77.20 | 1465.76 | 1465.85 | 1465.95 | 83.11 | 87.49 | 91.85 | 0.05581 |
| 14 | 2 | 12 | 2 | 77.10 | 1467.86 | 1467.96 | 1468.05 | 177.99 | 182.34 | 186.69 | 0.05581 |
| 15 | 2 | 17 | 2 | 74.26 | 1532.94 | 1533.08 | 1533.22 | 58.74 | 64.30 | 69.86 | 0.07798 |
| 15 | 7 | 17 | 2 | 74.01 | 1539.02 | 1539.16 | 1539.30 | 304.77 | 310.31 | 315.86 | 0.07798 |
| 15 | 7 | 14 | 2 | 73.74 | 1545.41 | 1545.52 | 1545.64 | 202.75 | 207.29 | 211.83 | 0.06422 |
| 15 | 2 | 14 | 2 | 73.65 | 1547.80 | 1547.92 | 1548.03 | 298.76 | 303.27 | 307.78 | 0.06422 |
| 15 | 7 | 14 | 2 | 73.44 | 1552.84 | 1552.96 | 1553.07 | 138.22 | 142.65 | 147.08 | 0.06422 |
| 15 | 2 | 13 | 2 | 73.34 | 1555.23 | 1555.34 | 1555.44 | 231.63 | 235.73 | 239.81 | 0.05963 |
| 15 | 7 | 14 | 2 | 73.13 | 1560.25 | 1560.36 | 1560.47 | 64.78 | 69.07 | 73.36 | 0.06422 |
| 15 | 2 | 14 | 2 | 73.04 | 1562.67 | 1562.79 | 1562.90 | 156.57 | 160.81 | 165.05 | 0.06422 |
| 15 | 7 | 14 | 2 | 72.83 | 1567.78 | 1567.89 | 1568.00 | 346.22 | 350.37 | 354.52 | 0.06422 |
| 15 | 2 | 14 | 2 | 72.73 | 1570.23 | 1570.34 | 1570.46 | 75.99 | 80.12 | 84.25 | 0.06422 |
| 15 | 7 | 14 | 2 | 72.52 | 1575.39 | 1575.50 | 1575.62 | 264.53 | 268.63 | 272.73 | 0.06422 |
| 15 | 2 | 14 | 2 | 72.42 | 1577.90 | 1578.01 | 1578.13 | 354.83 | 358.85 | 2.88 | 0.06422 |
| 15 | 7 | 15 | 2 | 72.21 | 1583.12 | 1583.25 | 1583.37 | 180.39 | 184.68 | 188.97 | 0.06881 |
| 15 | 2 | 15 | 2 | 72.11 | 1585.66 | 1585.78 | 1585.90 | 269.96 | 274.22 | 278.49 | 0.06881 |
| 15 | 7 | 15 | 2 | 71.90 | 1590.95 | 1591.08 | 1591.20 | 95.23 | 99.45 | 103.66 | 0.06881 |
| 16 | 7 | 16 | 2 | 70.55 | 1625.87 | 1626.00 | 1626.13 | 244.14 | 248.68 | 253.21 | 0.06897 |
| 16 | 2 | 16 | 2 | 70.45 | 1628.58 | 1628.70 | 1628.83 | 338.94 | 343.46 | 347.99 | 0.06897 |
| 16 | 7 | 16 | 2 | 70.24 | 1634.32 | 1634.45 | 1634.58 | 180.00 | 184.52 | 189.03 | 0.06897 |
| 16 | 2 | 16 | 2 | 70.14 | 1637.04 | 1637.17 | 1637.30 | 275.22 | 279.70 | 284.18 | 0.06897 |
| 16 | 7 | 17 | 2 | 69.92 | 1642.87 | 1643.00 | 1643.14 | 115.54 | 120.26 | 124.98 | 0.07328 |
| 16 | 2 | 16 | 2 | 69.82 | 1645.62 | 1645.75 | 1645.88 | 210.60 | 215.05 | 219.51 | 0.06897 |
| 16 | 7 | 17 | 2 | 69.61 | 1651.44 | 1651.57 | 1651.71 | 52.40 | 57.19 | 61.98 | 0.07328 |
| 16 | 2 | 17 | 2 | 69.50 | 1654.27 | 1654.41 | 1654.54 | 150.66 | 155.41 | 160.16 | 0.07328 |
| 16 | 7 | 17 | 2 | 69.29 | 1660.18 | 1660.32 | 1660.45 | 350.82 | 355.30 | 359.75 | 0.07328 |
| 16 | 2 | 16 | 2 | 69.19 | 1662.98 | 1663.11 | 1663.24 | 81.02 | 85.17 | 89.32 | 0.06897 |
| 16 | 7 | 17 | 2 | 68.98 | 1668.87 | 1669.01 | 1669.15 | 271.79 | 276.23 | 280.67 | 0.07328 |
| 16 | 2 | 17 | 2 | 68.88 | 1671.74 | 1671.88 | 1672.02 | 4.80 | 9.24 | 13.69 | 0.07328 |
| 16 | 7 | 17 | 2 | 68.66 | 1677.75 | 1677.89 | 1678.02 | 199.88 | 204.34 | 208.81 | 0.07328 |
| 16 | 2 | 17 | 2 | 68.56 | 1680.71 | 1680.85 | 1680.98 | 296.30 | 300.78 | 305.25 | 0.07328 |
| 17 | 7 | 18 | 2 | 67.39 | 1714.79 | 1714.94 | 1715.08 | 331.29 | 335.99 | 340.69 | 0.06977 |
| 17 | 2 | 17 | 2 | 67.28 | 1717.81 | 1717.95 | 1718.09 | 68.83 | 73.24 | 77.63 | 0.06589 |
| 17 | 7 | 18 | 2 | 67.07 | 1724.17 | 1724.31 | 1724.46 | 270.05 | 274.55 | 279.04 | 0.06977 |
| 17 | 2 | 18 | 2 | 66.97 | 1727.17 | 1727.31 | 1727.46 | 1.84 | 6.22 | 10.59 | 0.06977 |
| 17 | 7 | 18 | 2 | 66.76 | 1733.54 | 1733.68 | 1733.83 | 193.58 | 197.94 | 202.30 | 0.06977 |
| 17 | 2 | 18 | 2 | 66.66 | 1736.64 | 1736.79 | 1736.93 | 287.55 | 291.99 | 296.44 | 0.06977 |
| 17 | 7 | 19 | 2 | 66.45 | 1743.12 | 1743.27 | 1743.43 | 128.11 | 132.86 | 137.60 | 0.07364 |
| 17 | 2 | 18 | 2 | 66.35 | 1746.18 | 1746.33 | 1746.47 | 222.98 | 227.44 | 231.90 | 0.06977 |
| 17 | 7 | 19 | 2 | 66.14 | 1752.64 | 1752.80 | 1752.95 | 56.65 | 61.16 | 65.66 | 0.07364 |
| 17 | 2 | 19 | 2 | 66.04 | 1755.84 | 1756.00 | 1756.15 | 150.56 | 155.04 | 159.53 | 0.07364 |
| 17 | 7 | 19 | 2 | 65.84 | 1762.42 | 1762.58 | 1762.73 | 342.02 | 346.46 | 350.90 | 0.07364 |
| 17 | 2 | 19 | 2 | 65.73 | 1765.65 | 1765.80 | 1765.96 | 75.19 | 79.59 | 84.00 | 0.07364 |
| 17 | 7 | 19 | 2 | 65.53 | 1772.33 | 1772.49 | 1772.64 | 265.23 | 269.58 | 273.92 | 0.07364 |
| 17 | 2 | 19 | 2 | 65.43 | 1775.58 | 1775.73 | 1775.88 | 357.20 | 1.55 | 5.90 | 0.07364 |
| 18 | 7 | 20 | 2 | 64.55 | 1804.55 | 1804.71 | 1804.88 | 114.28 | 118.88 | 123.47 | 0.08032 |
| 18 | 2 | 20 | 2 | 64.45 | 1807.87 | 1808.03 | 1808.19 | 209.17 | 213.80 | 218.42 | 0.08032 |
| 18 | 7 | 20 | 2 | 64.24 | 1814.90 | 1815.06 | 1815.22 | 47.69 | 52.20 | 56.70 | 0.08032 |
| 18 | 2 | 20 | 2 | 64.15 | 1818.21 | 1818.37 | 1818.53 | 140.42 | 144.96 | 149.50 | 0.08032 |
| 18 | 7 | 21 | 2 | 63.94 | 1825.24 | 1825.41 | 1825.58 | 339.46 | 344.31 | 349.15 | 0.08434 |
| 18 | 2 | 20 | 2 | 63.84 | 1828.74 | 1828.90 | 1829.06 | 79.53 | 84.13 | 88.72 | 0.08032 |
| 18 | 7 | 21 | 2 | 63.64 | 1835.89 | 1836.06 | 1836.23 | 282.20 | 286.98 | 291.76 | 0.08434 |
| 18 | 2 | 21 | 2 | 63.54 | 1839.39 | 1839.56 | 1839.73 | 20.98 | 25.75 | 30.52 | 0.08434 |
| 18 | 7 | 21 | 2 | 63.34 | 1846.68 | 1846.85 | 1847.02 | 224.27 | 228.94 | 233.61 | 0.08434 |
| 18 | 2 | 21 | 2 | 63.24 | 1850.25 | 1850.42 | 1850.59 | 322.08 | 326.67 | 331.26 | 0.08434 |
| 18 | 7 | 22 | 2 | 63.03 | 1857.67 | 1857.85 | 1858.02 | 163.53 | 168.41 | 173.30 | 0.08835 |
| 18 | 2 | 22 | 2 | 62.93 | 1861.36 | 1861.54 | 1861.72 | 265.53 | 270.45 | 275.36 | 0.08835 |
| 22 | 7 | 29 | 2 | 59.15 | 2010.97 | 2011.20 | 2011.43 | 305.51 | 311.43 | 317.32 | 0.11694 |
| 22 | 2 | 22 | 2 | 59.04 | 2015.49 | 2015.67 | 2015.84 | 58.30 | 62.69 | 67.08 | 0.08871 |
| 22 | 7 | 25 | 2 | 58.84 | 2023.92 | 2024.12 | 2024.32 | 266.89 | 271.89 | 276.89 | 0.10081 |
| 22 | 2 | 24 | 2 | 58.75 | 2027.94 | 2028.13 | 2028.33 | 7.03 | 11.85 | 16.69 | 0.09677 |
| 22 | 7 | 25 | 2 | 58.55 | 2036.42 | 2036.63 | 2036.83 | 220.49 | 225.61 | 230.72 | 0.10081 |
| 22 | 2 | 24 | 2 | 58.46 | 2040.44 | 2040.63 | 2040.83 | 322.97 | 327.95 | 332.93 | 0.09677 |
| 22 | 7 | 25 | 2 | 58.27 | 2048.92 | 2049.12 | 2049.33 | 178.26 | 183.22 | 188.20 | 0.10081 |
| 22 | 2 | 24 | 2 | 58.17 | 2052.94 | 2053.13 | 2053.33 | 277.95 | 282.83 | 287.71 | 0.09677 |
| 22 | 7 | 25 | 2 | 57.98 | 2061.42 | 2061.62 | 2061.83 | 134.66 | 139.80 | 144.93 | 0.10081 |
| 22 | 2 | 25 | 2 | 57.89 | 2065.45 | 2065.66 | 2065.86 | 237.12 | 242.22 | 247.32 | 0.10081 |
| 23 | 7 | 26 | 2 | 57.13 | 2099.62 | 2099.83 | 2100.04 | 349.16 | 354.14 | 359.12 | 0.10156 |
| 23 | 2 | 25 | 2 | 57.04 | 2103.71 | 2103.91 | 2104.11 | 85.76 | 90.49 | 95.23 | 0.09766 |
| 23 | 7 | 26 | 2 | 56.85 | 2112.39 | 2112.60 | 2112.81 | 287.51 | 292.37 | 297.24 | 0.10156 |
| 23 | 2 | 25 | 2 | 56.75 | 2117.16 | 2117.36 | 2117.56 | 38.18 | 42.87 | 47.55 | 0.09766 |
| 23 | 7 | 26 | 2 | 56.56 | 2125.84 | 2126.05 | 2126.26 | 242.30 | 247.26 | 252.20 | 0.10156 |
| 23 | 2 | 25 | 2 | 56.48 | 2129.93 | 2130.14 | 2130.34 | 338.22 | 342.91 | 347.60 | 0.09766 |
| 23 | 7 | 26 | 2 | 56.29 | 2138.61 | 2138.82 | 2139.03 | 180.22 | 185.11 | 190.00 | 0.10156 |
| 23 | 2 | 26 | 2 | 56.14 | 2145.84 | 2146.05 | 2146.26 | 347.95 | 352.79 | 357.63 | 0.10156 |
| 23 | 7 | 26 | 2 | 55.95 | 2154.72 | 2154.93 | 2155.14 | 186.70 | 191.24 | 195.79 | 0.10156 |
| 23 | 2 | 25 | 2 | 55.86 | 2158.91 | 2159.11 | 2159.31 | 277.10 | 281.45 | 285.81 | 0.09766 |
| 24 | 7 | 28 | 2 | 54.30 | 2235.29 | 2235.51 | 2235.74 | 42.66 | 46.94 | 51.23 | 0.08642 |
| 24 | 2 | 29 | 2 | 54.21 | 2239.84 | 2240.07 | 2240.30 | 128.94 | 133.37 | 137.81 | 0.08951 |
| 24 | 7 | 29 | 2 | 54.02 | 2249.55 | 2249.79 | 2250.02 | 312.25 | 316.59 | 320.92 | 0.08951 |
| 24 | 2 | 28 | 2 | 53.93 | 2254.13 | 2254.36 | 2254.58 | 36.84 | 41.01 | 45.18 | 0.08642 |
| 24 | 7 | 29 | 2 | 53.74 | 2263.83 | 2264.07 | 2264.30 | 218.17 | 222.57 | 226.96 | 0.08951 |
| 24 | 2 | 29 | 2 | 53.65 | 2268.71 | 2268.95 | 2269.18 | 309.63 | 314.00 | 318.36 | 0.08951 |
| 24 | 7 | 30 | 2 | 53.45 | 2278.66 | 2278.91 | 2279.15 | 136.81 | 141.47 | 146.12 | 0.09259 |
| 24 | 2 | 30 | 2 | 53.36 | 2283.63 | 2283.87 | 2284.11 | 232.34 | 236.97 | 241.57 | 0.09259 |
| 24 | 7 | 30 | 2 | 53.16 | 2293.83 | 2294.07 | 2294.32 | 61.30 | 65.71 | 70.12 | 0.09259 |
| 24 | 7 | 31 | 2 | 52.98 | 2303.25 | 2303.50 | 2303.75 | 231.89 | 236.35 | 240.80 | 0.09568 |
| 24 | 2 | 31 | 2 | 52.89 | 2308.17 | 2308.42 | 2308.67 | 318.19 | 322.50 | 326.81 | 0.09568 |
| 25 | 7 | 32 | 2 | 52.59 | 2323.78 | 2324.03 | 2324.29 | 223.69 | 228.07 | 232.48 | 0.12121 |
| 25 | 2 | 31 | 2 | 52.50 | 2328.84 | 2329.09 | 2329.34 | 311.41 | 315.82 | 320.25 | 0.11742 |
| 25 | 7 | 32 | 2 | 52.30 | 2339.57 | 2339.82 | 2340.08 | 140.71 | 145.19 | 149.67 | 0.12121 |
| 25 | 2 | 33 | 2 | 52.20 | 2344.90 | 2345.16 | 2345.43 | 232.56 | 237.11 | 241.64 | 0.12500 |
| 25 | 7 | 33 | 2 | 51.99 | 2355.93 | 2356.19 | 2356.46 | 57.71 | 62.10 | 66.48 | 0.12500 |
| 25 | 2 | 34 | 2 | 51.90 | 2361.31 | 2361.59 | 2361.86 | 145.86 | 150.30 | 154.75 | 0.12879 |
| 25 | 7 | 34 | 2 | 51.69 | 2372.65 | 2372.93 | 2373.20 | 326.99 | 331.30 | 335.60 | 0.12879 |
| 25 | 2 | 35 | 2 | 51.59 | 2378.14 | 2378.43 | 2378.71 | 52.75 | 57.13 | 61.51 | 0.13258 |
| 26 | 2 | 35 | 2 | 50.97 | 2412.43 | 2412.72 | 2413.00 | 229.05 | 233.57 | 238.07 | 0.12281 |
| 26 | 7 | 35 | 2 | 50.76 | 2424.09 | 2424.38 | 2424.66 | 55.48 | 60.01 | 64.52 | 0.12281 |
| 26 | 2 | 36 | 2 | 50.66 | 2429.97 | 2430.26 | 2430.55 | 149.79 | 154.46 | 159.13 | 0.12632 |
| 26 | 7 | 36 | 2 | 50.45 | 2441.98 | 2442.27 | 2442.56 | 342.37 | 347.02 | 351.66 | 0.12632 |
| 26 | 2 | 36 | 2 | 50.35 | 2447.62 | 2447.91 | 2448.20 | 71.89 | 76.45 | 81.00 | 0.12632 |
| 26 | 7 | 36 | 2 | 50.14 | 2459.63 | 2459.92 | 2460.21 | 256.62 | 260.99 | 265.34 | 0.12632 |
| 26 | 2 | 35 | 2 | 50.04 | 2465.29 | 2465.58 | 2465.86 | 340.26 | 344.37 | 348.47 | 0.12281 |
| 26 | 7 | 36 | 2 | 49.84 | 2477.28 | 2477.57 | 2477.86 | 156.87 | 161.22 | 165.57 | 0.12632 |
| 30 | 7 | 35 | 2 | 47.50 | 2618.83 | 2619.11 | 2619.39 | 225.04 | 229.43 | 233.81 | 0.12681 |
| 30 | 2 | 34 | 2 | 47.41 | 2624.32 | 2624.59 | 2624.87 | 311.25 | 315.62 | 319.98 | 0.12319 |
| 30 | 7 | 35 | 2 | 47.22 | 2635.97 | 2636.25 | 2636.54 | 136.26 | 140.69 | 145.13 | 0.12681 |
| 30 | 2 | 34 | 2 | 47.14 | 2641.47 | 2641.75 | 2642.02 | 222.29 | 226.55 | 230.82 | 0.12319 |
| 30 | 7 | 35 | 2 | 46.95 | 2653.12 | 2653.40 | 2653.68 | 41.17 | 45.45 | 49.74 | 0.12681 |
| 30 | 2 | 36 | 2 | 46.86 | 2658.84 | 2659.13 | 2659.42 | 127.92 | 132.31 | 136.69 | 0.13043 |
| 30 | 7 | 36 | 2 | 46.68 | 2670.85 | 2671.14 | 2671.43 | 307.27 | 311.52 | 315.79 | 0.13043 |
| 30 | 2 | 31 | 2 | 46.59 | 2676.49 | 2676.74 | 2676.99 | 29.77 | 33.41 | 37.06 | 0.11232 |
| 31 | 7 | 37 | 2 | 46.02 | 2713.86 | 2714.16 | 2714.46 | 249.23 | 254.22 | 259.19 | 0.16017 |
| 31 | 2 | 32 | 2 | 45.93 | 2719.66 | 2719.91 | 2720.17 | 344.83 | 349.04 | 353.24 | 0.13853 |
| 31 | 7 | 36 | 2 | 45.75 | 2731.57 | 2731.86 | 2732.15 | 176.18 | 180.77 | 185.36 | 0.15584 |
| 31 | 7 | 38 | 2 | 45.66 | 2737.24 | 2737.55 | 2737.86 | 265.67 | 270.51 | 275.34 | 0.16450 |
| 31 | 2 | 44 | 2 | 45.55 | 2745.05 | 2745.41 | 2745.76 | 27.85 | 33.41 | 38.96 | 0.19048 |
| 31 | 7 | 44 | 2 | 45.33 | 2759.52 | 2759.87 | 2760.23 | 258.66 | 264.42 | 270.19 | 0.19048 |
| 32 | 2 | 33 | 2 | 44.78 | 2796.80 | 2797.07 | 2797.33 | 167.19 | 171.84 | 176.49 | 0.12840 |
| 32 | 7 | 34 | 2 | 44.62 | 2808.11 | 2808.39 | 2808.66 | 3.24 | 7.94 | 12.65 | 0.13230 |
| 32 | 2 | 33 | 2 | 44.54 | 2813.47 | 2813.73 | 2814.00 | 94.09 | 98.52 | 102.96 | 0.12840 |
| 32 | 7 | 34 | 2 | 44.37 | 2824.78 | 2825.06 | 2825.33 | 279.92 | 284.46 | 289.02 | 0.13230 |
| 32 | 2 | 22 | 2 | 44.30 | 2830.05 | 2830.22 | 2830.40 | 7.70 | 10.66 | 13.62 | 0.08560 |
| 32 | 7 | 33 | 2 | 44.14 | 2840.81 | 2841.08 | 2841.34 | 186.73 | 191.13 | 195.52 | 0.12840 |
| 32 | 2 | 34 | 2 | 44.06 | 2846.44 | 2846.72 | 2846.99 | 278.78 | 283.21 | 287.63 | 0.13230 |
| 32 | 7 | 34 | 2 | 43.90 | 2857.78 | 2858.05 | 2858.33 | 98.46 | 102.75 | 107.01 | 0.13230 |
| 33 | 7 | 46 | 2 | 43.41 | 2892.60 | 2892.97 | 2893.34 | 280.09 | 285.84 | 291.60 | 0.21005 |
| 33 | 2 | 51 | 2 | 43.29 | 2901.35 | 2901.76 | 2902.17 | 56.21 | 62.66 | 69.14 | 0.23288 |
| 33 | 7 | 53 | 2 | 43.05 | 2918.43 | 2918.86 | 2919.29 | 333.11 | 340.30 | 347.50 | 0.24201 |
| 33 | 7 | 36 | 2 | 42.82 | 2935.08 | 2935.38 | 2935.67 | 252.34 | 257.05 | 261.73 | 0.16438 |
| 33 | 2 | 33 | 2 | 42.74 | 2940.75 | 2941.01 | 2941.28 | 341.87 | 345.98 | 350.11 | 0.15068 |
| 34 | 2 | 36 | 2 | 42.27 | 2975.84 | 2976.13 | 2976.42 | 144.78 | 149.02 | 153.27 | 0.13187 |
| 34 | 7 | 36 | 2 | 42.10 | 2987.86 | 2988.15 | 2988.44 | 326.86 | 331.49 | 336.13 | 0.13187 |
| 34 | 2 | 36 | 2 | 42.02 | 2993.78 | 2994.07 | 2994.36 | 62.75 | 67.48 | 72.21 | 0.13187 |
| 34 | 7 | 36 | 2 | 41.86 | 3005.79 | 3006.08 | 3006.37 | 261.58 | 266.37 | 271.17 | 0.13187 |
| 34 | 2 | 31 | 2 | 41.79 | 3011.43 | 3011.68 | 3011.93 | 354.26 | 358.36 | 2.46 | 0.11355 |
| 34 | 7 | 35 | 2 | 41.63 | 3023.00 | 3023.28 | 3023.56 | 182.22 | 186.76 | 191.32 | 0.12821 |
| 34 | 2 | 31 | 2 | 41.53 | 3030.83 | 3031.08 | 3031.33 | 307.49 | 311.44 | 315.41 | 0.11355 |
| 34 | 7 | 32 | 2 | 41.39 | 3041.58 | 3041.84 | 3042.09 | 117.49 | 121.57 | 125.66 | 0.11722 |
| 35 | 2 | 33 | 2 | 40.40 | 3117.81 | 3118.07 | 3118.34 | 206.22 | 210.10 | 213.98 | 0.12268 |
| 35 | 7 | 33 | 2 | 40.26 | 3128.92 | 3129.18 | 3129.45 | 9.01 | 12.92 | 16.83 | 0.12268 |
| 35 | 2 | 33 | 2 | 40.20 | 3134.25 | 3134.51 | 3134.78 | 87.13 | 91.04 | 94.95 | 0.12268 |
| 35 | 7 | 34 | 2 | 40.05 | 3145.44 | 3145.72 | 3145.99 | 250.53 | 254.31 | 258.08 | 0.12639 |
| 35 | 2 | 34 | 2 | 39.99 | 3150.79 | 3151.07 | 3151.34 | 320.98 | 324.45 | 327.92 | 0.12639 |
| 35 | 7 | 34 | 2 | 39.84 | 3162.06 | 3162.33 | 3162.61 | 105.91 | 109.67 | 113.43 | 0.12639 |
| 35 | 2 | 34 | 2 | 39.78 | 3167.44 | 3167.71 | 3167.99 | 181.12 | 185.00 | 188.90 | 0.12639 |
| 35 | 7 | 34 | 2 | 39.64 | 3178.78 | 3179.05 | 3179.33 | 338.17 | 341.89 | 345.62 | 0.12639 |
| 36 | 7 | 35 | 2 | 39.22 | 3212.54 | 3212.82 | 3213.11 | 82.27 | 86.20 | 90.14 | 0.12500 |
| 36 | 2 | 35 | 2 | 39.15 | 3218.01 | 3218.29 | 3218.57 | 158.41 | 162.33 | 166.26 | 0.12500 |
| 36 | 7 | 35 | 2 | 39.01 | 3229.56 | 3229.84 | 3230.12 | 319.54 | 323.49 | 327.46 | 0.12500 |
| 36 | 2 | 35 | 2 | 38.94 | 3235.09 | 3235.37 | 3235.65 | 37.02 | 40.96 | 44.88 | 0.12500 |
| 36 | 7 | 35 | 2 | 38.80 | 3246.71 | 3246.99 | 3247.27 | 197.31 | 201.13 | 204.94 | 0.12500 |
| 36 | 2 | 35 | 2 | 38.73 | 3252.26 | 3252.54 | 3252.82 | 270.70 | 274.31 | 277.91 | 0.12500 |
| 36 | 7 | 35 | 2 | 38.59 | 3263.95 | 3264.23 | 3264.51 | 55.24 | 58.75 | 62.25 | 0.12500 |
| 36 | 2 | 35 | 2 | 38.53 | 3269.53 | 3269.81 | 3270.09 | 125.16 | 128.68 | 132.21 | 0.12500 |
| 37 | 7 | 36 | 2 | 38.18 | 3298.79 | 3299.08 | 3299.37 | 181.27 | 185.67 | 190.06 | 0.12500 |
| 37 | 2 | 36 | 2 | 38.11 | 3304.41 | 3304.70 | 3304.99 | 264.23 | 268.35 | 272.45 | 0.12500 |
| 37 | 7 | 36 | 2 | 37.97 | 3316.33 | 3316.62 | 3316.91 | 66.32 | 70.21 | 74.11 | 0.12500 |
| 37 | 2 | 36 | 2 | 37.90 | 3322.01 | 3322.30 | 3322.59 | 141.96 | 145.78 | 149.61 | 0.12500 |
| 37 | 7 | 36 | 2 | 37.76 | 3333.99 | 3334.28 | 3334.57 | 300.82 | 304.71 | 308.59 | 0.12500 |
| 37 | 2 | 36 | 2 | 37.69 | 3339.72 | 3340.01 | 3340.30 | 17.97 | 21.92 | 25.86 | 0.12500 |
| 37 | 7 | 36 | 2 | 37.55 | 3351.78 | 3352.07 | 3352.36 | 184.92 | 189.03 | 193.16 | 0.12500 |
| 37 | 2 | 36 | 2 | 37.48 | 3357.53 | 3357.82 | 3358.11 | 267.98 | 272.28 | 276.59 | 0.12500 |
| 38 | 2 | 37 | 2 | 37.18 | 3384.16 | 3384.46 | 3384.76 | 317.89 | 322.53 | 327.18 | 0.12458 |
| 38 | 7 | 37 | 2 | 37.03 | 3396.40 | 3396.69 | 3396.99 | 147.69 | 152.32 | 156.94 | 0.12458 |
| 38 | 2 | 37 | 2 | 36.97 | 3402.25 | 3402.55 | 3402.85 | 237.28 | 241.74 | 246.19 | 0.12458 |
| 38 | 7 | 37 | 2 | 36.83 | 3414.57 | 3414.87 | 3415.17 | 54.23 | 58.37 | 62.49 | 0.12458 |
| 38 | 2 | 37 | 2 | 36.76 | 3420.47 | 3420.76 | 3421.06 | 134.40 | 138.37 | 142.34 | 0.12458 |
| 38 | 7 | 37 | 2 | 36.62 | 3432.87 | 3433.17 | 3433.47 | 298.43 | 302.40 | 306.35 | 0.12458 |
| 38 | 2 | 37 | 2 | 36.55 | 3438.79 | 3439.09 | 3439.38 | 17.00 | 20.95 | 24.91 | 0.12458 |
| 38 | 7 | 38 | 2 | 36.41 | 3451.25 | 3451.56 | 3451.87 | 180.01 | 183.99 | 187.97 | 0.12795 |
| 39 | 2 | 38 | 2 | 36.14 | 3475.76 | 3476.07 | 3476.38 | 153.48 | 157.90 | 162.30 | 0.14179 |
| 39 | 7 | 38 | 2 | 36.00 | 3488.38 | 3488.68 | 3488.99 | 338.21 | 342.80 | 347.38 | 0.14179 |
| 39 | 7 | 38 | 2 | 35.81 | 3505.46 | 3505.76 | 3506.07 | 236.71 | 241.28 | 245.85 | 0.14179 |
| 39 | 2 | 38 | 2 | 35.74 | 3511.53 | 3511.83 | 3512.14 | 326.36 | 330.85 | 335.32 | 0.14179 |
| 39 | 7 | 39 | 2 | 35.60 | 3524.31 | 3524.62 | 3524.94 | 144.18 | 148.32 | 152.45 | 0.14552 |
| 39 | 2 | 38 | 2 | 35.49 | 3534.53 | 3534.84 | 3535.15 | 275.50 | 279.32 | 283.13 | 0.14179 |
| 39 | 7 | 39 | 2 | 35.34 | 3547.40 | 3547.72 | 3548.03 | 74.14 | 78.17 | 82.22 | 0.14552 |
| 40 | 7 | 39 | 2 | 35.14 | 3566.58 | 3566.90 | 3567.21 | 328.11 | 332.16 | 336.20 | 0.12342 |
| 40 | 2 | 39 | 2 | 35.07 | 3572.72 | 3573.03 | 3573.35 | 46.92 | 50.97 | 55.02 | 0.12342 |
| 40 | 7 | 39 | 2 | 34.93 | 3585.76 | 3586.07 | 3586.39 | 215.38 | 219.53 | 223.71 | 0.12342 |
| 40 | 2 | 39 | 2 | 34.86 | 3591.99 | 3592.31 | 3592.62 | 300.71 | 305.17 | 309.64 | 0.12342 |
| 40 | 7 | 40 | 2 | 34.72 | 3605.10 | 3605.43 | 3605.75 | 133.63 | 138.54 | 143.48 | 0.12658 |
| 40 | 2 | 40 | 2 | 34.65 | 3611.39 | 3611.71 | 3612.03 | 229.29 | 234.12 | 238.93 | 0.12658 |
| 40 | 7 | 40 | 2 | 34.51 | 3624.58 | 3624.90 | 3625.22 | 58.19 | 62.84 | 67.50 | 0.12658 |
| 40 | 2 | 40 | 2 | 34.45 | 3630.90 | 3631.22 | 3631.55 | 151.63 | 156.48 | 161.32 | 0.12658 |
| 41 | 7 | 41 | 2 | 34.14 | 3659.87 | 3660.20 | 3660.53 | 211.01 | 215.60 | 220.20 | 0.13443 |
| 41 | 2 | 40 | 2 | 34.07 | 3666.22 | 3666.54 | 3666.87 | 301.25 | 305.94 | 310.64 | 0.13115 |
| 41 | 7 | 41 | 2 | 33.93 | 3679.64 | 3679.97 | 3680.30 | 132.40 | 136.86 | 141.32 | 0.13443 |
| 41 | 2 | 41 | 2 | 33.87 | 3686.07 | 3686.40 | 3686.74 | 219.09 | 223.59 | 228.09 | 0.13443 |
| 41 | 7 | 41 | 2 | 33.72 | 3699.59 | 3699.92 | 3700.25 | 48.96 | 53.66 | 58.35 | 0.13443 |
| 41 | 2 | 41 | 2 | 33.66 | 3706.05 | 3706.38 | 3706.71 | 140.52 | 145.20 | 149.88 | 0.13443 |
| 41 | 7 | 41 | 2 | 33.52 | 3719.65 | 3719.98 | 3720.31 | 331.37 | 335.94 | 340.50 | 0.13443 |
| 41 | 2 | 19 | 2 | 33.45 | 3725.89 | 3726.05 | 3726.20 | 57.84 | 59.99 | 62.13 | 0.06230 |
| 42 | 7 | 42 | 2 | 32.69 | 3800.41 | 3800.75 | 3801.09 | 2.48 | 6.70 | 10.90 | 0.16406 |
| 42 | 2 | 42 | 2 | 32.63 | 3807.03 | 3807.37 | 3807.71 | 82.69 | 86.62 | 90.53 | 0.16406 |
| 42 | 7 | 43 | 2 | 32.49 | 3821.06 | 3821.40 | 3821.75 | 238.50 | 242.57 | 246.65 | 0.16797 |
| 42 | 2 | 43 | 2 | 32.42 | 3827.77 | 3828.12 | 3828.47 | 322.68 | 327.28 | 331.89 | 0.16797 |
| 42 | 7 | 43 | 2 | 32.28 | 3841.89 | 3842.23 | 3842.58 | 147.97 | 152.40 | 156.82 | 0.16797 |
| 42 | 2 | 43 | 2 | 32.21 | 3848.64 | 3848.99 | 3849.34 | 233.92 | 238.32 | 242.72 | 0.16797 |
| 43 | 2 | 43 | 2 | 31.79 | 3890.78 | 3891.13 | 3891.47 | 353.06 | 356.15 | 359.25 | 0.16350 |
| 43 | 7 | 44 | 2 | 31.65 | 3905.16 | 3905.51 | 3905.87 | 137.36 | 141.43 | 145.52 | 0.16730 |
| 43 | 2 | 44 | 2 | 31.59 | 3912.03 | 3912.38 | 3912.74 | 220.57 | 225.14 | 229.73 | 0.16730 |
| 43 | 7 | 44 | 2 | 31.44 | 3926.49 | 3926.85 | 3927.20 | 78.41 | 83.74 | 89.03 | 0.16730 |
| 43 | 2 | 44 | 2 | 31.38 | 3933.43 | 3933.78 | 3934.14 | 179.99 | 185.09 | 190.17 | 0.16730 |
| 43 | 7 | 44 | 2 | 31.24 | 3947.99 | 3948.35 | 3948.70 | 18.43 | 23.16 | 27.93 | 0.16730 |
| 44 | 2 | 45 | 2 | 30.96 | 3976.61 | 3976.97 | 3977.33 | 88.85 | 94.33 | 99.78 | 0.19912 |
| 44 | 7 | 45 | 2 | 30.82 | 3991.34 | 3991.70 | 3992.07 | 317.64 | 323.82 | 330.03 | 0.19912 |
| 44 | 2 | 1 | 2 | 30.76 | 3997.95 | 3997.95 | 3997.96 | 72.06 | 72.20 | 72.33 | 0.00442 |
| 44 | 7 | 45 | 2 | 30.63 | 4011.72 | 4012.08 | 4012.44 | 314.08 | 320.36 | 326.65 | 0.19912 |
| 44 | 2 | 45 | 2 | 30.56 | 4018.81 | 4019.18 | 4019.54 | 74.93 | 80.96 | 86.98 | 0.19912 |
| 44 | 7 | 45 | 2 | 30.42 | 4033.73 | 4034.10 | 4034.46 | 313.36 | 318.96 | 324.55 | 0.19912 |
| 45 | 7 | 46 | 2 | 30.00 | 4078.17 | 4078.55 | 4078.92 | 220.91 | 226.01 | 231.11 | 0.16547 |
| 45 | 2 | 46 | 2 | 29.93 | 4085.40 | 4085.77 | 4086.14 | 321.26 | 326.46 | 331.68 | 0.16547 |
| 45 | 7 | 46 | 2 | 29.79 | 4100.59 | 4100.96 | 4101.33 | 170.14 | 175.12 | 180.10 | 0.16547 |
| 45 | 2 | 46 | 2 | 29.72 | 4107.87 | 4108.25 | 4108.62 | 267.31 | 272.24 | 277.19 | 0.16547 |
| 45 | 7 | 47 | 2 | 29.58 | 4123.16 | 4123.53 | 4123.91 | 106.61 | 111.42 | 116.22 | 0.16906 |
| 45 | 2 | 47 | 2 | 29.52 | 4130.49 | 4130.86 | 4131.24 | 198.14 | 202.76 | 207.37 | 0.16906 |
| 46 | 7 | 47 | 2 | 29.17 | 4168.74 | 4169.12 | 4169.50 | 240.39 | 244.24 | 248.13 | 0.16491 |
| 46 | 2 | 47 | 2 | 29.10 | 4176.15 | 4176.52 | 4176.90 | 320.13 | 324.42 | 328.73 | 0.16491 |
| 46 | 7 | 47 | 2 | 28.96 | 4191.73 | 4192.10 | 4192.48 | 145.28 | 150.02 | 154.74 | 0.16491 |
| 46 | 2 | 48 | 2 | 28.89 | 4199.18 | 4199.57 | 4199.96 | 239.05 | 243.90 | 248.74 | 0.16842 |
| 46 | 7 | 48 | 2 | 28.75 | 4214.86 | 4215.25 | 4215.63 | 76.41 | 81.44 | 86.46 | 0.16842 |
| 46 | 2 | 48 | 2 | 28.68 | 4222.38 | 4222.76 | 4223.15 | 176.13 | 181.30 | 186.46 | 0.16842 |
| 47 | 2 | 48 | 2 | 28.47 | 4245.71 | 4246.09 | 4246.48 | 108.69 | 113.44 | 118.21 | 0.16495 |
| 47 | 7 | 48 | 2 | 28.33 | 4261.58 | 4261.96 | 4262.35 | 299.26 | 303.68 | 308.12 | 0.16495 |
| 47 | 2 | 48 | 2 | 28.26 | 4269.20 | 4269.58 | 4269.97 | 26.08 | 30.50 | 34.92 | 0.16495 |
| 47 | 7 | 49 | 2 | 28.12 | 4285.17 | 4285.57 | 4285.96 | 212.82 | 217.47 | 222.15 | 0.16838 |
| 47 | 2 | 49 | 2 | 28.05 | 4292.82 | 4293.21 | 4293.61 | 304.27 | 309.11 | 313.98 | 0.16838 |
| 47 | 7 | 49 | 2 | 27.91 | 4308.89 | 4309.28 | 4309.68 | 121.98 | 125.68 | 129.36 | 0.16838 |
| 48 | 2 | 49 | 2 | 27.64 | 4340.51 | 4340.91 | 4341.30 | 67.59 | 71.74 | 75.89 | 0.16388 |
| 48 | 7 | 50 | 2 | 27.49 | 4356.78 | 4357.19 | 4357.59 | 245.95 | 250.28 | 254.61 | 0.16722 |
| 48 | 2 | 50 | 2 | 27.43 | 4364.58 | 4364.99 | 4365.39 | 328.80 | 333.05 | 337.28 | 0.16722 |
| 48 | 7 | 50 | 2 | 27.28 | 4380.95 | 4381.36 | 4381.76 | 136.74 | 140.75 | 144.74 | 0.16722 |
| 48 | 2 | 50 | 2 | 27.22 | 4388.80 | 4389.20 | 4389.60 | 214.96 | 219.05 | 223.12 | 0.16722 |
| 48 | 7 | 50 | 2 | 27.07 | 4405.28 | 4405.68 | 4406.09 | 29.28 | 33.63 | 37.99 | 0.16722 |
| 50 | 7 | 70 | 2 | 23.20 | 4880.35 | 4880.91 | 4881.47 | 52.41 | 58.36 | 64.27 | 0.24911 |
| 50 | 2 | 69 | 2 | 23.12 | 4891.12 | 4891.68 | 4892.23 | 160.54 | 165.95 | 171.35 | 0.24555 |
| 50 | 7 | 71 | 2 | 22.94 | 4914.02 | 4914.59 | 4915.17 | 30.45 | 36.25 | 42.03 | 0.25267 |
| 50 | 2 | 71 | 2 | 22.86 | 4925.05 | 4925.62 | 4926.20 | 142.29 | 148.11 | 153.91 | 0.25267 |
| 51 | 2 | 72 | 2 | 20.81 | 5199.82 | 5200.40 | 5200.98 | 83.84 | 89.91 | 95.99 | 0.24828 |
| 51 | 7 | 72 | 2 | 20.63 | 5223.35 | 5223.93 | 5224.52 | 322.70 | 328.36 | 334.01 | 0.24828 |
| 51 | 2 | 73 | 2 | 20.55 | 5234.65 | 5235.24 | 5235.83 | 70.02 | 75.44 | 80.84 | 0.25172 |
| 51 | 7 | 73 | 2 | 20.38 | 5258.27 | 5258.86 | 5259.45 | 279.18 | 284.36 | 289.52 | 0.25172 |
| 54 | 2 | 72 | 2 | 17.14 | 5727.77 | 5728.35 | 5728.93 | 218.43 | 223.84 | 229.24 | 0.28800 |
| 54 | 7 | 72 | 2 | 16.98 | 5751.12 | 5751.70 | 5752.28 | 78.31 | 83.72 | 89.12 | 0.28800 |
| 54 | 2 | 38 | 2 | 16.91 | 5761.84 | 5762.15 | 5762.45 | 177.76 | 180.59 | 183.44 | 0.15200 |
| 54 | 7 | 68 | 2 | 16.77 | 5783.34 | 5783.89 | 5784.44 | 18.13 | 23.28 | 28.45 | 0.27200 |
| 55 | 2 | 64 | 2 | 14.93 | 6070.07 | 6070.59 | 6071.10 | 221.08 | 225.70 | 230.33 | 0.24903 |
| 55 | 7 | 64 | 2 | 14.79 | 6091.02 | 6091.54 | 6092.05 | 54.60 | 59.44 | 64.27 | 0.24903 |
| 55 | 2 | 64 | 2 | 14.73 | 6101.07 | 6101.59 | 6102.10 | 148.29 | 153.13 | 157.95 | 0.24903 |
| 55 | 7 | 65 | 2 | 14.60 | 6122.07 | 6122.59 | 6123.12 | 351.51 | 356.73 | 1.97 | 0.25292 |
| 56 | 7 | 68 | 2 | 12.84 | 6409.51 | 6410.05 | 6410.60 | 94.99 | 100.12 | 105.25 | 0.25000 |
| 56 | 2 | 68 | 2 | 12.78 | 6419.98 | 6420.53 | 6421.08 | 193.71 | 198.85 | 203.98 | 0.25000 |
| 56 | 7 | 68 | 2 | 12.64 | 6442.15 | 6442.70 | 6443.25 | 35.15 | 39.98 | 44.80 | 0.25000 |
| 56 | 2 | 68 | 2 | 12.58 | 6452.73 | 6453.28 | 6453.83 | 127.49 | 132.23 | 136.98 | 0.25000 |
| 57 | 7 | 72 | 2 | 10.87 | 6744.36 | 6744.94 | 6745.52 | 355.12 | 359.82 | 4.53 | 0.25000 |
| 57 | 2 | 72 | 2 | 10.81 | 6755.35 | 6755.93 | 6756.51 | 85.23 | 90.10 | 94.98 | 0.25000 |
| 57 | 7 | 72 | 2 | 10.68 | 6778.65 | 6779.23 | 6779.82 | 279.21 | 283.86 | 288.50 | 0.25000 |
| 57 | 2 | 72 | 2 | 10.61 | 6789.69 | 6790.27 | 6790.85 | 6.38 | 10.89 | 15.40 | 0.25000 |
| 58 | 2 | 74 | 2 | 9.04 | 7070.39 | 7070.99 | 7071.59 | 129.20 | 133.60 | 138.03 | 0.24749 |
| 58 | 7 | 75 | 2 | 8.91 | 7094.57 | 7095.17 | 7095.78 | 305.09 | 309.38 | 313.65 | 0.25084 |
| 58 | 2 | 75 | 2 | 8.84 | 7106.23 | 7106.83 | 7107.44 | 28.50 | 32.95 | 37.42 | 0.25084 |
| 58 | 7 | 75 | 2 | 8.71 | 7130.50 | 7131.11 | 7131.71 | 222.31 | 227.22 | 232.11 | 0.25084 |
| 59 | 7 | 77 | 2 | 7.33 | 7386.26 | 7386.88 | 7387.50 | 321.25 | 325.82 | 330.38 | 0.24839 |
| 59 | 2 | 77 | 2 | 7.27 | 7398.12 | 7398.74 | 7399.36 | 47.86 | 52.36 | 56.84 | 0.24839 |
| 59 | 7 | 78 | 2 | 7.13 | 7423.24 | 7423.87 | 7424.50 | 227.48 | 232.03 | 236.61 | 0.25161 |
| 59 | 2 | 78 | 2 | 7.07 | 7435.23 | 7435.86 | 7436.49 | 315.85 | 320.55 | 325.25 | 0.25161 |
| 60 | 7 | 80 | 2 | 5.56 | 7725.33 | 7725.97 | 7726.62 | 313.17 | 317.71 | 322.28 | 0.25000 |
| 60 | 2 | 80 | 2 | 5.49 | 7737.61 | 7738.25 | 7738.90 | 39.02 | 43.48 | 47.92 | 0.25000 |
| 60 | 7 | 80 | 2 | 5.36 | 7763.56 | 7764.21 | 7764.85 | 212.63 | 216.93 | 221.23 | 0.25000 |
| 60 | 2 | 80 | 2 | 5.30 | 7775.91 | 7776.56 | 7777.20 | 296.02 | 300.41 | 304.80 | 0.25000 |
| 1 | 6 | 1 | 10 | 142.26 | 145.54 | 145.58 | 145.62 | 144.48 | 143.29 | 142.10 | 0.01471 |
| 1 | 6 | 1 | 10 | 142.25 | 145.70 | 145.74 | 145.78 | 139.72 | 138.53 | 137.34 | 0.01471 |
| 1 | 6 | 1 | 10 | 142.24 | 145.86 | 145.90 | 145.95 | 134.96 | 133.78 | 132.59 | 0.01471 |
| 1 | 6 | 1 | 10 | 142.23 | 146.03 | 146.07 | 146.11 | 130.21 | 129.02 | 127.83 | 0.01471 |
| 1 | 3 | 1 | 12 | 141.96 | 151.24 | 151.29 | 151.34 | 337.38 | 335.98 | 334.57 | 0.01471 |
| 1 | 3 | 1 | 12 | 141.95 | 151.44 | 151.49 | 151.53 | 331.76 | 330.35 | 328.94 | 0.01471 |
| 1 | 3 | 1 | 12 | 141.94 | 151.63 | 151.68 | 151.73 | 326.13 | 324.73 | 323.32 | 0.01471 |
| 1 | 3 | 1 | 12 | 141.93 | 151.82 | 151.87 | 151.92 | 320.51 | 319.10 | 317.70 | 0.01471 |
| 1 | 6 | 1 | 12 | 141.78 | 154.70 | 154.75 | 154.80 | 237.19 | 235.79 | 234.40 | 0.01471 |
| 1 | 6 | 1 | 12 | 141.77 | 154.90 | 154.94 | 154.99 | 231.61 | 230.21 | 228.82 | 0.01471 |
| 1 | 6 | 1 | 12 | 141.76 | 155.09 | 155.14 | 155.19 | 226.03 | 224.64 | 223.24 | 0.01471 |
| 1 | 6 | 1 | 12 | 141.75 | 155.28 | 155.33 | 155.38 | 220.46 | 219.06 | 217.67 | 0.01471 |
| 1 | 3 | 1 | 10 | 141.47 | 160.64 | 160.68 | 160.72 | 67.29 | 66.15 | 65.00 | 0.01471 |
| 1 | 3 | 1 | 10 | 141.46 | 160.80 | 160.84 | 160.88 | 62.71 | 61.57 | 60.42 | 0.01471 |
| 1 | 3 | 1 | 10 | 141.46 | 160.96 | 161.00 | 161.04 | 58.13 | 56.99 | 55.84 | 0.01471 |
| 1 | 3 | 1 | 10 | 141.45 | 161.12 | 161.16 | 161.20 | 53.55 | 52.41 | 51.27 | 0.01471 |
| 1 | 3 | 1 | 10 | 141.44 | 161.28 | 161.32 | 161.36 | 48.98 | 47.84 | 46.69 | 0.01471 |
| 1 | 6 | 1 | 10 | 141.29 | 164.21 | 164.25 | 164.29 | 326.22 | 325.08 | 323.94 | 0.01471 |
| 1 | 6 | 1 | 10 | 141.28 | 164.37 | 164.41 | 164.45 | 321.67 | 320.54 | 319.40 | 0.01471 |
| 1 | 6 | 1 | 10 | 141.27 | 164.53 | 164.57 | 164.61 | 317.13 | 316.00 | 314.86 | 0.01471 |
| 1 | 6 | 1 | 10 | 141.26 | 164.69 | 164.73 | 164.77 | 312.59 | 311.46 | 310.32 | 0.01471 |
| 1 | 6 | 1 | 10 | 141.26 | 164.86 | 164.90 | 164.94 | 308.06 | 306.92 | 305.79 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.97 | 170.38 | 170.42 | 170.46 | 153.75 | 152.63 | 151.51 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.96 | 170.54 | 170.58 | 170.62 | 149.27 | 148.16 | 147.04 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.95 | 170.70 | 170.74 | 170.78 | 144.80 | 143.68 | 142.57 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.95 | 170.86 | 170.90 | 170.94 | 140.33 | 139.21 | 138.10 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.94 | 171.02 | 171.06 | 171.10 | 135.86 | 134.75 | 133.63 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.78 | 174.07 | 174.11 | 174.15 | 51.72 | 50.62 | 49.51 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.77 | 174.23 | 174.27 | 174.31 | 47.29 | 46.18 | 45.07 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.77 | 174.39 | 174.44 | 174.48 | 42.86 | 41.75 | 40.64 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.76 | 174.56 | 174.60 | 174.64 | 38.43 | 37.32 | 36.21 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.75 | 174.72 | 174.76 | 174.80 | 34.00 | 32.89 | 31.79 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.46 | 180.41 | 180.45 | 180.49 | 238.75 | 237.66 | 236.57 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.45 | 180.57 | 180.61 | 180.65 | 234.39 | 233.30 | 232.21 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.44 | 180.73 | 180.77 | 180.81 | 230.03 | 228.94 | 227.85 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.43 | 180.89 | 180.93 | 180.98 | 225.67 | 224.58 | 223.49 | 0.01471 |
| 1 | 3 | 1 | 10 | 140.43 | 181.06 | 181.10 | 181.14 | 221.31 | 220.22 | 219.13 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.26 | 184.22 | 184.27 | 184.31 | 135.98 | 134.90 | 133.82 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.26 | 184.39 | 184.43 | 184.47 | 131.66 | 130.58 | 129.50 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.25 | 184.55 | 184.59 | 184.63 | 127.34 | 126.26 | 125.18 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.24 | 184.71 | 184.75 | 184.79 | 123.01 | 121.93 | 120.86 | 0.01471 |
| 1 | 6 | 1 | 10 | 140.23 | 184.87 | 184.91 | 184.95 | 118.70 | 117.62 | 116.54 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.94 | 190.70 | 190.74 | 190.78 | 323.81 | 322.75 | 321.69 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.93 | 190.86 | 190.90 | 190.94 | 319.56 | 318.50 | 317.44 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.92 | 191.02 | 191.06 | 191.10 | 315.31 | 314.25 | 313.19 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.91 | 191.18 | 191.22 | 191.26 | 311.07 | 310.01 | 308.95 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.91 | 191.35 | 191.39 | 191.43 | 306.82 | 305.76 | 304.70 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.74 | 194.58 | 194.62 | 194.66 | 222.13 | 221.07 | 220.02 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.73 | 194.74 | 194.78 | 194.82 | 217.92 | 216.87 | 215.82 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.73 | 194.90 | 194.94 | 194.98 | 213.72 | 212.67 | 211.62 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.72 | 195.06 | 195.10 | 195.14 | 209.52 | 208.47 | 207.42 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.71 | 195.22 | 195.26 | 195.30 | 205.32 | 204.27 | 203.22 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.40 | 201.35 | 201.39 | 201.43 | 47.26 | 46.23 | 45.20 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.39 | 201.51 | 201.55 | 201.59 | 43.14 | 42.11 | 41.08 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.39 | 201.68 | 201.72 | 201.76 | 39.02 | 37.99 | 36.96 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.38 | 201.84 | 201.88 | 201.92 | 34.91 | 33.88 | 32.85 | 0.01471 |
| 1 | 3 | 1 | 10 | 139.37 | 202.00 | 202.04 | 202.08 | 30.79 | 29.76 | 28.74 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.20 | 205.38 | 205.42 | 205.46 | 305.11 | 304.09 | 303.07 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.19 | 205.54 | 205.58 | 205.62 | 301.04 | 300.02 | 299.01 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.18 | 205.70 | 205.74 | 205.78 | 296.97 | 295.96 | 294.94 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.18 | 205.86 | 205.90 | 205.94 | 292.91 | 291.90 | 290.88 | 0.01471 |
| 1 | 6 | 1 | 10 | 139.17 | 206.02 | 206.06 | 206.10 | 288.85 | 287.84 | 286.82 | 0.01471 |
| 1 | 3 | 1 | 10 | 138.86 | 212.22 | 212.26 | 212.30 | 134.51 | 133.52 | 132.52 | 0.01471 |
| 1 | 3 | 1 | 10 | 138.85 | 212.38 | 212.42 | 212.46 | 130.53 | 129.54 | 128.55 | 0.01471 |
| 1 | 3 | 1 | 10 | 138.84 | 212.54 | 212.58 | 212.62 | 126.56 | 125.56 | 124.57 | 0.01471 |
| 1 | 3 | 1 | 10 | 138.84 | 212.70 | 212.74 | 212.78 | 122.58 | 121.59 | 120.59 | 0.01471 |
| 1 | 3 | 1 | 10 | 138.83 | 212.86 | 212.90 | 212.94 | 118.61 | 117.61 | 116.62 | 0.01471 |
| 1 | 3 | 1 | 10 | 138.82 | 213.02 | 213.06 | 213.10 | 114.64 | 113.64 | 112.65 | 0.01471 |
| 2 | 6 | 16 | 2 | 135.21 | 287.26 | 287.39 | 287.52 | 275.18 | 272.65 | 270.12 | 0.05316 |
| 2 | 6 | 16 | 2 | 135.19 | 287.78 | 287.91 | 288.04 | 265.06 | 262.53 | 260.00 | 0.05316 |
| 2 | 6 | 21 | 2 | 134.88 | 294.21 | 294.38 | 294.55 | 140.81 | 137.59 | 134.37 | 0.06977 |
| 2 | 6 | 21 | 2 | 134.85 | 294.89 | 295.06 | 295.23 | 127.94 | 124.73 | 121.52 | 0.06977 |
| 2 | 3 | 19 | 2 | 134.27 | 307.25 | 307.40 | 307.56 | 260.63 | 257.91 | 255.18 | 0.06312 |
| 2 | 3 | 19 | 2 | 134.25 | 307.86 | 308.02 | 308.17 | 249.74 | 247.03 | 244.31 | 0.06312 |
| 2 | 3 | 19 | 2 | 134.22 | 308.48 | 308.63 | 308.78 | 238.89 | 236.19 | 233.48 | 0.06312 |
| 2 | 6 | 20 | 2 | 133.90 | 315.19 | 315.35 | 315.51 | 122.58 | 119.84 | 117.10 | 0.06645 |
| 2 | 6 | 20 | 2 | 133.87 | 315.83 | 315.99 | 316.15 | 111.63 | 108.89 | 106.17 | 0.06645 |
| 2 | 6 | 20 | 2 | 133.84 | 316.48 | 316.64 | 316.80 | 100.72 | 98.00 | 95.28 | 0.06645 |
| 2 | 3 | 18 | 2 | 133.35 | 327.23 | 327.37 | 327.52 | 285.30 | 283.01 | 280.72 | 0.05980 |
| 2 | 3 | 18 | 2 | 133.32 | 327.81 | 327.95 | 328.10 | 276.15 | 273.87 | 271.59 | 0.05980 |
| 2 | 6 | 18 | 2 | 133.08 | 332.99 | 333.14 | 333.28 | 196.04 | 193.83 | 191.63 | 0.05980 |
| 2 | 6 | 18 | 2 | 133.05 | 333.57 | 333.72 | 333.86 | 187.23 | 185.04 | 182.84 | 0.05980 |
| 2 | 3 | 19 | 2 | 132.62 | 343.07 | 343.22 | 343.38 | 48.00 | 45.83 | 43.65 | 0.06312 |
| 2 | 3 | 19 | 2 | 132.59 | 343.68 | 343.84 | 343.99 | 39.32 | 37.15 | 34.99 | 0.06312 |
| 3 | 3 | 23 | 2 | 128.20 | 442.88 | 443.06 | 443.25 | 99.85 | 98.59 | 97.34 | 0.04618 |
| 3 | 3 | 23 | 2 | 128.17 | 443.62 | 443.80 | 443.99 | 94.84 | 93.59 | 92.35 | 0.04618 |
| 3 | 3 | 23 | 2 | 128.14 | 444.36 | 444.54 | 444.73 | 89.88 | 88.65 | 87.42 | 0.04618 |
| 3 | 6 | 23 | 2 | 127.80 | 452.25 | 452.44 | 452.62 | 39.87 | 38.76 | 37.64 | 0.04618 |
| 3 | 6 | 23 | 2 | 127.77 | 453.00 | 453.18 | 453.37 | 35.43 | 34.32 | 33.22 | 0.04618 |
| 3 | 6 | 23 | 2 | 127.74 | 453.74 | 453.92 | 454.11 | 31.03 | 29.94 | 28.85 | 0.04618 |
| 3 | 3 | 21 | 2 | 127.23 | 465.73 | 465.90 | 466.07 | 326.02 | 325.19 | 324.35 | 0.04217 |
| 3 | 3 | 21 | 2 | 127.20 | 466.41 | 466.58 | 466.74 | 322.69 | 321.87 | 321.04 | 0.04217 |
| 3 | 3 | 21 | 2 | 127.17 | 467.08 | 467.25 | 467.42 | 319.40 | 318.58 | 317.77 | 0.04217 |
| 3 | 3 | 21 | 2 | 127.14 | 467.76 | 467.93 | 468.10 | 316.14 | 315.33 | 314.53 | 0.04217 |
| 3 | 6 | 21 | 2 | 126.78 | 476.43 | 476.60 | 476.77 | 277.47 | 276.77 | 276.07 | 0.04217 |
| 3 | 6 | 21 | 2 | 126.75 | 477.11 | 477.28 | 477.45 | 274.68 | 273.99 | 273.29 | 0.04217 |
| 3 | 6 | 21 | 2 | 126.72 | 477.78 | 477.95 | 478.12 | 271.92 | 271.23 | 270.55 | 0.04217 |
| 3 | 6 | 21 | 2 | 126.69 | 478.46 | 478.63 | 478.80 | 269.19 | 268.51 | 267.84 | 0.04217 |
| 3 | 3 | 24 | 2 | 125.91 | 497.27 | 497.47 | 497.66 | 206.66 | 206.14 | 205.63 | 0.04819 |
| 3 | 3 | 24 | 2 | 125.88 | 498.05 | 498.24 | 498.44 | 204.60 | 204.10 | 203.59 | 0.04819 |
| 3 | 3 | 24 | 2 | 125.85 | 498.82 | 499.02 | 499.21 | 202.59 | 202.10 | 201.60 | 0.04819 |
| 3 | 3 | 24 | 2 | 125.81 | 499.60 | 499.79 | 499.98 | 200.63 | 200.14 | 199.66 | 0.04819 |
| 3 | 6 | 24 | 2 | 125.39 | 509.77 | 509.97 | 510.16 | 178.33 | 177.97 | 177.62 | 0.04819 |
| 3 | 6 | 24 | 2 | 125.36 | 510.55 | 510.74 | 510.94 | 176.91 | 176.56 | 176.21 | 0.04819 |
| 3 | 6 | 24 | 2 | 125.33 | 511.32 | 511.52 | 511.71 | 175.53 | 175.19 | 174.85 | 0.04819 |
| 3 | 6 | 24 | 2 | 125.30 | 512.10 | 512.29 | 512.48 | 174.18 | 173.85 | 173.52 | 0.04819 |
| 4 | 3 | 23 | 2 | 124.45 | 532.97 | 533.15 | 533.34 | 151.27 | 151.18 | 151.10 | 0.04656 |
| 4 | 3 | 23 | 2 | 124.42 | 533.71 | 533.90 | 534.08 | 150.93 | 150.85 | 150.77 | 0.04656 |
| 4 | 3 | 23 | 2 | 124.39 | 534.45 | 534.64 | 534.82 | 150.61 | 150.54 | 150.47 | 0.04656 |
| 4 | 3 | 23 | 2 | 124.36 | 535.19 | 535.38 | 535.57 | 150.33 | 150.26 | 150.19 | 0.04656 |
| 4 | 3 | 23 | 2 | 124.33 | 535.94 | 536.12 | 536.31 | 150.07 | 150.01 | 149.95 | 0.04656 |
| 4 | 6 | 23 | 2 | 123.88 | 547.03 | 547.22 | 547.40 | 149.97 | 150.03 | 150.09 | 0.04656 |
| 4 | 6 | 23 | 2 | 123.85 | 547.77 | 547.96 | 548.15 | 150.22 | 150.29 | 150.35 | 0.04656 |
| 4 | 6 | 23 | 2 | 123.82 | 548.52 | 548.70 | 548.89 | 150.50 | 150.57 | 150.65 | 0.04656 |
| 4 | 6 | 23 | 2 | 123.79 | 549.26 | 549.44 | 549.63 | 150.81 | 150.90 | 150.98 | 0.04656 |
| 4 | 6 | 23 | 2 | 123.76 | 550.00 | 550.19 | 550.37 | 151.16 | 151.25 | 151.34 | 0.04656 |
| 4 | 3 | 22 | 2 | 122.78 | 574.42 | 574.59 | 574.77 | 180.18 | 180.52 | 180.86 | 0.04453 |
| 4 | 3 | 22 | 2 | 122.76 | 575.13 | 575.30 | 575.48 | 181.55 | 181.89 | 182.24 | 0.04453 |
| 4 | 3 | 22 | 2 | 122.73 | 575.84 | 576.01 | 576.19 | 182.95 | 183.30 | 183.66 | 0.04453 |
| 4 | 3 | 22 | 2 | 122.70 | 576.54 | 576.72 | 576.90 | 184.38 | 184.74 | 185.11 | 0.04453 |
| 4 | 3 | 22 | 2 | 122.67 | 577.25 | 577.43 | 577.61 | 185.84 | 186.21 | 186.59 | 0.04453 |
| 4 | 3 | 22 | 2 | 122.64 | 577.96 | 578.14 | 578.32 | 187.34 | 187.72 | 188.10 | 0.04453 |
| 4 | 6 | 22 | 2 | 122.15 | 590.49 | 590.66 | 590.84 | 218.51 | 219.01 | 219.53 | 0.04453 |
| 4 | 6 | 22 | 2 | 122.12 | 591.20 | 591.37 | 591.55 | 220.55 | 221.07 | 221.59 | 0.04453 |
| 4 | 6 | 22 | 2 | 122.09 | 591.91 | 592.08 | 592.26 | 222.64 | 223.16 | 223.69 | 0.04453 |
| 4 | 6 | 22 | 2 | 122.07 | 592.62 | 592.79 | 592.97 | 224.74 | 225.27 | 225.80 | 0.04453 |
| 4 | 6 | 22 | 2 | 122.04 | 593.33 | 593.50 | 593.68 | 226.87 | 227.41 | 227.95 | 0.04453 |
| 4 | 6 | 22 | 2 | 122.01 | 594.04 | 594.21 | 594.39 | 229.03 | 229.58 | 230.13 | 0.04453 |
| 8 | 3 | 22 | 2 | 112.80 | 844.52 | 844.70 | 844.88 | 99.71 | 102.50 | 105.29 | 0.05914 |
| 8 | 3 | 22 | 2 | 112.77 | 845.23 | 845.41 | 845.59 | 110.87 | 113.66 | 116.46 | 0.05914 |
| 8 | 3 | 22 | 2 | 112.75 | 845.94 | 846.12 | 846.30 | 122.05 | 124.85 | 127.65 | 0.05914 |
| 8 | 6 | 22 | 2 | 112.48 | 853.73 | 853.91 | 854.09 | 246.47 | 249.33 | 252.20 | 0.05914 |
| 8 | 6 | 22 | 2 | 112.46 | 854.44 | 854.62 | 854.80 | 257.93 | 260.80 | 263.68 | 0.05914 |
| 8 | 6 | 22 | 2 | 112.43 | 855.15 | 855.33 | 855.51 | 269.42 | 272.30 | 275.18 | 0.05914 |
| 8 | 3 | 21 | 2 | 112.00 | 868.01 | 868.18 | 868.35 | 121.71 | 124.55 | 127.40 | 0.05645 |
| 8 | 3 | 21 | 2 | 111.97 | 868.69 | 868.86 | 869.03 | 133.10 | 135.95 | 138.80 | 0.05645 |
| 8 | 3 | 21 | 2 | 111.95 | 869.37 | 869.54 | 869.71 | 144.51 | 147.37 | 150.23 | 0.05645 |
| 8 | 6 | 21 | 2 | 111.70 | 876.67 | 876.83 | 877.00 | 268.75 | 271.66 | 274.58 | 0.05645 |
| 8 | 6 | 21 | 2 | 111.68 | 877.34 | 877.51 | 877.68 | 280.41 | 283.33 | 286.25 | 0.05645 |
| 8 | 6 | 21 | 2 | 111.66 | 878.02 | 878.19 | 878.36 | 292.09 | 295.01 | 297.93 | 0.05645 |
| 8 | 3 | 19 | 2 | 111.26 | 890.11 | 890.26 | 890.41 | 143.94 | 146.67 | 149.39 | 0.05108 |
| 8 | 3 | 19 | 2 | 111.24 | 890.72 | 890.87 | 891.03 | 154.86 | 157.59 | 160.32 | 0.05108 |
| 8 | 3 | 19 | 2 | 111.21 | 891.33 | 891.49 | 891.64 | 165.79 | 168.52 | 171.26 | 0.05108 |
| 8 | 6 | 19 | 2 | 110.99 | 898.14 | 898.29 | 898.45 | 288.34 | 291.12 | 293.90 | 0.05108 |
| 8 | 6 | 19 | 2 | 110.97 | 898.75 | 898.91 | 899.06 | 299.48 | 302.26 | 305.05 | 0.05108 |
| 8 | 6 | 19 | 2 | 110.95 | 899.37 | 899.52 | 899.67 | 310.63 | 313.42 | 316.21 | 0.05108 |
| 9 | 3 | 19 | 2 | 107.23 | 965.21 | 965.37 | 965.52 | 272.84 | 276.74 | 280.66 | 0.07336 |
| 9 | 3 | 19 | 2 | 107.19 | 965.83 | 965.98 | 966.13 | 288.50 | 292.43 | 296.37 | 0.07336 |
| 9 | 6 | 19 | 2 | 106.80 | 971.29 | 971.45 | 971.60 | 72.05 | 76.16 | 80.27 | 0.07336 |
| 9 | 6 | 19 | 2 | 106.76 | 971.91 | 972.06 | 972.21 | 88.52 | 92.64 | 96.77 | 0.07336 |
| 9 | 6 | 15 | 2 | 106.22 | 979.34 | 979.46 | 979.58 | 292.94 | 296.34 | 299.75 | 0.05792 |
| 9 | 6 | 21 | 2 | 105.50 | 989.17 | 989.34 | 989.51 | 219.71 | 224.86 | 230.00 | 0.08108 |
| 9 | 3 | 22 | 2 | 104.92 | 997.07 | 997.24 | 997.42 | 105.79 | 111.45 | 117.11 | 0.08494 |
| 9 | 6 | 21 | 2 | 104.58 | 1001.57 | 1001.74 | 1001.91 | 251.01 | 256.52 | 262.04 | 0.08108 |
| 9 | 3 | 15 | 2 | 104.11 | 1008.09 | 1008.21 | 1008.33 | 105.65 | 109.69 | 113.72 | 0.05792 |
| 9 | 3 | 15 | 2 | 104.07 | 1008.57 | 1008.70 | 1008.82 | 121.79 | 125.83 | 129.87 | 0.05792 |
| 9 | 6 | 15 | 2 | 103.75 | 1012.99 | 1013.11 | 1013.23 | 270.12 | 274.21 | 278.31 | 0.05792 |
| 9 | 6 | 15 | 2 | 103.71 | 1013.47 | 1013.59 | 1013.71 | 286.51 | 290.61 | 294.72 | 0.05792 |
| 9 | 3 | 22 | 2 | 103.17 | 1020.75 | 1020.93 | 1021.11 | 175.89 | 182.03 | 188.17 | 0.08494 |
| 9 | 6 | 22 | 2 | 102.84 | 1025.34 | 1025.52 | 1025.69 | 335.48 | 341.69 | 347.90 | 0.08494 |
| 10 | 3 | 19 | 2 | 100.75 | 1054.36 | 1054.51 | 1054.66 | 305.97 | 311.72 | 317.47 | 0.08190 |
| 10 | 6 | 19 | 2 | 100.47 | 1058.30 | 1058.45 | 1058.61 | 94.55 | 100.36 | 106.17 | 0.08190 |
| 10 | 3 | 18 | 2 | 100.02 | 1064.62 | 1064.77 | 1064.91 | 336.19 | 341.79 | 347.39 | 0.07759 |
| 10 | 6 | 18 | 2 | 99.75 | 1068.44 | 1068.58 | 1068.73 | 124.06 | 129.72 | 135.38 | 0.07759 |
| 10 | 3 | 17 | 2 | 99.33 | 1074.55 | 1074.69 | 1074.82 | 4.59 | 10.04 | 15.49 | 0.07328 |
| 10 | 6 | 17 | 2 | 99.07 | 1078.23 | 1078.37 | 1078.51 | 151.74 | 157.24 | 162.76 | 0.07328 |
| 10 | 3 | 15 | 2 | 98.56 | 1085.59 | 1085.71 | 1085.83 | 90.76 | 95.74 | 100.72 | 0.06466 |
| 10 | 6 | 17 | 2 | 98.33 | 1089.02 | 1089.15 | 1089.29 | 232.63 | 238.34 | 244.06 | 0.07328 |
| 10 | 3 | 16 | 2 | 97.93 | 1094.73 | 1094.86 | 1094.99 | 113.10 | 118.58 | 124.05 | 0.06897 |
| 10 | 6 | 16 | 2 | 97.70 | 1098.14 | 1098.27 | 1098.40 | 258.76 | 264.30 | 269.84 | 0.06897 |
| 10 | 3 | 15 | 2 | 97.32 | 1103.72 | 1103.85 | 1103.97 | 140.65 | 145.94 | 151.24 | 0.06466 |
| 10 | 6 | 15 | 2 | 97.10 | 1107.04 | 1107.16 | 1107.28 | 286.34 | 291.69 | 297.04 | 0.06466 |
| 10 | 3 | 15 | 2 | 96.74 | 1112.45 | 1112.57 | 1112.69 | 167.46 | 172.90 | 178.33 | 0.06466 |
| 10 | 6 | 15 | 2 | 96.52 | 1115.71 | 1115.83 | 1115.95 | 314.51 | 319.99 | 325.48 | 0.06466 |
| 11 | 6 | 14 | 2 | 94.89 | 1140.51 | 1140.62 | 1140.73 | 38.70 | 44.18 | 49.66 | 0.06829 |
| 11 | 3 | 14 | 2 | 94.57 | 1145.48 | 1145.59 | 1145.71 | 281.74 | 287.28 | 292.84 | 0.06829 |
| 11 | 6 | 14 | 2 | 94.38 | 1148.44 | 1148.55 | 1148.67 | 67.75 | 73.34 | 78.93 | 0.06829 |
| 11 | 3 | 13 | 2 | 94.07 | 1153.32 | 1153.42 | 1153.53 | 310.92 | 316.19 | 321.45 | 0.06341 |
| 11 | 6 | 13 | 2 | 93.88 | 1156.25 | 1156.35 | 1156.46 | 98.31 | 103.62 | 108.92 | 0.06341 |
| 11 | 3 | 13 | 2 | 93.42 | 1163.56 | 1163.67 | 1163.77 | 111.11 | 116.50 | 121.89 | 0.06341 |
| 11 | 6 | 13 | 2 | 93.24 | 1166.41 | 1166.51 | 1166.62 | 257.93 | 263.36 | 268.79 | 0.06341 |
| 11 | 3 | 13 | 2 | 92.95 | 1171.02 | 1171.13 | 1171.23 | 137.87 | 143.35 | 148.84 | 0.06341 |
| 11 | 6 | 13 | 2 | 92.78 | 1173.79 | 1173.90 | 1174.00 | 283.41 | 288.93 | 294.45 | 0.06341 |
| 11 | 3 | 12 | 2 | 92.49 | 1178.39 | 1178.49 | 1178.58 | 166.67 | 171.82 | 176.97 | 0.05854 |
| 11 | 6 | 13 | 2 | 92.32 | 1181.12 | 1181.22 | 1181.33 | 312.21 | 317.84 | 323.46 | 0.06341 |
| 11 | 3 | 12 | 2 | 92.05 | 1185.57 | 1185.66 | 1185.76 | 192.01 | 197.26 | 202.49 | 0.05854 |
| 11 | 6 | 12 | 2 | 91.88 | 1188.25 | 1188.34 | 1188.44 | 337.59 | 342.87 | 348.14 | 0.05854 |
| 11 | 3 | 12 | 2 | 91.62 | 1192.60 | 1192.70 | 1192.79 | 216.40 | 221.74 | 227.07 | 0.05854 |
| 11 | 6 | 12 | 2 | 91.46 | 1195.21 | 1195.31 | 1195.41 | 0.96 | 6.33 | 11.70 | 0.05854 |
| 11 | 3 | 12 | 2 | 91.20 | 1199.49 | 1199.58 | 1199.68 | 239.35 | 244.78 | 250.20 | 0.05854 |
| 12 | 6 | 10 | 2 | 89.34 | 1230.73 | 1230.81 | 1230.89 | 238.58 | 243.31 | 248.02 | 0.03012 |
| 12 | 3 | 10 | 2 | 89.12 | 1234.48 | 1234.56 | 1234.64 | 98.17 | 102.90 | 107.62 | 0.03012 |
| 12 | 6 | 10 | 2 | 88.99 | 1236.75 | 1236.83 | 1236.91 | 230.90 | 235.63 | 240.35 | 0.03012 |
| 12 | 3 | 21 | 2 | 88.76 | 1240.57 | 1240.74 | 1240.91 | 94.67 | 104.59 | 114.49 | 0.06325 |
| 12 | 6 | 21 | 2 | 88.63 | 1242.81 | 1242.98 | 1243.15 | 225.77 | 235.68 | 245.57 | 0.06325 |
| 12 | 3 | 10 | 2 | 88.41 | 1246.70 | 1246.79 | 1246.87 | 93.22 | 97.93 | 102.63 | 0.03012 |
| 12 | 6 | 10 | 2 | 88.28 | 1249.02 | 1249.10 | 1249.18 | 228.09 | 232.78 | 237.47 | 0.03012 |
| 12 | 3 | 10 | 2 | 88.06 | 1252.81 | 1252.89 | 1252.97 | 88.40 | 93.08 | 97.76 | 0.03012 |
| 12 | 6 | 10 | 2 | 87.93 | 1255.11 | 1255.19 | 1255.27 | 221.64 | 226.31 | 230.98 | 0.03012 |
| 12 | 3 | 21 | 2 | 87.71 | 1258.94 | 1259.11 | 1259.28 | 83.13 | 92.91 | 102.68 | 0.06325 |
| 12 | 6 | 21 | 2 | 87.58 | 1261.19 | 1261.36 | 1261.53 | 212.89 | 222.65 | 232.41 | 0.06325 |
| 12 | 3 | 21 | 2 | 87.36 | 1264.94 | 1265.11 | 1265.28 | 68.67 | 78.40 | 88.13 | 0.06325 |
| 12 | 6 | 21 | 2 | 87.24 | 1267.19 | 1267.36 | 1267.53 | 197.84 | 207.55 | 217.28 | 0.06325 |
| 12 | 3 | 21 | 2 | 87.03 | 1270.87 | 1271.04 | 1271.21 | 49.05 | 58.75 | 68.44 | 0.06325 |
| 12 | 6 | 21 | 2 | 86.91 | 1273.10 | 1273.27 | 1273.44 | 176.38 | 186.06 | 195.75 | 0.06325 |
| 12 | 3 | 21 | 2 | 86.70 | 1276.74 | 1276.91 | 1277.08 | 24.62 | 34.29 | 43.95 | 0.06325 |
| 12 | 6 | 21 | 2 | 86.58 | 1278.94 | 1279.11 | 1279.28 | 150.19 | 159.86 | 169.51 | 0.06325 |
| 12 | 3 | 21 | 2 | 86.38 | 1282.60 | 1282.77 | 1282.94 | 358.34 | 7.98 | 17.62 | 0.06325 |
| 12 | 6 | 21 | 2 | 86.26 | 1284.78 | 1284.95 | 1285.12 | 122.67 | 132.30 | 141.92 | 0.06325 |
| 12 | 3 | 10 | 2 | 85.95 | 1290.53 | 1290.61 | 1290.69 | 89.15 | 93.73 | 98.29 | 0.03012 |
| 13 | 3 | 20 | 2 | 84.38 | 1319.36 | 1319.52 | 1319.68 | 277.21 | 286.29 | 295.38 | 0.04651 |
| 13 | 6 | 20 | 2 | 84.27 | 1321.50 | 1321.66 | 1321.82 | 37.57 | 46.64 | 55.72 | 0.04651 |
| 13 | 6 | 20 | 2 | 84.10 | 1324.68 | 1324.84 | 1325.00 | 216.37 | 225.44 | 234.51 | 0.04651 |
| 13 | 3 | 22 | 2 | 83.87 | 1329.05 | 1329.22 | 1329.40 | 101.91 | 111.85 | 121.81 | 0.05116 |
| 13 | 6 | 22 | 2 | 83.75 | 1331.39 | 1331.57 | 1331.75 | 233.46 | 243.40 | 253.33 | 0.05116 |
| 13 | 3 | 20 | 2 | 83.54 | 1335.36 | 1335.52 | 1335.68 | 95.27 | 104.27 | 113.26 | 0.04651 |
| 13 | 6 | 20 | 2 | 83.43 | 1337.50 | 1337.66 | 1337.82 | 214.40 | 223.37 | 232.34 | 0.04651 |
| 13 | 3 | 20 | 2 | 83.24 | 1341.07 | 1341.23 | 1341.39 | 52.83 | 61.77 | 70.70 | 0.04651 |
| 13 | 6 | 20 | 2 | 83.13 | 1343.22 | 1343.38 | 1343.54 | 171.54 | 180.45 | 189.36 | 0.04651 |
| 13 | 3 | 20 | 2 | 82.95 | 1346.79 | 1346.95 | 1347.11 | 8.37 | 17.22 | 26.07 | 0.04651 |
| 13 | 6 | 20 | 2 | 82.84 | 1348.93 | 1349.09 | 1349.26 | 125.95 | 134.76 | 143.59 | 0.04651 |
| 13 | 3 | 20 | 2 | 82.66 | 1352.50 | 1352.66 | 1352.82 | 320.29 | 329.05 | 337.80 | 0.04651 |
| 13 | 6 | 20 | 2 | 82.55 | 1354.64 | 1354.80 | 1354.96 | 76.52 | 85.23 | 93.94 | 0.04651 |
| 13 | 3 | 20 | 2 | 82.36 | 1358.25 | 1358.42 | 1358.58 | 270.97 | 279.60 | 288.24 | 0.04651 |
| 13 | 6 | 21 | 2 | 82.25 | 1360.42 | 1360.58 | 1360.75 | 26.44 | 35.47 | 44.49 | 0.04884 |
| 13 | 3 | 21 | 2 | 82.07 | 1364.02 | 1364.19 | 1364.36 | 217.54 | 226.47 | 235.40 | 0.04884 |
| 13 | 6 | 20 | 2 | 81.96 | 1366.19 | 1366.35 | 1366.51 | 331.59 | 340.04 | 348.49 | 0.04651 |
| 13 | 3 | 21 | 2 | 81.78 | 1369.81 | 1369.98 | 1370.15 | 160.44 | 169.22 | 178.00 | 0.04884 |
| 13 | 6 | 21 | 2 | 81.67 | 1372.00 | 1372.17 | 1372.33 | 273.38 | 282.10 | 290.82 | 0.04884 |
| 13 | 3 | 21 | 2 | 81.49 | 1375.63 | 1375.80 | 1375.97 | 99.84 | 108.48 | 117.11 | 0.04884 |
| 13 | 6 | 21 | 2 | 81.38 | 1377.82 | 1377.99 | 1378.16 | 210.91 | 219.49 | 228.07 | 0.04884 |
| 14 | 3 | 12 | 2 | 79.70 | 1412.14 | 1412.24 | 1412.33 | 88.85 | 93.47 | 98.10 | 0.02878 |
| 14 | 6 | 12 | 2 | 79.57 | 1414.89 | 1414.99 | 1415.08 | 220.19 | 224.80 | 229.42 | 0.02878 |
| 14 | 3 | 11 | 2 | 79.39 | 1418.57 | 1418.65 | 1418.74 | 35.56 | 39.79 | 44.02 | 0.02638 |
| 14 | 6 | 22 | 2 | 79.29 | 1420.61 | 1420.78 | 1420.96 | 132.75 | 141.19 | 149.64 | 0.05276 |
| 14 | 3 | 22 | 2 | 79.11 | 1424.43 | 1424.61 | 1424.78 | 314.62 | 323.06 | 331.49 | 0.05276 |
| 14 | 6 | 22 | 2 | 79.00 | 1426.73 | 1426.90 | 1427.08 | 63.86 | 72.28 | 80.71 | 0.05276 |
| 14 | 3 | 22 | 2 | 78.81 | 1430.62 | 1430.80 | 1430.98 | 248.75 | 257.16 | 265.57 | 0.05276 |
| 14 | 6 | 22 | 2 | 78.71 | 1432.94 | 1433.12 | 1433.30 | 358.82 | 7.22 | 15.62 | 0.05276 |
| 14 | 3 | 22 | 2 | 78.52 | 1436.85 | 1437.03 | 1437.21 | 183.77 | 192.15 | 200.53 | 0.05276 |
| 14 | 6 | 22 | 2 | 78.41 | 1439.20 | 1439.38 | 1439.56 | 294.51 | 302.87 | 311.23 | 0.05276 |
| 14 | 3 | 22 | 2 | 78.23 | 1443.12 | 1443.30 | 1443.48 | 118.83 | 127.15 | 135.48 | 0.05276 |
| 14 | 6 | 23 | 2 | 78.12 | 1445.48 | 1445.67 | 1445.85 | 229.56 | 238.24 | 246.91 | 0.05516 |
| 14 | 3 | 23 | 2 | 77.94 | 1449.43 | 1449.62 | 1449.81 | 54.00 | 62.63 | 71.27 | 0.05516 |
| 14 | 6 | 22 | 2 | 77.83 | 1451.81 | 1451.98 | 1452.16 | 164.15 | 172.37 | 180.59 | 0.05276 |
| 14 | 3 | 23 | 2 | 77.64 | 1455.81 | 1455.99 | 1456.18 | 349.00 | 357.55 | 6.09 | 0.05516 |
| 14 | 6 | 23 | 2 | 77.53 | 1458.20 | 1458.39 | 1458.57 | 99.10 | 107.61 | 116.11 | 0.05516 |
| 14 | 3 | 23 | 2 | 77.35 | 1462.26 | 1462.45 | 1462.64 | 284.74 | 293.17 | 301.61 | 0.05516 |
| 14 | 6 | 23 | 2 | 77.24 | 1464.68 | 1464.87 | 1465.05 | 34.58 | 42.99 | 51.38 | 0.05516 |
| 14 | 3 | 23 | 2 | 77.06 | 1468.76 | 1468.95 | 1469.14 | 218.58 | 226.90 | 235.23 | 0.05516 |
| 14 | 6 | 23 | 2 | 76.94 | 1471.21 | 1471.39 | 1471.58 | 328.00 | 336.29 | 344.57 | 0.05516 |
| 15 | 3 | 16 | 2 | 74.22 | 1534.04 | 1534.17 | 1534.30 | 103.17 | 108.39 | 113.62 | 0.07583 |
| 15 | 6 | 16 | 2 | 74.08 | 1537.39 | 1537.52 | 1537.65 | 238.89 | 244.11 | 249.32 | 0.07583 |
| 15 | 3 | 10 | 2 | 73.89 | 1541.85 | 1541.93 | 1542.01 | 59.32 | 62.59 | 65.85 | 0.04739 |
| 15 | 6 | 13 | 2 | 73.80 | 1544.08 | 1544.18 | 1544.29 | 149.15 | 153.38 | 157.61 | 0.06161 |
| 15 | 3 | 13 | 2 | 73.61 | 1548.71 | 1548.81 | 1548.92 | 334.75 | 338.92 | 343.09 | 0.06161 |
| 15 | 6 | 13 | 2 | 73.49 | 1551.49 | 1551.59 | 1551.70 | 84.90 | 89.03 | 93.16 | 0.06161 |
| 15 | 3 | 13 | 2 | 73.30 | 1556.12 | 1556.22 | 1556.33 | 266.17 | 270.24 | 274.31 | 0.06161 |
| 15 | 6 | 13 | 2 | 73.19 | 1558.90 | 1559.01 | 1559.11 | 13.38 | 17.39 | 21.40 | 0.06161 |
| 15 | 3 | 13 | 2 | 73.00 | 1563.60 | 1563.71 | 1563.81 | 191.33 | 195.25 | 199.17 | 0.06161 |
| 15 | 6 | 13 | 2 | 72.88 | 1566.42 | 1566.52 | 1566.63 | 295.94 | 299.82 | 303.70 | 0.06161 |
| 15 | 3 | 13 | 2 | 72.69 | 1571.17 | 1571.28 | 1571.38 | 110.49 | 114.32 | 118.16 | 0.06161 |
| 15 | 6 | 13 | 2 | 72.58 | 1574.02 | 1574.13 | 1574.23 | 214.62 | 218.44 | 222.26 | 0.06161 |
| 15 | 3 | 13 | 2 | 72.38 | 1578.85 | 1578.96 | 1579.06 | 28.75 | 32.48 | 36.22 | 0.06161 |
| 15 | 6 | 13 | 2 | 72.27 | 1581.74 | 1581.84 | 1581.95 | 131.21 | 134.93 | 138.64 | 0.06161 |
| 15 | 3 | 13 | 2 | 72.07 | 1586.63 | 1586.74 | 1586.84 | 304.29 | 307.98 | 311.66 | 0.06161 |
| 15 | 6 | 13 | 2 | 71.96 | 1589.56 | 1589.66 | 1589.77 | 46.64 | 50.29 | 53.94 | 0.06161 |
| 16 | 6 | 15 | 2 | 70.61 | 1624.35 | 1624.47 | 1624.59 | 190.52 | 194.78 | 199.03 | 0.07109 |
| 16 | 3 | 15 | 2 | 70.41 | 1629.61 | 1629.73 | 1629.85 | 15.12 | 19.35 | 23.59 | 0.07109 |
| 16 | 6 | 15 | 2 | 70.29 | 1632.78 | 1632.90 | 1633.02 | 126.04 | 130.28 | 134.52 | 0.07109 |
| 16 | 3 | 15 | 2 | 70.10 | 1638.10 | 1638.22 | 1638.34 | 311.86 | 316.03 | 320.21 | 0.07109 |
| 16 | 6 | 15 | 2 | 69.98 | 1641.31 | 1641.43 | 1641.55 | 62.12 | 66.27 | 70.41 | 0.07109 |
| 16 | 3 | 15 | 2 | 69.78 | 1646.66 | 1646.78 | 1646.91 | 246.52 | 250.71 | 254.89 | 0.07109 |
| 16 | 6 | 15 | 2 | 69.66 | 1649.88 | 1650.00 | 1650.12 | 358.16 | 2.38 | 6.60 | 0.07109 |
| 16 | 3 | 15 | 2 | 69.47 | 1655.35 | 1655.47 | 1655.59 | 188.11 | 192.30 | 196.49 | 0.07109 |
| 16 | 6 | 15 | 2 | 69.35 | 1658.61 | 1658.73 | 1658.85 | 299.15 | 303.17 | 307.19 | 0.07109 |
| 16 | 3 | 15 | 2 | 69.15 | 1664.04 | 1664.16 | 1664.28 | 115.28 | 119.19 | 123.10 | 0.07109 |
| 16 | 6 | 15 | 2 | 69.04 | 1667.31 | 1667.43 | 1667.55 | 221.10 | 225.02 | 228.94 | 0.07109 |
| 16 | 3 | 15 | 2 | 68.84 | 1672.85 | 1672.97 | 1673.09 | 40.61 | 44.53 | 48.45 | 0.07109 |
| 16 | 6 | 15 | 2 | 68.72 | 1676.15 | 1676.27 | 1676.40 | 147.95 | 151.88 | 155.82 | 0.07109 |
| 16 | 3 | 16 | 2 | 68.52 | 1681.83 | 1681.96 | 1682.09 | 332.88 | 337.09 | 341.30 | 0.07583 |
| 17 | 6 | 17 | 2 | 67.44 | 1713.08 | 1713.22 | 1713.35 | 275.93 | 280.37 | 284.82 | 0.07173 |
| 17 | 3 | 16 | 2 | 67.25 | 1718.96 | 1719.09 | 1719.22 | 105.49 | 109.61 | 113.72 | 0.06751 |
| 17 | 6 | 16 | 2 | 67.13 | 1722.48 | 1722.60 | 1722.73 | 217.07 | 221.13 | 225.18 | 0.06751 |
| 17 | 3 | 17 | 2 | 66.93 | 1728.31 | 1728.45 | 1728.59 | 36.30 | 40.44 | 44.57 | 0.07173 |
| 17 | 6 | 17 | 2 | 66.82 | 1731.83 | 1731.97 | 1732.10 | 142.12 | 146.26 | 150.39 | 0.07173 |
| 17 | 3 | 17 | 2 | 66.62 | 1737.81 | 1737.95 | 1738.09 | 323.45 | 327.67 | 331.91 | 0.07173 |
| 17 | 6 | 17 | 2 | 66.51 | 1741.39 | 1741.52 | 1741.66 | 74.31 | 78.57 | 82.84 | 0.07173 |
| 17 | 3 | 17 | 2 | 66.31 | 1747.34 | 1747.48 | 1747.62 | 258.60 | 262.79 | 266.98 | 0.07173 |
| 17 | 6 | 17 | 2 | 66.20 | 1750.91 | 1751.05 | 1751.18 | 5.56 | 9.62 | 13.67 | 0.07173 |
| 17 | 3 | 17 | 2 | 66.01 | 1757.05 | 1757.18 | 1757.32 | 185.68 | 189.68 | 193.68 | 0.07173 |
| 17 | 6 | 17 | 2 | 65.89 | 1760.67 | 1760.81 | 1760.95 | 291.28 | 295.25 | 299.23 | 0.07173 |
| 17 | 3 | 17 | 2 | 65.70 | 1766.88 | 1767.01 | 1767.15 | 110.34 | 114.26 | 118.19 | 0.07173 |
| 17 | 6 | 17 | 2 | 65.58 | 1770.56 | 1770.70 | 1770.83 | 215.03 | 218.91 | 222.79 | 0.07173 |
| 17 | 3 | 18 | 2 | 65.39 | 1776.80 | 1776.95 | 1777.09 | 32.02 | 36.15 | 40.28 | 0.07595 |
| 18 | 6 | 18 | 2 | 64.61 | 1802.68 | 1802.83 | 1802.97 | 61.22 | 65.32 | 69.44 | 0.08036 |
| 18 | 3 | 18 | 2 | 64.41 | 1809.15 | 1809.29 | 1809.44 | 245.83 | 249.97 | 254.11 | 0.08036 |
| 18 | 6 | 18 | 2 | 64.30 | 1813.03 | 1813.17 | 1813.32 | 355.40 | 359.47 | 3.53 | 0.08036 |
| 18 | 3 | 18 | 2 | 64.11 | 1819.50 | 1819.64 | 1819.79 | 176.76 | 180.85 | 184.94 | 0.08036 |
| 18 | 6 | 18 | 2 | 64.00 | 1823.37 | 1823.52 | 1823.66 | 286.41 | 290.53 | 294.65 | 0.08036 |
| 18 | 3 | 19 | 2 | 63.81 | 1830.03 | 1830.18 | 1830.33 | 116.27 | 120.63 | 124.98 | 0.08482 |
| 18 | 6 | 19 | 2 | 63.70 | 1833.97 | 1834.12 | 1834.28 | 227.99 | 232.32 | 236.65 | 0.08482 |
| 18 | 3 | 19 | 2 | 63.50 | 1840.72 | 1840.87 | 1841.03 | 58.39 | 62.69 | 66.99 | 0.08482 |
| 18 | 6 | 19 | 2 | 63.39 | 1844.74 | 1844.89 | 1845.04 | 170.47 | 174.72 | 178.97 | 0.08482 |
| 18 | 3 | 19 | 2 | 63.20 | 1851.61 | 1851.77 | 1851.92 | 358.94 | 3.07 | 7.19 | 0.08482 |
| 18 | 6 | 19 | 2 | 63.09 | 1855.70 | 1855.85 | 1856.01 | 109.59 | 113.78 | 117.96 | 0.08482 |
| 18 | 3 | 20 | 2 | 62.89 | 1862.74 | 1862.90 | 1863.06 | 303.74 | 308.21 | 312.68 | 0.08929 |
| 22 | 6 | 26 | 2 | 59.21 | 2008.39 | 2008.60 | 2008.80 | 239.13 | 244.59 | 250.04 | 0.11207 |
| 22 | 3 | 22 | 2 | 59.01 | 2016.97 | 2017.15 | 2017.33 | 95.08 | 99.48 | 103.89 | 0.09483 |
| 22 | 6 | 23 | 2 | 58.90 | 2021.64 | 2021.83 | 2022.01 | 210.60 | 215.18 | 219.75 | 0.09914 |
| 22 | 3 | 23 | 2 | 58.71 | 2029.46 | 2029.64 | 2029.83 | 44.90 | 49.54 | 54.19 | 0.09914 |
| 22 | 6 | 23 | 2 | 58.61 | 2034.14 | 2034.33 | 2034.51 | 162.71 | 167.40 | 172.08 | 0.09914 |
| 22 | 3 | 23 | 2 | 58.43 | 2041.96 | 2042.14 | 2042.33 | 2.10 | 6.89 | 11.67 | 0.09914 |
| 22 | 6 | 23 | 2 | 58.32 | 2046.64 | 2046.83 | 2047.01 | 121.67 | 126.30 | 130.93 | 0.09914 |
| 22 | 3 | 23 | 2 | 58.14 | 2054.46 | 2054.64 | 2054.83 | 316.35 | 321.08 | 325.81 | 0.09914 |
| 22 | 6 | 23 | 2 | 58.03 | 2059.14 | 2059.33 | 2059.51 | 76.42 | 81.17 | 85.91 | 0.09914 |
| 22 | 3 | 23 | 2 | 57.86 | 2066.99 | 2067.18 | 2067.37 | 276.00 | 280.67 | 285.34 | 0.09914 |
| 23 | 6 | 23 | 2 | 57.18 | 2097.31 | 2097.49 | 2097.68 | 293.80 | 298.27 | 302.72 | 0.09871 |
| 23 | 3 | 23 | 2 | 57.01 | 2105.28 | 2105.47 | 2105.65 | 122.62 | 126.96 | 131.29 | 0.09871 |
| 23 | 6 | 23 | 2 | 56.90 | 2110.07 | 2110.26 | 2110.44 | 233.82 | 238.13 | 242.43 | 0.09871 |
| 23 | 3 | 23 | 2 | 56.72 | 2118.73 | 2118.92 | 2119.10 | 74.78 | 79.11 | 83.45 | 0.09871 |
| 23 | 6 | 23 | 2 | 56.61 | 2123.52 | 2123.71 | 2123.89 | 187.58 | 191.97 | 196.37 | 0.09871 |
| 23 | 3 | 23 | 2 | 56.44 | 2131.51 | 2131.69 | 2131.88 | 14.80 | 19.11 | 23.42 | 0.09871 |
| 23 | 6 | 23 | 2 | 56.34 | 2136.30 | 2136.48 | 2136.67 | 126.23 | 130.56 | 134.88 | 0.09871 |
| 23 | 3 | 24 | 2 | 56.10 | 2147.45 | 2147.64 | 2147.84 | 24.89 | 29.33 | 33.77 | 0.10300 |
| 23 | 6 | 24 | 2 | 56.00 | 2152.33 | 2152.53 | 2152.72 | 134.67 | 138.93 | 143.17 | 0.10300 |
| 23 | 3 | 24 | 2 | 55.83 | 2160.49 | 2160.68 | 2160.87 | 311.22 | 315.38 | 319.55 | 0.10300 |
| 24 | 6 | 25 | 2 | 54.35 | 2232.76 | 2232.97 | 2233.17 | 354.73 | 358.55 | 2.37 | 0.08532 |
| 24 | 3 | 26 | 2 | 54.18 | 2241.60 | 2241.81 | 2242.02 | 162.48 | 166.46 | 170.45 | 0.08874 |
| 24 | 6 | 26 | 2 | 54.07 | 2246.96 | 2247.17 | 2247.38 | 263.81 | 267.75 | 271.68 | 0.08874 |
| 24 | 3 | 26 | 2 | 53.90 | 2255.88 | 2256.09 | 2256.30 | 69.21 | 73.11 | 77.00 | 0.08874 |
| 24 | 6 | 26 | 2 | 53.79 | 2261.25 | 2261.46 | 2261.66 | 169.51 | 173.45 | 177.40 | 0.08874 |
| 24 | 3 | 27 | 2 | 53.61 | 2270.51 | 2270.73 | 2270.95 | 343.19 | 347.24 | 351.31 | 0.09215 |
| 24 | 6 | 27 | 2 | 53.50 | 2275.99 | 2276.21 | 2276.43 | 85.94 | 90.05 | 94.19 | 0.09215 |
| 24 | 3 | 27 | 2 | 53.32 | 2285.49 | 2285.70 | 2285.92 | 267.51 | 271.59 | 275.67 | 0.09215 |
| 24 | 6 | 27 | 2 | 53.21 | 2291.11 | 2291.33 | 2291.55 | 11.70 | 15.69 | 19.67 | 0.09215 |
| 24 | 6 | 28 | 2 | 53.04 | 2300.45 | 2300.68 | 2300.90 | 181.66 | 185.73 | 189.81 | 0.09556 |
| 24 | 3 | 28 | 2 | 52.85 | 2310.07 | 2310.30 | 2310.52 | 350.92 | 354.79 | 358.64 | 0.09556 |
| 25 | 6 | 29 | 2 | 52.65 | 2320.91 | 2321.14 | 2321.37 | 175.07 | 179.01 | 182.96 | 0.12033 |
| 25 | 3 | 29 | 2 | 52.46 | 2330.77 | 2331.00 | 2331.24 | 345.61 | 349.76 | 353.92 | 0.12033 |
| 25 | 6 | 29 | 2 | 52.35 | 2336.70 | 2336.93 | 2337.16 | 90.75 | 94.83 | 98.91 | 0.12033 |
| 25 | 3 | 30 | 2 | 52.16 | 2346.90 | 2347.14 | 2347.38 | 266.61 | 270.71 | 274.81 | 0.12448 |
| 25 | 6 | 30 | 2 | 52.05 | 2352.98 | 2353.22 | 2353.46 | 8.79 | 12.82 | 16.84 | 0.12448 |
| 25 | 3 | 31 | 2 | 51.86 | 2363.36 | 2363.61 | 2363.86 | 179.00 | 183.04 | 187.08 | 0.12863 |
| 25 | 6 | 31 | 2 | 51.74 | 2369.61 | 2369.86 | 2370.11 | 279.01 | 282.97 | 286.94 | 0.12863 |
| 25 | 3 | 32 | 2 | 51.55 | 2380.26 | 2380.51 | 2380.77 | 85.47 | 89.45 | 93.44 | 0.13278 |
| 26 | 3 | 31 | 2 | 50.93 | 2414.56 | 2414.81 | 2415.06 | 263.08 | 267.07 | 271.07 | 0.11969 |
| 26 | 6 | 32 | 2 | 50.82 | 2420.97 | 2421.23 | 2421.49 | 5.52 | 9.65 | 13.77 | 0.12355 |
| 26 | 3 | 32 | 2 | 50.62 | 2432.17 | 2432.42 | 2432.68 | 185.06 | 189.21 | 193.35 | 0.12355 |
| 26 | 6 | 33 | 2 | 50.51 | 2438.76 | 2439.03 | 2439.30 | 290.82 | 295.09 | 299.34 | 0.12741 |
| 26 | 3 | 33 | 2 | 50.31 | 2449.79 | 2450.06 | 2450.33 | 105.93 | 110.06 | 114.20 | 0.12741 |
| 26 | 6 | 32 | 2 | 50.20 | 2456.43 | 2456.69 | 2456.95 | 208.17 | 212.10 | 216.02 | 0.12355 |
| 26 | 3 | 33 | 2 | 50.01 | 2467.44 | 2467.71 | 2467.97 | 11.47 | 15.34 | 19.20 | 0.12741 |
| 26 | 6 | 33 | 2 | 49.89 | 2474.06 | 2474.33 | 2474.59 | 108.74 | 112.71 | 116.67 | 0.12741 |
| 30 | 6 | 64 | 2 | 47.54 | 2615.97 | 2616.49 | 2617.01 | 180.96 | 188.92 | 196.88 | 0.15202 |
| 30 | 3 | 65 | 2 | 47.37 | 2626.68 | 2627.21 | 2627.73 | 348.93 | 357.30 | 5.68 | 0.15439 |
| 30 | 6 | 65 | 2 | 47.26 | 2633.11 | 2633.63 | 2634.16 | 91.22 | 99.47 | 107.71 | 0.15439 |
| 30 | 3 | 65 | 2 | 47.10 | 2643.83 | 2644.35 | 2644.88 | 258.83 | 266.92 | 275.02 | 0.15439 |
| 30 | 6 | 65 | 2 | 47.00 | 2650.25 | 2650.78 | 2651.30 | 357.57 | 5.58 | 13.57 | 0.15439 |
| 30 | 3 | 32 | 2 | 46.83 | 2661.03 | 2661.29 | 2661.55 | 161.00 | 164.89 | 168.78 | 0.07601 |
| 30 | 6 | 33 | 2 | 46.73 | 2667.63 | 2667.90 | 2668.16 | 259.74 | 263.69 | 267.64 | 0.07838 |
| 30 | 3 | 32 | 2 | 46.56 | 2678.57 | 2678.82 | 2679.08 | 59.99 | 63.75 | 67.51 | 0.07601 |
| 31 | 6 | 69 | 2 | 46.06 | 2710.83 | 2711.38 | 2711.94 | 197.91 | 207.39 | 216.87 | 0.18449 |
| 31 | 3 | 67 | 2 | 45.89 | 2722.00 | 2722.54 | 2723.08 | 23.02 | 31.79 | 40.54 | 0.17914 |
| 31 | 6 | 67 | 2 | 45.79 | 2728.62 | 2729.16 | 2729.70 | 129.53 | 138.12 | 146.69 | 0.17914 |
| 31 | 3 | 81 | 2 | 45.50 | 2747.99 | 2748.65 | 2749.30 | 74.12 | 84.46 | 94.82 | 0.21658 |
| 31 | 6 | 81 | 2 | 45.38 | 2755.98 | 2756.63 | 2757.28 | 201.44 | 211.95 | 222.48 | 0.21658 |
| 31 | 3 | 9 | 2 | 45.24 | 2766.07 | 2766.15 | 2766.22 | 6.29 | 7.48 | 8.69 | 0.02406 |
| 32 | 3 | 63 | 2 | 44.74 | 2799.09 | 2799.60 | 2800.11 | 207.17 | 216.02 | 224.86 | 0.12600 |
| 32 | 6 | 63 | 2 | 44.65 | 2805.33 | 2805.84 | 2806.35 | 315.31 | 324.10 | 332.87 | 0.12600 |
| 32 | 3 | 63 | 2 | 44.50 | 2815.75 | 2816.26 | 2816.77 | 131.87 | 140.21 | 148.55 | 0.12600 |
| 32 | 6 | 63 | 2 | 44.41 | 2822.00 | 2822.51 | 2823.02 | 234.03 | 242.37 | 250.73 | 0.12600 |
| 32 | 3 | 61 | 2 | 44.27 | 2832.03 | 2832.52 | 2833.01 | 40.80 | 48.99 | 57.19 | 0.12200 |
| 32 | 6 | 61 | 2 | 44.18 | 2838.11 | 2838.60 | 2839.10 | 141.89 | 150.07 | 158.24 | 0.12200 |
| 32 | 3 | 63 | 2 | 44.03 | 2848.76 | 2849.26 | 2849.77 | 315.99 | 324.11 | 332.22 | 0.12600 |
| 32 | 6 | 63 | 2 | 43.94 | 2855.01 | 2855.51 | 2856.02 | 55.01 | 62.99 | 70.97 | 0.12600 |
| 33 | 6 | 84 | 2 | 43.46 | 2888.88 | 2889.56 | 2890.23 | 222.26 | 232.78 | 243.32 | 0.19400 |
| 33 | 3 | 94 | 2 | 43.23 | 2904.77 | 2905.53 | 2906.29 | 110.45 | 122.51 | 134.60 | 0.21709 |
| 33 | 6 | 97 | 2 | 43.10 | 2914.18 | 2914.96 | 2915.74 | 262.22 | 275.11 | 288.07 | 0.22402 |
| 33 | 3 | 26 | 2 | 42.94 | 2926.51 | 2926.72 | 2926.93 | 109.14 | 112.68 | 116.22 | 0.06005 |
| 33 | 6 | 67 | 2 | 42.86 | 2932.14 | 2932.68 | 2933.22 | 203.93 | 212.90 | 221.84 | 0.15473 |
| 33 | 3 | 65 | 2 | 42.71 | 2943.08 | 2943.61 | 2944.13 | 17.96 | 26.00 | 34.04 | 0.15012 |
| 34 | 3 | 66 | 2 | 42.23 | 2978.29 | 2978.83 | 2979.36 | 180.71 | 188.56 | 196.41 | 0.12865 |
| 34 | 6 | 66 | 2 | 42.14 | 2984.91 | 2985.45 | 2985.98 | 280.47 | 288.78 | 297.11 | 0.12865 |
| 34 | 3 | 66 | 2 | 41.99 | 2996.24 | 2996.77 | 2997.30 | 102.77 | 111.52 | 120.32 | 0.12865 |
| 34 | 6 | 67 | 2 | 41.90 | 3002.85 | 3003.39 | 3003.93 | 212.74 | 221.73 | 230.71 | 0.13060 |
| 34 | 3 | 65 | 2 | 41.75 | 3013.71 | 3014.23 | 3014.76 | 31.61 | 40.17 | 48.74 | 0.12671 |
| 34 | 6 | 64 | 2 | 41.67 | 3020.15 | 3020.66 | 3021.18 | 136.15 | 144.49 | 152.83 | 0.12476 |
| 34 | 3 | 59 | 2 | 41.50 | 3033.00 | 3033.47 | 3033.95 | 341.81 | 349.33 | 356.84 | 0.11501 |
| 34 | 6 | 60 | 2 | 41.42 | 3038.93 | 3039.42 | 3039.90 | 75.62 | 83.27 | 90.93 | 0.11696 |
| 35 | 3 | 62 | 2 | 40.37 | 3120.06 | 3120.56 | 3121.06 | 239.02 | 246.32 | 253.64 | 0.12450 |
| 35 | 6 | 62 | 2 | 40.29 | 3126.19 | 3126.69 | 3127.19 | 328.99 | 336.32 | 343.66 | 0.12450 |
| 35 | 3 | 62 | 2 | 40.16 | 3136.52 | 3137.02 | 3137.52 | 120.70 | 128.12 | 135.57 | 0.12450 |
| 35 | 6 | 62 | 2 | 40.08 | 3142.70 | 3143.20 | 3143.70 | 211.81 | 218.99 | 226.12 | 0.12450 |
| 35 | 3 | 62 | 2 | 39.95 | 3153.09 | 3153.59 | 3154.09 | 349.91 | 356.18 | 2.44 | 0.12450 |
| 35 | 6 | 62 | 2 | 39.88 | 3159.30 | 3159.80 | 3160.30 | 68.85 | 75.44 | 82.08 | 0.12450 |
| 35 | 3 | 63 | 2 | 39.74 | 3169.75 | 3170.25 | 3170.76 | 213.74 | 220.90 | 228.03 | 0.12651 |
| 35 | 6 | 63 | 2 | 39.67 | 3176.00 | 3176.50 | 3177.01 | 300.29 | 307.22 | 314.15 | 0.12651 |
| 36 | 6 | 64 | 2 | 39.25 | 3209.73 | 3210.24 | 3210.76 | 42.86 | 50.09 | 57.33 | 0.12427 |
| 36 | 3 | 64 | 2 | 39.12 | 3220.36 | 3220.88 | 3221.40 | 191.17 | 198.35 | 205.53 | 0.12427 |
| 36 | 6 | 64 | 2 | 39.04 | 3226.73 | 3227.24 | 3227.76 | 279.87 | 287.08 | 294.32 | 0.12427 |
| 36 | 3 | 64 | 2 | 38.91 | 3237.45 | 3237.97 | 3238.48 | 69.81 | 76.94 | 84.07 | 0.12427 |
| 36 | 6 | 64 | 2 | 38.83 | 3243.86 | 3244.38 | 3244.89 | 158.40 | 165.49 | 172.57 | 0.12427 |
| 36 | 3 | 65 | 2 | 38.70 | 3254.63 | 3255.16 | 3255.68 | 300.78 | 307.31 | 313.80 | 0.12621 |
| 36 | 6 | 65 | 2 | 38.62 | 3261.08 | 3261.60 | 3262.13 | 19.95 | 26.35 | 32.76 | 0.12621 |
| 36 | 3 | 65 | 2 | 38.49 | 3271.93 | 3272.46 | 3272.98 | 155.31 | 161.93 | 168.56 | 0.12621 |
| 37 | 6 | 65 | 2 | 38.21 | 3295.89 | 3296.42 | 3296.94 | 137.09 | 145.10 | 153.11 | 0.12241 |
| 37 | 3 | 66 | 2 | 38.08 | 3306.83 | 3307.37 | 3307.90 | 298.25 | 305.60 | 312.90 | 0.12429 |
| 37 | 6 | 66 | 2 | 38.00 | 3313.41 | 3313.94 | 3314.47 | 27.30 | 34.40 | 41.51 | 0.12429 |
| 37 | 3 | 67 | 2 | 37.87 | 3324.45 | 3324.99 | 3325.53 | 173.99 | 181.10 | 188.21 | 0.12618 |
| 37 | 6 | 66 | 2 | 37.79 | 3331.06 | 3331.59 | 3332.12 | 261.52 | 268.64 | 275.75 | 0.12429 |
| 37 | 3 | 67 | 2 | 37.66 | 3342.17 | 3342.71 | 3343.25 | 51.38 | 58.79 | 66.21 | 0.12618 |
| 37 | 6 | 67 | 2 | 37.58 | 3348.82 | 3349.36 | 3349.90 | 143.35 | 150.90 | 158.46 | 0.12618 |
| 37 | 3 | 67 | 2 | 37.45 | 3360.00 | 3360.54 | 3361.08 | 304.93 | 313.13 | 321.35 | 0.12618 |
| 38 | 3 | 68 | 2 | 37.14 | 3386.65 | 3387.19 | 3387.74 | 356.52 | 5.04 | 13.57 | 0.12409 |
| 38 | 6 | 68 | 2 | 37.07 | 3393.40 | 3393.94 | 3394.49 | 101.15 | 109.66 | 118.16 | 0.12409 |
| 38 | 3 | 68 | 2 | 36.94 | 3404.76 | 3405.31 | 3405.85 | 274.42 | 282.44 | 290.43 | 0.12409 |
| 38 | 6 | 68 | 2 | 36.86 | 3411.56 | 3412.10 | 3412.65 | 12.07 | 19.79 | 27.48 | 0.12409 |
| 38 | 3 | 69 | 2 | 36.73 | 3422.99 | 3423.55 | 3424.10 | 167.87 | 175.22 | 182.55 | 0.12591 |
| 38 | 6 | 69 | 2 | 36.65 | 3429.82 | 3430.38 | 3430.93 | 257.94 | 265.32 | 272.71 | 0.12591 |
| 38 | 3 | 69 | 2 | 36.52 | 3441.33 | 3441.88 | 3442.44 | 50.58 | 57.92 | 65.24 | 0.12591 |
| 38 | 6 | 69 | 2 | 36.44 | 3448.21 | 3448.76 | 3449.32 | 140.53 | 147.73 | 154.93 | 0.12591 |
| 39 | 3 | 70 | 2 | 36.10 | 3478.34 | 3478.90 | 3479.47 | 190.54 | 198.72 | 206.92 | 0.12367 |
| 39 | 6 | 70 | 2 | 36.03 | 3485.29 | 3485.85 | 3486.42 | 292.21 | 300.60 | 308.98 | 0.12367 |
| 39 | 3 | 70 | 2 | 35.91 | 3495.35 | 3495.92 | 3496.48 | 83.10 | 91.60 | 100.15 | 0.12367 |
| 39 | 6 | 71 | 2 | 35.84 | 3502.34 | 3502.91 | 3503.49 | 189.68 | 198.38 | 207.08 | 0.12544 |
| 39 | 3 | 71 | 2 | 35.71 | 3514.13 | 3514.70 | 3515.28 | 4.20 | 12.41 | 20.60 | 0.12544 |
| 39 | 6 | 71 | 2 | 35.63 | 3521.17 | 3521.74 | 3522.32 | 102.43 | 110.12 | 117.78 | 0.12544 |
| 39 | 3 | 71 | 2 | 35.45 | 3537.15 | 3537.73 | 3538.30 | 307.88 | 314.87 | 321.86 | 0.12544 |
| 39 | 6 | 72 | 2 | 35.38 | 3544.25 | 3544.83 | 3545.41 | 34.52 | 41.73 | 48.96 | 0.12721 |
| 40 | 3 | 72 | 2 | 35.04 | 3575.37 | 3575.95 | 3576.53 | 81.19 | 88.71 | 96.20 | 0.12329 |
| 40 | 6 | 72 | 2 | 34.96 | 3582.57 | 3583.15 | 3583.73 | 173.73 | 181.24 | 188.77 | 0.12329 |
| 40 | 3 | 73 | 2 | 34.83 | 3594.66 | 3595.25 | 3595.84 | 338.95 | 347.49 | 356.07 | 0.12500 |
| 40 | 6 | 73 | 2 | 34.75 | 3601.89 | 3602.48 | 3603.06 | 85.12 | 93.91 | 102.73 | 0.12500 |
| 40 | 3 | 73 | 2 | 34.62 | 3614.07 | 3614.66 | 3615.25 | 268.91 | 277.44 | 285.92 | 0.12500 |
| 40 | 6 | 73 | 2 | 34.55 | 3621.35 | 3621.94 | 3622.53 | 12.29 | 20.60 | 28.95 | 0.12500 |
| 40 | 3 | 74 | 2 | 34.41 | 3633.61 | 3634.21 | 3634.80 | 192.41 | 201.40 | 210.36 | 0.12671 |
| 40 | 6 | 74 | 2 | 34.34 | 3640.93 | 3641.53 | 3642.13 | 301.40 | 310.22 | 319.03 | 0.12671 |
| 41 | 6 | 74 | 2 | 34.17 | 3656.60 | 3657.20 | 3657.79 | 165.90 | 174.13 | 182.37 | 0.12292 |
| 41 | 3 | 75 | 2 | 34.04 | 3668.95 | 3669.56 | 3670.16 | 341.17 | 350.02 | 358.86 | 0.12458 |
| 41 | 6 | 75 | 2 | 33.96 | 3676.35 | 3676.95 | 3677.56 | 87.40 | 95.77 | 104.08 | 0.12458 |
| 41 | 3 | 75 | 2 | 33.83 | 3688.83 | 3689.44 | 3690.04 | 256.86 | 265.26 | 273.71 | 0.12458 |
| 41 | 6 | 75 | 2 | 33.76 | 3696.28 | 3696.89 | 3697.49 | 1.89 | 10.50 | 19.11 | 0.12458 |
| 41 | 3 | 76 | 2 | 33.63 | 3708.82 | 3709.44 | 3710.05 | 179.81 | 188.49 | 197.17 | 0.12625 |
| 41 | 6 | 76 | 2 | 33.55 | 3716.31 | 3716.93 | 3717.54 | 285.14 | 293.65 | 302.15 | 0.12625 |
| 41 | 3 | 76 | 2 | 33.43 | 3728.14 | 3728.76 | 3729.37 | 89.44 | 98.11 | 106.80 | 0.12625 |
| 42 | 6 | 77 | 2 | 32.72 | 3796.99 | 3797.61 | 3798.24 | 319.41 | 327.28 | 335.14 | 0.16383 |
| 42 | 3 | 78 | 2 | 32.59 | 3809.89 | 3810.52 | 3811.15 | 115.07 | 122.05 | 128.98 | 0.16596 |
| 42 | 6 | 78 | 2 | 32.52 | 3817.62 | 3818.25 | 3818.88 | 199.61 | 206.58 | 213.62 | 0.16596 |
| 42 | 3 | 79 | 2 | 32.39 | 3830.65 | 3831.29 | 3831.93 | 1.21 | 9.78 | 18.37 | 0.16809 |
| 42 | 6 | 79 | 2 | 32.31 | 3838.43 | 3839.07 | 3839.71 | 103.75 | 111.93 | 120.09 | 0.16809 |
| 42 | 3 | 79 | 2 | 32.18 | 3851.55 | 3852.18 | 3852.82 | 270.80 | 278.93 | 287.05 | 0.16809 |
| 43 | 3 | 80 | 2 | 31.76 | 3893.71 | 3894.35 | 3895.00 | 19.42 | 25.35 | 31.32 | 0.16563 |
| 43 | 6 | 80 | 2 | 31.68 | 3901.63 | 3902.28 | 3902.92 | 98.15 | 105.17 | 112.26 | 0.16563 |
| 43 | 3 | 81 | 2 | 31.55 | 3914.98 | 3915.63 | 3916.29 | 259.56 | 268.57 | 277.80 | 0.16770 |
| 43 | 6 | 80 | 2 | 31.48 | 3922.95 | 3923.60 | 3924.24 | 22.94 | 33.39 | 43.71 | 0.16563 |
| 43 | 3 | 81 | 2 | 31.35 | 3936.39 | 3937.05 | 3937.70 | 222.31 | 231.51 | 240.69 | 0.16770 |
| 43 | 6 | 81 | 2 | 31.27 | 3944.42 | 3945.07 | 3945.73 | 331.30 | 339.91 | 348.48 | 0.16770 |
| 44 | 3 | 82 | 2 | 30.93 | 3979.62 | 3980.28 | 3980.94 | 133.90 | 143.75 | 153.59 | 0.16566 |
| 44 | 6 | 82 | 2 | 30.85 | 3987.74 | 3988.40 | 3989.06 | 257.86 | 268.58 | 279.42 | 0.16566 |
| 44 | 3 | 83 | 2 | 30.73 | 3999.91 | 4000.58 | 4001.25 | 106.44 | 118.22 | 130.06 | 0.16768 |
| 44 | 6 | 82 | 2 | 30.66 | 4008.09 | 4008.75 | 4009.41 | 251.02 | 262.55 | 274.05 | 0.16566 |
| 44 | 3 | 83 | 2 | 30.53 | 4021.86 | 4022.52 | 4023.19 | 125.01 | 135.90 | 146.78 | 0.16768 |
| 44 | 6 | 83 | 2 | 30.45 | 4030.08 | 4030.75 | 4031.42 | 256.50 | 266.99 | 277.44 | 0.16768 |
| 45 | 3 | 84 | 2 | 30.11 | 4066.15 | 4066.83 | 4067.51 | 56.69 | 65.92 | 75.17 | 0.16535 |
| 45 | 6 | 84 | 2 | 30.03 | 4074.48 | 4075.16 | 4075.84 | 170.33 | 179.59 | 188.87 | 0.16535 |
| 45 | 3 | 85 | 2 | 29.90 | 4088.49 | 4089.17 | 4089.86 | 4.56 | 14.14 | 23.70 | 0.16732 |
| 45 | 6 | 85 | 2 | 29.82 | 4096.87 | 4097.55 | 4098.24 | 119.89 | 129.17 | 138.42 | 0.16732 |
| 45 | 3 | 85 | 2 | 29.69 | 4110.99 | 4111.67 | 4112.36 | 308.65 | 317.74 | 326.82 | 0.16732 |
| 45 | 6 | 85 | 2 | 29.62 | 4119.41 | 4120.10 | 4120.78 | 58.95 | 67.73 | 76.49 | 0.16732 |
| 46 | 3 | 86 | 2 | 29.28 | 4156.41 | 4157.11 | 4157.80 | 124.00 | 130.51 | 136.96 | 0.16507 |
| 46 | 6 | 86 | 2 | 29.20 | 4164.94 | 4165.64 | 4166.33 | 203.38 | 209.97 | 216.61 | 0.16507 |
| 46 | 3 | 87 | 2 | 29.07 | 4179.31 | 4180.02 | 4180.72 | 356.50 | 4.69 | 12.89 | 0.16699 |
| 46 | 6 | 87 | 2 | 28.99 | 4187.90 | 4188.60 | 4189.31 | 98.26 | 106.79 | 115.36 | 0.16699 |
| 46 | 3 | 88 | 2 | 28.86 | 4202.38 | 4203.09 | 4203.80 | 278.93 | 287.79 | 296.65 | 0.16891 |
| 46 | 6 | 87 | 2 | 28.78 | 4211.02 | 4211.72 | 4212.43 | 27.33 | 36.24 | 45.16 | 0.16699 |
| 47 | 3 | 89 | 2 | 28.44 | 4248.94 | 4249.66 | 4250.38 | 148.43 | 157.24 | 166.02 | 0.16604 |
| 47 | 6 | 89 | 2 | 28.36 | 4257.69 | 4258.41 | 4259.13 | 254.08 | 262.51 | 270.90 | 0.16604 |
| 47 | 3 | 89 | 2 | 28.23 | 4272.46 | 4273.17 | 4273.89 | 63.47 | 71.80 | 80.16 | 0.16604 |
| 47 | 6 | 89 | 2 | 28.15 | 4281.26 | 4281.98 | 4282.70 | 166.64 | 175.11 | 183.57 | 0.16604 |
| 47 | 3 | 90 | 2 | 28.02 | 4296.09 | 4296.82 | 4297.54 | 344.39 | 353.14 | 1.83 | 0.16791 |
| 47 | 6 | 90 | 2 | 27.94 | 4304.95 | 4305.68 | 4306.40 | 83.86 | 91.09 | 98.23 | 0.16791 |
| 48 | 3 | 91 | 2 | 27.60 | 4343.83 | 4344.56 | 4345.29 | 102.87 | 110.85 | 118.87 | 0.16606 |
| 48 | 6 | 91 | 2 | 27.53 | 4352.79 | 4353.53 | 4354.26 | 202.23 | 210.34 | 218.44 | 0.16606 |
| 48 | 3 | 91 | 2 | 27.39 | 4367.92 | 4368.65 | 4369.39 | 3.78 | 11.43 | 19.05 | 0.16606 |
| 48 | 6 | 91 | 2 | 27.32 | 4376.94 | 4377.68 | 4378.41 | 96.67 | 104.06 | 111.42 | 0.16606 |
| 48 | 3 | 92 | 2 | 27.18 | 4392.16 | 4392.90 | 4393.65 | 249.31 | 257.00 | 264.74 | 0.16788 |
| 48 | 6 | 92 | 2 | 27.11 | 4401.24 | 4401.98 | 4402.73 | 345.65 | 353.66 | 1.66 | 0.16788 |
| 50 | 3 | 128 | 2 | 23.34 | 4862.15 | 4863.18 | 4864.21 | 212.80 | 223.86 | 234.95 | 0.24903 |
| 50 | 6 | 128 | 2 | 23.24 | 4874.74 | 4875.77 | 4876.81 | 351.40 | 2.85 | 14.21 | 0.24903 |
| 50 | 3 | 129 | 2 | 23.08 | 4895.76 | 4896.80 | 4897.84 | 205.74 | 215.94 | 226.20 | 0.25097 |
| 50 | 6 | 129 | 2 | 22.98 | 4908.40 | 4909.44 | 4910.48 | 333.07 | 343.72 | 354.38 | 0.25097 |
| 51 | 3 | 132 | 2 | 20.77 | 5204.62 | 5205.69 | 5206.75 | 133.95 | 145.02 | 156.06 | 0.24906 |
| 51 | 6 | 132 | 2 | 20.67 | 5217.59 | 5218.65 | 5219.72 | 265.96 | 276.50 | 287.01 | 0.24906 |
| 51 | 3 | 133 | 2 | 20.52 | 5239.47 | 5240.54 | 5241.61 | 113.63 | 123.20 | 132.69 | 0.25094 |
| 51 | 6 | 133 | 2 | 20.42 | 5252.49 | 5253.56 | 5254.63 | 228.20 | 237.66 | 247.14 | 0.25094 |
| 54 | 3 | 131 | 2 | 17.10 | 5732.54 | 5733.59 | 5734.65 | 263.02 | 272.99 | 282.96 | 0.25586 |
| 54 | 6 | 131 | 2 | 17.02 | 5745.41 | 5746.46 | 5747.52 | 24.63 | 34.60 | 44.54 | 0.25586 |
| 54 | 3 | 125 | 2 | 16.88 | 5765.64 | 5766.65 | 5767.65 | 213.00 | 222.36 | 231.72 | 0.24414 |
| 54 | 6 | 125 | 2 | 16.80 | 5777.89 | 5778.90 | 5779.90 | 327.06 | 336.47 | 345.88 | 0.24414 |
| 55 | 3 | 117 | 2 | 14.90 | 6074.35 | 6075.29 | 6076.23 | 259.62 | 268.24 | 276.89 | 0.24894 |
| 55 | 6 | 117 | 2 | 14.82 | 6085.89 | 6086.84 | 6087.78 | 6.60 | 15.44 | 24.29 | 0.24894 |
| 55 | 3 | 118 | 2 | 14.70 | 6105.34 | 6106.30 | 6107.25 | 188.44 | 197.40 | 206.44 | 0.25106 |
| 55 | 6 | 118 | 2 | 14.63 | 6116.92 | 6117.88 | 6118.83 | 300.30 | 309.73 | 319.17 | 0.25106 |
| 56 | 3 | 124 | 2 | 12.75 | 6424.50 | 6425.50 | 6426.50 | 235.72 | 244.92 | 254.06 | 0.24900 |
| 56 | 6 | 124 | 2 | 12.67 | 6436.71 | 6437.71 | 6438.71 | 346.77 | 355.74 | 4.67 | 0.24900 |
| 56 | 3 | 125 | 2 | 12.55 | 6457.27 | 6458.28 | 6459.28 | 166.42 | 174.95 | 183.42 | 0.25100 |
| 56 | 6 | 125 | 2 | 12.48 | 6469.56 | 6470.57 | 6471.57 | 266.25 | 273.97 | 281.62 | 0.25100 |
| 57 | 6 | 130 | 2 | 10.90 | 6738.68 | 6739.73 | 6740.78 | 309.27 | 317.71 | 326.18 | 0.24857 |
| 57 | 3 | 131 | 2 | 10.78 | 6760.10 | 6761.16 | 6762.21 | 125.26 | 134.23 | 143.20 | 0.25048 |
| 57 | 6 | 131 | 2 | 10.71 | 6772.95 | 6774.00 | 6775.06 | 233.10 | 241.68 | 250.26 | 0.25048 |
| 57 | 3 | 131 | 2 | 10.59 | 6794.45 | 6795.51 | 6796.56 | 43.18 | 51.25 | 59.30 | 0.25048 |
| 58 | 3 | 136 | 2 | 9.01 | 7075.32 | 7076.42 | 7077.51 | 165.63 | 173.70 | 181.78 | 0.25000 |
| 58 | 6 | 136 | 2 | 8.94 | 7088.64 | 7089.74 | 7090.83 | 262.99 | 270.83 | 278.65 | 0.25000 |
| 58 | 3 | 136 | 2 | 8.81 | 7111.18 | 7112.28 | 7113.37 | 65.64 | 74.10 | 82.66 | 0.25000 |
| 58 | 6 | 136 | 2 | 8.74 | 7124.56 | 7125.65 | 7126.75 | 173.53 | 182.56 | 191.60 | 0.25000 |
| 59 | 3 | 141 | 2 | 7.24 | 7403.24 | 7404.38 | 7405.52 | 84.79 | 92.93 | 101.05 | 0.24912 |
| 59 | 6 | 141 | 2 | 7.16 | 7417.09 | 7418.23 | 7419.36 | 183.33 | 191.45 | 199.58 | 0.24912 |
| 59 | 3 | 142 | 2 | 7.04 | 7440.38 | 7441.52 | 7442.67 | 354.32 | 2.88 | 11.43 | 0.25088 |
| 59 | 6 | 142 | 2 | 6.97 | 7454.28 | 7455.42 | 7456.57 | 97.99 | 106.59 | 115.19 | 0.25088 |
| 60 | 3 | 146 | 2 | 5.46 | 7742.90 | 7744.08 | 7745.25 | 75.22 | 83.19 | 91.10 | 0.25000 |
| 60 | 6 | 146 | 2 | 5.39 | 7757.20 | 7758.37 | 7759.55 | 170.42 | 178.23 | 186.02 | 0.25000 |
| 60 | 3 | 146 | 2 | 5.27 | 7781.21 | 7782.39 | 7783.57 | 332.15 | 340.18 | 348.21 | 0.25000 |
| 60 | 6 | 146 | 2 | 5.19 | 7795.54 | 7796.72 | 7797.90 | 69.03 | 76.84 | 84.61 | 0.25000 |
| 1 | 4 | 1 | 10 | 142.37 | 143.32 | 143.36 | 143.40 | 210.36 | 209.16 | 207.97 | 0.01515 |
| 1 | 4 | 1 | 10 | 142.36 | 143.48 | 143.52 | 143.56 | 205.57 | 204.38 | 203.18 | 0.01515 |
| 1 | 4 | 1 | 10 | 142.35 | 143.64 | 143.68 | 143.72 | 200.79 | 199.59 | 198.40 | 0.01515 |
| 1 | 4 | 1 | 10 | 142.35 | 143.80 | 143.84 | 143.88 | 196.01 | 194.81 | 193.61 | 0.01515 |
| 1 | 5 | 1 | 10 | 142.31 | 144.42 | 144.46 | 144.50 | 177.61 | 176.42 | 175.22 | 0.01515 |
| 1 | 5 | 1 | 10 | 142.31 | 144.58 | 144.62 | 144.66 | 172.84 | 171.64 | 170.45 | 0.01515 |
| 1 | 5 | 1 | 10 | 142.30 | 144.74 | 144.78 | 144.82 | 168.07 | 166.87 | 165.68 | 0.01515 |
| 1 | 5 | 1 | 10 | 142.29 | 144.90 | 144.95 | 144.99 | 163.30 | 162.10 | 160.91 | 0.01515 |
| 1 | 4 | 1 | 12 | 141.90 | 152.40 | 152.44 | 152.49 | 303.89 | 302.49 | 301.08 | 0.01515 |
| 1 | 4 | 1 | 12 | 141.89 | 152.59 | 152.64 | 152.69 | 298.28 | 296.88 | 295.47 | 0.01515 |
| 1 | 4 | 1 | 12 | 141.88 | 152.78 | 152.83 | 152.88 | 292.67 | 291.27 | 289.87 | 0.01515 |
| 1 | 4 | 1 | 12 | 141.87 | 152.98 | 153.03 | 153.07 | 287.06 | 285.66 | 284.26 | 0.01515 |
| 1 | 5 | 1 | 12 | 141.84 | 153.55 | 153.60 | 153.65 | 270.49 | 269.09 | 267.69 | 0.01515 |
| 1 | 5 | 1 | 12 | 141.83 | 153.74 | 153.79 | 153.84 | 264.90 | 263.50 | 262.10 | 0.01515 |
| 1 | 5 | 1 | 12 | 141.82 | 153.94 | 153.99 | 154.03 | 259.30 | 257.91 | 256.51 | 0.01515 |
| 1 | 5 | 1 | 12 | 141.81 | 154.13 | 154.18 | 154.23 | 253.71 | 252.32 | 250.92 | 0.01515 |
| 1 | 4 | 1 | 10 | 141.41 | 161.83 | 161.87 | 161.91 | 33.44 | 32.29 | 31.15 | 0.01515 |
| 1 | 4 | 1 | 10 | 141.40 | 161.99 | 162.03 | 162.07 | 28.87 | 27.73 | 26.59 | 0.01515 |
| 1 | 4 | 1 | 10 | 141.39 | 162.15 | 162.19 | 162.23 | 24.30 | 23.16 | 22.02 | 0.01515 |
| 1 | 4 | 1 | 10 | 141.39 | 162.32 | 162.36 | 162.40 | 19.74 | 18.60 | 17.45 | 0.01515 |
| 1 | 4 | 1 | 10 | 141.38 | 162.48 | 162.52 | 162.56 | 15.17 | 14.03 | 12.89 | 0.01515 |
| 1 | 5 | 1 | 10 | 141.35 | 163.02 | 163.06 | 163.10 | 359.90 | 358.76 | 357.62 | 0.01515 |
| 1 | 5 | 1 | 10 | 141.34 | 163.18 | 163.22 | 163.26 | 355.34 | 354.20 | 353.06 | 0.01515 |
| 1 | 5 | 1 | 10 | 141.33 | 163.34 | 163.38 | 163.42 | 350.78 | 349.65 | 348.51 | 0.01515 |
| 1 | 5 | 1 | 10 | 141.32 | 163.50 | 163.54 | 163.58 | 346.23 | 345.09 | 343.96 | 0.01515 |
| 1 | 5 | 1 | 10 | 141.32 | 163.66 | 163.70 | 163.74 | 341.68 | 340.54 | 339.40 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.91 | 171.61 | 171.65 | 171.69 | 119.57 | 118.45 | 117.34 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.90 | 171.77 | 171.81 | 171.85 | 115.11 | 113.99 | 112.88 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.89 | 171.94 | 171.98 | 172.02 | 110.65 | 109.54 | 108.42 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.88 | 172.10 | 172.14 | 172.18 | 106.19 | 105.08 | 103.97 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.88 | 172.26 | 172.30 | 172.34 | 101.74 | 100.62 | 99.51 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.85 | 172.84 | 172.88 | 172.92 | 85.71 | 84.60 | 83.49 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.84 | 173.00 | 173.04 | 173.08 | 81.26 | 80.15 | 79.04 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.83 | 173.16 | 173.20 | 173.24 | 76.82 | 75.71 | 74.59 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.82 | 173.32 | 173.36 | 173.40 | 72.37 | 71.26 | 70.15 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.81 | 173.48 | 173.52 | 173.56 | 67.93 | 66.82 | 65.71 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.39 | 181.68 | 181.73 | 181.77 | 204.33 | 203.25 | 202.16 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.39 | 181.85 | 181.89 | 181.93 | 199.99 | 198.90 | 197.82 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.38 | 182.01 | 182.05 | 182.09 | 195.64 | 194.55 | 193.47 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.37 | 182.17 | 182.21 | 182.25 | 191.29 | 190.21 | 189.12 | 0.01515 |
| 1 | 4 | 1 | 10 | 140.36 | 182.33 | 182.37 | 182.41 | 186.95 | 185.86 | 184.77 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.33 | 182.95 | 182.99 | 183.03 | 170.22 | 169.14 | 168.06 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.32 | 183.11 | 183.15 | 183.19 | 165.89 | 164.80 | 163.72 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.31 | 183.27 | 183.31 | 183.35 | 161.55 | 160.46 | 159.38 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.31 | 183.43 | 183.48 | 183.52 | 157.21 | 156.13 | 155.04 | 0.01515 |
| 1 | 5 | 1 | 10 | 140.30 | 183.60 | 183.64 | 183.68 | 152.87 | 151.79 | 150.71 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.87 | 192.00 | 192.04 | 192.08 | 289.65 | 288.60 | 287.54 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.86 | 192.16 | 192.20 | 192.24 | 285.42 | 284.36 | 283.30 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.86 | 192.32 | 192.36 | 192.40 | 281.19 | 280.13 | 279.07 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.85 | 192.48 | 192.52 | 192.56 | 276.96 | 275.90 | 274.84 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.84 | 192.64 | 192.68 | 192.72 | 272.73 | 271.67 | 270.61 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.81 | 193.29 | 193.33 | 193.37 | 255.83 | 254.78 | 253.72 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.80 | 193.45 | 193.49 | 193.53 | 251.61 | 250.56 | 249.50 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.79 | 193.61 | 193.65 | 193.69 | 247.39 | 246.34 | 245.29 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.78 | 193.77 | 193.81 | 193.85 | 243.18 | 242.12 | 241.07 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.77 | 193.93 | 193.97 | 194.01 | 238.96 | 237.91 | 236.86 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.33 | 202.69 | 202.73 | 202.77 | 13.13 | 12.10 | 11.08 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.33 | 202.85 | 202.89 | 202.93 | 9.03 | 8.00 | 6.98 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.32 | 203.01 | 203.05 | 203.09 | 4.93 | 3.90 | 2.88 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.31 | 203.18 | 203.22 | 203.26 | 0.83 | 359.81 | 358.78 | 0.01515 |
| 1 | 4 | 1 | 10 | 139.30 | 203.34 | 203.38 | 203.42 | 356.73 | 355.71 | 354.69 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.27 | 204.03 | 204.07 | 204.11 | 339.15 | 338.13 | 337.11 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.26 | 204.19 | 204.23 | 204.27 | 335.06 | 334.04 | 333.02 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.25 | 204.35 | 204.39 | 204.43 | 330.98 | 329.96 | 328.94 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.24 | 204.51 | 204.55 | 204.59 | 326.90 | 325.88 | 324.86 | 0.01515 |
| 1 | 5 | 1 | 10 | 139.24 | 204.68 | 204.72 | 204.76 | 322.82 | 321.80 | 320.78 | 0.01515 |
| 2 | 4 | 16 | 2 | 135.37 | 283.94 | 284.07 | 284.20 | 340.86 | 338.29 | 335.73 | 0.06324 |
| 2 | 4 | 16 | 2 | 135.35 | 284.46 | 284.58 | 284.71 | 330.60 | 328.04 | 325.48 | 0.06324 |
| 2 | 5 | 16 | 2 | 135.29 | 285.60 | 285.73 | 285.86 | 307.91 | 305.36 | 302.81 | 0.06324 |
| 2 | 5 | 16 | 2 | 135.27 | 286.12 | 286.25 | 286.37 | 297.72 | 295.17 | 292.63 | 0.06324 |
| 2 | 5 | 3 | 2 | 134.91 | 293.87 | 293.90 | 293.92 | 147.26 | 146.80 | 146.34 | 0.01186 |
| 2 | 4 | 19 | 2 | 134.15 | 309.87 | 310.02 | 310.18 | 214.33 | 211.65 | 208.96 | 0.07510 |
| 2 | 4 | 19 | 2 | 134.12 | 310.48 | 310.64 | 310.79 | 203.60 | 200.93 | 198.25 | 0.07510 |
| 2 | 4 | 19 | 2 | 134.09 | 311.10 | 311.25 | 311.40 | 192.91 | 190.25 | 187.58 | 0.07510 |
| 2 | 5 | 19 | 2 | 134.03 | 312.53 | 312.69 | 312.84 | 168.02 | 165.38 | 162.74 | 0.07510 |
| 2 | 5 | 19 | 2 | 134.00 | 313.15 | 313.30 | 313.45 | 157.47 | 154.83 | 152.20 | 0.07510 |
| 2 | 5 | 19 | 2 | 133.97 | 313.76 | 313.91 | 314.06 | 146.95 | 144.33 | 141.71 | 0.07510 |
| 2 | 4 | 18 | 2 | 133.26 | 329.14 | 329.29 | 329.43 | 255.21 | 252.95 | 250.69 | 0.07115 |
| 2 | 4 | 18 | 2 | 133.23 | 329.72 | 329.87 | 330.02 | 246.18 | 243.93 | 241.68 | 0.07115 |
| 2 | 5 | 18 | 2 | 133.17 | 331.07 | 331.22 | 331.36 | 225.38 | 223.15 | 220.92 | 0.07115 |
| 2 | 5 | 18 | 2 | 133.14 | 331.65 | 331.80 | 331.94 | 216.46 | 214.24 | 212.02 | 0.07115 |
| 3 | 4 | 23 | 2 | 128.07 | 446.00 | 446.18 | 446.37 | 79.12 | 77.91 | 76.71 | 0.04582 |
| 3 | 4 | 23 | 2 | 128.03 | 446.74 | 446.92 | 447.11 | 74.30 | 73.11 | 71.91 | 0.04582 |
| 3 | 4 | 23 | 2 | 128.00 | 447.48 | 447.67 | 447.85 | 69.53 | 68.34 | 67.16 | 0.04582 |
| 3 | 5 | 23 | 2 | 127.93 | 449.12 | 449.31 | 449.50 | 59.10 | 57.94 | 56.78 | 0.04582 |
| 3 | 5 | 23 | 2 | 127.90 | 449.87 | 450.05 | 450.24 | 54.47 | 53.32 | 52.17 | 0.04582 |
| 3 | 5 | 23 | 2 | 127.87 | 450.61 | 450.79 | 450.98 | 49.88 | 48.74 | 47.61 | 0.04582 |
| 3 | 4 | 21 | 2 | 127.08 | 469.29 | 469.46 | 469.63 | 308.89 | 308.10 | 307.31 | 0.04183 |
| 3 | 4 | 21 | 2 | 127.05 | 469.97 | 470.14 | 470.31 | 305.74 | 304.96 | 304.18 | 0.04183 |
| 3 | 4 | 21 | 2 | 127.02 | 470.65 | 470.82 | 470.99 | 302.63 | 301.86 | 301.09 | 0.04183 |
| 3 | 4 | 21 | 2 | 126.99 | 471.32 | 471.49 | 471.66 | 299.56 | 298.80 | 298.04 | 0.04183 |
| 3 | 5 | 21 | 2 | 126.93 | 472.87 | 473.03 | 473.20 | 292.69 | 291.94 | 291.19 | 0.04183 |
| 3 | 5 | 21 | 2 | 126.90 | 473.54 | 473.71 | 473.88 | 289.71 | 288.97 | 288.23 | 0.04183 |
| 3 | 5 | 21 | 2 | 126.87 | 474.22 | 474.39 | 474.56 | 286.76 | 286.04 | 285.31 | 0.04183 |
| 3 | 5 | 21 | 2 | 126.84 | 474.90 | 475.07 | 475.24 | 283.87 | 283.16 | 282.44 | 0.04183 |
| 3 | 4 | 24 | 2 | 125.74 | 501.44 | 501.64 | 501.83 | 196.09 | 195.62 | 195.16 | 0.04781 |
| 3 | 4 | 24 | 2 | 125.70 | 502.22 | 502.41 | 502.61 | 194.25 | 193.79 | 193.34 | 0.04781 |
| 3 | 4 | 24 | 2 | 125.67 | 502.99 | 503.19 | 503.38 | 192.44 | 192.00 | 191.56 | 0.04781 |
| 3 | 4 | 24 | 2 | 125.64 | 503.77 | 503.96 | 504.15 | 190.68 | 190.24 | 189.81 | 0.04781 |
| 3 | 5 | 20 | 2 | 125.57 | 505.56 | 505.72 | 505.88 | 186.77 | 186.42 | 186.08 | 0.03984 |
| 3 | 5 | 20 | 2 | 125.54 | 506.20 | 506.36 | 506.52 | 185.41 | 185.07 | 184.74 | 0.03984 |
| 3 | 5 | 20 | 2 | 125.52 | 506.85 | 507.01 | 507.17 | 184.07 | 183.74 | 183.41 | 0.03984 |
| 3 | 5 | 20 | 2 | 125.49 | 507.49 | 507.65 | 507.81 | 182.75 | 182.42 | 182.10 | 0.03984 |
| 3 | 5 | 20 | 2 | 125.46 | 508.14 | 508.30 | 508.46 | 181.46 | 181.14 | 180.83 | 0.03984 |
| 4 | 4 | 23 | 2 | 124.26 | 537.65 | 537.84 | 538.03 | 149.61 | 149.58 | 149.54 | 0.04656 |
| 4 | 4 | 23 | 2 | 124.23 | 538.40 | 538.58 | 538.77 | 149.47 | 149.43 | 149.40 | 0.04656 |
| 4 | 4 | 23 | 2 | 124.20 | 539.14 | 539.32 | 539.51 | 149.34 | 149.31 | 149.28 | 0.04656 |
| 4 | 4 | 23 | 2 | 124.17 | 539.88 | 540.07 | 540.25 | 149.23 | 149.21 | 149.20 | 0.04656 |
| 4 | 4 | 23 | 2 | 124.14 | 540.62 | 540.81 | 540.99 | 149.17 | 149.16 | 149.15 | 0.04656 |
| 4 | 5 | 23 | 2 | 124.07 | 542.35 | 542.53 | 542.72 | 149.16 | 149.17 | 149.18 | 0.04656 |
| 4 | 5 | 23 | 2 | 124.04 | 543.09 | 543.27 | 543.46 | 149.21 | 149.23 | 149.25 | 0.04656 |
| 4 | 5 | 23 | 2 | 124.01 | 543.83 | 544.02 | 544.20 | 149.30 | 149.33 | 149.36 | 0.04656 |
| 4 | 5 | 23 | 2 | 123.98 | 544.57 | 544.76 | 544.94 | 149.41 | 149.44 | 149.48 | 0.04656 |
| 4 | 5 | 23 | 2 | 123.95 | 545.32 | 545.50 | 545.69 | 149.54 | 149.58 | 149.62 | 0.04656 |
| 4 | 4 | 22 | 2 | 122.57 | 579.77 | 579.95 | 580.13 | 191.26 | 191.66 | 192.05 | 0.04453 |
| 4 | 4 | 22 | 2 | 122.54 | 580.48 | 580.66 | 580.83 | 192.86 | 193.26 | 193.67 | 0.04453 |
| 4 | 4 | 22 | 2 | 122.52 | 581.19 | 581.37 | 581.54 | 194.48 | 194.89 | 195.31 | 0.04453 |
| 4 | 4 | 22 | 2 | 122.49 | 581.90 | 582.08 | 582.25 | 196.14 | 196.56 | 196.98 | 0.04453 |
| 4 | 4 | 22 | 2 | 122.46 | 582.61 | 582.79 | 582.96 | 197.83 | 198.25 | 198.68 | 0.04453 |
| 4 | 4 | 22 | 2 | 122.43 | 583.32 | 583.50 | 583.67 | 199.54 | 199.97 | 200.40 | 0.04453 |
| 4 | 5 | 22 | 2 | 122.36 | 585.13 | 585.31 | 585.49 | 204.05 | 204.50 | 204.96 | 0.04453 |
| 4 | 5 | 22 | 2 | 122.33 | 585.84 | 586.02 | 586.20 | 205.87 | 206.33 | 206.79 | 0.04453 |
| 4 | 5 | 22 | 2 | 122.30 | 586.55 | 586.73 | 586.91 | 207.72 | 208.19 | 208.65 | 0.04453 |
| 4 | 5 | 22 | 2 | 122.28 | 587.26 | 587.44 | 587.62 | 209.59 | 210.07 | 210.54 | 0.04453 |
| 4 | 5 | 22 | 2 | 122.25 | 587.97 | 588.15 | 588.33 | 211.49 | 211.98 | 212.46 | 0.04453 |
| 4 | 5 | 22 | 2 | 122.22 | 588.68 | 588.86 | 589.04 | 213.43 | 213.92 | 214.41 | 0.04453 |
| 8 | 4 | 22 | 2 | 112.69 | 847.59 | 847.77 | 847.95 | 148.21 | 151.03 | 153.85 | 0.05914 |
| 8 | 4 | 22 | 2 | 112.67 | 848.30 | 848.48 | 848.66 | 159.49 | 162.31 | 165.14 | 0.05914 |
| 8 | 4 | 22 | 2 | 112.64 | 849.01 | 849.19 | 849.37 | 170.79 | 173.62 | 176.46 | 0.05914 |
| 8 | 5 | 22 | 2 | 112.58 | 850.67 | 850.84 | 851.02 | 197.22 | 200.06 | 202.90 | 0.05914 |
| 8 | 5 | 22 | 2 | 112.56 | 851.38 | 851.55 | 851.73 | 208.59 | 211.43 | 214.28 | 0.05914 |
| 8 | 5 | 22 | 2 | 112.54 | 852.09 | 852.26 | 852.44 | 219.97 | 222.83 | 225.68 | 0.05914 |
| 8 | 4 | 21 | 2 | 111.90 | 870.89 | 871.06 | 871.23 | 170.26 | 173.12 | 175.99 | 0.05645 |
| 8 | 4 | 21 | 2 | 111.88 | 871.57 | 871.74 | 871.91 | 181.72 | 184.59 | 187.46 | 0.05645 |
| 8 | 4 | 21 | 2 | 111.85 | 872.25 | 872.42 | 872.58 | 193.21 | 196.09 | 198.97 | 0.05645 |
| 8 | 5 | 21 | 2 | 111.80 | 873.78 | 873.95 | 874.12 | 219.30 | 222.20 | 225.09 | 0.05645 |
| 8 | 5 | 21 | 2 | 111.78 | 874.46 | 874.62 | 874.79 | 230.87 | 233.77 | 236.67 | 0.05645 |
| 8 | 5 | 21 | 2 | 111.76 | 875.13 | 875.30 | 875.47 | 242.46 | 245.36 | 248.27 | 0.05645 |
| 8 | 4 | 19 | 2 | 111.17 | 892.79 | 892.94 | 893.09 | 191.75 | 194.50 | 197.24 | 0.05108 |
| 8 | 4 | 19 | 2 | 111.15 | 893.40 | 893.55 | 893.70 | 202.74 | 205.49 | 208.24 | 0.05108 |
| 8 | 4 | 19 | 2 | 111.13 | 894.01 | 894.16 | 894.32 | 213.75 | 216.50 | 219.26 | 0.05108 |
| 8 | 5 | 19 | 2 | 111.08 | 895.46 | 895.62 | 895.77 | 239.88 | 242.65 | 245.41 | 0.05108 |
| 8 | 5 | 19 | 2 | 111.06 | 896.08 | 896.23 | 896.38 | 250.95 | 253.71 | 256.48 | 0.05108 |
| 8 | 5 | 19 | 2 | 111.04 | 896.69 | 896.84 | 896.99 | 262.03 | 264.80 | 267.57 | 0.05108 |
| 9 | 4 | 19 | 2 | 107.09 | 967.24 | 967.39 | 967.54 | 324.93 | 328.92 | 332.91 | 0.07170 |
| 9 | 4 | 19 | 2 | 107.05 | 967.85 | 968.00 | 968.16 | 340.92 | 344.92 | 348.93 | 0.07170 |
| 9 | 5 | 19 | 2 | 106.95 | 969.26 | 969.42 | 969.57 | 17.98 | 22.03 | 26.08 | 0.07170 |
| 9 | 5 | 19 | 2 | 106.90 | 969.88 | 970.03 | 970.18 | 34.20 | 38.27 | 42.34 | 0.07170 |
| 9 | 4 | 21 | 2 | 105.71 | 986.25 | 986.41 | 986.58 | 131.96 | 136.98 | 142.01 | 0.07925 |
| 9 | 5 | 21 | 2 | 105.61 | 987.70 | 987.87 | 988.04 | 175.46 | 180.54 | 185.63 | 0.07925 |
| 9 | 4 | 22 | 2 | 104.81 | 998.57 | 998.74 | 998.92 | 153.79 | 159.49 | 165.19 | 0.08302 |
| 9 | 5 | 21 | 2 | 104.69 | 1000.08 | 1000.25 | 1000.42 | 202.63 | 208.10 | 213.58 | 0.07925 |
| 9 | 4 | 15 | 2 | 103.99 | 1009.72 | 1009.84 | 1009.96 | 160.11 | 164.17 | 168.23 | 0.05660 |
| 9 | 4 | 15 | 2 | 103.95 | 1010.20 | 1010.32 | 1010.45 | 176.35 | 180.41 | 184.47 | 0.05660 |
| 9 | 5 | 15 | 2 | 103.87 | 1011.35 | 1011.47 | 1011.59 | 214.84 | 218.92 | 222.99 | 0.05660 |
| 9 | 5 | 15 | 2 | 103.83 | 1011.83 | 1011.95 | 1012.07 | 231.15 | 235.23 | 239.31 | 0.05660 |
| 9 | 4 | 22 | 2 | 103.06 | 1022.28 | 1022.45 | 1022.63 | 228.72 | 234.88 | 241.04 | 0.08302 |
| 9 | 5 | 22 | 2 | 102.95 | 1023.81 | 1023.98 | 1024.16 | 282.01 | 288.19 | 294.38 | 0.08302 |
| 10 | 4 | 19 | 2 | 100.65 | 1055.67 | 1055.82 | 1055.98 | 355.33 | 1.09 | 6.86 | 0.08190 |
| 10 | 5 | 19 | 2 | 100.56 | 1056.99 | 1057.14 | 1057.29 | 44.85 | 50.64 | 56.43 | 0.08190 |
| 10 | 4 | 18 | 2 | 99.93 | 1065.90 | 1066.04 | 1066.19 | 25.43 | 31.05 | 36.67 | 0.07759 |
| 10 | 5 | 18 | 2 | 99.84 | 1067.17 | 1067.32 | 1067.46 | 74.82 | 80.45 | 86.09 | 0.07759 |
| 10 | 4 | 17 | 2 | 99.24 | 1075.77 | 1075.91 | 1076.05 | 53.36 | 58.82 | 64.29 | 0.07328 |
| 10 | 5 | 17 | 2 | 99.15 | 1077.01 | 1077.15 | 1077.28 | 102.62 | 108.10 | 113.59 | 0.07328 |
| 10 | 4 | 16 | 2 | 98.49 | 1086.69 | 1086.82 | 1086.95 | 136.31 | 141.64 | 146.97 | 0.06897 |
| 10 | 5 | 16 | 2 | 98.40 | 1087.86 | 1087.99 | 1088.12 | 184.70 | 190.05 | 195.40 | 0.06897 |
| 10 | 4 | 16 | 2 | 97.86 | 1095.87 | 1096.00 | 1096.13 | 161.46 | 166.96 | 172.46 | 0.06897 |
| 10 | 5 | 16 | 2 | 97.78 | 1097.01 | 1097.14 | 1097.27 | 210.02 | 215.54 | 221.06 | 0.06897 |
| 10 | 4 | 15 | 2 | 97.25 | 1104.83 | 1104.95 | 1105.07 | 189.04 | 194.35 | 199.67 | 0.06466 |
| 10 | 5 | 15 | 2 | 97.18 | 1105.93 | 1106.05 | 1106.18 | 237.60 | 242.94 | 248.27 | 0.06466 |
| 10 | 4 | 15 | 2 | 96.67 | 1113.54 | 1113.66 | 1113.78 | 216.45 | 221.91 | 227.36 | 0.06466 |
| 10 | 5 | 15 | 2 | 96.59 | 1114.63 | 1114.75 | 1114.87 | 265.59 | 271.06 | 276.53 | 0.06466 |
| 11 | 5 | 14 | 2 | 94.96 | 1139.51 | 1139.62 | 1139.73 | 350.25 | 355.71 | 1.18 | 0.06897 |
| 11 | 4 | 14 | 2 | 94.51 | 1146.47 | 1146.58 | 1146.69 | 330.15 | 335.72 | 341.28 | 0.06897 |
| 11 | 5 | 14 | 2 | 94.44 | 1147.45 | 1147.56 | 1147.68 | 18.68 | 24.26 | 29.84 | 0.06897 |
| 11 | 4 | 13 | 2 | 94.01 | 1154.30 | 1154.40 | 1154.51 | 359.93 | 5.20 | 10.48 | 0.06404 |
| 11 | 5 | 13 | 2 | 93.94 | 1155.27 | 1155.38 | 1155.48 | 49.06 | 54.34 | 59.63 | 0.06404 |
| 11 | 4 | 13 | 2 | 93.36 | 1164.51 | 1164.62 | 1164.72 | 160.07 | 165.48 | 170.88 | 0.06404 |
| 11 | 5 | 13 | 2 | 93.30 | 1165.46 | 1165.57 | 1165.67 | 209.16 | 214.57 | 219.99 | 0.06404 |
| 11 | 4 | 13 | 2 | 92.89 | 1171.95 | 1172.05 | 1172.16 | 186.41 | 191.90 | 197.40 | 0.06404 |
| 11 | 5 | 13 | 2 | 92.83 | 1172.88 | 1172.98 | 1173.08 | 235.07 | 240.58 | 246.09 | 0.06404 |
| 11 | 4 | 12 | 2 | 92.43 | 1179.31 | 1179.41 | 1179.50 | 215.65 | 220.81 | 225.97 | 0.05911 |
| 11 | 5 | 12 | 2 | 92.38 | 1180.22 | 1180.32 | 1180.42 | 264.31 | 269.48 | 274.66 | 0.05911 |
| 11 | 4 | 12 | 2 | 91.99 | 1186.46 | 1186.56 | 1186.66 | 240.56 | 245.81 | 251.07 | 0.05911 |
| 11 | 5 | 12 | 2 | 91.94 | 1187.36 | 1187.45 | 1187.55 | 289.24 | 294.51 | 299.78 | 0.05911 |
| 11 | 4 | 12 | 2 | 91.57 | 1193.47 | 1193.57 | 1193.66 | 264.49 | 269.84 | 275.19 | 0.05911 |
| 11 | 5 | 12 | 2 | 91.51 | 1194.34 | 1194.44 | 1194.54 | 312.68 | 318.03 | 323.40 | 0.05911 |
| 11 | 4 | 11 | 2 | 91.15 | 1200.36 | 1200.45 | 1200.53 | 288.22 | 293.21 | 298.20 | 0.05419 |
| 12 | 4 | 10 | 2 | 89.43 | 1229.18 | 1229.26 | 1229.34 | 147.99 | 152.71 | 157.43 | 0.02915 |
| 12 | 5 | 10 | 2 | 89.38 | 1229.96 | 1230.04 | 1230.12 | 193.27 | 197.99 | 202.71 | 0.02915 |
| 12 | 4 | 10 | 2 | 89.08 | 1235.24 | 1235.32 | 1235.40 | 142.57 | 147.29 | 152.02 | 0.02915 |
| 12 | 5 | 21 | 2 | 89.02 | 1236.08 | 1236.25 | 1236.41 | 191.70 | 201.62 | 211.54 | 0.06122 |
| 12 | 4 | 21 | 2 | 88.72 | 1241.32 | 1241.49 | 1241.66 | 138.55 | 148.46 | 158.36 | 0.06122 |
| 12 | 5 | 21 | 2 | 88.67 | 1242.06 | 1242.23 | 1242.40 | 181.93 | 191.84 | 201.74 | 0.06122 |
| 12 | 4 | 10 | 2 | 88.37 | 1247.48 | 1247.56 | 1247.64 | 138.38 | 143.07 | 147.78 | 0.02915 |
| 12 | 5 | 10 | 2 | 88.32 | 1248.24 | 1248.33 | 1248.41 | 183.02 | 187.71 | 192.41 | 0.02915 |
| 12 | 4 | 10 | 2 | 88.02 | 1253.57 | 1253.66 | 1253.74 | 132.85 | 137.53 | 142.20 | 0.02915 |
| 12 | 5 | 10 | 2 | 87.97 | 1254.34 | 1254.42 | 1254.50 | 177.27 | 181.94 | 186.61 | 0.02915 |
| 12 | 4 | 21 | 2 | 87.66 | 1259.69 | 1259.86 | 1260.03 | 126.41 | 136.18 | 145.95 | 0.06122 |
| 12 | 5 | 21 | 2 | 87.62 | 1260.44 | 1260.61 | 1260.78 | 169.65 | 179.42 | 189.19 | 0.06122 |
| 12 | 4 | 21 | 2 | 87.32 | 1265.69 | 1265.86 | 1266.03 | 111.74 | 121.46 | 131.19 | 0.06122 |
| 12 | 5 | 21 | 2 | 87.28 | 1266.44 | 1266.61 | 1266.78 | 154.81 | 164.52 | 174.25 | 0.06122 |
| 12 | 4 | 21 | 2 | 86.99 | 1271.61 | 1271.78 | 1271.95 | 91.51 | 101.20 | 110.89 | 0.06122 |
| 12 | 5 | 21 | 2 | 86.95 | 1272.36 | 1272.53 | 1272.69 | 133.94 | 143.63 | 153.32 | 0.06122 |
| 12 | 4 | 21 | 2 | 86.66 | 1277.48 | 1277.65 | 1277.82 | 66.50 | 76.15 | 85.82 | 0.06122 |
| 12 | 5 | 21 | 2 | 86.62 | 1278.21 | 1278.38 | 1278.55 | 108.34 | 118.00 | 127.66 | 0.06122 |
| 12 | 4 | 21 | 2 | 86.34 | 1283.32 | 1283.49 | 1283.66 | 39.64 | 49.28 | 58.92 | 0.06122 |
| 12 | 5 | 21 | 2 | 86.30 | 1284.06 | 1284.23 | 1284.40 | 81.39 | 91.02 | 100.65 | 0.06122 |
| 13 | 4 | 20 | 2 | 84.35 | 1320.07 | 1320.23 | 1320.39 | 317.20 | 326.28 | 335.36 | 0.04796 |
| 13 | 5 | 20 | 2 | 84.31 | 1320.79 | 1320.95 | 1321.11 | 357.61 | 6.69 | 15.77 | 0.04796 |
| 13 | 5 | 7 | 2 | 84.13 | 1324.28 | 1324.34 | 1324.39 | 194.14 | 197.32 | 200.49 | 0.01679 |
| 13 | 4 | 22 | 2 | 83.83 | 1329.83 | 1330.01 | 1330.18 | 145.78 | 155.73 | 165.67 | 0.05276 |
| 13 | 5 | 22 | 2 | 83.79 | 1330.61 | 1330.79 | 1330.97 | 189.63 | 199.57 | 209.51 | 0.05276 |
| 13 | 4 | 20 | 2 | 83.51 | 1336.07 | 1336.23 | 1336.39 | 134.86 | 143.85 | 152.83 | 0.04796 |
| 13 | 5 | 20 | 2 | 83.47 | 1336.79 | 1336.95 | 1337.11 | 174.86 | 183.85 | 192.83 | 0.04796 |
| 13 | 4 | 20 | 2 | 83.21 | 1341.79 | 1341.95 | 1342.11 | 92.58 | 101.51 | 110.43 | 0.04796 |
| 13 | 5 | 20 | 2 | 83.17 | 1342.50 | 1342.66 | 1342.82 | 131.86 | 140.77 | 149.69 | 0.04796 |
| 13 | 4 | 20 | 2 | 82.91 | 1347.50 | 1347.66 | 1347.82 | 47.31 | 56.16 | 65.00 | 0.04796 |
| 13 | 5 | 20 | 2 | 82.88 | 1348.21 | 1348.38 | 1348.54 | 86.65 | 95.48 | 104.31 | 0.04796 |
| 13 | 4 | 20 | 2 | 82.62 | 1353.21 | 1353.38 | 1353.54 | 359.24 | 7.99 | 16.73 | 0.04796 |
| 13 | 5 | 20 | 2 | 82.58 | 1353.93 | 1354.09 | 1354.25 | 38.13 | 46.86 | 55.58 | 0.04796 |
| 13 | 4 | 21 | 2 | 82.32 | 1358.97 | 1359.14 | 1359.31 | 309.38 | 318.44 | 327.49 | 0.05036 |
| 13 | 5 | 20 | 2 | 82.29 | 1359.70 | 1359.86 | 1360.02 | 348.17 | 356.78 | 5.38 | 0.04796 |
| 13 | 4 | 20 | 2 | 82.03 | 1364.75 | 1364.91 | 1365.07 | 255.78 | 264.26 | 272.74 | 0.04796 |
| 13 | 5 | 21 | 2 | 82.00 | 1365.46 | 1365.63 | 1365.80 | 293.51 | 302.41 | 311.30 | 0.05036 |
| 13 | 4 | 21 | 2 | 81.74 | 1370.54 | 1370.71 | 1370.87 | 198.04 | 206.80 | 215.55 | 0.05036 |
| 13 | 5 | 21 | 2 | 81.70 | 1371.26 | 1371.43 | 1371.60 | 235.55 | 244.29 | 253.02 | 0.05036 |
| 13 | 4 | 21 | 2 | 81.45 | 1376.36 | 1376.53 | 1376.70 | 136.82 | 145.43 | 154.04 | 0.05036 |
| 13 | 5 | 21 | 2 | 81.41 | 1377.09 | 1377.26 | 1377.43 | 174.11 | 182.70 | 191.30 | 0.05036 |
| 14 | 4 | 13 | 2 | 79.65 | 1413.03 | 1413.14 | 1413.24 | 131.61 | 136.62 | 141.63 | 0.03030 |
| 14 | 5 | 12 | 2 | 79.61 | 1413.97 | 1414.07 | 1414.16 | 176.29 | 180.91 | 185.53 | 0.02797 |
| 14 | 4 | 22 | 2 | 79.36 | 1419.07 | 1419.25 | 1419.43 | 59.77 | 68.22 | 76.67 | 0.05128 |
| 14 | 5 | 22 | 2 | 79.33 | 1419.84 | 1420.02 | 1420.19 | 96.26 | 104.71 | 113.16 | 0.05128 |
| 14 | 4 | 22 | 2 | 79.07 | 1425.19 | 1425.37 | 1425.55 | 351.05 | 359.48 | 7.91 | 0.05128 |
| 14 | 5 | 22 | 2 | 79.03 | 1425.96 | 1426.14 | 1426.32 | 27.45 | 35.88 | 44.32 | 0.05128 |
| 14 | 4 | 22 | 2 | 78.78 | 1431.40 | 1431.57 | 1431.75 | 285.47 | 293.88 | 302.29 | 0.05128 |
| 14 | 5 | 22 | 2 | 78.74 | 1432.17 | 1432.35 | 1432.52 | 322.17 | 330.57 | 338.97 | 0.05128 |
| 14 | 4 | 22 | 2 | 78.49 | 1437.64 | 1437.81 | 1437.99 | 220.71 | 229.09 | 237.46 | 0.05128 |
| 14 | 5 | 22 | 2 | 78.45 | 1438.42 | 1438.60 | 1438.77 | 257.62 | 265.99 | 274.36 | 0.05128 |
| 14 | 4 | 22 | 2 | 78.19 | 1443.91 | 1444.09 | 1444.27 | 155.91 | 164.23 | 172.54 | 0.05128 |
| 14 | 5 | 23 | 2 | 78.16 | 1444.69 | 1444.88 | 1445.06 | 192.57 | 201.26 | 209.94 | 0.05361 |
| 14 | 4 | 23 | 2 | 77.90 | 1450.22 | 1450.41 | 1450.60 | 90.77 | 99.39 | 108.01 | 0.05361 |
| 14 | 5 | 22 | 2 | 77.86 | 1451.02 | 1451.19 | 1451.37 | 127.48 | 135.71 | 143.94 | 0.05128 |
| 14 | 4 | 23 | 2 | 77.61 | 1456.60 | 1456.79 | 1456.97 | 25.76 | 34.29 | 42.82 | 0.05361 |
| 14 | 5 | 23 | 2 | 77.57 | 1457.40 | 1457.59 | 1457.77 | 62.46 | 70.98 | 79.50 | 0.05361 |
| 14 | 4 | 23 | 2 | 77.31 | 1463.07 | 1463.26 | 1463.44 | 321.40 | 329.83 | 338.26 | 0.05361 |
| 14 | 5 | 23 | 2 | 77.28 | 1463.88 | 1464.06 | 1464.25 | 358.03 | 6.44 | 14.86 | 0.05361 |
| 14 | 4 | 23 | 2 | 77.02 | 1469.58 | 1469.76 | 1469.95 | 255.11 | 263.43 | 271.74 | 0.05361 |
| 14 | 5 | 23 | 2 | 76.98 | 1470.39 | 1470.58 | 1470.76 | 291.58 | 299.88 | 308.18 | 0.05361 |
| 15 | 4 | 16 | 2 | 74.17 | 1535.16 | 1535.29 | 1535.42 | 148.55 | 153.76 | 158.98 | 0.07512 |
| 15 | 5 | 16 | 2 | 74.12 | 1536.27 | 1536.40 | 1536.53 | 193.55 | 198.77 | 203.99 | 0.07512 |
| 15 | 4 | 13 | 2 | 73.88 | 1542.25 | 1542.35 | 1542.46 | 75.30 | 79.54 | 83.77 | 0.06103 |
| 15 | 5 | 12 | 2 | 73.84 | 1543.18 | 1543.28 | 1543.38 | 113.07 | 116.97 | 120.87 | 0.05634 |
| 15 | 4 | 13 | 2 | 73.57 | 1549.63 | 1549.74 | 1549.84 | 11.57 | 15.72 | 19.87 | 0.06103 |
| 15 | 5 | 13 | 2 | 73.53 | 1550.56 | 1550.67 | 1550.77 | 48.28 | 52.42 | 56.57 | 0.06103 |
| 15 | 4 | 13 | 2 | 73.27 | 1557.05 | 1557.15 | 1557.26 | 302.07 | 306.12 | 310.17 | 0.06103 |
| 15 | 5 | 13 | 2 | 73.23 | 1557.97 | 1558.08 | 1558.18 | 337.80 | 341.83 | 345.86 | 0.06103 |
| 15 | 4 | 13 | 2 | 72.96 | 1564.54 | 1564.64 | 1564.75 | 226.24 | 230.14 | 234.04 | 0.06103 |
| 15 | 5 | 13 | 2 | 72.92 | 1565.47 | 1565.58 | 1565.68 | 261.01 | 264.89 | 268.78 | 0.06103 |
| 15 | 4 | 13 | 2 | 72.65 | 1572.12 | 1572.22 | 1572.33 | 144.98 | 148.82 | 152.66 | 0.06103 |
| 15 | 5 | 13 | 2 | 72.61 | 1573.07 | 1573.17 | 1573.28 | 179.79 | 183.63 | 187.47 | 0.06103 |
| 15 | 4 | 13 | 2 | 72.34 | 1579.82 | 1579.92 | 1580.03 | 63.12 | 66.84 | 70.55 | 0.06103 |
| 15 | 5 | 13 | 2 | 72.30 | 1580.78 | 1580.88 | 1580.99 | 97.17 | 100.90 | 104.62 | 0.06103 |
| 15 | 4 | 13 | 2 | 72.03 | 1587.61 | 1587.71 | 1587.82 | 338.52 | 342.19 | 345.86 | 0.06103 |
| 15 | 5 | 13 | 2 | 71.99 | 1588.58 | 1588.69 | 1588.79 | 12.63 | 16.29 | 19.95 | 0.06103 |
| 16 | 4 | 14 | 2 | 70.69 | 1622.29 | 1622.41 | 1622.52 | 118.28 | 122.24 | 126.20 | 0.06731 |
| 16 | 5 | 14 | 2 | 70.65 | 1623.33 | 1623.44 | 1623.55 | 154.51 | 158.48 | 162.45 | 0.06731 |
| 16 | 4 | 15 | 2 | 70.37 | 1630.66 | 1630.79 | 1630.91 | 52.11 | 56.34 | 60.57 | 0.07212 |
| 16 | 5 | 15 | 2 | 70.33 | 1631.72 | 1631.84 | 1631.96 | 89.06 | 93.29 | 97.52 | 0.07212 |
| 16 | 4 | 15 | 2 | 70.06 | 1639.16 | 1639.29 | 1639.41 | 348.54 | 352.69 | 356.85 | 0.07212 |
| 16 | 5 | 15 | 2 | 70.02 | 1640.24 | 1640.36 | 1640.48 | 25.33 | 29.49 | 33.64 | 0.07212 |
| 16 | 4 | 15 | 2 | 69.74 | 1647.74 | 1647.86 | 1647.98 | 283.62 | 287.81 | 292.00 | 0.07212 |
| 16 | 5 | 15 | 2 | 69.70 | 1648.81 | 1648.93 | 1649.05 | 320.83 | 325.04 | 329.24 | 0.07212 |
| 16 | 4 | 15 | 2 | 69.43 | 1656.44 | 1656.56 | 1656.68 | 225.68 | 229.82 | 233.96 | 0.07212 |
| 16 | 5 | 15 | 2 | 69.39 | 1657.53 | 1657.65 | 1657.77 | 262.81 | 266.92 | 271.03 | 0.07212 |
| 16 | 4 | 15 | 2 | 69.11 | 1665.13 | 1665.25 | 1665.37 | 150.51 | 154.43 | 158.36 | 0.07212 |
| 16 | 5 | 15 | 2 | 69.07 | 1666.22 | 1666.34 | 1666.46 | 185.83 | 189.75 | 193.67 | 0.07212 |
| 16 | 4 | 15 | 2 | 68.80 | 1673.94 | 1674.07 | 1674.19 | 76.16 | 80.09 | 84.02 | 0.07212 |
| 16 | 5 | 15 | 2 | 68.76 | 1675.05 | 1675.17 | 1675.29 | 112.04 | 115.97 | 119.90 | 0.07212 |
| 17 | 5 | 16 | 2 | 67.48 | 1711.94 | 1712.06 | 1712.19 | 238.79 | 242.97 | 247.15 | 0.06809 |
| 17 | 4 | 17 | 2 | 67.21 | 1720.11 | 1720.25 | 1720.39 | 142.22 | 146.58 | 150.94 | 0.07234 |
| 17 | 5 | 16 | 2 | 67.17 | 1721.31 | 1721.43 | 1721.56 | 180.12 | 184.21 | 188.29 | 0.06809 |
| 17 | 4 | 16 | 2 | 66.90 | 1729.51 | 1729.64 | 1729.76 | 72.19 | 76.06 | 79.93 | 0.06809 |
| 17 | 5 | 16 | 2 | 66.86 | 1730.68 | 1730.81 | 1730.93 | 107.34 | 111.24 | 115.12 | 0.06809 |
| 17 | 4 | 17 | 2 | 66.58 | 1739.01 | 1739.14 | 1739.28 | 0.35 | 4.61 | 8.87 | 0.07234 |
| 17 | 5 | 17 | 2 | 66.55 | 1740.20 | 1740.34 | 1740.47 | 37.48 | 41.74 | 46.00 | 0.07234 |
| 17 | 4 | 17 | 2 | 66.28 | 1748.53 | 1748.67 | 1748.80 | 294.62 | 298.75 | 302.87 | 0.07234 |
| 17 | 5 | 17 | 2 | 66.24 | 1749.72 | 1749.86 | 1750.00 | 330.36 | 334.43 | 338.51 | 0.07234 |
| 17 | 4 | 17 | 2 | 65.97 | 1758.26 | 1758.39 | 1758.53 | 220.95 | 224.94 | 228.93 | 0.07234 |
| 17 | 5 | 17 | 2 | 65.93 | 1759.46 | 1759.60 | 1759.74 | 256.14 | 260.13 | 264.11 | 0.07234 |
| 17 | 4 | 17 | 2 | 65.66 | 1768.10 | 1768.24 | 1768.38 | 145.33 | 149.24 | 153.14 | 0.07234 |
| 17 | 5 | 17 | 2 | 65.62 | 1769.33 | 1769.47 | 1769.61 | 180.35 | 184.24 | 188.11 | 0.07234 |
| 17 | 4 | 18 | 2 | 65.35 | 1778.05 | 1778.20 | 1778.34 | 67.62 | 71.78 | 75.94 | 0.07660 |
| 18 | 4 | 18 | 2 | 64.38 | 1810.44 | 1810.59 | 1810.73 | 282.55 | 286.66 | 290.77 | 0.07965 |
| 18 | 5 | 18 | 2 | 64.34 | 1811.74 | 1811.88 | 1812.03 | 319.20 | 323.28 | 327.35 | 0.07965 |
| 18 | 4 | 18 | 2 | 64.07 | 1820.79 | 1820.93 | 1821.08 | 213.16 | 217.26 | 221.36 | 0.07965 |
| 18 | 5 | 18 | 2 | 64.04 | 1822.08 | 1822.22 | 1822.37 | 249.62 | 253.72 | 257.83 | 0.07965 |
| 18 | 4 | 19 | 2 | 63.77 | 1831.34 | 1831.50 | 1831.65 | 153.58 | 157.93 | 162.27 | 0.08407 |
| 18 | 5 | 19 | 2 | 63.73 | 1832.66 | 1832.81 | 1832.96 | 190.81 | 195.15 | 199.48 | 0.08407 |
| 18 | 4 | 19 | 2 | 63.47 | 1842.06 | 1842.21 | 1842.37 | 95.90 | 100.19 | 104.47 | 0.08407 |
| 18 | 5 | 19 | 2 | 63.43 | 1843.40 | 1843.55 | 1843.70 | 133.25 | 137.52 | 141.78 | 0.08407 |
| 18 | 4 | 19 | 2 | 63.16 | 1852.98 | 1853.13 | 1853.28 | 35.69 | 39.83 | 43.96 | 0.08407 |
| 18 | 5 | 19 | 2 | 63.13 | 1854.34 | 1854.49 | 1854.64 | 72.50 | 76.66 | 80.82 | 0.08407 |
| 18 | 4 | 20 | 2 | 62.86 | 1864.13 | 1864.29 | 1864.45 | 342.21 | 346.68 | 351.16 | 0.08850 |
| 18 | 5 | 20 | 2 | 62.82 | 1865.52 | 1865.68 | 1865.84 | 20.74 | 25.23 | 29.72 | 0.08850 |
| 22 | 5 | 26 | 2 | 59.25 | 2006.62 | 2006.83 | 2007.04 | 192.71 | 198.24 | 203.76 | 0.11304 |
| 22 | 4 | 23 | 2 | 58.97 | 2018.51 | 2018.70 | 2018.89 | 133.29 | 137.88 | 142.48 | 0.10000 |
| 22 | 5 | 23 | 2 | 58.93 | 2020.08 | 2020.26 | 2020.45 | 171.97 | 176.55 | 181.13 | 0.10000 |
| 22 | 4 | 22 | 2 | 58.68 | 2031.04 | 2031.21 | 2031.39 | 84.51 | 88.96 | 93.42 | 0.09565 |
| 22 | 5 | 23 | 2 | 58.64 | 2032.58 | 2032.76 | 2032.95 | 123.22 | 127.89 | 132.57 | 0.10000 |
| 22 | 4 | 22 | 2 | 58.39 | 2043.54 | 2043.71 | 2043.89 | 42.82 | 47.39 | 51.96 | 0.09565 |
| 22 | 5 | 23 | 2 | 58.35 | 2045.08 | 2045.26 | 2045.45 | 82.21 | 86.92 | 91.62 | 0.10000 |
| 22 | 4 | 22 | 2 | 58.10 | 2056.04 | 2056.21 | 2056.39 | 356.81 | 1.36 | 5.90 | 0.09565 |
| 22 | 5 | 23 | 2 | 58.07 | 2057.58 | 2057.76 | 2057.95 | 36.33 | 41.10 | 45.87 | 0.10000 |
| 22 | 4 | 23 | 2 | 57.82 | 2068.56 | 2068.74 | 2068.93 | 315.36 | 320.03 | 324.70 | 0.10000 |
| 23 | 5 | 23 | 2 | 57.22 | 2095.71 | 2095.90 | 2096.08 | 255.04 | 259.58 | 264.12 | 0.09871 |
| 23 | 4 | 23 | 2 | 56.97 | 2106.88 | 2107.06 | 2107.25 | 159.83 | 164.14 | 168.44 | 0.09871 |
| 23 | 5 | 23 | 2 | 56.94 | 2108.48 | 2108.66 | 2108.85 | 196.85 | 201.14 | 205.44 | 0.09871 |
| 23 | 4 | 23 | 2 | 56.68 | 2120.33 | 2120.51 | 2120.70 | 112.16 | 116.52 | 120.88 | 0.09871 |
| 23 | 5 | 23 | 2 | 56.65 | 2121.93 | 2122.11 | 2122.30 | 149.80 | 154.19 | 158.57 | 0.09871 |
| 23 | 4 | 23 | 2 | 56.41 | 2133.10 | 2133.29 | 2133.47 | 51.90 | 56.22 | 60.52 | 0.09871 |
| 23 | 5 | 23 | 2 | 56.37 | 2134.70 | 2134.88 | 2135.07 | 89.02 | 93.35 | 97.67 | 0.09871 |
| 23 | 4 | 24 | 2 | 56.07 | 2149.08 | 2149.27 | 2149.46 | 62.02 | 66.40 | 70.77 | 0.10300 |
| 23 | 5 | 24 | 2 | 56.03 | 2150.71 | 2150.90 | 2151.09 | 98.64 | 102.94 | 107.25 | 0.10300 |
| 23 | 4 | 24 | 2 | 55.80 | 2162.12 | 2162.32 | 2162.51 | 346.44 | 350.60 | 354.75 | 0.10300 |
| 24 | 4 | 25 | 2 | 54.42 | 2229.27 | 2229.47 | 2229.67 | 288.90 | 292.68 | 296.46 | 0.08651 |
| 24 | 5 | 25 | 2 | 54.39 | 2231.01 | 2231.22 | 2231.42 | 321.64 | 325.44 | 329.24 | 0.08651 |
| 24 | 4 | 26 | 2 | 54.14 | 2243.38 | 2243.59 | 2243.80 | 196.31 | 200.29 | 204.27 | 0.08997 |
| 24 | 5 | 26 | 2 | 54.11 | 2245.17 | 2245.38 | 2245.59 | 230.23 | 234.20 | 238.16 | 0.08997 |
| 24 | 4 | 26 | 2 | 53.86 | 2257.67 | 2257.88 | 2258.09 | 102.55 | 106.46 | 110.38 | 0.08997 |
| 24 | 5 | 26 | 2 | 53.83 | 2259.46 | 2259.66 | 2259.87 | 135.89 | 139.83 | 143.77 | 0.08997 |
| 24 | 4 | 27 | 2 | 53.58 | 2272.33 | 2272.55 | 2272.77 | 17.21 | 21.29 | 25.36 | 0.09343 |
| 24 | 5 | 27 | 2 | 53.54 | 2274.16 | 2274.38 | 2274.60 | 51.47 | 55.55 | 59.65 | 0.09343 |
| 24 | 4 | 27 | 2 | 53.29 | 2287.36 | 2287.58 | 2287.80 | 302.62 | 306.66 | 310.69 | 0.09343 |
| 24 | 5 | 27 | 2 | 53.25 | 2289.24 | 2289.45 | 2289.67 | 337.20 | 341.21 | 345.21 | 0.09343 |
| 24 | 5 | 27 | 2 | 53.07 | 2298.60 | 2298.82 | 2299.03 | 148.02 | 151.97 | 155.93 | 0.09343 |
| 25 | 4 | 29 | 2 | 52.43 | 2332.74 | 2332.98 | 2333.21 | 20.81 | 24.97 | 29.14 | 0.11934 |
| 25 | 5 | 29 | 2 | 52.39 | 2334.72 | 2334.95 | 2335.19 | 55.95 | 60.10 | 64.22 | 0.11934 |
| 25 | 4 | 30 | 2 | 52.12 | 2348.92 | 2349.16 | 2349.40 | 300.84 | 304.92 | 309.00 | 0.12346 |
| 25 | 5 | 30 | 2 | 52.09 | 2350.95 | 2351.19 | 2351.44 | 335.00 | 339.05 | 343.10 | 0.12346 |
| 25 | 4 | 31 | 2 | 51.82 | 2365.44 | 2365.69 | 2365.94 | 212.51 | 216.52 | 220.53 | 0.12757 |
| 25 | 5 | 31 | 2 | 51.78 | 2367.53 | 2367.78 | 2368.03 | 245.91 | 249.89 | 253.87 | 0.12757 |
| 25 | 4 | 32 | 2 | 51.51 | 2382.40 | 2382.66 | 2382.92 | 118.52 | 122.47 | 126.42 | 0.13169 |
| 25 | 5 | 31 | 2 | 51.47 | 2384.56 | 2384.81 | 2385.06 | 151.51 | 155.31 | 159.12 | 0.12757 |
| 26 | 4 | 32 | 2 | 50.90 | 2416.68 | 2416.94 | 2417.20 | 296.94 | 301.06 | 305.19 | 0.12355 |
| 26 | 5 | 32 | 2 | 50.86 | 2418.83 | 2419.09 | 2419.34 | 331.24 | 335.36 | 339.48 | 0.12355 |
| 26 | 4 | 33 | 2 | 50.58 | 2434.35 | 2434.62 | 2434.88 | 220.17 | 224.44 | 228.71 | 0.12741 |
| 26 | 5 | 32 | 2 | 50.54 | 2436.58 | 2436.84 | 2437.09 | 255.85 | 259.98 | 264.11 | 0.12355 |
| 26 | 4 | 32 | 2 | 50.27 | 2452.02 | 2452.28 | 2452.54 | 140.46 | 144.45 | 148.44 | 0.12355 |
| 26 | 5 | 33 | 2 | 50.24 | 2454.21 | 2454.47 | 2454.74 | 174.16 | 178.24 | 182.32 | 0.12741 |
| 26 | 4 | 32 | 2 | 49.97 | 2469.66 | 2469.92 | 2470.18 | 43.88 | 47.65 | 51.42 | 0.12355 |
| 26 | 5 | 33 | 2 | 49.93 | 2471.85 | 2472.12 | 2472.38 | 75.95 | 79.87 | 83.80 | 0.12741 |
| 30 | 4 | 64 | 2 | 47.33 | 2628.83 | 2629.34 | 2629.86 | 23.19 | 31.43 | 39.67 | 0.14066 |
| 30 | 5 | 64 | 2 | 47.30 | 2630.97 | 2631.49 | 2632.00 | 57.39 | 65.58 | 73.77 | 0.14066 |
| 30 | 4 | 64 | 2 | 47.06 | 2645.97 | 2646.49 | 2647.00 | 291.92 | 299.87 | 307.80 | 0.14066 |
| 30 | 5 | 65 | 2 | 47.03 | 2648.11 | 2648.63 | 2649.16 | 324.77 | 332.80 | 340.82 | 0.14286 |
| 30 | 4 | 33 | 2 | 46.80 | 2663.22 | 2663.49 | 2663.75 | 193.89 | 197.89 | 201.88 | 0.07253 |
| 30 | 5 | 66 | 2 | 46.75 | 2665.70 | 2666.24 | 2666.77 | 231.08 | 239.00 | 246.92 | 0.14505 |
| 30 | 4 | 33 | 2 | 46.52 | 2680.75 | 2681.02 | 2681.28 | 91.83 | 95.71 | 99.57 | 0.07253 |
| 30 | 5 | 66 | 2 | 46.48 | 2683.24 | 2683.77 | 2684.30 | 127.83 | 135.48 | 143.14 | 0.14505 |
| 31 | 4 | 68 | 2 | 46.13 | 2706.29 | 2706.83 | 2707.38 | 120.93 | 130.12 | 139.36 | 0.15741 |
| 31 | 5 | 69 | 2 | 46.09 | 2708.55 | 2709.11 | 2709.67 | 159.18 | 168.64 | 178.13 | 0.15972 |
| 31 | 4 | 67 | 2 | 45.86 | 2724.21 | 2724.75 | 2725.29 | 58.80 | 67.52 | 76.22 | 0.15509 |
| 31 | 5 | 66 | 2 | 45.82 | 2726.42 | 2726.95 | 2727.49 | 94.36 | 102.89 | 111.40 | 0.15278 |
| 31 | 4 | 81 | 2 | 45.46 | 2750.65 | 2751.31 | 2751.96 | 116.35 | 126.75 | 137.15 | 0.18750 |
| 31 | 5 | 81 | 2 | 45.42 | 2753.32 | 2753.97 | 2754.62 | 158.79 | 169.24 | 179.70 | 0.18750 |
| 32 | 4 | 63 | 2 | 44.71 | 2801.17 | 2801.68 | 2802.19 | 243.34 | 252.15 | 260.95 | 0.12600 |
| 32 | 5 | 63 | 2 | 44.68 | 2803.25 | 2803.76 | 2804.27 | 279.35 | 288.12 | 296.89 | 0.12600 |
| 32 | 4 | 63 | 2 | 44.47 | 2817.84 | 2818.35 | 2818.85 | 166.10 | 174.40 | 182.69 | 0.12600 |
| 32 | 5 | 63 | 2 | 44.44 | 2819.92 | 2820.43 | 2820.94 | 200.02 | 208.32 | 216.62 | 0.12600 |
| 32 | 4 | 61 | 2 | 44.24 | 2834.05 | 2834.55 | 2835.04 | 74.50 | 82.68 | 90.85 | 0.12200 |
| 32 | 5 | 61 | 2 | 44.21 | 2836.09 | 2836.58 | 2837.07 | 108.22 | 116.41 | 124.58 | 0.12200 |
| 32 | 4 | 63 | 2 | 44.00 | 2850.84 | 2851.34 | 2851.85 | 349.15 | 357.21 | 5.26 | 0.12600 |
| 32 | 5 | 63 | 2 | 43.97 | 2852.92 | 2853.42 | 2853.93 | 22.09 | 30.11 | 38.11 | 0.12600 |
| 33 | 5 | 84 | 2 | 43.50 | 2886.10 | 2886.77 | 2887.45 | 179.05 | 189.56 | 200.09 | 0.17797 |
| 33 | 4 | 95 | 2 | 43.19 | 2907.89 | 2908.65 | 2909.42 | 160.16 | 172.47 | 184.81 | 0.20127 |
| 33 | 5 | 96 | 2 | 43.15 | 2911.01 | 2911.79 | 2912.56 | 210.64 | 223.20 | 235.82 | 0.20339 |
| 33 | 4 | 67 | 2 | 42.92 | 2927.73 | 2928.27 | 2928.81 | 129.69 | 138.82 | 147.94 | 0.14195 |
| 33 | 5 | 66 | 2 | 42.89 | 2929.94 | 2930.47 | 2931.00 | 167.00 | 175.96 | 184.90 | 0.13983 |
| 33 | 4 | 64 | 2 | 42.68 | 2945.23 | 2945.75 | 2946.26 | 50.83 | 58.72 | 66.61 | 0.13559 |
| 34 | 4 | 67 | 2 | 42.20 | 2980.49 | 2981.03 | 2981.57 | 213.27 | 221.34 | 229.44 | 0.12984 |
| 34 | 5 | 67 | 2 | 42.17 | 2982.70 | 2983.24 | 2983.78 | 246.51 | 254.73 | 263.01 | 0.12984 |
| 34 | 4 | 67 | 2 | 41.96 | 2998.44 | 2998.98 | 2999.52 | 139.19 | 148.16 | 157.15 | 0.12984 |
| 34 | 5 | 66 | 2 | 41.93 | 3000.65 | 3001.18 | 3001.71 | 175.97 | 184.87 | 193.77 | 0.12791 |
| 34 | 4 | 65 | 2 | 41.73 | 3015.85 | 3016.38 | 3016.90 | 66.58 | 75.10 | 83.63 | 0.12597 |
| 34 | 5 | 65 | 2 | 41.70 | 3018.00 | 3018.52 | 3019.05 | 101.45 | 109.93 | 118.43 | 0.12597 |
| 34 | 4 | 60 | 2 | 41.47 | 3034.97 | 3035.46 | 3035.94 | 13.03 | 20.68 | 28.32 | 0.11628 |
| 34 | 5 | 59 | 2 | 41.45 | 3036.96 | 3037.43 | 3037.91 | 44.38 | 51.90 | 59.42 | 0.11434 |
| 35 | 4 | 62 | 2 | 40.35 | 3122.10 | 3122.60 | 3123.10 | 268.96 | 276.30 | 283.62 | 0.12475 |
| 35 | 5 | 61 | 2 | 40.32 | 3124.15 | 3124.64 | 3125.14 | 299.01 | 306.23 | 313.46 | 0.12274 |
| 35 | 4 | 62 | 2 | 40.14 | 3138.58 | 3139.08 | 3139.58 | 151.30 | 158.72 | 166.15 | 0.12475 |
| 35 | 5 | 62 | 2 | 40.11 | 3140.63 | 3141.13 | 3141.63 | 181.75 | 189.07 | 196.39 | 0.12475 |
| 35 | 4 | 62 | 2 | 39.93 | 3155.16 | 3155.66 | 3156.16 | 15.80 | 22.08 | 28.39 | 0.12475 |
| 35 | 5 | 62 | 2 | 39.90 | 3157.23 | 3157.73 | 3158.23 | 42.02 | 48.42 | 54.85 | 0.12475 |
| 35 | 4 | 63 | 2 | 39.72 | 3171.83 | 3172.33 | 3172.84 | 242.92 | 249.98 | 257.02 | 0.12676 |
| 35 | 5 | 63 | 2 | 39.69 | 3173.91 | 3174.42 | 3174.93 | 271.79 | 278.76 | 285.72 | 0.12676 |
| 36 | 4 | 63 | 2 | 39.30 | 3205.52 | 3206.02 | 3206.53 | 343.75 | 350.88 | 358.01 | 0.12281 |
| 36 | 5 | 64 | 2 | 39.28 | 3207.61 | 3208.13 | 3208.65 | 13.19 | 20.43 | 27.68 | 0.12476 |
| 36 | 4 | 64 | 2 | 39.09 | 3222.48 | 3223.00 | 3223.52 | 220.69 | 227.88 | 235.07 | 0.12476 |
| 36 | 5 | 64 | 2 | 39.07 | 3224.60 | 3225.12 | 3225.64 | 250.25 | 257.45 | 264.65 | 0.12476 |
| 36 | 4 | 64 | 2 | 38.88 | 3239.59 | 3240.10 | 3240.62 | 99.35 | 106.48 | 113.62 | 0.12476 |
| 36 | 5 | 64 | 2 | 38.86 | 3241.72 | 3242.24 | 3242.76 | 128.91 | 136.05 | 143.17 | 0.12476 |
| 36 | 4 | 65 | 2 | 38.68 | 3256.79 | 3257.31 | 3257.84 | 327.42 | 333.86 | 340.29 | 0.12671 |
| 36 | 5 | 65 | 2 | 38.65 | 3258.93 | 3259.46 | 3259.98 | 353.72 | 0.14 | 6.54 | 0.12671 |
| 37 | 4 | 66 | 2 | 38.05 | 3309.03 | 3309.56 | 3310.09 | 328.33 | 335.55 | 342.76 | 0.12406 |
| 37 | 5 | 66 | 2 | 38.03 | 3311.21 | 3311.75 | 3312.28 | 357.89 | 5.05 | 12.19 | 0.12406 |
| 37 | 4 | 66 | 2 | 37.84 | 3326.66 | 3327.19 | 3327.72 | 203.08 | 210.11 | 217.13 | 0.12406 |
| 37 | 5 | 66 | 2 | 37.82 | 3328.86 | 3329.39 | 3329.92 | 232.20 | 239.27 | 246.36 | 0.12406 |
| 37 | 4 | 67 | 2 | 37.64 | 3344.38 | 3344.92 | 3345.46 | 81.79 | 89.25 | 96.72 | 0.12594 |
| 37 | 5 | 67 | 2 | 37.61 | 3346.60 | 3347.14 | 3347.68 | 112.48 | 119.99 | 127.50 | 0.12594 |
| 37 | 4 | 67 | 2 | 37.43 | 3362.23 | 3362.77 | 3363.31 | 339.01 | 347.32 | 355.66 | 0.12594 |
| 37 | 5 | 67 | 2 | 37.40 | 3364.46 | 3365.00 | 3365.54 | 13.38 | 21.75 | 30.12 | 0.12594 |
| 38 | 4 | 68 | 2 | 37.12 | 3388.90 | 3389.44 | 3389.99 | 31.50 | 40.01 | 48.52 | 0.12409 |
| 38 | 5 | 68 | 2 | 37.09 | 3391.15 | 3391.69 | 3392.24 | 66.36 | 74.82 | 83.29 | 0.12409 |
| 38 | 4 | 68 | 2 | 36.91 | 3407.02 | 3407.57 | 3408.12 | 307.39 | 315.30 | 323.18 | 0.12409 |
| 38 | 5 | 68 | 2 | 36.88 | 3409.29 | 3409.84 | 3410.39 | 339.92 | 347.73 | 355.52 | 0.12409 |
| 38 | 4 | 69 | 2 | 36.70 | 3425.26 | 3425.82 | 3426.38 | 197.84 | 205.15 | 212.48 | 0.12591 |
| 38 | 5 | 69 | 2 | 36.68 | 3427.55 | 3428.10 | 3428.66 | 227.88 | 235.22 | 242.57 | 0.12591 |
| 38 | 4 | 69 | 2 | 36.49 | 3443.62 | 3444.17 | 3444.73 | 80.73 | 88.04 | 95.33 | 0.12591 |
| 38 | 5 | 69 | 2 | 36.47 | 3445.92 | 3446.47 | 3447.03 | 110.80 | 118.04 | 125.28 | 0.12591 |
| 39 | 4 | 70 | 2 | 36.08 | 3480.65 | 3481.21 | 3481.78 | 224.18 | 232.42 | 240.67 | 0.12367 |
| 39 | 5 | 70 | 2 | 36.05 | 3482.96 | 3483.53 | 3484.09 | 258.02 | 266.29 | 274.58 | 0.12367 |
| 39 | 4 | 71 | 2 | 35.89 | 3497.67 | 3498.25 | 3498.82 | 118.31 | 127.05 | 135.80 | 0.12544 |
| 39 | 5 | 70 | 2 | 35.86 | 3500.01 | 3500.58 | 3501.14 | 154.09 | 162.72 | 171.34 | 0.12367 |
| 39 | 4 | 71 | 2 | 35.68 | 3516.48 | 3517.05 | 3517.62 | 37.66 | 45.73 | 53.74 | 0.12544 |
| 39 | 5 | 71 | 2 | 35.66 | 3518.83 | 3519.40 | 3519.97 | 70.43 | 78.31 | 86.14 | 0.12544 |
| 39 | 4 | 72 | 2 | 35.43 | 3539.52 | 3540.10 | 3540.68 | 336.71 | 343.77 | 350.84 | 0.12721 |
| 39 | 5 | 71 | 2 | 35.40 | 3541.89 | 3542.46 | 3543.03 | 5.55 | 12.54 | 19.56 | 0.12544 |
| 40 | 4 | 72 | 2 | 35.01 | 3577.77 | 3578.35 | 3578.93 | 112.07 | 119.53 | 126.98 | 0.12308 |
| 40 | 5 | 73 | 2 | 34.99 | 3580.16 | 3580.75 | 3581.34 | 142.81 | 150.37 | 157.94 | 0.12479 |
| 40 | 4 | 73 | 2 | 34.80 | 3597.06 | 3597.65 | 3598.24 | 14.00 | 22.65 | 31.32 | 0.12479 |
| 40 | 5 | 73 | 2 | 34.78 | 3599.48 | 3600.06 | 3600.65 | 49.50 | 58.19 | 66.89 | 0.12479 |
| 40 | 4 | 73 | 2 | 34.60 | 3616.50 | 3617.09 | 3617.68 | 303.76 | 312.11 | 320.44 | 0.12479 |
| 40 | 5 | 73 | 2 | 34.57 | 3618.93 | 3619.51 | 3620.10 | 338.09 | 346.38 | 354.66 | 0.12479 |
| 40 | 4 | 74 | 2 | 34.39 | 3636.05 | 3636.64 | 3637.24 | 228.97 | 237.86 | 246.74 | 0.12650 |
| 40 | 5 | 74 | 2 | 34.36 | 3638.49 | 3639.09 | 3639.68 | 265.28 | 274.10 | 282.93 | 0.12650 |
| 41 | 4 | 74 | 2 | 34.02 | 3671.42 | 3672.02 | 3672.61 | 17.18 | 25.86 | 34.48 | 0.12292 |
| 41 | 5 | 75 | 2 | 33.99 | 3673.88 | 3674.49 | 3675.09 | 52.62 | 61.23 | 69.77 | 0.12458 |
| 41 | 4 | 75 | 2 | 33.81 | 3691.32 | 3691.92 | 3692.52 | 291.57 | 300.09 | 308.62 | 0.12458 |
| 41 | 5 | 75 | 2 | 33.78 | 3693.80 | 3694.40 | 3695.01 | 326.62 | 335.20 | 343.77 | 0.12458 |
| 41 | 4 | 76 | 2 | 33.60 | 3711.32 | 3711.94 | 3712.55 | 215.18 | 223.83 | 232.46 | 0.12625 |
| 41 | 5 | 75 | 2 | 33.57 | 3713.82 | 3714.43 | 3715.03 | 250.33 | 258.80 | 267.25 | 0.12458 |
| 41 | 4 | 76 | 2 | 33.40 | 3730.65 | 3731.26 | 3731.88 | 125.07 | 133.84 | 142.63 | 0.12625 |
| 41 | 5 | 76 | 2 | 33.37 | 3733.17 | 3733.78 | 3734.39 | 161.13 | 169.90 | 178.68 | 0.12625 |
| 42 | 4 | 78 | 2 | 32.57 | 3812.47 | 3813.10 | 3813.73 | 143.42 | 150.27 | 157.09 | 0.16561 |
| 42 | 5 | 78 | 2 | 32.54 | 3815.05 | 3815.68 | 3816.31 | 171.45 | 178.30 | 185.16 | 0.16561 |
| 42 | 4 | 78 | 2 | 32.36 | 3833.25 | 3833.88 | 3834.51 | 36.09 | 44.48 | 52.82 | 0.16561 |
| 42 | 5 | 79 | 2 | 32.33 | 3835.84 | 3836.47 | 3837.11 | 70.27 | 78.55 | 86.78 | 0.16773 |
| 42 | 4 | 79 | 2 | 32.15 | 3854.15 | 3854.79 | 3855.42 | 304.07 | 312.23 | 320.38 | 0.16773 |
| 42 | 5 | 79 | 2 | 32.13 | 3856.76 | 3857.40 | 3858.04 | 337.48 | 345.59 | 353.68 | 0.16773 |
| 43 | 4 | 80 | 2 | 31.74 | 3896.35 | 3897.00 | 3897.64 | 44.07 | 50.28 | 56.59 | 0.16563 |
| 43 | 5 | 80 | 2 | 31.71 | 3898.99 | 3899.63 | 3900.28 | 70.16 | 76.85 | 83.65 | 0.16563 |
| 43 | 4 | 80 | 2 | 31.53 | 3917.64 | 3918.29 | 3918.93 | 297.85 | 307.73 | 317.78 | 0.16563 |
| 43 | 5 | 81 | 2 | 31.50 | 3920.29 | 3920.95 | 3921.60 | 339.58 | 350.17 | 0.85 | 0.16770 |
| 43 | 4 | 81 | 2 | 31.32 | 3939.07 | 3939.73 | 3940.38 | 259.76 | 268.75 | 277.66 | 0.16770 |
| 43 | 5 | 81 | 2 | 31.29 | 3941.74 | 3942.39 | 3943.05 | 295.95 | 304.62 | 313.28 | 0.16770 |
| 44 | 4 | 82 | 2 | 30.90 | 3982.32 | 3982.98 | 3983.64 | 174.17 | 184.11 | 194.10 | 0.16532 |
| 44 | 5 | 82 | 2 | 30.88 | 3985.03 | 3985.69 | 3986.35 | 215.26 | 225.46 | 235.77 | 0.16532 |
| 44 | 4 | 83 | 2 | 30.71 | 4002.64 | 4003.31 | 4003.98 | 154.68 | 166.56 | 178.47 | 0.16734 |
| 44 | 5 | 83 | 2 | 30.68 | 4005.36 | 4006.03 | 4006.70 | 203.08 | 214.94 | 226.71 | 0.16734 |
| 44 | 4 | 83 | 2 | 30.50 | 4024.60 | 4025.27 | 4025.94 | 169.43 | 180.19 | 190.92 | 0.16734 |
| 44 | 5 | 83 | 2 | 30.47 | 4027.34 | 4028.01 | 4028.68 | 213.27 | 223.87 | 234.43 | 0.16734 |
| 45 | 4 | 84 | 2 | 30.08 | 4068.92 | 4069.60 | 4070.28 | 94.56 | 103.80 | 113.04 | 0.16535 |
| 45 | 5 | 84 | 2 | 30.06 | 4071.71 | 4072.38 | 4073.06 | 132.51 | 141.75 | 150.99 | 0.16535 |
| 45 | 4 | 85 | 2 | 29.88 | 4091.29 | 4091.97 | 4092.66 | 43.54 | 53.04 | 62.49 | 0.16732 |
| 45 | 5 | 85 | 2 | 29.85 | 4094.08 | 4094.76 | 4095.45 | 81.94 | 91.30 | 100.63 | 0.16732 |
| 45 | 4 | 85 | 2 | 29.67 | 4113.79 | 4114.48 | 4115.16 | 345.78 | 354.81 | 3.81 | 0.16732 |
| 45 | 5 | 85 | 2 | 29.64 | 4116.61 | 4117.29 | 4117.98 | 22.69 | 31.60 | 40.47 | 0.16732 |
| 46 | 4 | 86 | 2 | 29.25 | 4159.25 | 4159.94 | 4160.64 | 150.42 | 156.81 | 163.22 | 0.16538 |
| 46 | 5 | 86 | 2 | 29.22 | 4162.10 | 4162.79 | 4163.48 | 176.73 | 183.19 | 189.66 | 0.16538 |
| 46 | 4 | 87 | 2 | 29.04 | 4182.18 | 4182.88 | 4183.58 | 30.07 | 38.35 | 46.67 | 0.16731 |
| 46 | 5 | 87 | 2 | 29.02 | 4185.04 | 4185.74 | 4186.44 | 64.01 | 72.36 | 80.72 | 0.16731 |
| 46 | 4 | 87 | 2 | 28.83 | 4205.27 | 4205.97 | 4206.67 | 315.01 | 323.76 | 332.53 | 0.16731 |
| 46 | 5 | 87 | 2 | 28.81 | 4208.14 | 4208.85 | 4209.55 | 351.01 | 359.84 | 8.67 | 0.16731 |
| 47 | 4 | 88 | 2 | 28.41 | 4251.86 | 4252.57 | 4253.28 | 184.14 | 192.76 | 201.35 | 0.16449 |
| 47 | 5 | 89 | 2 | 28.39 | 4254.77 | 4255.49 | 4256.21 | 219.33 | 227.93 | 236.49 | 0.16636 |
| 47 | 4 | 89 | 2 | 28.21 | 4275.39 | 4276.11 | 4276.83 | 97.64 | 106.03 | 114.43 | 0.16636 |
| 47 | 5 | 89 | 2 | 28.18 | 4278.33 | 4279.04 | 4279.76 | 132.08 | 140.50 | 148.96 | 0.16636 |
| 47 | 4 | 90 | 2 | 28.00 | 4299.05 | 4299.78 | 4300.50 | 19.58 | 27.97 | 36.22 | 0.16822 |
| 47 | 5 | 90 | 2 | 27.97 | 4302.00 | 4302.73 | 4303.45 | 52.87 | 60.72 | 68.42 | 0.16822 |
| 48 | 4 | 91 | 2 | 27.58 | 4346.82 | 4347.55 | 4348.29 | 135.68 | 143.82 | 152.00 | 0.16606 |
| 48 | 5 | 91 | 2 | 27.55 | 4349.81 | 4350.54 | 4351.28 | 168.97 | 177.17 | 185.35 | 0.16606 |
| 48 | 4 | 91 | 2 | 27.37 | 4370.93 | 4371.66 | 4372.40 | 35.01 | 42.62 | 50.21 | 0.16606 |
| 48 | 5 | 91 | 2 | 27.34 | 4373.94 | 4374.67 | 4375.40 | 66.06 | 73.56 | 81.06 | 0.16606 |
| 48 | 4 | 92 | 2 | 27.16 | 4395.19 | 4395.93 | 4396.67 | 280.91 | 288.77 | 296.65 | 0.16788 |
| 48 | 5 | 92 | 2 | 27.13 | 4398.22 | 4398.96 | 4399.70 | 313.14 | 321.09 | 329.07 | 0.16788 |
| 50 | 4 | 128 | 2 | 23.31 | 4866.34 | 4867.37 | 4868.40 | 257.99 | 269.31 | 280.71 | 0.24903 |
| 50 | 5 | 128 | 2 | 23.28 | 4870.54 | 4871.57 | 4872.60 | 304.49 | 316.03 | 327.57 | 0.24903 |
| 50 | 4 | 129 | 2 | 23.05 | 4899.97 | 4901.01 | 4902.05 | 247.38 | 257.82 | 268.32 | 0.25097 |
| 50 | 5 | 129 | 2 | 23.02 | 4904.19 | 4905.23 | 4906.27 | 290.02 | 300.62 | 311.26 | 0.25097 |
| 51 | 4 | 132 | 2 | 20.74 | 5208.94 | 5210.01 | 5211.07 | 178.70 | 189.62 | 200.49 | 0.24906 |
| 51 | 5 | 132 | 2 | 20.71 | 5213.27 | 5214.33 | 5215.39 | 222.71 | 233.43 | 244.12 | 0.24906 |
| 51 | 4 | 133 | 2 | 20.48 | 5243.80 | 5244.88 | 5245.95 | 152.00 | 161.42 | 170.84 | 0.25094 |
| 51 | 5 | 133 | 2 | 20.45 | 5248.14 | 5249.21 | 5250.29 | 190.03 | 199.43 | 208.84 | 0.25094 |
| 54 | 4 | 131 | 2 | 17.08 | 5736.83 | 5737.88 | 5738.94 | 303.53 | 313.52 | 323.53 | 0.25586 |
| 54 | 5 | 131 | 2 | 17.05 | 5741.12 | 5742.17 | 5743.23 | 344.12 | 354.10 | 4.09 | 0.25586 |
| 54 | 4 | 125 | 2 | 16.86 | 5769.73 | 5770.73 | 5771.74 | 251.06 | 260.46 | 269.85 | 0.24414 |
| 54 | 5 | 125 | 2 | 16.83 | 5773.81 | 5774.82 | 5775.82 | 289.05 | 298.43 | 307.81 | 0.24414 |
| 55 | 4 | 117 | 2 | 14.87 | 6078.19 | 6079.14 | 6080.08 | 294.88 | 303.61 | 312.35 | 0.24841 |
| 55 | 5 | 118 | 2 | 14.85 | 6082.04 | 6082.99 | 6083.94 | 330.58 | 339.46 | 348.33 | 0.25053 |
| 55 | 4 | 118 | 2 | 14.68 | 6109.21 | 6110.16 | 6111.11 | 225.15 | 234.27 | 243.42 | 0.25053 |
| 55 | 5 | 118 | 2 | 14.65 | 6113.06 | 6114.01 | 6114.96 | 262.35 | 271.65 | 280.98 | 0.25053 |
| 56 | 4 | 124 | 2 | 12.72 | 6428.57 | 6429.57 | 6430.57 | 272.99 | 282.11 | 291.23 | 0.24900 |
| 56 | 5 | 124 | 2 | 12.70 | 6432.64 | 6433.64 | 6434.64 | 310.05 | 319.09 | 328.13 | 0.24900 |
| 56 | 4 | 125 | 2 | 12.53 | 6461.37 | 6462.37 | 6463.38 | 200.74 | 209.05 | 217.31 | 0.25100 |
| 56 | 5 | 125 | 2 | 12.50 | 6465.46 | 6466.47 | 6467.48 | 234.10 | 242.13 | 250.09 | 0.25100 |
| 57 | 4 | 131 | 2 | 10.76 | 6764.38 | 6765.44 | 6766.49 | 161.61 | 170.54 | 179.46 | 0.25000 |
| 57 | 5 | 131 | 2 | 10.73 | 6768.66 | 6769.72 | 6770.78 | 197.68 | 206.47 | 215.23 | 0.25000 |
| 57 | 4 | 131 | 2 | 10.56 | 6798.75 | 6799.81 | 6800.86 | 75.79 | 83.69 | 91.55 | 0.25000 |
| 57 | 5 | 131 | 2 | 10.54 | 6803.04 | 6804.10 | 6805.15 | 107.57 | 115.29 | 122.99 | 0.25000 |
| 58 | 4 | 136 | 2 | 8.99 | 7079.76 | 7080.86 | 7081.96 | 198.31 | 206.37 | 214.43 | 0.25000 |
| 58 | 5 | 136 | 2 | 8.96 | 7084.20 | 7085.29 | 7086.39 | 230.80 | 238.80 | 246.79 | 0.25000 |
| 58 | 4 | 136 | 2 | 8.79 | 7115.64 | 7116.74 | 7117.83 | 100.61 | 109.41 | 118.28 | 0.25000 |
| 58 | 5 | 136 | 2 | 8.76 | 7120.10 | 7121.20 | 7122.29 | 136.77 | 145.77 | 154.79 | 0.25000 |
| 59 | 4 | 141 | 2 | 7.21 | 7407.85 | 7408.99 | 7410.13 | 117.70 | 125.77 | 133.86 | 0.24912 |
| 59 | 5 | 141 | 2 | 7.19 | 7412.48 | 7413.61 | 7414.75 | 150.51 | 158.60 | 166.67 | 0.24912 |
| 59 | 4 | 142 | 2 | 7.01 | 7445.01 | 7446.15 | 7447.30 | 28.87 | 37.41 | 45.94 | 0.25088 |
| 59 | 5 | 142 | 2 | 6.99 | 7449.64 | 7450.79 | 7451.93 | 63.35 | 71.85 | 80.42 | 0.25088 |
| 60 | 4 | 146 | 2 | 5.44 | 7747.67 | 7748.84 | 7750.02 | 107.25 | 115.09 | 122.93 | 0.25000 |
| 60 | 5 | 146 | 2 | 5.41 | 7752.43 | 7753.61 | 7754.79 | 138.93 | 146.71 | 154.49 | 0.25000 |
| 60 | 4 | 146 | 2 | 5.24 | 7785.99 | 7787.16 | 7788.34 | 4.66 | 12.67 | 20.67 | 0.25000 |
| 60 | 5 | 146 | 2 | 5.22 | 7790.77 | 7791.94 | 7793.12 | 37.07 | 44.99 | 52.89 | 0.25000 |
| 59 | 8 | 141 | 2 | 7.31 | 7389.32 | 7390.46 | 7391.60 | 343.73 | 352.02 | 0.36 | 0.50000 |
| 59 | 8 | 141 | 2 | 7.11 | 7426.32 | 7427.46 | 7428.60 | 249.96 | 258.29 | 266.64 | 0.50000 |
| 60 | 8 | 146 | 2 | 5.54 | 7728.50 | 7729.67 | 7730.85 | 335.54 | 343.78 | 352.03 | 0.50000 |
| 60 | 8 | 146 | 2 | 5.34 | 7766.73 | 7767.91 | 7769.08 | 233.81 | 241.71 | 249.66 | 0.50000 |

- Fringe pattern at three temperatures

The table below contains the fringe pattern recorded from the three descent calibrations of the IR spectrometer.

| Pixel | T=264.322 | T=259.399 | T=251.661 |
| --- | --- | --- | --- |
| 7 | 0.985 | 0.984 | 0.982 |
| 8 | 0.990 | 0.991 | 0.991 |
| 9 | 0.996 | 1.000 | 1.002 |
| 10 | 1.003 | 1.009 | 1.013 |
| 11 | 1.012 | 1.014 | 1.017 |
| 12 | 1.022 | 1.016 | 1.013 |
| 13 | 1.016 | 1.012 | 1.010 |
| 14 | 1.011 | 1.006 | 1.003 |
| 15 | 1.015 | 1.011 | 1.009 |
| 16 | 0.997 | 0.997 | 0.998 |
| 17 | 0.995 | 0.993 | 0.992 |
| 18 | 0.976 | 0.977 | 0.979 |
| 19 | 0.986 | 0.987 | 0.989 |
| 20 | 0.991 | 0.995 | 0.998 |
| 21 | 1.000 | 1.002 | 1.001 |
| 22 | 0.999 | 1.000 | 0.996 |
| 23 | 0.982 | 0.989 | 0.993 |
| 24 | 1.000 | 1.000 | 1.001 |
| 25 | 1.032 | 1.022 | 1.012 |
| 26 | 1.009 | 1.008 | 1.007 |
| 27 | 0.997 | 0.998 | 0.999 |
| 28 | 0.982 | 0.985 | 0.987 |
| 29 | 0.987 | 0.988 | 0.988 |
| 30 | 1.007 | 1.004 | 1.001 |
| 31 | 1.010 | 1.010 | 1.010 |
| 32 | 1.000 | 1.003 | 1.005 |
| 33 | 0.980 | 0.988 | 0.993 |
| 34 | 0.975 | 0.981 | 0.986 |
| 35 | 0.992 | 0.992 | 0.994 |
| 36 | 1.013 | 1.009 | 1.009 |
| 37 | 1.022 | 1.018 | 1.018 |
| 38 | 1.009 | 1.010 | 1.011 |
| 39 | 1.000 | 0.999 | 0.996 |
| 40 | 0.974 | 0.978 | 0.982 |
| 41 | 0.980 | 0.981 | 0.984 |
| 42 | 1.014 | 1.007 | 1.004 |
| 43 | 1.026 | 1.020 | 1.017 |
| 44 | 1.021 | 1.019 | 1.017 |
| 45 | 0.999 | 1.001 | 1.005 |
| 46 | 0.982 | 0.988 | 0.992 |
| 47 | 0.973 | 0.977 | 0.981 |
| 48 | 0.988 | 0.987 | 0.988 |
| 49 | 1.021 | 1.015 | 1.011 |
| 50 | 1.037 | 1.032 | 1.027 |
| 51 | 1.034 | 1.031 | 1.028 |
| 52 | 1.015 | 1.017 | 1.015 |
| 53 | 0.990 | 0.997 | 0.999 |
| 54 | 0.975 | 0.983 | 0.986 |
| 55 | 1.000 | 0.995 | 0.986 |
| 56 | 1.002 | 0.997 | 0.994 |
| 57 | 1.019 | 1.013 | 1.010 |
| 58 | 1.042 | 1.033 | 1.025 |
| 59 | 1.029 | 1.025 | 1.022 |
| 60 | 1.006 | 1.007 | 1.008 |
| 61 | 0.977 | 0.985 | 0.990 |
| 62 | 0.971 | 0.977 | 0.980 |
| 63 | 0.951 | 0.960 | 0.970 |
| 64 | 0.984 | 0.982 | 0.982 |
| 65 | 0.987 | 0.987 | 0.988 |
| 66 | 1.016 | 1.009 | 1.005 |
| 67 | 1.018 | 1.014 | 1.011 |
| 68 | 1.013 | 1.011 | 1.010 |
| 69 | 0.997 | 0.999 | 1.000 |
| 70 | 0.982 | 0.986 | 0.990 |
| 71 | 0.972 | 0.977 | 0.981 |
| 72 | 0.973 | 0.977 | 0.980 |
| 73 | 0.981 | 0.982 | 0.983 |
| 74 | 0.992 | 0.990 | 0.990 |
| 75 | 1.002 | 0.999 | 0.998 |
| 76 | 1.017 | 1.013 | 1.009 |
| 77 | 1.015 | 1.012 | 1.009 |
| 78 | 1.014 | 1.012 | 1.009 |
| 79 | 0.997 | 0.999 | 0.998 |
| 80 | 0.980 | 0.986 | 0.987 |
| 81 | 0.966 | 0.973 | 0.977 |
| 82 | 0.967 | 0.972 | 0.977 |
| 83 | 0.983 | 0.988 | 0.991 |
| 84 | 1.012 | 1.014 | 1.014 |
| 85 | 1.031 | 1.032 | 1.032 |
| 86 | 1.040 | 1.034 | 1.030 |
| 87 | 1.026 | 1.026 | 1.026 |
| 88 | 1.045 | 1.039 | 1.033 |
| 89 | 1.045 | 1.045 | 1.043 |
| 90 | 1.027 | 1.028 | 1.031 |
| 91 | 1.026 | 1.029 | 1.030 |
| 92 | 0.999 | 1.005 | 1.008 |
| 93 | 0.980 | 0.987 | 0.992 |
| 94 | 0.969 | 0.972 | 0.975 |
| 95 | 0.972 | 0.971 | 0.972 |
| 96 | 0.992 | 0.984 | 0.973 |
| 97 | 0.964 | 0.964 | 0.966 |
| 98 | 0.969 | 0.965 | 0.964 |
| 99 | 0.971 | 0.967 | 0.967 |
| 100 | 0.985 | 0.980 | 0.978 |
| 101 | 1.011 | 1.004 | 0.998 |
| 102 | 1.017 | 1.008 | 1.004 |
| 103 | 1.022 | 1.017 | 1.014 |
| 104 | 1.016 | 1.013 | 1.011 |
| 105 | 1.014 | 1.014 | 1.011 |
| 106 | 0.985 | 0.994 | 1.001 |
| 107 | 1.004 | 1.007 | 1.010 |
| 108 | 1.014 | 1.011 | 1.013 |
| 109 | 1.028 | 1.021 | 1.018 |
| 110 | 1.028 | 1.022 | 1.019 |
| 111 | 0.904 | 0.954 | 0.979 |
| 112 | 1.009 | 1.007 | 1.007 |
| 113 | 1.004 | 1.000 | 1.000 |
| 114 | 0.998 | 0.993 | 0.991 |
| 115 | 1.006 | 1.000 | 0.996 |
| 116 | 1.010 | 1.004 | 0.999 |

For pixels beyond 116, the fringe pattern is small, and factors of unity in the above table are to be used.  Interpolation or extrapolation from the above table to the optics temperature of interest gives the expected fringe pattern.  Division of the intensity spectrum from the data by the fringe pattern gives the corrected intensity spectrum without fringes.